## TOC
| No. | Title | rows | Implementation
Within 7-4-2021 Rev 04 (for public gpt2) | 1 | Fabian brochure | 51 | Y
| 2 | user need and market requirement | 85 | Y
| 3 | product requirement specification | 748 | Y
| 4 | IFU | 598 | N
| 5 | System architecture | 182 | Y
| 6 | Hardware subsystem(gas, pneumatic) | 85 | Y
| 7 | Hardware subsystem(others) | 200 | N
| 8 | Software subsystem (9+1) | 292+ | N | HFO GUI: 1030
| 9 | Viewer (CIR, DIR, Market, clinical_workflow, clinical_HA) | 244 | Y
12-8-2021 Rev 05 (for public gpt2 + register simple transformers) | 10 | Viewer (CIR, DIR/SHA, Market, clinical_workflow, clinical_HA) | 244 | half Y | add full predications for register users
| 11 | SRS | 200 | half Y | wait for Alex inputs for cause, mitigations, update harm, severity
| 12 | FDA Recalls, FDA Maude (Fuji film, GE, Simens SYNGO imaging system, IBM watson, AGFA) | 549
| 13 | iPlatform (FW, App, Lap-App) | 401
| 14 | Zeus
| 15 | Ares
| 16 | Athena
| 17 | User Manual

## TrainingData_simpleTransformer
Context | Failure Mode | Hazard | Hazardous Situation (HS) | Failure Causes | Harm | Severity | Probability | Safety Design | Protective Measure | Information for Safety | Comment
Subsystem: The system shall render the crucial information from the DICOM data on every viewing panel. The information shall include . Modality Acquisition date Patient name Patient date of birth Patient gender | missing data | patient data issue | Loss of Patient Data during during intra-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
If the DICOM image contains scale information, the viewer will render one ruler with marks near the left edge of the 2D viewing panel. Depending on the zoom level, the ruler length will be adjusted. For each ruler length, there are either 5 or 10 marks. The total ruler length is displayed on the bottom end of the ruler. | fail to adjust size | application program issue | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
If the DICOM image contains information about patient orientation in space, there is a three-dimensional indicator displayed in the bottom right corner on each viewing panel. There are also characters (H, F, L, R, A, P) displayed for patient orientation on the edges of each 2D view. | incorrect display format | no display or display failure | Inappropriate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall support different DICOM modalities: MRI, CT, and X-Ray. | support failed | application program issue | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to open a DICOM series | difficult to open | improper or incorrect procedure or method | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall use the DCMTK library for opening the DICOM series | unable to download library | -defective component -improper or incorrect procedure or method -inadequate storage | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall be able to interface with cloud storage to retrieve DICOM image data. | fail to retrieve data | data issue | Unable to Generate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
A software component (in the following: DSP Agent) that interfaces with the cloud storage system and retrieves a DICOM compliant study to local storage (target folder). | fail to retrieve data | data back-up problem | Unable to Generate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The DSP Agent verifies complete retrieval from the cloud storage. Once the retrieval is completed, the DSP agent changes the name of the downloaded object from a temporary name to the final name (e.g. for single files “.image1.dcm” to “image1.dcm”, or folders “.study1” to “study1”). | fail to retrieve data | data issue | Inaccessible Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall test the validity of the DICOM image data when loading the medical imagery. | validation failed | protective measure issue | Unable to Generate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall display a loading screen when loading and/or processing data | fail to monitor loading images | no display or display failure | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display a loading screen while loading datasets into memory. This screen shall show an animation to indicate activity. | fail to monitor loading images | no display or display failure | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Each overlay shall have a distinct color which matches the color displayed on the 3D view. | missing color code | incorrect display | Inappropriate Patient 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall provide the option to visualize the overlay as either a semi-transparent infill or contour visualization. For the semi-transparent infill visualization, users can modify the opacity for each structure separately. | fail to adjust opacity | device difficult to setup or prepare | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Users can change the opacity for the semi-transparent infill visualization in the “3D Model” toolbar drawer. | fail to adjust opacity | optical issue | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall allow for toggling the visibility of the pre-operative image data in the 3D view by left-clicking (or tapping) the “eye ” icon in the 3D viewing panel. This allows the user to view the combination of pre-operative image data and 3D model, or just 3D model. The transparency of the 3D model objects can be changed using the list in the control panel. | fail to toggle | device operates differently than expected | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The “eye” button/icon allows the user to toggle the pre-operative image data on and off. | fail to toggle off 2D images | data issue | unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display the Multi-Planar Resection (MPR) cut planes for each visible 2D view. | unable to advance mpr cut plane | device operates differently than expected | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The amount of image that is visible within one view shall be indicated as a portion of the cut plane on the other views. | unable to advance cut plane | application program issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Patient PHI including acquisition date shall be able to be hidden or displayed based on user selection. | fail to hide display | incorrect display | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display the DICOM meta data in the 2D views - Patient name - DOB - sex - study acquisition date. - Series name | fail to display DICOM data | image display error | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to save the display adjustments made for a 2D view. | fail to save configuration | use of incorrect control settings | Unable to Save View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to restore previously saved display adjustments for 2D view | fail to restore previous adjustments | use of incorrect control settings | Unable to Retrieve Previous View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall render multiple 2D views of the preoperative medical image. | unable to render 2D views | application program issue | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to display three MPRs and a 3D rendering view. | unable to render 3D views | application program issue | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall provide a visible representation of the orientation of the patient in the 2D view. | incorrect patient orientation | failure to align | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall allow the user to define a beginning and ending point for measurement. | fail to pick end point | incorrect measurement | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user activates the measurement mode by left-clicking or tapping on the “Distance” button in the control panel. Image manipulation interactions are now locked (e.g. pan, zoom, rotation). Then the user can left-click or tap to place the start point followed by the end point within the same viewing panel (regardless which data is shown). | fail to deactivate mode | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The distance mode automatically ends after completing the measurement, and the previously selected interaction mode (e.g. scroll) continues. The user can enter the measurement mode using the “Distance” button repeatedly. | loss of measurement | incorrect measurement | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The maximum number of distance measurements shall be limited to 30. | distance out of range | incorrect measurement | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Measurement can be performed on DICOM-compliant data sets without calibration (e.g. no relationship between physical size and voxel/pixel provided) (e.g. X-ray imag es). The ruler overlay/measurement tool is not shown in this case. | fail to measure | measurement system incompatibility | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: After performing a measurement (see 3.2), the measured length (see 3.1 for units and formatting) will be shown next to the end point of the measurement in the viewing panel, and in the list of all measurements in the control panel. The list of all measurements in the control panel will further include an automatically assigned descriptor to the measurement element (e.g. “Distance 1”). | fail to display measurement value | no display or display failure | Inappropriate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall use millimeters for all measurements. The unit is abbreviated by mm. The accuracy is limited to one digit after the decimal. The decimal symbol must be localized (e.g. “12.3 mm” in the US and “12,3 mm” in France and Germany). There shall be one space between the number and the unit. | incorrect unit format | data issue | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall accept various HID inputs for performing zooming operations. | unable to zoom operation | application program issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The mouse wheel scroll shall zoom in and out. | unable to zoom in views | application program issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Middle-click and drag shall zoom in and out. | fail to drag operation | human-device interface issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Double right-click and drag up or down will zoom in 3D view | fail to drag operation | human-device interface issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom shall be done in real-time as input is made | zoom slow | device sensing issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom step shall be configurable via a configuration file | inappropriate zoom step | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The zoom out shall be limited to the bounding box for the 3D models | inappropriate zoom out limit | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The zoom out shall be limited to the bounding box for the 2D slice | inappropriate zoom out limit | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom in for 3D models shall be limited to 2-5mm or 1.2% of the bounding box. | inappropriate zoom in limit | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom out shall be limited by the dimensions of the DICOM image volume. The anatomical structure is not used in determining image extents. | inappropriate zoom out limit | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom in for 2D models shall be limited to 2% of the bounding box. | inappropriate zoom in limit | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The zoom Step size shall define the magnification made per step. | inappropriate zoom step | use of incorrect control settings | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom shall be performed by moving the camera closer to the 3D volume. | inappropriate zoom operation | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom shall be performed by magnification for a 2D image. | inappropriate zoom operation | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom to extents for 3D views are performed when an axis shortcut is pressed. | inappropriate zoom operation | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The ruler that is part of the 2D view will change measurement reading as zoom is performed. | incorrect readings | incorrect measurement | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Zoom shall apply to MPRs and the 3D view independently. | inappropriate zoom operation | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
For MPRs, the default zoom level is the same for each panel, and is defined to allow the image to fill the panel, but no part of the image is clipped in either panel. The default zoom level of the 3D view is defined such that the bounding box of all visible structures/meshes fits the camera viewport. | inappropriate zoom level | use of incorrect control settings | Inappropriate Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall accept various HID inputs for performing panning operations | inappropriate pan operation | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Right click and drag moves the visible area within the viewing panel. | fail to drag operation | use of device issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Right click and drag the image in the direction of mouse movement. | fail to drag operation | use of device issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
For 3D, right click and drag move the camera in the opposite direction of the mouse movement | fail to drag operation | use of device issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Movement of the image is proportional to the mouse movement. | unproportional movement | unintended movement | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Movement is allowed both vertically and horizontally. | inappropriate movement direction | unintended movement | Unable to Manipulate View during intra-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
MPR cut plane indicator shall be updated in the other views when the image is panned | fail to update indicator | no display or display failure | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The pan interaction is implemented as right-click and drag with mouse. | unable to pan image | human-device interface issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The pan is independently used for each MPR and 3D view. | fail to pan independently | use of device issue | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The default pan is set so that the images are centered in the panels. | fail to retrieve default setting | use of incorrect control settings | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The "Re-Center" button in the control panel resets the pan functionality. | fail to reset pan | reset issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
There is no max and min pan, the user can drag the image outside the view and continue dragging it. The user can drag the image back into view by dragging the background. | fail to drag operation | use of device issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The patient orientation is changed in coordination with the displayed anatomical structure. | incorrect patient orientation | failure to align | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Left click and drag rotates the 3D model in the view | fail to rotate 3D model | use of device issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Center of rotation is the center of the bounding box of the displayed segmentations. Those that are visible | rotation out of range | incorrect display | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall allow for rotation of the 3D view by left-clicking (or tapping) and dragging across the viewing panel. The interaction metaphor follows the virtual sphere rotation. The user can reset the orientation of the 3D view using the reset button in the control panel. | unable to rotate 3D view | application program issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall update in real-time the movement of the MPR cut planes. | fail to provide real time movement | unintended movement | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to generate an 2D image from a cut plane passing through a 3D model. | fail to generate 2D images | application program issue | unable to Generate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall allow users to turn on and off the visibility of segmentations. | fail to toggle visibility | no display or display failure | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Visibility of a segmentation is indicated by the check box within the U/I | fail to indicate visibility | no display or display failure | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display the name of each segmentation. | inadequate message display | incorrect display | Inappropriate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display each segmentation in its own color. | display color out of specification | incorrect display | Inappropriate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The opacity of a segmentation can be adjustable | fail to adjust opacity | optical issue | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Opacity shall be indicated by a percentage | fail to provide opacity percentage | optical issue | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Opacity can be changed via a slider | fail to adjust opacity | optical issue | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The volume of the segmentation structure shall be displayed | fail to display volume | no display or display failure | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Volumes shall be displayed by the software in SI units | inappropriate volume unit | incorrect measurement | Unable to Measure 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Each segmentation has its own color, though there may be duplications | inappropriate color scheme | incorrect display | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall allow the user to reset the parameters for visibility and opacity to defaults as provided by the vendor. | fail to reset default parameters | reset issue | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
If the segmentation is visible, the segmentation shall be shown in the MPR views with the same opacity as set for the 3D view. | incorrect opacity | optical issue | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Segmentation displayed in the MPR views can be viewed as a solid with opacity or as an outline of the segmentation. | fail to view as solid opacity | incorrect display | unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Segmentations can be hidden or shown within the MPR views. | fail to hide segmentation | incorrect display | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall visualize all structures provided as part of the 3D model metadata in the list of volumes in the toolbar. | fail to review 3D model | application program issue | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Note that the names of the segmentations are provided as part of the 3D model metadata. Changes to names (e.g. capitalization) shall not be performed by the Image Viewer application. | undesired name changes | patient data issue | Inaccessible Patient Data during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can select an item in the list. After the user selects an element of the list, the cut planes of the MPR automatically advance to intersect with the center of the object, and the translation of the image (pan) automatically positions the object at the center of the viewing panel. | unable to advance cut plane | application program issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall visualize the volumes if provided as part of the 3D model metadata as part of the list of volumes in the control panel. | fail to visualize volume | no display or display failure | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Following 3.1, all units are provided and displayed in cubic centimeters (“cc”). Two digits after the decimal shall be shown. The volumes will be displayed as “12.30 cc” or “12,30 cc”. The list of objects (segmentations) shall read at least: Organ Volume arteries 15.20 cc portal vein 27.86 cc liver 945.70 cc | incorrect unit format | data issue | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The volumes are provided as part of the 3D model. Should the accuracy (digits after the decimal) be higher than two decimal places, the Image Viewer application shall round to the nearest hundredth (number with two decimal places). For instance, 0.3161 → 0.32; or 9.991 → 10.00; or 0.0019 → 0.00 | low volume accuracy | volume accuracy issue | Inappropriate Patient 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall render the objects/segmentations of the 3D model in the Visible Patient model container as part of the control panel. This information includes (i) object name, (ii) object volume in cc, (iii) object color, (iv) object transparency setting. See 2.17 through 2.19. | inadequate 3D model data | data issue | Inappropriate Patient 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The webviewer instances of the Medical Image Viewer will not allow the user to render the modelling report. The user can access the modelling report through the webportal | fail to generate report | data issue | Inaccessible Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
A system that enables measurements and annotations within datasets | unable to measure | measurement system incompatibility | Unable to Measure 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The user can set the transparency for each object (segmentation) of the 3D model in the control panel. Each 3D model object is listed in the 3D Model drawer, and for each object the user can use a slider to change the transparency from min (transparent = 0.0) to max (opaque = 1.0). No detailed transparency percentage is provided. The default transparency levels are loaded from the 3D model meta information | fail to set transparency | optical issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can reset the transparency levels to the default transparency levels using the reset button inside the 3D Model drawer in the control panel. The list of objects (segmentations) shall contain following information: Display on/off Organ Volume Color/Transparency checkbox arteries 15.20 cc Color picker button checkbox portal vein 27.86 cc Color picker button checkbox liver 945.70 cc Color picker button | fail to reset transparency level | optical issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Note that the column “Organ” is described in 2.17, and “Volume” in 2.18. Note that “Volume” can only be displayed if the metadata is provided. | fail to visualize volume | no display or display failure | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Note: Change of color was descoped from R1 and moved to later release. | inappropriate color scheme | incorrect display | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall provide a visible representation of the orientation of the patient in the 3D view. | fail to visualize orientation | no display or display failure | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The orientation of the patient representation shall follow the rotation of the camera around the 3D model. | incorrect patient orientation | failure to align | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to display a volumetric rendering of the CT or MRI data. | fail to provide volumetric rendering | application program issue | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study. | unable to render 3d views | application program issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to display the volumetric rendering of the CT and MRI data in conjunction with the 3D segmentations visible within the same view. | fail to display volumetric rendering | no display or display failure | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to control the window width and level values for a volumetric rendering. | unable to adjust window width | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall define presets that are available for different window width and level values, e.g., CT Bone, CT Lungs, etc. | inappropriate preset value | use of incorrect control settings | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall provide the ability to automatically set the window width and level values to show the entire intensity range of the image. | fail to auto-set window | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Presets for window width and level shall be able to be defined in a configuration file. | fail to define presets | device difficult to setup or prepare | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to save the display adjustments made for a 3D view. | fail to save adjustments | device difficult to setup or prepare | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to restore previously saved display adjustments for 3D view. | fail to restore previous adjustments | reset issue | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall allow patient PHI to be hidden or displayed based on user selection. - Patient name - DOB - Sex - and study acquisition date. | fail to hide display | incorrect display | Unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall display DICOM meta data in the 3D view. | fail to display DICOM data | no display or display failure | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall have the capability to render the pre-operative image data and the Visible Patient dataset (“3D models”) in a 3D viewing panel. | fail to render image data | data issue | Unable to Display Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
If there is no 3D dataset available, the 3D view shall be empty and the system shall display an informative message. | inadequate message display | incorrect display | Inappropriate Patient 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
If the pre-operative volumetric image data and/or 3D model is available, the system shall provide the ability to display a 3D view of the data set. | fail to display 3D view | no display or display failure | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The volumetric rendering and 3D reconstruction datasets shall be displayed in the same coordinate system. | fail to display in the same coordinate system | incorrect display | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Users shall be able to switch on or off the volumetric rendering based on DICOM images in the 3D view. | fail to toggle visibility | incorrect display | Unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
using the “Reconstruction” check box in the “Viewing” toolbar drawer | inadequate check box | human-device interface issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall provide the option of changing the window level and window width for the pre-operative volumetric image data. | unable to adjust window width | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall allow users to modify window level and window width. | unable to adjust window level | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Users shall be able to choose between presets and manual windowing, see DIGITAL-DIR_DH-128. | fail to toggle | use of device issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
For the 3D model data, the system shall render the objects using the color and transparency settings for each object. | inappropriate render color | use of incorrect control settings | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall provide the ability to place virtual clips to simulate resection and to view volume simulation results. | fail to place virtual clip | application program issue | Delayed Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can start the volume simulation of resection by selecting the “Clip Apply Model” toolbar drawer. Once the toolbar drawer is active, just 3D viewing panel is shown and other MPR views are hidden. Interactions with the 3D view stay the same as outlined before. The clip applying mode remains active until the user deactivates the clip applying mode by either deselecting the “Clip Apply Model” toolbar drawer, selecting another toolbar drawer except “Annotation” or “My Views”, or clicking the “X” button overlaid on the 3D view. | fail to simulate resection | use of device issue | Delayed Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can hover over a specific structure of the organ, called “network mesh” from now on, visualized in the 3D viewing panel. The network mesh for different datasets are the following; for a lung dataset the bronchial branches, for liver the portal vein network, and for kidney the renal arteries. | fail to generate network mesh | application program issue | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
On hovering, colored highlights are shown dynamically indicating which parts of the network mesh are available to place clips on. The interaction point defines the origin of a ray that intersects with the network mesh. | fail to highlight color on | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can apply a clip by right-clicking on the network mesh. The system shall render the simulation of the resection for the given clip. The color highlight on the network mesh then stays without hovering over it. | fail to click on the network mesh | application program issue | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can start the volume simulation using the “Clip” button in the control panel. | fail to generate volume simulation | application program issue | Inappropriate Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
After at least one clip is applied, the system shall provide the following information. -Targeted organ, -Simulated resected part per clip, -Simulated total resected part (= sum of “Simulated resected part per clip”), -Simulated remaining part -Target organ nodules, -Healthy organ, -Simulated remaining healthy | inadequate clip info | data issue | Loss of Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
None of these volumes are calculated by the Image Viewer application. All volumes are provided through the 3D model metadata. The Image Viewer simply performs a user-based data selection, aggregation and visualization. | incorrect volume calculation | parameter calculation error due to software problem | Inappropriate Patient 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Following 3.1, all units are provided and displayed in cubic centimeters (“cc”). The volumes will be displayed as “12.30 cc” or “12,30 cc”. | incorrect unit format | data issue | Incorrect 3D Model Measurement during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The user can advance (“scroll”) the MPR cut plane through the image volume using left-click and drag mouse while placing the mouse cursor in the target view panel, or switch into scroll mode using the scroll button in the control panel and tap + swipe in the target view panel. The user then can advance the MPR cut plane through the volume for each view panel individually. | unable to advance cut plane | application program issue | unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
the total number of image elements in the respective direction and the current position of the MPR cut plane are shown. For instance: “Slice 100/212”. The step size of the advancement is defined by the spacing of the native scanning/imaging technology. Slice numbers will only be displayed for the principle orientation. | fail to display cut plane | no display or display failure | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall restrict the movement of the MPR to the extents of the 3D model volume. | uncontrolled movement | unintended movement | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall list all annotations in the annotation (or measurement) drawer in the control panel. The user can select an annotation by left-clicking or tapping on the line item. When selecting an annotation, the cut planes for each MPR advance to intersect with annotation and the translation (pan) is set to center the annotation on the screen. After selecting a line item, the system shall show a “x” button in the selected line to allow for deletion of the annotation. (Compare to 3.3) | fail to place annotations | use of device issue | Unable to Annotate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
A system that enables the simulation of a resection by clip applying to approximated segments. | fail to provide simulation resection | application program issue | Loss of Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
I need to view volumes of the target anatomy and pathology and of organ segments to be resected based on virtual, simulated clips I placed , AND | fail to view resected target anatomy | application program issue | Delayed Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology | fail to manipulate visualization parameters | poor quality image | Unable to Manipulate 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The system shall provide the ability to place an arrow as an annotation object. | fail to place arrow annotation | application program issue | Unable to Annotate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can enter the arrow annotation mode by left-clicking or tapping the “arrow” button in the control panel. The system freezes/pauses all image manipulations when in arrow annotation mode. The user then places the start point followed by the end point. The system draws a line from start point to end point and places an arrow head at the end point. The arrow annotation mode is completed after the user places the end point. | fail to place arrow annotation | application program issue | Unable to Annotate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can enter the text annotation mode by left-clicking or tapping the “text” button in the control panel. The system freezes/pauses all image manipulations when in text annotation mode. The user then places an anchor point in which the system places a text box. The user can type or paste text by left-clicking inside the text box. | fail to add any notes | human-device interface issue | Unable to Annotate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
All annotations are automatically listed in the annotation (or measurement) drawer in the control panel. Arrow annotations are automatically numbers and displayed with the index, e.g. “Arrow 1”. Text annotations are displayed using the text entered by the user, e.g.: “Text: User input”. | fail to number all annotations | application program issue | Unable to Annotate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall be able to run on hardware running Ubuntu Linux. | fail to run ubuntu linux | computer operating system issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The software shall be able to be hosted as a web-based application accessible through a web browser. | fail to access through web browser | computer system security issue | Inaccessible Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to run on a virtual machine deployed on an Azure server. | fail to run on virtual machine | application program issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The software shall support the following browsers: Microsoft Edge version 93 or greater, Google Chrome version 93 or greater, Apple Safari Version 14.0 or greater. | fail to support browsers | application program issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The system shall have the capability to stream the rendered view of the application from the rendering engine to the user. | fail to stream rendered view | incorrect display | Unable to Generate Stream during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall have a thin browser-side client interface, which allows for display of the video stream provided by the system. | fail to display video stream | incorrect display | Unable to Generate Stream during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The stream shall always be rendered as 1080p for an aspect ratio of 16:9. | low rendered resolution | Image Resolution Poor | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The stream shall be compressed and transmitted using the following codecs, supporting following combinations of browser applications and operating systems: | fail to transmit stream | data issue | Unable to Transmit Stream during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall use a GPU assigned to the virtual machine to perform off-screen rendering of the entire UI. | fail to run GPU | operating system becomes nonfunctional | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The minimum specification of the server shall be: • Ubuntu 18.04 LTS x64 • NVIDIA GPU with minimum 4GB of GPU memory that allows for rendering and encoding and decoding using the NVIDIA H.264 codec • 256 GB SSD drive • Minimum of 8 vCPUs • 16 GB RAM memory | low server configuration | use of incorrect control settings | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
User inputs shall be captured and forwarded to the Image Viewer app using a thin browser-side client interface. | fail to capture inputs | computer software issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
This interface shall display the rendered view provided by the system and capture interactions within the display area. | fail to display view | no display or display failure | Unable to Display View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Interactions shall include mouse events (position, buttons, and scroll), touch events (tap, drag, zoom via 2 finger pinch), and keyboard inputs (keyboard inputs may be deferred to R2). | lack of events | Communication or Transmission Issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The interactions shall be sent as JSON messages with the ability to determine order of interactions and time of event. | fail to send JSON messages | Communication or Transmission Issue | Unable to Transmit Stream during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall indicate the recommended resolution of the user’s browser window | low browser window resolution | Image Resolution Poor | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall inform the user that the target client resolution is 1080p if the browser-side client interface detects a resolution of the display area lower than 720p in one or both directions (1280 x 720px). | fail to detect low resolution | Image Resolution Poor | Poor Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be deployed on the Digital Hub and viewable through devices connected to the Digital Hub | fail to deploy software on hardware | computer software issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The hardware shall use a GPU perform on- and off-screen rendering of the entire UI. The minimum specification shall be: • Ubuntu 18.04 LTS x64 • NVIDIA GPU with minimum 4GB of GPU memory that allows for rendering and encoding and decoding using the NVIDIA H.264 codec • 256 GB SSD drive • Minimum of 8 CPU cores • 16 GB RAM memory | low hardware configuration | use of incorrect control settings | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
The hardware shall enable user inputs using mouse and keyboard , or a touchscreen. | lack of mouse input | human-device interface issue | Software Unresponsive during intra-operative |  | No harm | 0 |  |  |  |  | SW_FMEA
For each interaction modality, the interactions are defined in the respective interaction requirement. Text input can be performed using real or virtual keyboards. The virtual keyboard is rendered by the operating system or Hub infrastructure. The viewer app expects one single keyboard stream, and the operating system is responsible for the localization of the virtual keyboard. | lack of interaction requirement | data issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall create a rendered stream of the UI that is subsequently distributed by the Hub infrastructure. | fail to create rendered stream | data issue | Unable to Generate Stream during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The viewer app shall render one single stream at 1080p and up to 60 frames per second.No other streams will be rendered. | low rendered frames | Image Resolution Poor | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Close button overlay in top right corner of each viewing panel | fail to indicate button status | no display or display failure | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Switch in toolbar for toggling visibility of 3D view and 2D views. If switch for 2D views is on, visibility of individual MPR view panels can be set via icons. | fail to toggle visibility | incorrect display | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: To hide a view, the user can lift-click with a mouse or tap the (X) button in the top right corner of each view (MPR and 3D) to remove this panel from the layout. | fail to hide view | incorrect display | unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The Window Center and Window Width (see DICOM PS3.3 2020e, 0028,1050 and 0028,1051) changes which data is visualized in all MPR and 3D views simultaneously. | unable to adjust window width | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The Window Center defines the center point on the histogram of all gray scale values in the image around which values are selected. | fail to define window center | dislodged or dislocated | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The Window Width defines how close values need to be to the Window Center to be selected. | fail to define window width | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The default Window Center is centered to the histogram, the default width is set to the full width of the histogram. Therefore, the default Window Width and window center include all gray scale values in the image, and shows the entire intensity range of the image. | fail to define default window center | dislodged or dislocated | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can reset Window Width and Center using the “Reset” button in the toolbar. | fail to reset window width | failure to reset | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
For CT images only, the user can select predefined Window Width and Center from a predefined list. The presets include: CT Liver (W:150 C:30), CT Abdomen (W:400 C:50), CT Lungs (W:1500 C:-600), CT Mediastinum (W:350 C:50), CT Bone (W:1800 C:400) | fail to retrieve preset window width | application program issue | Unable to Display Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
After/during changing the Window Center and Width, the newly selected part of the histogram is mapped linearly to the visualizable range from white to black. | fail to map linearly | data issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The window width and window center values can be manually adjusted. Adjustment via sliders in the toolbar “Windowing” drawer. | fail to adjust window manually | application program issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The software shall be able to reset the window with and window center to the initial default values. | fail to reset window | failure to reset | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Resetting via the button “Reset” in the toolbar “Windowing” drawer. | fail to reset | failure to reset | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The software shall be able to re-center all view panels on user request. | fail to reset view | failure to reset | unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The user can use the “Re-Center” button in the toolbar to change following viewing properties for all hidden and visible view panels simultaneously | fail to center view | reset issue | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Reset positioning (pan) to default values | fail to reset position | failure to reset | unable to Manipulate Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Reset cut plane to their default (sagittal, longitudinal, axial) | fail to reset default settings | failure to reset | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The zoom level stays the same | unstable zoom level | unstable | Inappropriate View Manipulation during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The software shall be able to toggle the visibility of 3D view and MPR views (axial, sagittal, coronal projection) on user request. The view panels are laid out automatically based on predetermined layouts. | fail to toggle visibility | incorrect display | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Switch in toolbar for toggling visibility of 3D view and 2D views. If switch for 2D views is on, visibility of individual MPR view panels can be set via icons. | fail to toggle visibility | no display or display failure | unable to Display 3D Model during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The layouts include (I) the default 2x2 view (three MPR and one 3D view on the top left), (II) a three view panel view, (III) a 1x1 (side-by-side), and (IV) a maximized “focus-view”. | inadequate panel view | incorrect display | Unable to Display View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The software shall be able to change how much screen space each view takes up, that is, resize the view panels. | unable to resize viewing panel | reset issue | Unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Resizing is done by left clicking on a view border and drag. | fail to resize image | reset issue | unable to Manipulate Patient Image during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall provide a method for saving the current workspace and restoring the workspace to a previously stored version. | fail to restore previous workspace | loss of data | Unable to Display Window during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Subsystem: The Viewer software shall be able to save the current layout of views to a workspace. | fail to save workspace | data back-up problem | Unable to Save View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
This includes all view parameters of 2D and 3D views and changes to view parameters of 3D models as well as windowing and annotations. If we are in clip apply mode this includes placed clips. | fail to manipulate parameters | application program issue | Delayed Pre-Operative Plan during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
The system shall have the ability to rename a saved view | unable to rename view | device markings issue | unable to Manipulate View during pre-operative |  | Pre-operative Delayed | 0 |  |  |  |  | SW_FMEA
Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer. | fail to compatible with regular computer | Incompatibility Problem | [8 empty] | DIR
The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | fail to open images | image display error | [8 empty] | DIR
The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | inappropriate image format | data issue | [8 empty] | DIR
The system shall have the capability to render 2D views of the preoperative patient images. | unable to render 2D views | application program issue | [8 empty] | DIR
The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study. | unable to render 3D views | application program issue | [8 empty] | DIR
The system shall have the capability to render a view of 3D model segmentations | unable to render 3D views | application program issue | [8 empty] | DIR
The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view | unable to render 2D overlay on 3D segmentation | application program issue | [8 empty] | DIR
The system shall provide a method to zoom in and out on the views. | unable to zoom in views | application program issue | [8 empty] | DIR
The system shall provide a method to pan a viewing panel. | unable to pan viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to rotate a 2D viewing panel. | unable to rotate 2D viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to rotate a 3D viewing panel. | unable to rotate 3D viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel. | fail to toggle image visibility | application program issue | [8 empty] | DIR
The system shall provide a method to rotate the cut plane within an MPR view. | unable to rotate cut plane | application program issue | [8 empty] | DIR
The system shall provide a method to undo and redo user interactions. | fail to undo user interaction | use of device issue | [8 empty] | DIR
The system shall provide a method to undo and redo user interactions. | fail to redo user interaction | use of device issue | [8 empty] | DIR
The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size. | unable to resize viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method for hiding a viewing panel. | unable to hide viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to adjust the Window Center and Window Width adjustments using presets. | unable to adjust window center | application program issue | [8 empty] | DIR
The system shall provide a method for manually adjusting window width and center values for the intensity range | unable to adjust window width | application program issue | [8 empty] | DIR
The system shall provide a method for a reset all viewing panels simultaneously | unable to reset viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to select the layout for the viewing panels | unable to select layout for viewing panel | application program issue | [8 empty] | DIR
The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume. | unable to advance MPR cut plane | application program issue | [8 empty] | DIR
The system shall be able to display metadata contained in the DICOM data. | fail to display DICOM data | image display error | [8 empty] | DIR
The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects. | fail to list 3D models | application program issue | [8 empty] | DIR
The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model. | fail to display volumetric sizing information | no display or display failure | [8 empty] | DIR
The system shall be able to display metadata contained in the 3D model container | fail to display metadata | no display or display failure | [8 empty] | DIR
The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels | fail to display semi-transparent overlay | application program issue | [8 empty] | DIR
The system shall use SI units and their derivatives | inappropriate measurement unit | incorrect measurement | [8 empty] | DIR
The system shall provide a method for performing one or more measurements within one of the viewing panels | unable to perform multiple measurements | application program issue | [8 empty] | DIR
The system shall label and list measurements automatically | fail to label measurement automatically | application program issue | [8 empty] | DIR
The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object. | fail to place annotation on 3D model | application program issue | [8 empty] | DIR
The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel. | fail to edit notes | application program issue | [8 empty] | DIR
The system shall allow for simulation of resection by applying a clip to approximated segments. | fail to simulate resection | application program issue | [8 empty] | DIR
The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip. | fail to display volumetric sizing information | no display or display failure | [8 empty] | DIR
The system shall provide the ability to support individual language packs. | inadequate language pack | inadequate user interface | [8 empty] | DIR
The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language. | inadequate language pack | inadequate user interface | [8 empty] | DIR
The system shall, when appropriate and available, use standard internationally recognized symbols. | unstandardized symbol | device markings issue | [8 empty] | DIR
The system shall, when appropriate and available, use standard internationally recognized icons. | unstandardized icon | device markings issue | [8 empty] | DIR
The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System. | fail to record operative data | data back-up problem | [8 empty] | DIR
The System shall record all functions performed by the System pre, intra, and post operatively. | fail to record operative data | loss of data | [8 empty] | DIR
The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel. | fail to record monitoring data | data back-up problem | [8 empty] | DIR
The System shall record system functions and operational events generated by the System. | fail to record operational events | data back-up problem | [8 empty] | DIR
The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning. | fail to create pre-op planning | application program issue | [8 empty] | CIR
The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images. | incomplete 3D model | application program issue | [8 empty] | CIR
All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated. | incorrect revision | device markings issue | [8 empty] | CIR
3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator. | fail to create 3D Model | poor quality image | [8 empty] | CIR
3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon. | fail to provide patient diagnosis info | use of device issue | [8 empty] | CIR
Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components. | software corrupted | computer software issue | [8 empty] | CIR
DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data. | image not in DICOM format | image display error | [8 empty] | CIR
PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images. | fail to process medical imaging | application program issue | [8 empty] | CIR
Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format. | image not in DICOM format | data issue | [8 empty] | CIR
Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively | fail to refer pre-op plan intraoperatively | application program issue | [8 empty] | CIR
3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models | fail to upload 3D model | application program issue | [8 empty] | CIR
Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | inadequate IFU | inadequate instructions for healthcare professional | [8 empty] | CIR
Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | lack of video | nadequate instructions for non-healthcare professional | [8 empty] | CIR
Marketing and Sales Representative: Internal users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies. | fail to monitor product use | Market Issue | [8 empty] | CIR
View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to manipulate visualization parameters | application program issue | [8 empty] | CIR
Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings. | unable to measure point to point | application program issue | [8 empty] | CIR
2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings | unable to develop surgical plan | application program issue | [8 empty] | CIR
View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to review 3D model of pathology | application program issue | [8 empty] | CIR
Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings | fail to manipulate visualization parameters | application program issue | [8 empty] | CIR
Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | fail to provide simulation resection | application program issue | [8 empty] | CIR
Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | incorrect volume calculation | parameter calculation error due to software problem | [8 empty] | CIR
Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure image display | application program issue | [8 empty] | CIR
Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure 3D model display | application program issue | [8 empty] | CIR
Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display 3D model at OR monitor | incompatibility problem | [8 empty] | CIR
Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display patient image at OR monitor | incompatibility problem | [8 empty] | CIR
Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device. | clinical workflow interrupted | disinfection or sterilization issue at user location | [8 empty] | CIR
Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice. | patient data breached | unauthorized access to computer system | [8 empty] | CIR
Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently. | lack of local language | human-device interface issue | [8 empty] | CIR
Indications for Use for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party. | Indications for Use misalignment | failure to align | [8 empty] | CIR
Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain regulatory clerance | compliance issue | [8 empty] | CIR
Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain registration | compliance issue | [8 empty] | CIR
Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to capture data | data issue | [8 empty] | CIR
Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to store data | data issue | [8 empty] | CIR
Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to follow privacy regulation | compliance issue | [8 empty] | CIR
Branding: As company marketing representative, I need the application to follow company defined branding, so that I can maximize consistency in the look and feel of company digital products that may be used by the same user base and reduce cognitive load on the users when using company digital products. | branding misaligned | compliance issue | [8 empty] | CIR
Demo Capabilities: As company marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively. | unclear workflow | data issue | [8 empty] | CIR
Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body. | incorrect procedure | bacterial contamination of device | [8 empty] | CHA
Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | [8 empty] | CHA
3D projections may be incorrect when the pair of images used is a secondary capture generated from a series containing more than two images, where one of the images has been manually flipped. | image flipped | image reversal | [8 empty] | FDA1
Projection errors may lead to errors in calculation of 3D clinical parameters when adjusting the position, size and shape of the 3D objects used. | projection error | image display error | [8 empty] | FDA1
Projection errors may lead to errors in calculation of 3D clinical parameters when adjusting the position, size and shape of the 3D objects used. | incorrect parameter | use of incorrect control settings | [8 empty] | FDA1
Incorrect clinical parameters can be a contributing factor in incorrect clinical decisions, or sub-optimal surgical interventions. | incorrect clinical decision | poor quality image | [8 empty] | FDA1
A customer experienced when using IMPAX CV Reporting software, specifically, when building a NIV report, the NIV Cardio report was showing incorrect findings for Aneurysm. | false output | computer software issue | [8 empty] | FDA1
A database handling error could occur during the image acquisition process affecting the completeness of acquired images with Centricity PACS-IW. | incomplete image | image display error | [8 empty] | FDA1
There is a potential that one or more image series (i.e. all images within an image set) may be missing from an exam without indication to the user. While this is rare, this can occur with imaging studies that consist of a very small number of images per series. E.g. CR Thorax exam with 1 image per series. | missing image series | image display error | [8 empty] | FDA1
A defect component of the Compano reader power supply can cause the power supply to burn and some smoke to leave the housing. | power supply burned | Power Source Issue | [8 empty] | FDA1
A dentist found implant positions are not correctly exported from the implant planning software SICAT IMPLANT V2.0 - for the specific export format CMG.DXD. | incorrect position calculation | computer software issue | [8 empty] | FDA1
A migrated study that has annotations will display the annotations, but when a prior study is loaded that should have annotations, the second study will not display annotations. | fail to display annotations | no display or display failure | [8 empty] | FDA1
A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | incorrect image measurement | incorrect measurement | [8 empty] | FDA1
A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | fail to display calibration | image display error | [8 empty] | FDA1
A product issues happens for RF projection images. If user measures on RF image, the .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | fail to display label | image display error | [8 empty] | FDA1
A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | incorrect image measurement | incorrect measurement | [8 empty] | FDA1
A product issues happens for RF projection images. The .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | fail to display calibration | image display error | [8 empty] | FDA1
A report can be confirmed with the incorrect patient demographics resulting in it being saved under the incorrect patient record in the Electronic Medical Record (EMR). | incorrect patient record | patient data issue | [8 empty] | FDA1
A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | multiple accesses at once time | computer software issue | [8 empty] | FDA1
A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | missing read-only notification | device alarm system issue | [8 empty] | FDA1
A software defect negatively affects the quality of images taken during diagnostic imaging (mammograms) which may result in misdiagnosis. | poor quality image | therapeutic or diagnostic output failure | [8 empty] | FDA1
A software defect that has the potential to result in image misdiagnosis and incorrect treatment of a patient. | software defect | therapeutic or diagnostic output failure | [8 empty] | FDA1
A software defect was discovered that causes images to be out of context with clinical information. | missing clinical information | image display error | [8 empty] | FDA1
A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. | washed-out grey image | poor quality image | [8 empty] | FDA1
A software functionality in the report sections Findings Information and Summary of Measured Findings is not performing properly. The system will use the original values instead of displaying and saving corrected/modified values. Outdated information could be sent to the referring physician when creating the formal report in the reporting system. | improper report information | application program issue | [8 empty] | FDA1
A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. | inconsistent diagnosis | application program issue | [8 empty] | FDA1
Adding an additional comparison study to the Clinical Side Bar in the Enterprise Imaging Diagnostics Desktop, can result in a change of the sort order of the studies in the Clinical Side Bar. The thumbnail images of the studies are no longer chronologically sorted with the most recent study at the top. | inappropriate sort order | application program issue | [8 empty] | FDA1
Affected software product versions were found during internal audit to have been marketed for sale without sufficient documentation pertaining to certain required quality procedures. | insufficient documentation | device markings issue | [8 empty] | FDA1
After a period of time running Fusion Workstation, the Hounsfield measurement tool will report incorrect values. | incorrect measurement value | incorrect measurement | [8 empty] | FDA1
An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. | data calculated incorrectly | incorrect software programming calculations | [8 empty] | FDA1
An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. | decreased sensitivity | decreased sensitivity | [8 empty] | FDA1
An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | inappropriate labeling orientation | device markings issue | [8 empty] | FDA1
An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | image flipped | image reversal | [8 empty] | FDA1
An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | incorrect grid selected | use of device issue | [8 empty] | FDA1
An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | incorrect biopsy | unexpected therapeutic results | [8 empty] | FDA1
Any change made to current ultrasound machine software or hardware upgrades, adding or modifying custom fields, or adding a new Ultrasound Machine, must be tested prior to the use in production environment to ensure that ultrasound measurements are crossing over correctly to the AS system | lack of release test | false device output | [8 empty] | FDA1
Cardio study list does not show STAT studies without refreshing. | inappropriate list displayed | incorrect display | [8 empty] | FDA1
Carestream Health received a complaint related to CARESTREAM Image Suite 4 from a foreign hospital stating that the annotation on the overlay is displayed as "L (Left)", when it should be "R(Right)". | incorrect image annotation | image display error | [8 empty] | FDA1
Centricity Universal Viewer measurements saved into a DICOM Grayscale Presentation State are incorrect in subsequent views for exams containing series with different pixel sizes and may lead to a potential misdiagnosis. | inappropriate pixel size | image resolution poor | [8 empty] | FDA1
Centricity Universal Viewer measurements saved into a DICOM Grayscale Presentation State are incorrect in subsequent views for exams containing series with different pixel sizes and may lead to a potential misdiagnosis. | inappropriate pixel size | therapeutic or diagnostic output failure | [8 empty] | FDA1
Customers may experience an issue with the software study preferences when changes are made to the study protocol, resulting in incorrect patient follow-up. | inappropriate software release | application program version or upgrade problem | [8 empty] | FDA1
Cut lines on the image may present horizontally rather than vertically. | incorrect image cut line | image orientation incorrect | [8 empty] | FDA1
Digital templates were created with the incorrect files. | incorrect file applied | device markings issue | [8 empty] | FDA1
ON 07/14/2016, MERGE WAS NOTIFIED THAT THE STATION CRASHED WHILE DICTATING A REPORT, SUBSEQUENTLY, THE REPORT WAS MISSING. THE CUSTOMER REPORTED THAT THE SYSTEM STATED THERE WAS AN EXISTING REPORT WHEN THE USER TRIED TO RE-DICTATE THE REPORT, HOWEVER NO EXISTING REPORT COULD BE FOUND. AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT THE MOST PROBABLE CAUSE OF MISSING REPORT WAS DUE TO THE REPORT FILE NOT UPLOADING TO THE SERVER. | missing report | device markings issue | [8 empty] | FDA2
ON (B)(6) 2016 MERGE HEALTHCARE'S CUSTOMER, GE HEALTHCARE, REPORTED THAT ONE OF THEIR INSTALL SITES (B)(4) CONTACTED GE HEATHCARE TO NOTIFY THEM THAT NO LICENSE IS AVAILABLE WHEN LAUNCHING (B)(4). AS A RESULT OF THIS MALFUNCTION, USERS AT (B)(4) WERE NOT ABLE TO READ MAMMOGRAPHY IMAGES UNTIL THE MORNING OF (B)(6) 2016. THIS MALFUNCTION CAN RESULT IN DELAY IN DIAGNOSIS OR TREATMENT. | unable to read image | application program issue | [8 empty] | FDA2
ACCORDINGLY THE RESPONSE ON 5/19/2016 "...SITES HAVE BEEN UPGRADED FROM (B)(4) VERSION 3.7.3.7 TO (B)(4) VERSION 3.7.3.9SP3. THIS HAS ENABLED THE CUSTOMERS TO BYPASS THE ERROR, BECAUSE THE LICENSE ARCHITECTURE USING (B)(4) VERSION 3.7.3.9SP3 HAS A NEW DESIGN THAT USES WORKSTATION-BASED LICENSES. THE OLD SYSTEM USED SERVER-BASED LICENSES. (B)(4)'S SERVER UPGRADE WAS STARTED FRIDAY EVENING APRIL 22, AND BY MONDAY MORNING APRIL 25, 2016, ALL OF THEIR MAMMO STATIONS WERE ABLE TO VIEW MAMMO IMAGES ONCE AGAIN." THE PRODUCT IS FUNCTIONING AS EXPECTED AS OF 4/25/2016 AM ACCORDING TO GE HEALTHCARE. (B)(4). | incompatible license | application program issue | [8 empty] | FDA2
ON (B)(6) 2016 MERGE HEALTHCARE'S CUSTOMER, GE HEALTHCARE, REPORTED THAT ONE OF THEIR INSTALL SITES (B)(4) CALLED GE HEALTHCARE TO NOTIFY THAT THEY WERE GETTING LICENSE ERRORS WHEN OPENING STUDIES IN (B)(4). AS A RESULT OF THIS MALFUNCTION, USERS AT (B)(6) WERE NOT ABLE TO READ MAMMOGRAPHY IMAGES UNTIL THE EVENING OF (B)(6) 2016. THIS MALFUNCTION CAN RESULT IN DELAY IN DIAGNOSIS OR TREATMENT OF THE PATIENT. (B)(4). | license error | application program issue | [8 empty] | FDA2
Event Description: (B)(4) IS MEDICAL IMAGE REVIEW WORKSTATION SOFTWARE. THE PRODUCT CONSISTS OF FEATURES THAT ALLOW QUALIFIED MEDICAL PROFESSIONALS TO VIEW PATIENT MEDICAL IMAGES WITH THE DESIRED VIEWING PROTOCOL AND WORKFLOW IN ORDER TO OPTIMIZE THE EFFICIENT USE OF THEIR TIME. | unable to read image | application program issue | [8 empty] | FDA2
ON (B)(6)2017, MERGE HEALTHCARE TECHNICAL SUPPORT WAS NOTIFIED BY A (B)(4) CUSTOMER THAT PROCEDURES WERE NOT ABLE TO BE PERFORMED CAUSING A DELAY IN TREATMENT FOR 2 PATIENTS. PER THE CUSTOMER SITE, EMERGENCY DEPARTMENT STUDIES ARE TYPICALLY READ WITHIN 30 MINUTES HOWEVER, THERE WERE TWO STUDIES THAT WERE NOT READ UNTIL THE NEXT MORNING. WITH MERGE (B)(4) NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. THERE WAS NO REPORTED HARM TO EITHER PATIENT A OR PATIENT B AS A RESULT OF THE DELAY. (B)(4). | unable to read image | therapeutic or diagnostic output failure | [8 empty] | FDA2
Event Description: (B)(4). MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. | fail to create 2d images | application program issue | [8 empty] | FDA2
Event Description: (B)(4). MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. | fail to create 3d images | application program issue | [8 empty] | FDA2
ON (B)(6) 2017, A CUSTOMER REPORTED THAT WHEN TWO STUDIES WERE MERGED USING THE QC EDITOR, THE STUDY COMMENTS MADE ON THE SOURCE STUDY WERE NOT RETAINED AND MOVED TO THE TARGET STUDY. THE LOSS OF THE COMMENTS MAY LEAD TO A DELAY IN PATIENT CARE OR MISDIAGNOSIS. HOWEVER, THE IMAGES ARE AVAILABLE FOR DICTATION AND THERE HAS BEEN NO REPORT OF HARM TO A PATIENT DUE TO THIS ISSUE. (B)(4). | lost of annotation | image display error | [8 empty] | FDA2
Event Description: (B)(6) 2017 CUSTOMER REPORTED THAT THEY WERE UNABLE TO OPEN ALL IMAGES OF A CERTAIN STUDY IN THE SYNGO.PLAZA VIEWER. THE AFFECTED STUDY CONTAINED 240 IMAGES, HOWEVER, ONLY 52 WERE READABLE. ACCORDING TO THE SYNGO.PLAZA'S LOG FILES, THE IMAGE FILES WERE WRITTEN SUCCESSFULLY TO THE SHORT TERM STORAGE (STS) DISKS, HOWEVER THE STS'S FILESYSTEM CONTAINED ONLY EMPTY FILES (NO INFORMATION INCLUDING DICOM HEADER IS PRESENT). THE SYNGO.PLAZA SYSTEM RELIES ON WINDOWS CONFIRMATION TO DETERMINE WHETHER A FILE WAS SUCCESSFULLY STORED. SINCE NO ERRORS HAVE BEEN LOGGED IN THE SYNGO.PLAZA DURING THE STORAGE PROCESS TO THE STS AND THERE ARE KNOWN INTERMITTENT HARD DRIVE ISSUES IDENTIFIED IN WINDOWS EVENT LOGS, THEREFORE IT IS ASSUMED THAT THE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE CORRUPT STS FILES. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THIS INCIDENT OCCURRED IN (B)(6). | file corrupted | application program issue | [8 empty] | FDA2
Event Description: (B)(6) WAS ENROLLED ON PROTOCOL (B)(6) OF RESPONSE ADAPTED THERAPY BASED ON POSITRON EMISSION TOMOGRAPHY (PET) FOR BULKY STATE I AND STAGE II CLASSICAL HODGKIN LYMPHOMA ON (B)(6) 2015. THE PT PRESENTED TO CLINIC FEELING WELL ON (B)(6) 2015 FOR CONSIDERATION TO BEGIN CYCLE 3. THE LAB RESULTS DEMONSTRATED GRADE 4 NEUTROPENIA. (B)(6), 2015 ANEUT 0.33; (B)(6) 2015 WBC 1.71 (L), (B)(6) 2015 RBC 4.44, (B)(6) 2015 HGB 12.1, (B)(6) 2015 HGB 12.1, (B)(6) 2015 HCT 36.7, ON (B)(6) 2015 CYCLE 3, NEUTROPENIA REDUCED TO GRADE 1. THE PT IS FEELING WELL AND WILL PROCEED WITH PROTOCOL TREATMENT. (B)(6) 2015 ANEUT 1.51, (B)(6) 2015 WBC 5.23, (B)(6) 2015 RBC 4.18, (B)(6) 2015 HGB 11.5 (L), (B)(6) 2015 HCT 34.8 (L).ATTRIBUTE TO GR, 4 NEUTROPHIL COUNT DECREASED. (B)(6). PHASE II TRIAL OF RESPONSE-ADAPTED THERAPY BASED ON POSITRON EMISSION TOMOGRAPHY (PET) FOR BULKY STAGE 1 AND STAGE II CLASSICAL HODGKIN LYMPHOMA (HL). | inappropriate therapy | therapeutic or diagnostic output failure | [8 empty] | FDA2
Event Description: A CUSTOMER REPORTED A PATIENT DATA ERROR WHILE USING THE TOMTEC-ARENA. THE SOFTWARE ALLOWED PATIENT DATA ACQUIRED FROM ONE DATA SET TO BE INCLUDED IN THE NEXT PATIENT STUDY. THERE WAS NO MISDIAGNOSIS OR INJURY ASSOCIATED WITH THIS EVENT. | incorrect patient information | patient data issue | [8 empty] | FDA2
Event Description: A CUSTOMER REPORTED AN EMERGENCY ROOM (ER) PATIENT WAS IMAGED USING A GENERIC NAME PRIOR TO SCHEDULING THE STUDY WITH THE PT'S ACTUAL NAME RESULTING IN TWO PATIENTS IN HMI. THE TWO STUDIES WERE MERGED IN HMI USING THE WA MANAGER TOOL. IN THIS INSTANCE, MERGING OF THE SCHEDULED STUDY AND PERFORMED STUDY WAS INCOMPLETE WITH THE SCHEDULED STUDY NOT BEING REMOVED FROM THE SCHEDULED STUDIES LIST. SUBSEQUENTLY, THE ER PT WAS INADVERTENTLY SUBJECTED TO A SECOND CHEST X-RAY IMAGING PROCEDURE. THE DATE OF THIS ADVERSE EVENT AND PT INFO WERE NOT PROVIDED BY THE REPORTING FACILITY. | fail to merge multiple images | application program issue | [8 empty] | FDA2
Event Description: A CUSTOMER REPORTED MEASUREMENT DATA CAN APPEAR FROM A PREVIOUS STUDY ONTO A NEW STUDY IF AN ERROR MESSAGE APPEARS DURING LAUNCH WHILE USING THE TOMTEC-ARENA TTA 2. THIS EVENT OCCURRED DURING A RESEARCH STUDY. THERE WAS NO INJURY OR CLINICAL USE ASSOCIATED WITH THIS EVENT. | inappropriate measurement data displayed | incorrect display | [8 empty] | FDA2
Event Description: A CUSTOMER REPORTED THAT WHEN WORKING WITH MULTIPLE STUDIES AND USING THUMBNAIL PANELS, THE THUMBNAIL TAB CHANGED UNEXPECTEDLY AND THE RADIOLOGIST ALMOST DICTATED A REPORT ON THE BASIS OF LOOKING AT THE WRONG STUDY. THE PROBLEM WAS DISCOVERED AND NO MISDIAGNOSIS OCCURRED. | tab changed unexpectedly | application program issue | [8 empty] | FDA2
Event Description: A CUSTOMER USED A SURGIGUIDE TO CREATE THE OSTEOTOMY FOR TWO KEYSTONE DENTAL IMPLANTS ON TOOTH POSITION #7 AND #10. THE CENTRAL AXIS OF THE OSTEOTOMY FOR BOTH IMPLANTS DID DEVIATE 12 DEGREES TO THE BUCCAL SIDE FROM THE ORIGINAL PLANNING. THE INCORRECT IMPLANT PLACEMENT LEADS TO THE FACT THAT THE APICAL PART OF BOTH IMPLANTS ARE PERFORATING THE BUCCAL PLATE AND NEED GRAFTING OR THE IMPLANTS NEED TO BE REMOVED AND REPLACED DURING A SECOND CORRECTIVE SURGERY. | incorrect implant placement | positioning issue | [8 empty] | FDA2
Event Description: A DATA LOSS IN THE SYNGO IMAGING APPLICATION OCCURRED FOR 62 OUT OF 601 IMAGES IN ONE SERIES ON (B)(6) 2017. NO DATA MIX-UP OCCURRED; NO PATIENT RESCAN WAS REQUIRED. DUE TO CHANGES ON THE POWER SUPPLY OF THE SYNGO IMAGING ENVIRONMENT PERFORMED BY THE CUSTOMER (B)(6) 2017, AN IMPROPER SHUTDOWN OF THE STORAGE TERM RAID AND THE SYNGO IMAGING APPLICATION OCCURRED. THIS RESULTED IN THE AFFECTED IMAGES BEING CREATED WITH 0 KB FILE SIZE IN THE SYNGO IMAGING'S SHORT TERM STORAGE PRIOR TO THEM BEING ARCHIVED IN THE LONG TERM ARCHIVE. THE AFFECTED SERIES WERE AVAILABLE IN THE SYNGO IMAGING VIEWER, HOWEVER, ONLY 539 IMAGES CAN BE DISPLAYED AND 62 IMAGES WERE CORRUPT. THESE CORRUPT IMAGES WERE FOUND DUE TO MAINTENANCE TASKS ON THE SYSTEM (B)(6) 2017. THE REPORTED INCIDENT OCCURRED IN (B)(6). | image with 0 kb size | data issue | [8 empty] | FDA2
Event Description: A DATA LOSS OF APPROXIMATELY 200 CLINICAL IMAGE STUDIES WHILE ARCHIVING WAS REPORTED FROM A SITE. SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE AND FOUND THAT USER ACTIONS CAUSED THE DATA LOSS AND THAT THE DATA LOSS WAS UNRECOVERABLE. THE HISTORICAL IMAGES ARE NO LONGER AVAILABLE FOR COMPARISON WITH NEWER IMAGES HOWEVER THE CLINICAL REPORTS GENERATED FROM THE READ OF THE STUDIES PRIOR TO ARCHIVAL ARE STILL PRESENT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | image loss | loss of data | [8 empty] | FDA2
Event Description: A HAZARD MAY ARISE WHEN NON-DICOM COMPLIANT JPEG-2000-COMPRESSED IMAGES ARE RECEIVED BY THE (B)(6) COMPONENT WITH DICOM BIT DEPTH (BITS STORED TAG) THAT DOES NOT MATCH THE BIT DEPTH IN THE PIXEL DATA CODESTREAM. THESE IMAGES ARE DECOMPRESSED WITH A LOSS OF PRECISION IN THE IMAGE DATA WHICH CAN LEAD TO FAILING IMAGE REGISTRATION. | non-dicom format image | data issue | [8 empty] | FDA2
Event Description: A HAZARD MAY ARISE WHEN NON-DICOM COMPLIANT JPEG-2000-COMPRESSED IMAGES ARE RECEIVED BY THE (B)(6) COMPONENT WITH DICOM BIT DEPTH (BITS STORED TAG) THAT DOES NOT MATCH THE BIT DEPTH IN THE PIXEL DATA CODESTREAM. THESE IMAGES ARE DECOMPRESSED WITH A LOSS OF PRECISION IN THE IMAGE DATA WHICH CAN LEAD TO FAILING IMAGE REGISTRATION. | loss of precision | data issue | [8 empty] | FDA2
Event Description: A LITHIUM ION BATTERY USED FOR OUR PORTABLE COMPUTERS SPARKED, OVERHEATED AND MELTED. AN ENTIRE FLOOR OF OUR HOSPITAL HAD TO BE EVACUATED IN ORDER TO PROTECT OUR PATIENTS FROM SMOKE INHALATION. | battery melted | battery issue | [8 empty] | FDA2
Event Description: A MERGE CARDIO CUSTOMER REPORTED THAT THEIR STUDY LIST DOES NOT SHOW NEW STUDIES UNLESS MANUALLY REFRESHED. THE CLIENTS ARE GETTING NOTIFICATIONS REGARDING REPORT AND STUDY STATUS UPDATES BUT IF THEY HAVE A FILTER SET FOR ALL QC'D REPORTS FOR TODAY AND DON'T MANUALLY REFRESH THE STUDY LIST BETWEEN REVIEWING STUDIES, THEY DON'T SEE STAT STUDIES UNLESS THE STUDY LIST IS MANUALLY REFRESHED. (B)(4). | inappropriate list displayed | incorrect display | [8 empty] | FDA2
Event Description: A MERGE ORTHOCASE CUSTOMER REPORTED THAT THE DISPLAYED VALUE OF THE MEASUREMENT IS CHANGING AFTER SAVING A PLAN WITH A MEASUREMENT RESULT. | meassurement changed | application program issue | [8 empty] | FDA2
Event Description: WHEN USING THE PERPENDICULAR LINE TOOL, THERE IS AN OPTION TO "SNAP" THE LINES TOGETHER, WHICH PREVENTS THE PERPENDICULAR LINES FROM EXTENDING PAST THE STRAIGHT/BASE LINE. THE SYSTEM DEFAULT IS TO SNAP THESE LINES. TO REMOVE THE SNAPPING FUNCTION FROM THE CURRENT ANNOTATION, SIMPLY RIGHT CLICK NEAR THE LINE AND DESELECT "SNAP TO." WHILE THE SYSTEM DEFAULT IS SET TO SNAP THE LINES TOGETHER, THE ACTUAL "SNAP" FUNCTION MAY NOT OCCUR UNTIL THE PLAN IS SAVED, AT WHICH TIME THE LINE END POINT WITHIN 20 PIXELS OF THE BASE LINE WILL ADJUST ITS LOCATION TO SNAP PERFECTLY TO THE BASE LINE. THIS MAY ADJUST THE LENGTH AND CORRESPONDING MEASUREMENT OF THE PERPENDICULAR LINE WHEN THE PLAN IS REOPENED. (B)(4). | snap function unexpected | application program issue | [8 empty] | FDA2
Event Description: A NEUROSURGEON INITIALLY MISINTERPRETED ANCHOR IMAGES DISPLAYED IN THE VIEWPORT OF THE CHANGE HEALTHCARE ENTERPRISE VIEWER (CHEV) AS A PRIOR STUDY. IT WAS REPORTED THIS CAUSED THE NEUROSURGEON TO ALMOST USE THE WRONG IMAGE FOR SURGERY. NO ADVERSE CONSEQUENCES WERE REPORTED. | image misinterpreted | use of device issue | [8 empty] | FDA2
Event Description: A POTENTIAL LOSS OF CLINICALLY-RELEVANT DATA WHICH MAY BE UNRECOVERABLE WAS REPORTED. THE INVESTIGATION INTO THE ROOT CAUSE AND RECOVERABILITY OF THE DATA IS ONGOING. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES ARE STILL AVAILABLE IN THE SITES HIS. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | missing historical study images | loss of data | [8 empty] | FDA2
Event Description: A RADIOLOGIST WAS REVIEWING A TUMOR CASE IN RADSUITE AND NOTICED THAT THE PIXEL VALUE TOOL WAS DISPLAYING DIFFERENT VALUES THEN THE PET MODALITY'S SOFTWARE. TO AID IN THE INTERPRETATION OF PET IMAGES, THE PIXEL VALUE TOOL WITHIN THE RADSUITE SOFTWARE CAN BE USED TO VIEW THE ABSORPTION OF AN INJECTED ISOTOPE IN BODY TISSUE BASED ON STANDARDIZED UPTAKE VALUES (SUV). THE PATIENT WEIGHT VALUE IS USED FOR CALCULATING THE SUV. IF THE PATIENT WEIGHT IS MISSING, THE VALUE IS NULL. IF A STUDY WITH A PATIENT WEIGHT VALUE OF NULL IS VIEWED AFTER A STUDY FOR A DIFFERENT PATIENT WITH A PATIENT WEIGHT VALUE GREATER THAN 0, THE WEIGHT FROM THE PREVIOUS STUDY WILL BE USED TO CALCULATE THE SUVS FOR THE STUDY WITH THE NULL WEIGHT VALUE. THIS MAY RESULT IN INCORRECT SUVS. IF OTHER CLINICAL FACTORS ARE NOT AVAILABLE OR TAKEN INTO DUE CONSIDERATION, INCORRECT SUVS COULD POTENTIALLY LEAD TO OVERSTATEMENT OR UNDERSTATEMENT OF THE EFFECTIVENESS OF A CANCER TREATMENT, OR TO A MISDIAGNOSIS OF A CANCER. | null weight value | patient data issue |  |  | misdiagnosis of a cancer | [5 empty] | FDA2
Event Description: A RADSUITE CUSTOMER REPORTED THAT WHEN THE RADIOLOGIST IS SCROLLING THROUGH THE IMAGES IN THE VIEWER THERE HAVE BEEN DELAYS OF BETWEEN 5 AND 10 MINUTES. USER ERROR IN SETTING UP A CUSTOM MACRO AND LACK OF VERIFICATION OF THAT MACRO RESULTED IN THIS ISSUE. THIS MACRO CONTAINS A PERSISTABLE THAT NEEDS TO QUERY THE DATABASE EVERYTIME IT IS CALLED ON. WHEN THE USER IS SCROLLING THRU IMAGES THE MACRO IS MAKING A DATABASE CALL AND NEEDS TO HAVE THE INFORMATION FROM THE DATABASE BEFORE IT CAN GO ANY FURTHER THUS CAUSING THE VIEWPORT TO "FREEZE" AS IT IS WAITING ON THE INFORMATION FROM THE DATABASE. MOST OF THE TIME IT IS NOT A PROBLEM BUT THERE ARE TIMES WHERE THERE SEEMS TO BE CONTENTION IN THE DATABASE OR A NETWORK OR STORAGE ISSUE THAT IS CAUSING A DELAY. (B)(4). | scroll image delayed | application network issue | [8 empty] | FDA2
Event Description: A SERVICE CALL WAS PLACED FOR THE SMART CR UNIT (S/N (B)(4)) AT ABOUT 2AM CST ON (B)(6) 2016. CUSTOMER ALLEGES SMOKE COMING FROM THE UNIT. HOSPITAL STAFF UNPLUGGED THE UNIT AND AWAITED SERVICE. SERVICE ENGINEER ARRIVED ON-SITE LATER THAT DAY AND DETERMINED THE POWER SUPPLY TO BE THE ISSUE. A NEW POWER SUPPLY WAS ORDERED (B)(4) AND REPLACED THE FOLLOWING DAY, (B)(6) 2016. THE UNIT WAS TURNED OVER TO THE HOSPITAL IN GOOD WORKING ORDER. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | device burned | smoking | [8 empty] | FDA2
Event Description: A SIMPLE REQUEST FOR SERVICE SUPPORT TO RECOVER DATA AFTER A HARDWARE FAILURE IN THE SITE'S LONG TERM ARCHIVE SYSTEM WAS ESCALATED TO SIEMENS. HOWEVER. IT WAS IDENTIFIED THAT AN UNKNOWN AMOUNT OF IMAGE STUDIES WERE MISSING IN THE SITES LONG TERM ARCHIVE. THE HISTORICAL STUDY IMAGES ARE NO LONGER . HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. | missing historical study images | loss of data | [8 empty] | FDA2
Event Description: A SINGLE IMAGE IN A CT SERIES WAS FLIPPED AND SAVED. WHEN THE RADIOLOGIST MADE HIS FINAL INTERPRETATION, THERE WAS NOT A SIGNIFICANT VISUAL INDICATION THAT THE IMAGE HAD BEEN MANIPULATED. | image flipped | image reversal | [8 empty] | FDA2
Event Description: A SITE USING MERGE PACS REPORTED THAT A MIGRATED STUDY THAT HAS ANNOTATIONS WILL DISPLAY THAT STUDY'S ANNOTATIONS, BUT WHEN A PRIOR IS LOADED THAT SHOULD HAVE ANNOTATIONS THEN THE SECOND STUDY WILL NOT DISPLAY ANNOTATIONS. SITE EXPRESSED CONCERNS THAT SEEING NO ANNOTATIONS COULD POSSIBLY INDICATE TO THE PROVIDER THAT NOTHING SIGNIFICANT WAS NOTED WHEN THE IMAGE WAS READ INITIALLY, AND THEREFORE COULD POTENTIALLY IMPACT THE TREATMENT PLAN OR COULD POTENTIALLY RESULT IN A FAILURE TO TREAT ACCORDINGLY. | missing annotation | incorrect display | delayed treatment plan |  | treatment failure | [5 empty] | FDA2
Event Description: A SOFTWARE CODING ERROR RESULTED IN AN INCORRECT REGISTRY SETTING WHERE THE UNITS THAT ARE DISPLAYED FOR ORGAN DOSE ARE MISMATCHED WITH THE NUMERICAL VALUE DISPLAYED. THE PROBLEM WAS NOT DISCOVERED FOR A LONG TIME, LIKELY BECAUSE CUSTOMERS HAD NOT CHANGED THE DEFAULT REGISTRY SETTING FOR RADIATION DOSAGE (WHICH IS CORRECT) TO DISPLAY ORGAN DOSE UNTIL NOW. IF A CLINICIAN DID NOT NOTICE THAT THE DISPLAYED INDIVIDUAL IMAGE DOSAGE WAS WELL OUTSIDE OF THE NORMAL RANGE, THEY COULD THINK THAT THE PATIENT HAD RECEIVED LESS RADIATION THAN THEY ACTUALLY HAD AND SO ANY ADDITIONAL IMAGING STUDIES WOULD EXPOSE THE PATIENT TO MORE CUMULATIVE RADIATION THAN EXPECTED. THE INCORRECT DOSAGE VALUE IS NOT PRESERVED OR SAVED UNLESS A SCREEN CAPTURE IS GENERATED OR THE "COPY IMAGE" OR "SAVE IMAGE" FEATURE IS USED. FOR ORGANIZATIONS THAT CALCULATE A CUMULATIVE RADIATION DOSAGE VALUE FOR PATIENTS, THE CORRECT DOSAGE DATA FROM ALL TYPES OF SCANS (NOT JUST MAMMOGRAPHY) SHOULD BE IN A PACS OR OTHER CENTRAL LOCATION OR SYSTEM, NOT IN THE MAMMO APPLICATION. THE ACTUAL DOSE VALUES SAVED IN THE PACS IN THE DICOM IMAGE HEADERS ARE CORRECT IN CASE THEY ARE USED TO CALCULATE CUMULATIVE DOSE. | incorrect dose displayed | dose calculation error due to software problem | under dose |  | Patient injury | [5 empty] | FDA2
Event Description: A SPECIFIC WORKFLOW IN ONE CUSTOMER'S ENVIRONMENT IS LIKELY CONTRIBUTING TO THE FILLING UP OF THE RADSUITE CACHE. THE SITE TYPICALLY CONFIGURES LESS STORAGE FOR THE LOCAL EA (DSS PARTITIONS) AND MORE STORAGE FOR THE RADSUITE CACHE (UV_DSS PARTITIONS). WITH THEIR CONFIGURATION, IT IS LIKELY THAT MANY OF THEIR STUDIES WILL NOT QUALIFY FOR WATERMARKING TO REMOVE THEM FROM RADSUITE'S STORAGE AND FREE UP SPACE THERE. WHEN THE RADSUITE CACHE IS FULL, USERS WILL OBSERVE STORAGE FAILURE MESSAGES FROM THE MODALITY DEVICES. THOSE STUDIES COULD BE VIEWED ON THE MODALITIES OR STORED ON ANOTHER RADSUITE NODE FOR VIEWING THERE. IF NOT, THEN THERE IS A POTENTIAL THAT SOME STUDIES MIGHT NOT BE VIEWED, ALTHOUGH WE HAVE NOT RECEIVED ANY REPORTS OF THAT ACTUALLY OCCURRING. THE PROBLEM WAS FOUND AT ONE CUSTOMER SITE BY MERGE'S SYSTEM MONITORING PROCESS. | insufficient memory size | data back-up problem | [8 empty] | FDA2
Event Description: A SRS TREATMENT FOR 2 METASTASES IN THE LEFT PARIETAL AND LEFT OCCIPITAL LOBE IN THE BRAIN, PERFORMED ON (B)(6) 2020, HAD BEEN PLANNED WITH THE AID OF BRAINLAB IMAGE FUSION 4.0 (AND BRAINLAB MULTIPLE BRAIN METS SRS 2.0). DURING PLANNING/TREATMENT THE USER: PERFORMED A RIGID FUSION OF A MR AND CT IMAGE SET (TO COMBINE THE INFORMATION CONTAINED WITHIN THE INDIVIDUAL IMAGE SETS). PROCEEDED WITH CALCULATION OF AN ELASTIC FUSION FOR THIS MR/CT FUSION PAIR (A NEW MR IMAGE SET IS GENERATED USING A DEFORMABLE REGISTRATION ALGORITHM, BY BRAINLAB CRANIAL DISTORTION CORRECTION). REVIEWED AND ACCEPTED THE RESULTING ELASTICALLY DEFORMED / DISTORTION CORRECTED MR (INCL. DISTORTION CORRECTED OBJECTS). CREATED A DOSE PLAN BASED ON THE DISTORTION CORRECTED MR AND CT IMAGE SET IRRADIATED THE METASTASES ACCORDING TO THE DOSE PLAN. SUSPECTED ON (B)(6) 2020, FOR ONE METASTASIS, A DEVIATION OF ACTUAL METASTASIS LOCATION/SIZE VS. IRRADIATED TARGET VOLUME, BY COMPARING AFOREMENTIONED IMAGE SETS AND A POST-TREATMENT MR VIA RIGID FUSION. ACCORDING TO THE HOSPITAL: THE TREATMENT WAS A SRS TREATMENT (1 FRACTION), 2 BRAIN METASTASES, 20-24 GY. ONE METASTASIS (LEFT PARIETAL) WAS NOT IRRADIATED AS INTENDED, THE SECOND METASTASIS (LEFT OCCIPITAL) TREATED WITH THIS SAME PLAN SEEMS TO BE RESOLVED. FOR THE AFFECTED METASTASIS, A LARGER VOLUME WAS TREATED THAN WHAT CORRESPONDED TO THE REAL METASTASIS SIZE (THE METASTASIS RECEIVED 50%-71% OF THE INTENDED DOSE DISTRIBUTION, AND APPROX. 0.35 CC OF HEALTHY BRAIN TISSUE RECEIVED UNINTENDED RADIATION DOSE). THE INCORRECT DOSAGE EXCEEDED CLINICALLY ACCEPTABLE LIMITS, AND THERE WAS A NEGATIVE CLINICAL EFFECT TO THE PATIENT (THE AFFECTED METASTASIS CONTINUED TO GROW). IF FURTHER REMEDIAL ACTIONS ARE NEEDED IS CURRENTLY UNDECIDED. 8 MORE METASTASES WERE TREATED ON (B)(6) 2020 AND THE PHYSICIAN WILL MONITOR THE PATIENT AND THE AFFECTED METASTASIS). | incorrect dose plan | dose calculation error due to software problem | [8 empty] | FDA2
Event Description: A USER REPORTED THAT APPROXIMATELY 1000 STUDIES WERE LOST FROM THE PACS STORAGE AND A REFERRING PHYSICIAN WAS UNABLE TO VIEW IMAGES. RADIOLOGIST REPORTS FOR THESE DELETED STUDIES REMAINED AVAILABLE. PRELIMINARY ANALYSIS DETERMINED THE PROBLEM IS DUE TO A CONFIGURATION KEY THAT WAS ACTIVATED IN COMBINATION WITH A SOFTWARE BUG. WHEN THE FEATURE FORCEGLOBALSTUDYDELETION IS ENABLED, ANY SERIES DELETIONS THAT OCCUR ON A SERVER ARE AUTOMATICALLY ELEVATED TO A SYSTEM-LEVEL STUDY DELETION. THIS SYSTEM-LEVEL DELETION BYPASSES EXISTING MITIGATION POLICIES (I.E., ARCHIVE POLICIES) ON OTHER SERVERS. | image deleted | loss of data |  | inappropriate configuration key activiated; system level deletion | studied delayed | [5 empty] | FDA2
Event Description: ACCORDING TO THE LITERATURE STUDY. A RETROSPECTIVE ANALYSIS OF A SINGLE CENTER EXPERIENCE IN ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY MICROWAVE ABLATION IN HYBRID OPERATING ROOM. A TOTAL OF 30 LUNG NODULES FROM 25 PATIENTS WERE TREATED. PNEUMOTHORAX OCCURRED IN 2 CASES (6.67%) AND BOTH REQUIRED CHEST DRAIN INSERTION. | lung nodule | therapeutic or diagnostic output failure |  | inappropriate microwave ablation | additional surgery procedure | [5 empty] | FDA2
Event Description: AFTER A PERIOD OF TIME RUNNING THE FUSION WORKSTATION (FW) 3.0.2 P3 AND LATER, THE HOUNSFIELD MEASUREMENT TOOL WILL REPORT INCORRECT VALUES. IF THE EXAM IS CLOSED, AND USER EXITS FW AND LOGS BACK IN TO REOPEN THE EXAM, THE MEASUREMENT WILL THEN BE REPORTED CORRECTLY. THERE APPEARS TO BE NO CONSISTENCY TO THE METHOD TO INVOKE THE ERROR. THE FACILITY REPORTED THE MEASUREMENTS TO BE CLEARLY AND OBVIOUSLY INCORRECT UPON VIEWING THE RESULTS, AND IT HAPPENS ABOUT 1 OUT OF EVERY 500 STUDIES. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS OR MISDIAGNOSES OF A PATIENT, HOWEVER THIS WAS DEEMED A POTENTIAL SAFETY ISSUE. | incorrect display value | incorrect display |  | no consistency method | no harm | [5 empty] | FDA2
Event Description: AFTER PERFORMING A SUCCESSFUL 5DSA DV 3D DURING A CEREBRAL ANGIOGRAM, THE IMAGES WERE SENT TO THE SYNGO WORKSTATION OF IPP 8 FOR INTERPRETATION, SOFTWARE VERSION (B)(4). THE 3D WAS UNABLE TO SUCCESSFULLY RECONSTRUCT. ERROR CODE "THE SETTINGS OF THE USED ACQUISITION PROGRAM ARE NOT SUITABLE FOR THE REQUESTED 3D RECONSTRUCTION." A DEFAULT 3D RECONSTRUCTION IS PERFORMED. SEVERAL ATTEMPTS WERE MADE TO SALVAGE THE DATA FOR USE BUT WERE UNSUCCESSFUL. A SECOND RECONSTRUCTION WAS PERFORMED UNDER THE 3D 5DSA HEAD AND THE IMAGES WERE SUCCESSFULLY RECONSTRUCTED. | fail to reconstruct 3D model | application program issue |  | inappropriate settings of the acquisition program | plan delayed | [5 empty] | FDA2
Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 12 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: A NEW FINDING WAS ADDED TO COMMENT ON THE LEFT GREATER SAPHENOUS VEIN, BUT THE SENTENCE PRODUCED READS "RIGHT GREATER SAPHENOUS VEIN (DIST CALF): VENOUS INCOMPETENCE NOTED." POTENTIAL IMPACT: WHEN THE USER GOES TO REPORT ON THE LEFT GREATER SAPHENOUS VEIN (DIST CALF), THE SENTENCE WILL REFER TO THE RIGHT SAPHENOUS VEIN. THIS COULD POTENTIALLY LEAD TO MISDIAGNOSIS AND INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE # Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | incorrect clinical data | data issue | [8 empty] | FDA2
Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 2 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, ONLY TWO REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: LABEL WAS CHANGED FROM TRUNCAL V REGURG P 1/2-TIME TO PULMONARY INSUFFICIENCY PRESSURE HALF TIME (PI PHT). POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED THOSE REPORTS PRIOR TO THE CHANGE COULD DISPLAY INCORRECT CLINICAL DATA IF VIEWED IN RM OR IF USED AS A TEMPLATE FOR A REPEAT STUDY ON THAT PATIENT - IN THIS EXAMPLE THE TWO DISEASE STATES ARE VERY DIFFERENT, LEADING TO POTENTIAL INCORRECT DIAGNOSIS. DESCRIPTION OF MEASUREMENT DOES NOT CORRELATE WHEN USING HOVER TOOL WHICH CAN CAUSE CONFUSION. TRUNCUS ARTERIOSUS IS A RARE HEART DEFECT THAT'S PRESENT AT BIRTH (CONGENITAL). IN TRUNCUS ARTERIOSUS, IT MEANS THAT ONE LARGE BLOOD VESSEL LEADS OUT OF THE HEART. | incorrect clinical data | data issue |  | user error | no harm | [5 empty] | FDA2
Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2016-00001, 1225058-2016-00002, AND 1225058-2016-00003 TO THE FDA ON (B)(6) 2016 FOR A SITE IN THE US. THIS REPORT IS FOR A 4TH OCCURRENCE OF THE SAME ISSUE AT A DIFFERENT SITE AND INCLUDES ONE (1) REPORT AFFECTED ON (B)(6) 2016, WHEN THE ERROR OCCURRED. THIS WAS AN INTERNAL DISCOVERY BY AGFA DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN AS REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA (B)(4) (1225058-12-22-2016-001-C). REPORT ID DATE ERROR OCCURRED. (B)(4), (B)(6) 2016. BACKGROUND ON NOVEMBER 22, 2016, A CUSTOMER REPORTED TO AGFA THAT WHEN USING IMPAX CV REPORTING FOR NON-INVASIVE VASCULAR ULTRASOUND (NIV), WHEN BUILDING A NIV REPORT, IN THE CONCLUSION SECTION OF THE UPPER ARTERIAL DUPLEX, IF YOU SELECT ANY CHECK BOXES UNDER THE LEFT ANEURYSM IT PUTS IN THE REPORT THAT THERE WAS NO ANEURYSM. ON NOVEMBER 22, 2016, AGFA TECHNICAL SUPPORT CONTACTED THE CUSTOMER AND UPDATED HIM THAT ONLY THREE (3) REPORTS WERE AFFECTED AND THOSE THREE (3) REPORTS WERE ALL LISTED AS TEST REPORTS. THE CUSTOMER WAS PROVIDED THE LIST OF THE AFFECTED REPORTS. AGFA TECHNICAL SUPPORT RE-CONFIGURED THE REPORTS TO PRODUCE CORRECT STATEMENTS IN BOTH THE TEST AND PRODUCTION ENVIRONMENT. A PROBLEM RECORD WAS CREATED ON NOVEMBER 23, 2016 TO ADDRESS THE ISSUE VIA (B)(4). ON (B)(6) 2016 AGFA HAS IMPLEMENTED A CORRECTION VIA (B)(4) (1225058-12-22-2016-001-C). AGFA SENT URGENT SAFETY NOTICES TO AFFECTED CONSIGNEES FOR THIS ISSUE DESCRIBED IN (B)(4) | incorrect clinical data | data issue |  | user error | no harm | [5 empty] | FDA2
Event Description: ALL IMAGES ON THE SYNGO.PLAZA SYSTEM WERE NOT ACCESSIBLE IN THE ONLINE STORAGE DISKS. THERE WERE MULTIPLE POWER BLACKOUTS AND NO UNINTERRUPTABLE POWER SUPPLY (UPS) IS INSTALLED, WHICH COULD WITHSTAND VOLTAGE VARIATIONS. THE DAMAGE ON THE STS ARRAYS WAS UNRECOVERABLE AND THE SYSTEM WAS REINSTALLED FROM SCRATCH TO MAKE STORING OF NEW IMAGES POSSIBLE AGAIN. DURING ATTEMPTS TO RESTORE DATA FROM SYNGO.PLAZA'S ARCHIVE ALSO A MAJOR DAMAGE ON ITS DISKS WAS IDENTIFIED BECAUSE OF THE SAME POWER OUTAGES. ALL DATA ON THE ARCHIVE (5375 STUDIES) CANNOT BE ACCESSED ON ITS FILESYSTEM AS WELL. THE RESTORE AND CHECKS ON THE ARCHIVE ARE ONGOING, IF THE FILESYSTEM CAN BE MADE ACCESSIBLE AGAIN AND THUS USE IT FOR THE RESTORE. THE CUSTOMER WAS INFORMED ABOUT AFFECTED PATIENTS/STUDIES. BASED ON QUANTUM, OF CURRENTLY NOT ACCESSIBLE DATA, THIS IS PRELIMINARY TREATED AS A DATA LOSS. | fail to access image | failure to service |  | power blackouts | [6 empty] | FDA2
Event Description: ALLEGEDLY, WHILE PREPARING AN OR ROOM FOR THE FIRST CASE OF THE DAY, A STAFF MEMBER HEARD A LOUD POP COMING FROM THE MONITOR. SHE WENT TO NOTIFY HER SUPERVISOR OF THE POP; HOWEVER, WHEN SHE RETURNED TO THE OR ROOM, FLAMES WERE COMING FROM THE BACK OF THE MONITOR. STAFF WAS ABLE TO EXTINGUISH THE FIRE USING TWO CO2 EXTINGUISHERS. | catch fire | fire |  | device pop | [6 empty] | FDA2
Event Description: ANNULUS MEASUREMENT WAS PERFORMED BY A RADIOLOGIST WITH ASSISTANCE OF ADVANTAGE WORKSTATION TAVI (TRANSCATHETER AORTIC VALVE REPLACEMENT) SOFTWARE ON A PATIENT CT SCAN OBTAINED (B)(6) 2018. A TAVI VALVE IMPLANTATION PROCEDURE WAS PERFORMED ON (B)(6) 2018 WHERE A MEDTRONIC EVOLUTE PRO 29MM WAS PLACED, BASED ON RADIOLOGIST MEASUREMENT OF 465.4 MM2. SEVERE PERIPROSTHETIC VASCULAR INSUFFICIENCY WAS IDENTIFIED FOLLOWING VALVE PLACEMENT. A TAVI VALVE IN VALVE PROCEDURE WAS THEN PERFORMED IN (B)() 2018 FOLLOWING THE SAME RADIOLOGIST MEASUREMENT OF ANNULUS USING A MORE RECENT CT IMAGE AND DIFFERENT METHOD THAT RESULTED IN A MEASUREMENT OF 547 MM2, WHERE AN EDWARDS SAPIEN3 29MM VALVE WAS PLACED WITH NORMAL POST IMPLANT OUTCOME. | inappropriate valve location | positioning issue | [8 empty] | FDA2
Event Description: ANNULUS MEASUREMENT WAS PERFORMED BY A RADIOLOGIST WITH ASSISTANCE OF ADVANTAGE WORKSTATION TAVI (TRANSCATHETER AORTIC VALVE REPLACEMENT) SOFTWARE ON A PATIENT CT SCAN OBTAINED (B)(6) 2018. A TAVI VALVE IMPLANTATION PROCEDURE WAS PERFORMED IN (B)(6) 2018 WHERE A MEDTRONIC EVOLUTE PRO 26MM WAS PLACED, BASED ON RADIOLOGIST MEASUREMENT OF 377 MM2. WITHIN 5 MINUTES OF VALVE PLACEMENT, THE VALVE MIGRATED TO ASCENDING AORTA. DURING THIS TAVI PROCEDURE, FOLLOWING THE SAME RADIOLOGIST RE-MEASUREMENT OF ANNULUS USING A DIFFERENT METHOD THAT RESULTED IN A MEASUREMENT OF 390 MM2, AN EDWARDS SAPIEN3 29MM VALVE WAS PLACED WITH NORMAL POST IMPLANT OUTCOME. | valve migrated | migration of device or device component | [8 empty] | FDA2
Event Description: AS PART OF A CHANGE OF EHR AT THIS INSTITUTION, A NEW PACS SYSTEM WAS INTRODUCED REPLACING ONE THAT WAS FUNCTIONING WELL. NONE OF THE HOSPITAL IT EXPERTS, NOR EXPERTS FROM THE VENDOR, CHECKED THE COMPUTERS TO DETERMINE IF THE IMAGING SYSTEM'S LAUNCHER COULD BE LOADED. AS IT TURNS OUT, INNUMERABLE COMPUTERS DO NOT HAVE THE STORAGE TO LOAD THE THIN CLIENT LAUNCHER AND THUS, NO RADIOGRAPH IMAGES COULD BE SEEN ON THESE SYSTEMS LINKED TO THE HOSPITAL'S EHR. THIS HAS ADVERSE IMPACT ESPECIALLY AT SITES THAT HAVE ONE OR TWO COMPUTERS CAUSING DELAYS IN CARE AND DIAGNOSIS. THIS IS ESPECIALLY DANGEROUS IN PTS WITH CARDIOPULMONARY DISEASE AND HEART FAILURE OR PNEUMONIA. THERE HAS BEEN NO REMEDY PUT FORTH AND NO ONE HAS ASSUMED LEADERSHIP TO SOLVE THE DYSFUNCTION. | fail to load image system | problem with software installation |  | insufficient storage | diagnostic delayed | [5 empty] | FDA2
Event Description: BEFORE GUIDED IMPLANT SURGERY WAS CONDUCTED, THE DOCTOR IDENTIFIED THAT THE IMPLANT LENGTHS BETWEEN THE REFERENCE CHART PROVIDED WITH THE GUIDE AND HIS TREATMENT PLANNING FILE DID NOT MATCH. THE REFERENCE CHART STATED THAT THE IMPLANT LENGTH WAS 6 MM BUT THE PLANNING FILE STATED IT WAS 8 MM. HE ALSO NOTICED THAT THERE WAS NO 6 MM OPTION AVAILABLE FOR HIS SELECTED IMPLANT TYPE. AS THE DOCTOR WAS UNCERTAIN WHICH IMPLANT LENGTH WAS CONSIDERED DURING GUIDE MANUFACTURE, HE CHOSE TO FREE-HAND THE SURGERY. THERE HAVE BEEN NO ADVERSE EVENTS REPORTED FOR THE PATIENT. ALTHOUGH THE ISSUE WAS FIRST REPORTED AS A SURGICAL GUIDE COMPLIANT, THE INVESTIGATION IDENTIFIED THAT THE CAUSE WAS RELATED TO A DISCREPANCY BETWEEN INVIVODENTAL/TXSTUDIO SOFTWARE VERSIONS WHICH WAS USED TO CREATE THE TREATMENT PLAN. THEREFORE, THE SUSPECT MEDICAL DEVICE SECTION REFLECTS THE SOFTWARE PRODUCT INFORMATION. | implant mismatched | component incompatible | incorrect treatment plan during surgery | discrepancy of software version | surgery delayed | [5 empty] | FDA2
Event Description: CADATREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING A STUDY WITH INCORRECT PATIENT LATERALITY. THE PATIENT'S LATERALITY WAS ENTERED AT THE SCANNER INCORRECTLY. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER: COMPLAINT-(B)(4). | incorrect patient laterality | patient data issue |  | user error | [6 empty] | FDA2
Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2017, A RADIOLOGY TECHNOLOGIST CONTACTED MERGE TECHNICAL SUPPORT FOR ASSISTANCE WITH STUDY PROCESSING. DURING TROUBLESHOOTING IT WAS DISCOVERED THAT THE PATIENT ORIENTATION HAD BEEN INCORRECTLY ENTERED AT THE MRI SCANNER. CADSTREAM IS UNABLE TO EDIT IMAGE DICOM DATA. PER THE CUSTOMER'S REQUEST, ADJUSTMENTS WERE MADE TO REACH A 'READY TO READ' STATUS WITH THE INCORRECT DICOM REFLECTED ON ALL IMAGES. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | incorrect patient orientation | patient data issue |  | user error | [6 empty] | FDA2
Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED ((B)(4)) THAT FOLLOWING A ROUTINE SYSTEM POWER DOWN, THE TIME WITHIN THE SYSTEM IS INCORRECT AND THE DRIVES ARE NOT AVAILABLE. | incorrect time | reset issue |  | power blackouts | [6 empty] | FDA2
Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED ((B)(6)) THAT IMAGES ARE PROCESSING FINE WITH THE EXCEPTION OF THE RIGHT AND LEFT DESIGNATION ON THE SUBTRACTION IMAGES. CUSTOMER RECENTLY UPGRADED TO CADSTREAM 6.0.1. WHEN CADSTREAM BREAKS UP THE BREAST SEQUENCES INTO RIGHT AND LEFT, THE SEQUENCES WERE LABELED AS R POST IMAGES, OR R SUB 2, ETC. USING VERSION 5.0.7 OF CADSTREAM. NOW ONLY THE POST IMAGES ARE LABELED. | incorrect image labeling | device markings issue | [8 empty] | FDA2
Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED (CASE (B)(4)) THEY ARE HAVING ISSUES SENDING IMAGES TO CADSTREAM. THIS COULD CAUSE A DELAY IN PATIENT CARE. | fail to send image | communication or transmission issue |  |  | diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT IMAGES WERE SLOW TO LOAD DURING A BIOPSY PROCEDURE. THE CUSTOMER REPORTED COMPLETING THE BIOPSY BY ESTIMATING THE TARGET LOCATION OF THE BIOPSY BASED ON LANDMARKS AND NOT VISUAL ENHANCEMENT ON IMAGES. ESTIMATING TARGETS HAS THE POTENTIAL TO LEAD TO AN INCORRECT FOLLOW UP WHICH MAY LEAD TO A DELAY IN DIAGNOSIS AND TREATMENT. | image load slow | improper device output | [8 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST ASSISTANCE WITH RE-PROCESSING IMAGES THAT WERE BELIEVED TO BE PROCESSED INCORRECTLY. SUPPORT REVIEWED THE INCOMING IMAGES AND SYSTEM REMOTELY. SUPPORT FOUND THAT THE IMAGES APPEARED TO BE SCANNED WITH INCORRECT PATIENT ORIENTATION. THE CUSTOMER CONFIRMED THE ORIENTATION OF THE IMAGES. CADSTREAM DOES NOT ALLOW FOR THE EDITING OR MODIFICATION OF THE ORIGINAL IMAGE. THE CUSTOMER HAS THE ABILITY TO MAKE ANNOTATIONS REGARDING THE ORIENTATION, HOWEVER, CADSTREAM WILL CONTINUE TO REFLECT THE ORIENTATION AS ORIGINALLY SCANNED ON THE PROCESSED IMAGES. SUPPORT INSTRUCTED THE CUSTOMER REGARDING THESE IMAGES AND ORIENTATION.. THE INCORRECT ORIENTATION IS NOT A MALFUNCTION OF THE CADSTREAM DEVICE NOR IS IT A FAILURE OF THE DEVICE TO MEET SPECIFICATION. THE ISSUE IS USE ERROR CAUSED ORIGINATING AT THE MRI SCANNER DUE TO IMPROPER DATA ENTRY. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect patient orientation | patient data issue |  |  | diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST ASSISTANCE WITH RE-PROCESSING IMAGES THAT WERE BELIEVED TO BE PROCESSED INCORRECTLY. SUPPORT REVIEWED THE INCOMING IMAGES AND SYSTEM REMOTELY. SUPPORT FOUND THAT THE STUDIES WERE PROCESSED AS EXPECTED; HOWEVER, IT WAS NOTED THAT THE IMAGES APPEARED TO BE SCANNED WITH INCORRECT PATIENT ORIENTATION. THE CUSTOMER CONFIRMED THE IMAGES WERE SCANNED SUPINE INSTEAD OF PRONE SO ALL OF THE IMAGES WERE UPSIDE DOWN. CADSTREAM DOES NOT ALLOW FOR THE EDITING OR MODIFICATION OF THE ORIGINAL IMAGE. THE CUSTOMER HAS THE ABILITY TO MAKE ANNOTATIONS REGARDING THE ORIENTATION, HOWEVER, CADSTREAM WILL CONTINUE TO REFLECT THE ORIENTATION AS ORIGINALLY SCANNED ON THE PROCESSED IMAGES. SUPPORT INSTRUCTED THE CUSTOMER REGARDING THESE IMAGES AND ORIENTATION. THE INCORRECT ORIENTATION IS NOT A MALFUNCTION OF THE CADSTREAM DEVICE NOR IS IT A FAILURE OF THE DEVICE TO MEET SPECIFICATION. THE ISSUE IS USE ERROR CAUSED ORIGINATING AT THE MRI SCANNER DUE TO IMPROPER DATA ENTRY. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | image scanned incorrectly | image reversal |  |  | diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2014 A CUSTOMER CALLED TO REPORT THAT A STUDY WOULD NOT PROCESS. THE ERROR WAS THAT CADSTREAM COULDN'T CREATE REFORMATS OF THE RIGHT/LEFT DUE TO SERIES BEING OBLIQUED. THE TECH USED A SKIN MARKER THAT WAS TOO LARGE, GENERALLY USED FOR EXTREMITIES, AND THE CADSTREAM SOFTWARE IDENTIFIED IT AS THE CHEST WALL. MERGE CUSTOMER SUPPORT TRIED TO TURN OFF EVERY PROCESSING WITHIN THE PROCESS TAB AND CREATED A MATCH REFORMAT WHICH ALLOWED THE CHEST WALL TO BE MOVED BACK, BUT WHEN THE CADSTREAM REPROCESSED THE STUDY THE 'FADED' AREA WAS STILL IN ITS' ORIGINAL LOCATION. MERGE CUSTOMER SERVICE WAS UNABLE TO PROCESS THE STUDY. THE PATIENT MAY NEED TO BE RESCHEDULED FOR A SECOND SCAN. (B)(4) Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | fail to create reformat | device operates differently than expected |  |  | reschedule scan | 2 |  |  |  |  | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED AN IMAGE EXPORTED FROM CADSTREAM WAS SHOWING UP UNDER A DIFFERENT PATIENT IN PACS. WHILE THE PACS SYSTEM FLAGGED THE INCOMING IMAGE AND ALERTED THE CUSTOMER, IT IS UNCLEAR AT THIS TIME IF THE INCORRECT IMAGE ASSIGNMENT IS ATTRIBUTED TO THE CADSTREAM SYSTEM OR THE PACS SYSTEM. A PATIENT'S IMAGE SHOWING UP UNDER A DIFFERENT PATIENT HAS THE POTENTIAL OF LEADING TO A DELAY IN CARE AND/OR AN INCORRECT TREATMENT OR DIAGNOSIS. HOWEVER, THERE IS NO INDICATION OF ANY HARM TO THE PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | incorrect patient data | patient data issue |  | fail to flag incorrect image | incorrect treatment | 3 |  |  |  |  | FDA2
Event Description: CALCIUM SCORING IN MERGE UNITY PACS IS A NON-INVASIVE WAY OF OBTAINING INFORMATION ABOUT THE PRESENCE, LOCATION, AND EXTENT OF CALCIFIED PLAQUE ON CORONARY ARTERIES. DURING DEVELOPMENT OF THE NEXT RELEASE OF UNITY (VERSION 11.0), THE DEVELOPMENT TEAM FOUND THAT THE CURRENT ALGORITHM OVERESTIMATES THE CALCIUM SCORE FOR SITES THAT OBTAIN CALCIUM SCORING SCANS AT SLICE THICKNESS OF <3MM AND A SLICE INTERVAL OF 0. IT WAS PRESUMED THAT THE SCORE FOR SITES WITH A SLICE THICKNESS OF >3MM AND A THICKNESS OF 0 WOULD BE UNDERESTIMATED, HOWEVER PERFORMING CORONARY CALCIUM SCORING CTS USING A SLICE THICKNESS OF >3MM IS UNLIKELY. THE ERROR IN THE CALCULATION MINIMALLY IMPACTS EVIDENCE AND MILD EVIDENCE OF CORONARY ARTERY DISEASE, AND WAS DEEMED UNLIKELY TO CAUSE ADVERSE HEALTH CONSEQUENCES. THERE WERE NO REPORTED PATIENTS WHO WERE AFFECTED BY THE ERROR. UNITY 11.0.2 WAS RELEASED ON 11/25/2015 TO ACCOUNT FOR SLICES NOT EQUAL TO 3MM AND TO DISPLAY A WARNING WHENEVER SLICE THICKNESS IS NOT 3MM. | scoring overestimated | false positive result |  | inappropriate algorithm | inappropriate diagnostic | 3 |  |  |  |  | FDA2
Event Description: CUSTOMER CALLED ABOUT A MERGE CARDIO REPORT FOR ONE PATIENT THAT IS ATTACHED TO THE STUDY OF ANOTHER. ONE PATIENT HAD US STUDY DONE ON (B)(6) 2012. IT HAS A CONFIRMED REPORT, HOWEVER, THE REPORT BELONGS TO ANOTHER PATIENT WHO HAD US STUDY ON (B)(6) 2012. WHEN TRYING TO VIEW THE REPORT, THE USER GETS THE SAME GLOBAL EXCEPTION HANDLER ERROR AS BEFORE: EXCEPTION INFORMATION: EXCEPTION TYPE: HTTPEXCEPTION. EXCEPTION MESSAGE: CANNOT HAVE MULTIPLE ITEMS SELECTED IN A DROPDOWN LIST. THE PDF GENERATION FUNCTIONALITY WAS NOT DESIGNED TO BE EXCLUSIVE PER PATIENT. IF TWO OR MORE USERS WERE REPORTING ON DIFFERENT PATIENTS AND REQUESTED THE GENERATION OF A PDF REPORT FILE AT THE SAME TIME, THE SYSTEM WOULD INCORRECTLY ATTACH THE FIRST PDF REPORT FILE GENERATED TO THE OTHER PATIENT(S). Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. | exception message error | device alarm system issue | [8 empty] | FDA2
Event Description: CUSTOMER REPORT ALLEGES SMOKE COMING FROM SMARTCR COMPUTED RADIOGRAPHY (CR) READER UNIT AROUND 23:00 AFTER FIRST NOTICING A BURNING SMELL AROUND 21:00. UNIT WAS REMOVED FROM NICU AND TAKEN TO PLANT OPS. THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | burning odor | smoking | [8 empty] | FDA2
Event Description: CUSTOMER REPORTED AN ISSUE OF HEMO DATA CREATING INCORRECT PATIENT STUDY WHEN IMPORTED TO CARDIO. IN THE VERY RARE COINCIDENCE WHERE TWO PATIENTS ARE ADDED TWO HEMO DATA SETS ARE IMPORTED INTO CARDIO DATABASE AT ALMOST EXACTLY THE SAME TIME, CARDIO WILL POTENTIALLY ASSOCIATE THE HEMO DATA WITH THE INCORRECT PATIENT. THE HEMO DATA WOULD NOT SHOW WITH THE EXPECTED PATIENT BUT WITH ANOTHER PATIENT. THIS ISSUE CAN POTENTIALLY RESULT IN ONE OF TWO RISKS: DELAY IN PATIENT CARE SINCE THE PHYSICIAN WILL NEED TO SEARCH FOR THE HEMO DATA THAT WAS EXPECTED OR ASK THAT DATA GET REIMPORTED BEFORE THE REPORT CAN BE FINALIZED. IN THE VERY RARE OCCURRENCE THAT THE HEMO DATA IS SAVED UNDER A PATIENT THAT HAS A HEMO STUDY SCHEDULED OR COMPLETED WITHIN THE SAME TIME FRAME. THE PHYSICIAN CAN LOOK AT THE HEMO DATA OF THE INCORRECT PATIENT AND ISSUE THE WRONG TREATMENT OR DIAGNOSIS. (B)(4). | incorrect patient information | patient data issue |  | user error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: CUSTOMER REPORTED FRAME SKIPPING ISSUES WHEN PLAYING CINE SERIES. THE RADIOLOGIST RECENTLY MISSED A BRAIN BLEED ON A PATIENT WHICH CAUSED A NEGATIVE OUTCOME, BUT DID NOT RESULT IN DEATH. THE BRAIN BLEED APPARENTLY SHOWED ON ONLY ONE IMAGE AND IT WAS THE IMAGE THAT WAS SKIPPED. INFORMATION REPORTED INDICATES A REPORTABLE MALFUNCTION DATA DISPLAY ISSUE (NON-IMAGE QUALITY), WHICH MAY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY REQUIRING MEDICAL INTERVENTION; THEREFORE, A REPORTABLE MALFUNCTION WILL BE FILED. | image skipped | data issue | missed brain bleed image | user error | serious injury | [5 empty] | FDA2
Event Description: CUSTOMER REPORTED THAT SWITCHING STUDIES CAUSES THE MEASUREMENT LIST TO KEEP MEASUREMENTS FROM THE PREVIOUS STUDY, AND MIXES THEM WITH THOSE FROM THE NEW ONE. MEASUREMENT LIST CONTAINS MEASUREMENTS FROM ANOTHER STUDY. | measurement lists mixed | incorrect measurement | [8 empty] | FDA2
Event Description: CUSTOMER REPORTED THAT SYNAPSE WAS UNABLE TO DISPLAY IMAGE FILES AND THE "IMAGE NOT LOADED" MESSAGE WAS DISPLAYED. SUBSEQUENT INSPECTION BY SERVICE PERSONNEL IDENTIFIED THE IMAGES WERE NO LONGER IN THE DATABASE. | fail to load image | application program issue | [8 empty] | FDA2
Event Description: CUSTOMER REPORTED THAT WHEN MEASURING A LESION ON AN UNMAGNIFIED MAMMOGRAPHY IMAGE USING MERGE PACS, AND THEN PERFORMING THE SAME MEASUREMENT ON THE IMAGE MAGNIFIED BY THE HOLOGIC IMAGER, THE MEASUREMENT WAS NOT THE SAME. IF DOING MEASUREMENTS ON MAGNIFIED AND UNMAGNIFIED IMAGES IT IS RELATIVELY OBVIOUS THAT THE MEASUREMENTS ARE NOT CONSISTENT. HOWEVER, IF PERFORMING MEASUREMENT ONLY ON A MAGNIFIED IMAGE IT MAY NOT BE OBVIOUS THAT THE MEASUREMENTS ARE NOT ACCURATE. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | inconsistent measurement | incorrect measurement | [8 empty] | FDA2
Event Description: CUSTOMER REPORTED WHEN MEASURING WITH THE MERGE CARDIO TOOLBAR, IT DISPLAYS CM ON THE WORKSTATION, BUT MM ON THE REPORT. EX: THE RVOT VALUE SENT FROM THE CART IS 1.5CM. THE UNITS OF MEASURE IN THE (B)(4) TABLE ARE CORRECT (CM) HOWEVER THE UNITS OF MEASURE IN THE 17 MEASUREMENTS. XSLT FILE ARE INCORRECT (MM). THIS CAUSES THE NUMBER IN THE REPORT TO DISPLAY AS 0.15CM. WHEN TAKING MEASUREMENTS FROM IMAGES ON THE CARDIO WORKSTATION OR FROM THE US CART, NUMBERS ARE NOT CROSSING TO THE REPORT IN THE CORRECT UNIT OF MEASURE. SPECIFICALLY, TWO MEASUREMENTS: RIGHT VENTRICULAR OUTFLOW TRACT (RVOT) DIAMETER AND MAIN PULMONARY ARTERY ED DIAMETER. SINCE THE TWO VALUES ARE SMALLER THAN THEY ARE WHICH CAN MISLEAD THE CARDIOLOGIST THAT THERE ARE ABNORMALITY. THE CARDIOLOGIST COULD POTENTIALLY PRESCRIBE UNNECESSARY MEDICATIONS OR TREATMENT. (B)(4). | inappropriate measurement unit | incorrect measurement |  |  | unnecessary medication or treatment | [5 empty] | FDA2
Event Description: DISTANCE MEASUREMENTS MAY NOT BE CALCULATED ACCURATELY BY THE UNIVERSAL VIEWER WHEN APPLIED TO IMAGES ACQUIRED WITH A MAGNIFICATION FACTOR ON SIEMEN'S XA MODALITY (ANGIOGRAPHY X-RAY IMAGES) . DURING A PRODUCT DEMONSTRATION OF THE UNIVERSAL VIEWER WEB 6.0 , A RADIOLOGIST MEASURED VALUES ONSCREEN EQUALING TO 1.36CM AND 1.31CM, WHEN THE ACTUAL MEASUREMENTS SHOULD HAVE BEEN AROUND 1CM. IF NOT NOTICED BY THE CAREGIVER, THIS COULD LEAD TO AN OVERESTIMATE OF THE ANATOMY OF INTEREST. WHEN A MODALITY SENDS IMAGES TO PACS THAT ONLY CONTAIN VALUES IN THE IMAGER PIXEL SPACING (0018, 1164) , AND MAGNIFICATION FACTOR (0018, 1114) DICOM TAGS, BUT DO NOT INCLUDE A VALUE IN THE PIXEL SPACING (0020,0030) DICOM TAG THE MEASUREMENTS MAY BE INCONSISTENT WITH THE MAGNIFICATION FACTOR RESULTING A DISCREPANCY IN THE MEASUREMENTS DISPLAYED ON THE VIEWER. GE RECOMMENDS THAT USERS INCLUDE AN APPROPRIATELY PLACED OBJECT OF KNOWN SIZE IN THE IMAGE TO BE USED FOR CALIBRATION THAT WILL ALLOW USERS TO OBTAIN ACCURATE DISTANCE MEASUREMENTS. UPON CALIBRATION THE MEASUREMENTS ARE CORRECT. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: INVESTIGATION REVEALED THERE ARE MULTIPLE PRIMARY ROOT CAUSES THAT LED TO THE REPORTED ISSUE: 1. DIFFERENCE IN DICOM IMPLEMENTATION ACROSS VENDORS AND PRODUCTS 2. LOGIC ERROR IN THE USAGE OF DICOM TAGS IN THE UNIVERSAL VIEWER PRODUCT FOR DIGITAL PROJECTION MODALITIES. 3. REQUIREMENTS ARE LACKING DETAILS. 4. CROSS-FUNCTIONAL COMMUNICATION. 5. INSUFFICIENT DICOM CONFORMANCE STATEMENT. | inaccuracy distance measurement | incorrect measurement | overestimate the anatomy of interest | difference in DICOM implementation across vendors and products; logic error in the usage of DICOM tags | inappropriate treatment or diagnostic |  |  |  | calibration |  | FDA2
Event Description: DOING A MR FUSION PROSTATE BIOPSY WITH ANESTHESIA AND DURING THE PLANNING STAGE OF THE STUDY, THE MACHINE RESET/TURNED OFF. THE ENTIRE MAPPING PROCESS/PROCEDURE HAD TO BE RESTARTED. | device reset | reset issue | restart mapping process |  | diagnostic delayed | [5 empty] | FDA2
Event Description: DUE TO AN IDENTIFIED HARD DRIVE FAILURE ON THE SYNGO.PLAZA SERVER, ALL STUDIES AVAILABLE IN SHORT TERM STORAGE (STS) WERE LOST. AS THERE IS NO ARCHIVE SOLUTION IN THIS SITE, IS NOT POSSIBLE TO RESTORE THE DATA IN THE STS. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED STUDIES AND THE POSSIBILITY TO RECOVER THE DELETED DATA; HOWEVER, THE AFFECTED DATA IS CONSIDERED AS CLINICALLY RELEVANT. THE REPORTED EVENT OCCURRED IN (B)(6). | hard drive failure | device inoperable | loss of data | hardware failure | diagnostic delayed | [5 empty] | FDA2
Event Description: DUE TO HARDWARE ISSUES WITH THE SYNGO.PLAZA'S HARD DRIVES, A YET UNIDENTIFIED AMOUNT OF MEDICAL IMAGES IS NOT ACCESSIBLE ANYMORE. THE SYSTEM WAS CONFIGURED AS A WORKSTATION AND NOT AS A PACS. NO LONG TERM STORAGE ARCHIVE IS CONFIGURED, THUS THERE IS NO POSSIBILITY TO RESTORE DATA FROM THERE. ONCE THE RESTORE ATTEMPTS OF THE AFFECTED HARD DRIVES ARE COMPLETED, AN UPDATED REPORT WILL FOLLOW CONTAINING THE DETAILED INFORMATION ABOUT THE DATA LOSS. | fail to access image | data issue | loss of data | hard drive issue | [6 empty] | FDA2
Event Description: DURING A BILATERAL TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 10 STEM FOR BOTH PATIENT'S LEFT AND RIGHT HIPS. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED SIZE 10 AND SIZE 12 STEMS. IT IS UNKNOWN WHICH SIDE RECEIVED WHICH SIZED STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | inconsistent size | size incorrect for patient | during hip arthroplasty |  | surgery delayed | [5 empty] | FDA2
Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THAT THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A STEM 3 TO 8 SIZES SMALLER THAN THE STEM SIZE THAT WAS IMPLANTED. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 11 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | inappropriate part size | size incorrect for patient | [8 empty] | FDA2
Event Description: DURING AN OPEN SURGERY, THE FACILITY USED THE SUBJECT OTV-S7PRO FOR OBSERVATION OF THE OPERATIONAL FIELD AND RECORDING THE IMAGES. UPON CONNECTING A CAMERA HEAD, OTV-S7PRO-HD-E12 TO OTV-S7PRO, A MALFUNCTION OF THE WHOLE IMAGE TURNING GREEN OCCURRED PERIODICALLY. THE FACILITY ATTEMPTED TO TURN THE POWER SWITCH ON AND OFF BUT THE DEVICE DID NOT COME TO NORMAL. THE PROCEDURE WAS COMPLETED BY REPLACING THE SUBJECT OTV-S7PRO TO CV-190. NO HEALTH DAMAGE TO THE PATIENT WAS REPORTED OTHER THAN REPLACING THE DEVICE. IN RESPONSE TO THE FAILURE, THEY HAD BEEN SUBSTITUTED FOR THE OLYMPUS-OWNED OTV-S7PRO AND OTV-S7PROH-HD-E12 LATER. HOWEVER, THE SIMILAR PHENOMENON OCCURRED AGAIN. | image turning color | image display error | during open surgery |  | surgery delayed | [5 empty] | FDA2
Event Description: DURING ARTHROSCOPY AND ORIF ORTHOPEDIC PROCEDURE, THE COMPUTER ON THE TOWER (WHERE YOU CAN ENTER THE PATIENT'S INFORMATION AND WHERE IT HAS THE CONTROL RECORDINGS) TURNED OFF SPONTANEOUSLY. IT HAD TO BE RESET AND TURNED BACK ON. WHEN IT WAS RESET, THE PATIENT'S DATA WAS MISSING FROM THE SYSTEM. NO INJURY TO THE PATIENT. | unexpected shutdown | failure to shut off | loss of patient data during intra-operative |  | surgery delayed | [5 empty] | FDA2
Event Description: DURING INTERNAL REVIEW IT WAS FOUND THAT IN THE ADVANCED REPORTING (AR) SYSTEM OF SYNAPSE CARDIOVASCULAR SOFTWARE, WHEN UPDATING A SAVED VOLUME MEASUREMENT ON AN ECHO REPORT, THE VALUES FOR AREA AND VOLUME APPEARED TO BE SWAPPED IN THE RESULTS PANEL. THIS ISSUE AFFECTS AR VERSIONS 6.0.4 THROUGH 6.2.0. THE IMPACT IS IN ECHO REPORTS AND NON-INVASIVE VASCULAR REPORTS AND AFFECTS 10 OF 17 AR SITES. THERE WAS NO ADVERSE EVENT OR INJURY REPORTED; THIS INCIDENT IS BEING REPORTED IN AN ABUNDANCE OF CAUTION. NO OTHER INFORMATION WAS PROVIDED. | incorrect measurement value | incorrect measurement | during internal review |  | no harm | [5 empty] | FDA2
Event Description: DURING MIGRATION FROM THE SIEMENS SYNGO IMAGING TO SECTRA PACS, IT WAS FOUND THAT 13 STUDIES HAD NOT MIGRATED TO SECTRA PACS. THE DATA LOSS WAS OF CLINICAL RELEVANCE; HOWEVER, NO PATIENT RESCANS WERE REQUIRED. THE LOST DATA IS OLD AND THE CLINICAL RELEVANCE OF LOST DATA IS ASSESSED BY SIEMENS EXPERTS AS VERY LOW. READINGS AND REPORTS FOR ALL STUDIES HAVE BEEN SUCCESSFULLY GENERATED PRIOR TO THE INCIDENT. ONLY ONE IMAGE DATED 2017 IS CONSIDERED CLINICALLY RELEVANT; IT WAS CORRUPTED DUE TO COMPRESSION AFTER RECEIVING. ALL OTHER AFFECTED STUDIES (9 IMAGES FROM 2008 AND ONE IMAGE FROM 2013) ARE BETWEEN 5-10 YEARS OLD AND THEIR CLINICAL RELEVANCE WAS ASSESSED AS LOW. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN GREAT BRITAIN. | data corrupted | patient data issue | loss of patient data during imaging migration | due to data compression after receiving | no harm | [5 empty] | FDA2
Event Description: DURING MIGRATION FROM THE SYNGO IMAGING PACS TO THE SYNGO.PLAZA PACS IT WAS IDENTIFIED THAT 297,808 IMAGES BELONGING TO 1,434 STUDIES COULD NOT BE TRANSFERRED TO THE NEW PACS. THEREFORE, THESE IMAGES ARE NOT AVAILABLE IN THE NEW ARCHIVE FOR COMPARISON WITH PRIOR STUDIES. THE AFFECTED DATA IS DATED WITHIN THE TIME FRAME 2003 TO 2017. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE INCIDENT OCCURRED IN (B)(6). | fail to transfer image | communication or transmission issue | loss of patient data during imaging migration | due to data compression after receiving | no harm | [5 empty] | FDA2
Event Description: DURING TWENTY-THREE TOTAL HIP PROCEDURES, THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A STEM 3 TO 8 SIZES SMALLER THAN THE STEM PLANNED TO BE IMPLANTED. IN EACH PROCEDURE, THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED THE PROPER STEM. THERE WERE NO PATIENT INJURIES OR DELAYS IN PROCEDURE. | incorrect implant size | size incorrect for patient | during hip procedure | software error | surgery delayed | [5 empty] | FDA2
Event Description: ECHO MACHINE PROBE LOCKED UP TWICE. MACHINE REBOOTED. A XSPROBE FROM ANOTHER MACHINE OBTAINED AND TESTING COMPLETED. FDA SAFETY REPORT ID # (B)(4). | probe locked | failure to sense | [6 empty] | device reboot |  | FDA2
Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). ON (B)(6) 2016, A CUSTOMER EMAILED SUPPORT ALLEGING THAT THE MEASUREMENTS ON A DIGITAL RADIOGRAPHY FLUOROSCOPY (RF) IMAGE WERE INACCURATE. MERGE HEALTHCARE INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT MEASUREMENTS ARE DISPLAYED IN CENTIMETERS, UNTIL CALIBRATION (CAL) IS PERFORMED. ONCE CALIBRATION IS PERFORMED ON THE IMAGE, THE IMAGE WILL DISPLAY IN CAL, NOT CENTIMETERS. THE RF IMAGE TYPE THAT THE CUSTOMER WAS USING IS MEASURING IN CENTIMETERS AFTER THE IMAGE IS CALIBRATED WHEN IT SHOULD BE MEASURED IN CAL. THE CUSTOMER CONFIRMED THAT NO PATIENTS WERE HARMED DUE TO THIS ISSUE. HOWEVER THERE IS A POTENTIAL FOR A MIS-TREATMENT OR MIS-DIAGNOSIS DUE TO UNEXPECTED MEASUREMENT UNITS ON AN IMAGE. (B)(4). | inaccuracy image measurement | incorrect measurement | unexpected measurement units | out of calibration | mis-treatment or mis-diagnostic |  |  |  | calibration |  | FDA2
Event Description: A CUSTOMER REPORTED ((B)(4)) THAT THEY WERE UNABLE TO SEND IMAGES TO THE SERVER. THE INVESTIGATION SHOWED THIS WAS AN ISSUE WITH THE PERMISSIONS THE USER SET WHEN THEY BUILT THE DATABASE. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | fail to transfer image | communication or transmission issue | during sending images to server | permission issue | surgery aborted | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED ((B)(4)) THEY WERE UNABLE TO RETRIEVE STUDIES FROM REMOTE DEVICES TO SEND TO THEIR PRIMARY WORKSTATION. THE CUSTOMER REPORTED THAT THEY WERE ABLE TO RESOLVE THE ISSUE. | fail to retrieve image | communication or transmission issue | during sending images to server | software issue | surgery aborted | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED ((B)(4)) THAT AFTER SETTING UP V 4.0.3 OF THE EFILM SERVER, THEY WERE UNABLE TO RECEIVE IMAGES UNLESS EFILM SERVICE IS RESTARTED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. THE INVESTIGATION DISCOVERED THAT IMAGES WERE BEING STORED ON A NETWORK SHARE. MERGE INFORMED THE CUSTOMER THAT THE SYSTEM DOES NOT SUPPORT NETWORK STORAGE. | fail to transfer image | communication or transmission issue | during sending images to server | system does not support network storage | surgery aborted | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED (CASE (B)(4)) IMAGES WERE FREEZING ON THE EFILM SYSTEM. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | image frozen | image display error | [8 empty] | FDA2
Event Description: ON (B)(6) 2017, A CUSTOMER WAS CONCERNED ABOUT A SOFTWARE DEFECT IDENTIFIED IN EFILM V. 3.4 WHICH THE CUSTOMER WAS CURRENTLY RUNNING. THE DEFECT WAS PART OF RECALL 2183926-09/20/2016-074-C (FDA CLASSIFICATION #: Z-0707-2017) REPORTED TO THE FDA AND HAS THE POTENTIAL TO LEAD TO INACCURATE MEASUREMENTS. THE ISSUES DOES NOT IMPACT ALL IMAGES, ONLY SPECIFIC PROJECTION IMAGES. CR, DX, MG AND XA IMAGES ARE THE MOST COMMON TYPES OF IMAGES AND THERE WERE NO ISSUES RELATED TO THESE. USE OF THE RECALLED VERSION OF THIS PRODUCT MAY RESULT IN INCORRECT MEASUREMENTS HOWEVER, A USER WILL LIKELY BE ABLE TO NOTICE THAT THE MEASUREMENTS ARE NOT AS EXPECTED. THE CUSTOMER WAS INFORMED THE DEFECT WAS CORRECTED IN EFILM VERSION 4.1.1 AND ABOVE. THE CUSTOMER IS UNABLE TO UPGRADE TO CURRENT VERSION OF EFILM 4.2 BECAUSE THE OPERATING SYSTEM THE CUSTOMER IS RUNNING IS NOT COMPATIBLE WITH THE CURRENT VERSION OF EFILM. WHEN THE CUSTOMER UPGRADES TO THE MINIMUM SYSTEM REQUIREMENTS, THE CUSTOMER WILL BE ABLE TO UPDATE THEIR SYSTEM TO THE FIXED VERSION OF SOFTWARE. THE CUSTOMER DID NOT INDICATE ANY INACCURATE MEASUREMENTS WERE OBTAINED WITH THIS PRODUCT. INACCURATE MEASUREMENTS MAY RESULT IN A DELAY IN PATIENT CARE AND/OR CHANGE IN RECOMMENDATION FOR TREATMENT, HOWEVER, THERE IS NO INDICATION THAT THE ISSUE, REPORTED BY THE CUSTOMER, RESULTED IN A DEATH OR SERIOUS INJURY. (B)(4). | inaccuracy image measurement | incorrect measurement | during treatment | incompatible software application with operating system | inappropriate treatment or diagnostic |  |  |  | bug fixed |  | FDA2
Event Description: ENDOSCOPIC CAMERA FAILED TO TURN ON DURING CASE; UNABLE TO STORE DATA. | fail to turn on | device inoperable | loss of patient data |  | procedure delayed | [5 empty] | FDA2
Event Description: E-PRESCRIPTIONS WERE SENT USING ARIA EHR SOFTWARE PRODUCED BY (B)(4). THEY WERE NEVER RECEIVED BY (B)(6) PHARMACY ON (B)(6). EHR VENDOR INFORMED, BUT INSISTS THAT PRESCRIPTIONS WERE SENT. PHARMACY INSISTS THEY NEVER RECEIVED THEM. | fail to receive prescription | communication or transmission issue | [8 empty] | FDA2
Event Description: ER CALLED AN AMI. PT WAS BROUGHT TO LAB AND LOADED ONTO TABLE. PT WAS CLEANED AND DRAPED PER PROTOCOL. ARTERIAL SHEATH WAS PLACED. WHEN TECH TRIED TO GUIDE C-ARM INTO PLACE TO OBTAIN IMAGES OF ARTERIES THE C-ARM DID NOT WORK. TECH RAN TO BACK AND REBOOTED MACHINE IN ATTEMPT TO CORRECT ERROR. MACHINE CAME BACK UP WITH SAME ERROR MESSAGE. PHYSICIAN DECIDED IT WAS BEST INTEREST OF PT TO CARE FLIGHT TO A HIGHER LEVEL OF CARE FACILITY TO FINISH HEART CATH. BIOMED AND APPROPRIATE PARTIES WERE NOTIFIED. | fail to run | device inoperable | during procedure |  | surgery canceled |  |  |  | reboot device |  | FDA2
Event Description: FOLLOWING AN UPGRADE TO CHRS 14.0 A CUSTOMER SITE, THE CUSTOMER REPORTED ISSUES ON (B)(6) 2020 WHERE SOME NEWLY IMPORTED ULTRASOUND IMAGES WEREN'T IN ORDER AND SOME OF THEM WERE CUT OFF AND BLACK. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. | image disordered | image display error |  | software error | [6 empty] | FDA2
Event Description: FOLLOWING THE UPDATE FROM V3.3 TO V3.5 OF THE DYNACAD SOFTWARE IT WAS NOTED THAT THE KTRANS MAP WAS NOT RENDERED CORRECTLY ON THE DYNACAD CLIENT. INVESTIGATION BY THE INVIVO FIELD TEAM DISCOVERED THAT THE KTRANS MAP WAS DISPLAYED CORRECTLY WHEN THE DYNACAD CLIENT WAS INSTALLED ON THE SERVER MACHINE, BUT NOT WHEN THE DYNACAD CLIENT WAS INSTALLED ON A REMOTE MACHINE. | incorrect map displayed | incorrect display | [8 empty] | FDA2
Event Description: FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) RECENTLY COMPLETED A RETROSPECTIVE REVIEW OF COMPLAINT FILES FROM 2013 THROUGH 2015 AND HAS DETERMINED THAT THIS COMPLAINT MAY MEET THE THRESHOLD FOR REPORTING BECAUSE IMAGES WERE REPORTED AS LOST. FMSU IS THEREFORE REPORTING IT AT THIS TIME. THE SITE (B)(6) REPORTS THAT THEY HAVE LOST TWO IMAGES DURING OPERATION OF A CLUSTERED (B)(4) CONSOLE RUNNING V6.2. THESE IMAGES WERE TAKEN BETWEEN 1:00 AND 1:10 PM ON (B)(6) 2013. THESE IMAGES HAD BECOME INACCESSIBLE AND THE (B)(4) WAS THEN REBOOTED. REBOOTING DELETES THE IMAGES FROM THE DATABASE. ONCE THE REBOOT PROCESS WAS COMPLETE THE IMAGES WERE NO LONGER PRESENT ON THE (B)(4) CONSOLE. THE FACILITY DID NOT PROVIDE ANY INFORMATION INDICATING THIS AS AN ADVERSE EVENT, I.E., THAT THE PATIENT WAS RE-IMAGED. FUJIFILM THEREFORE IS CONSIDERING THIS AS A POSSIBLE MALFUNCTION. | missing image series | loss of data | during operation | malfunction | re-imaged and diagnostic delayed | [5 empty] | FDA2
Event Description: AN EXAM AT (B)(6) HOSPITAL IN (B)(6), WAS ENTERED MANUALLY ON THE MODALITY AND THE ACCESSION NUMBER WAS ENTERED INCORRECTLY BY THE TECHNOLOGIST. AS A RESULT, THE EXAM WAS SENT FROM THE MODALITY TO THE PATIENT EXAM FOLDER IN SYNAPSE MATCHING THE ACCESSION NUMBER THAT WAS ENTERED. THIS PATIENT FOLDER HOWEVER WAS NOT FOR THE CORRECT PATIENT. SYNAPSE HAS FUNCTIONALITY DESIGNED TO MITIGATE THIS TYPE OF EVENT. WHEN STUDIES ARE RECEIVED FROM A MODALITY, SYNAPSE PULLS THE PATIENT INFORMATION FROM THE DICOM HEADER. SYNAPSE THEN TRIES TO MATCH THE PATIENT INFORMATION, E.G., THE PATIENT ID, PATIENT NAME, ETC. WITH THE ACCESSION NUMBER. WHEN SYNAPSE FINDS A POSSIBLE MISMATCH IT CREATES AN ANOMALY ALERT. ANYTIME THE EXAM IS OPENED AN ANOMALY ALERT WILL OPEN ON THE USER'S MONITOR SCREEN INDICATING THAT AN ANOMALY EXISTS FOR THIS EXAM. THE ANOMALY ALERT IDENTIFIED IN THE EVENT ABOVE WOULD INDICATE "MISMATCH PATIENT NAME." A SCREENSHOT OF AN ANOMALY ALERT IS PROVIDED FOR YOUR REFERENCE (THIS IS NOT FROM THE PARTICULAR EVENT). THE USER HAS THE ABILITY TO SELECT A CHECKBOX NOT TO SHOW ANOMALIES WHILE IN THEIR CURRENT SESSION WHICH WAS DONE IN THIS INSTANCE. AFTER BEING NOTIFIED OF THE ANOMALY AND TURNING OFF THE WARNING THE RADIOLOGIST STILL WENT AHEAD AND DICTATED THE RESULTS OF PATIENT A'S EXAM TO THE PATIENT B'S EXAM FOLDER; AS A RESULT THERE WAS AN ALLEGED MISDIAGNOSIS OF PATIENT B. FMSU IS NOT AWARE OF WHETHER TREATMENT WAS PROVIDED ON THE BASIS OF THE MISDIAGNOSIS OR OF THE OUTCOME. EXAMS WITH ANOMALIES ARE ALSO COPIED TO AN ANOMALY FOLDER. THE PACS ADMINISTRATOR SHOULD BE REVIEWING THE ANOMALY FOLDER PERIODICALLY TO RESOLVE ANOMALIES. OR, AS IN THIS SITUATION, THE RADIOLOGIST SHOULD HAVE ALERTED THE PACS ADMIN THAT AN ANOMALY NEEDED TO BE RESOLVED. A COPY OF THE PAGE FROM THE OPERATION MANUAL SHOWING THE ADMIN FOLDER AND ANOMALY SUB-FOLDER HAS BEEN ATTACHED. THE SPECIFIC ACTION OF MANUALLY ENTERING AN INCORRECT ACCESSION NUMBER IS CLEARLY USER ERROR AND SYNAPSE CANNOT PREVENT SUCH AN ERROR FROM OCCURING. SYNAPSE OPERATED EXACTLY AS DESIGNED. | enter incorrect patient data | patient data issue | during procedure | user error | misdiagnosis | [5 empty] | FDA2
Event Description: THE DATE OF OCCURRENCE IS UNKNOWN, HOWEVER ONE CAN PRESUME IT WAS CLOSE TO (B)(6) 2014, THE DATE FUJIFILM WAS NOTIFIED BY THE USER.. ORDER ENTERED IN THE (B)(4) WAS SENT TO THE (B)(4) AND THE (B)(4) CONSOLE WORKSTATION WORKLIST. TECHNOLOGIST IMAGED PATIENT AND SENT EXAM IMAGES TO (B)(4). TECHNOLOGIST WENT BACK INTO THE (B)(4) CONSOLE WORKSTATION AND CHANGED THE ACCESSION NUMBER OF THE ACQUIRED EXAM AND TRIED TO RESEND THE IMAGES TO (B)(4). THE IMAGES WITH THE NEW ACCESSION NUMBER WERE NOT ACCEPTED BY THE (B)(4) BECAUSE THE STUDY ID WAS ALREADY ATTACHED TO THE ACCESSION NUMBER ORIGINALLY ASSIGNED TO THE PATIENT (AS ENTERED IN THE (B)(4)). THEREFORE THE IMAGES BECAME STUCK IN THE QUEUE OF THE (B)(4) CONSOLE BECAUSE THEY WERE REJECTED BY THE (B)(4). TECHNOLOGIST MAY HAVE REINITIALIZED EITHER THE (B)(4) CONSOLE OR THE OUTPUT QUEUE. TECHNOLOGIST COULD NOT FIND THE IMAGES AND CONCLUDED THE IMAGES HAD BEEN ERASED FROM THE LOCAL DATABASE. THE PATIENT WAS REIMAGED AND THE TECHNOLOGIST REPORTED THE ORIGINAL IMAGES AS LOST. HOWEVER, THE IMAGES WERE ALREADY IN THE (B)(4) SO THE TECHNOLOGIST DID NOT NEED TO REIMAGE THE PATIENT. | fail to transfer image | communication or transmission issue | image stuck in the queue | user error | reimaged | [5 empty] | FDA2
Event Description: GE PACS UNIVERSAL VIEWER SOFTWARE IS SUBSTANTIALLY WORSE THAN THE PREVIOUS CLIENT CALLED CENTRICITY. THIS PT HAD A BREAST SONOGRAM AND MAMMOGRAM. THE NEW VIEWER CANNOT SHOW MAMMOGRAMS AND SONOGRAMS AT THE SAME TIME. THIS CREATES INTERPRETIVE DIFFICULTIES THAT DIDN'T EXIST WITH THE OLDER CLIENT CALLED CENTRICITY. | unable to display multiple views | incorrect display | during procedure | software error | inconvience | [5 empty] | FDA2
Event Description: HAD MAMMOGRAM AT (B)(6) HOSPITAL AND BECAME VERY ILL WITHIN 3 WEEKS. I HAD SILICONE BREAST IMPLANTS AND HAD NO IDEA THEY BOTH GOT RUPTURED. I STARTED IN (B)(6) GOING TO DOCTOR. UNABLE TO WALK. SEVERE PAIN IN FEET. MANY TEST TO FOLLOW. WAS SENT TO PODIATRIST AND RHEUMATOLOGIST. WAS IMMEDIATELY PUT ON MELOXICAM FOR INFLAMMATION. TESTED POSITIVE ANA FOR POSSIBLE LUPUS. HAD MANY BLOOD TEST AND I HAVE ALL OF THE RESULTS. AS OF(B)(6) I NO LONGER WAS ABLE TO WORK. DIAGNOSED WITH MID FOOT OSTEOARTHRITIS AND SURGERY WAS SUGGESTED. I WAS ALSO DIAGNOSED WITH LOWER BACK OSTEOARTHRITIS AND IN (B)(6) WITH FIBRO. I APPLIED FOR SOCIAL SECURITY DISABILITY AND AS APPROVED. I DID NOT CORRELATE THE BREAST IMPLANTS TILL APPROX (B)(6). I REQUESTED AN MRI AND IT WAS SHOWN BOTH IMPLANTS WERE RUPTURED. THEN UPON INVESTIGATING, I FOUND THAT ALL OF MY SYMPTOMS WERE RELATED TO SILICONE POISONING DUE TO RUPTURED IMPLANTS. I GOT SO SICK SO FAST AND DID NOT KNOW WHY. SINCE THE REMOVAL OF THE IMPLANTS, I HAVE HAD 2 ADD'L RECONSTRUCTION SURGERIES INCLUDING A LATISSIMUS FLAP. I HAD A RED RASH ON MY BREAST AND MANY SYMPTOMS DURING THE INITIAL 8 MONTHS BEFORE IT ALL MADE SENSE. I HAVE PICTURES, NOTES AND MANY BILLS. THE HOSPITAL HAS DENIED THE MAMMOGRAM CAUSED THE RUPTURE BUT BOTH AT THE SAME TIME IS SO OBVIOUS. THE IMPLANTS WERE SILICONE MADE BY MENTOR 600CC. | silicone breast implant ruptured | material rupture |  | mammogram | red rash; paint | [5 empty] | FDA2
Event Description: HOUNSFIELD MEASUREMENTS ARE USED TO MEASURE TISSUE DENSITY AND MAY BE USED AS AN INPUT TO CLINICAL DECISION MAKING. AFTER A PERIOD OF TIME RUNNING FUSION WORKSTATION 3.0.2 P3 AND LATER, THE HOUNSFIELD MEASUREMENT TOOL (AND REPORTEDLY OTHER MEASUREMENT TOOLS) WILL REPORT INCORRECT VALUES. IF THE EXAM IS CLOSED AND THE USER EXITS FUSION WORKSTATION AND THEN LOGS BACK IN AND REOPENS THE EXAM, THE MEASUREMENT WILL THEN BE REPORTED CORRECTLY. THE CUSTOMER REPORTED THAT MEASURING IN MATRIX IS NOT ALWAYS ACCURATE ON TWO DIFFERENT WORKSTATIONS. CLIENT DOES NOT CONSIDER IT A MAJOR ISSUE BUT DOES RESULT IN THE DOCTORS HAVING TO REBOOT THESE WORKSTATIONS. ONCE THIS IS DONE THEN THEY FUNCTION PROPERLY. | inconsistent measurement | incorrect measurement | during procedure | different workstations | diagnostic delayed |  |  |  | roboot workstation |  | FDA2
Event Description: ICONNECT ACCESS DISPLAYED INCORRECT PRIORS REPORT(S) IN THE VIEWPORT AREA ONLY WHEN MORE THAN ONE PRIOR STUDY (2 OR MORE) WAS VIEWED. NO INJURIES WERE REPORTED. ON THE STUDY HISTORY TAB WHERE THE PRIOR STUDIES ARE LISTED, THE CORRECT PRIORS REPORT WAS ALWAYS DISPLAYED ON CLICKING THE STUDY NAME. THE ISSUE HAPPENED ONLY IN THE VIEWER. THERE IS ONLY ONE COMPLAINT REPORTED SO FAR, OVER THE DEPLOYMENT PERIOD FROM (B)(6) 2012 (ICONNECT ACCESS 3.0 RELEASE) UNTIL THE PRESENT. THE LONE COMPLAINT WAS NOT REPORTED FROM AN ACTUAL CLINICAL USER. IT WAS ENCOUNTERED BY A PACS ADMINISTRATOR AT A MERGE CUSTOMER SITE WHILE EVALUATING THE SYSTEM IN PREPARATION FOR AN UPCOMING GO-LIVE. | incorrect report displayed | display misread | user confused | software error | no harm | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHEN LAUNCHING IMAGES FOR PATIENT A, PATIENT B'S IMAGES WERE DISPLAYED. THE RADIOLOGIST NOTICED THAT THE STUDY DESCRIPTION FOR PATIENT A'S EXAM DID NOT MATCH THE IMAGES THAT WERE DISPLAYED. DUE TO ICONNECT ACCESS NOT PRESENTING PATIENT DATA AS EXPECTED THERE IS A POTENTIAL FOR EITHER A DELAY IN TREATMENT OR AN INCORRECT TREATMENT OR DIAGNOSIS. THE ISSUE WAS DETECTABLE BY USER. THERE WAS NO INDICATION OF PATIENT HARM OR MISDIAGNOSIS DUE TO THIS ISSUE. (B)(4). | incorrect patient image displayed | image display error | during procedure |  | treatment delayed | [5 empty] | FDA2
Event Description: ICONNECT ENTERPRISE ARCHIVE A VENDOR NEUTRAL ARCHIVE FOR STORAGE AND COMMUNICATIONS OF MEDICAL IMAGES AND DATA. CUSTOMER REPORTED THAT THEIR PREFETCHER IS NOT WORKING AS EXPECTED. DURING A READ OF A STUDY ((B)(6) 2015: CT CHEST SCAN WITHOUT CONTRAST), THE RADIOLOGIST DID NOT HAVE A RELEVANT PRIOR ((B)(6) 2013: CT CHEST SCAN WITH CONTRAST) THAT HE WAS EXPECTING. PACS ADMINISTRATOR WAS ABLE TO MANUALLY RETRIEVE THE PRIOR NOT AVAILABLE LOCALLY. THE SITE REVIEWED MERGE PACS AND FOUND A THIRD RELEVANT PRIOR THAT WAS NOT RETRIEVED ((B)(6) 2012: CT SCAN). BOTH STUDIES WERE NOT RETRIEVED FALL UNDER THE "LAST 3 IDENTICAL MODALITY" RULE. (B)(4). | scan without contrast | poor quality image | [8 empty] | FDA2
Event Description: IMAGES FROM PHILIPS PET MODALITY DISPLAY INCORRECT SUV VALUE (STANDARDIZED UPTAKE VALUE), WHEN THE IMAGES ARE LOADED TO THE EMBEDDED TRUED APPLICATION FROM THE SYNGO IMAGING PATIENT LIST. IN THE INITIAL CASE, THE CORRECT VALUE IS '3', WHICH IS VALIDATED ON THE PHILIPS PET MODALITY WORKSTATION, HOWEVER, IN TRUED THE VALUE SHOWS AS '15'. THE INVESTIGATION IS ON-GOING. THE REPORTED INCIDENT OCCURRED IN (B)(6). | incorrect display value | image display error | [8 empty] | FDA2
Event Description: IMAGES WERE SENT FROM DAR - 8000I TO PACS AFTER FLUOROSCOPIC EXAMINATION, BUT ONLY ONE WAS STORED IN PACS. EVEN WHEN CHECKING THE DB OF THE DAR - 8000I DEVICE, THERE WAS ONLY ONE IMAGE. THEREFORE, RE-EXAMINATION WAS CARRIED OUT. | fail to store image | loss of data | during procedure |  | diagnostic delayed |  |  |  | re-examination |  | FDA2
Event Description: IN RESPONSE TO THE FAILURE WITH THE FACILITY'S OTV-S7PRO AND OTV-S7PROH-HD-E12 ON (B)(6), THEY HAD BEEN SUBSTITUTED FOR THE OLYMPUS-OWNED EQUIPMENT. DURING AN OPEN SURGERY, THE FACILITY USED THE SUBSTITUTED OTV-S7PRO, WHICH WAS THE SUBJECT DEVICE, AND THE SUBSTITUTED OTV-S7PROH-HD-E12 FOR OBSERVATION OF THE OPERATIONAL FIELD AND RECORDING THE IMAGES. HOWEVER, THE SIMILAR PHENOMENON OCCURRED AGAIN AS THE IMAGE MALFUNCTION TURNING GREEN PERIODICALLY WHICH WAS CAUSED BY THE FACILITY'S EQUIPMENT. THE PROCEDURE WAS COMPLETED BY REPLACING THE SUBJECT OTV-S7PRO TO CV-190. NO HEALTH DAMAGE TO THE PATIENT WAS REPORTED OTHER THAN REPLACING THE DEVICE. | image turning color | image display error | during an open surgery | caused by facility equipment | procedure delayed |  |  |  | substitute camera lists |  | FDA2
Event Description: IN THE PACS SUITE FOR MAMMOGRAPHY IMAGES, IN THE DIAGNOSTIC VIEW (LM VIEW)THE PACS WILL READ AND LABEL THE LEFT 3D TOMO INDICATOR BAR AS REVERSED IN ORIENTATION ON PACS ON BREAST TOMOS (I.E. LATERAL BREAST REGISTERS AS "M" FOR MEDIAL, AND VICE VERSA) FOR THE HOLOGIC AFFIRM PRONE BIOPSY SYSTEM AND THE GE ESSENTIAL WITH SENOCLAIRE. THIS HAS BEEN RECOGNIZED BY THE RADIOLOGISTS AND THEY ARE AWARE OF ITS OCCURRENCE. PHILIPS WAS FIRST CONTACTED ABOUT THIS OBSERVATION IN (B)(6) 2020. | image reversal | image reversal | during procedure | software error | diagnostic delayed | [5 empty] | FDA2
Event Description: IN THIS CASE A DENTIST UTILIZED A WAFER TO SUPPORT THE CORRECTION OF THE PATIENT'S OCCLUSAL SITUATION SURGICALLY. HOWEVER HE DID NOT ORDER THE CORRECT TYPE OF WAFER FOR THE PLANNED SURGERY. THE SURGERY RESULTED IN AN OPEN BITE AND THE PATIENT IS REFUSING ANOTHER SURGICAL INTERVENTION. | incorrect implant type | component incompatible | during the surgery | user order incorrect part | additional surgery procedure | [5 empty] | FDA2
Event Description: IN THIS EVENT IT WAS REPORTED THAT A SURGIGUIDE WAS USED FOR IMPLANT PLACEMENT IN REGION #20. THE IMPLANT WAS PLACED TOO DEEP AND TOO CLOSE TO THE ROOT OF THE NEIGHBOURING TOOTH. THE SURGERY WAS COMPLETED, NO ADDITIONAL TREATMENT WAS NECESSARY. | improper implant position | malposition of sevice | during teeth surgery | user operation issue | patient discomfortable | [5 empty] | FDA2
Event Description: IN THIS EVENT IT WAS REPORTED THAT A SURGIGUIDE WOULD NOT FIT A PATIENT PROPERLY. THE DENTIST STATED IT WAS OFF APPROXIMATELY 3MM, SHIFTED TO THE BUCCAL. THE PROBLEM WAS NOTICED DURING SURGERY. THE PATIENT WAS NOT SEDATED. BONE GRAFT MATERIAL AND A MEMBRANE WERE PLACED AT THE SURGICAL SITE. THE IMPLANTS WERE NOT PLACED, BECAUSE THE OSTEOTOMY WAS TOO FAR BUCCAL, THERE WOULD HAVE BEEN NO PRIMARY STABILITY. AN ADDITIONAL SURGERY IS NOW REQUIRED. | implant dislodged | dislodged or dislocated | during teeth surgery | low stability | additional surgery procedure | [5 empty] | FDA2
Event Description: IN THIS EVENT, IT WAS REPORTED THAT A DENTIST HAD ORDERED A MUCOSA-SUPPORTED SURGIGUIDE. HOWEVER, THE GUIDE HE RECEIVED WAS ALSO SUPPORTING ON THE TWO REMAINING DISTAL TEETH, CAUSING A GAP BETWEEN THE GUIDE AND THE MUCOSA SURFACE. THE GUIDE WAS UNSTABLE AS THERE WAS A SPACE OF 1-2MM BETWEEN THE FLANGE AND MUCOSA BOTH BETWEEN THE BUCCAL AND THE PALATAL SIDES. ALL THIS WAS NOTICED WHEN THE PATIENT WAS IN THE CHAIR. THE DOCTOR TRIED TO REMOVE THE PORTION OF THE GUIDE THAT SUPPORTED ON THE TEETH BUT THIS COMPROMISED THE FIT OF THE GUIDE. HE STILL TRIED TO USE IT BY DRILLING FOR IMPLANT ONE (AT TOOTH POSITION 16) BUT HE ENTERED THE SINUS, WHICH WAS NOT PLANNED. | inappropriate guide | failure to osseointegrate | during teeth surgery | compromised the fit | additional surgery procedure | [5 empty] | FDA2
Event Description: IN THIS EVENT, IT WAS REPORTED THAT A SURGIGUIDE WAS ROCKING AND NOT STABLE. CHANGING AND TRIMMING THE GUIDE DID NOT FIX THE ISSUE ENTIRELY. THE DOCTOR THEN MADE A FLAP TO POSITION THE GUIDE ON THE BONE, BUT THE GUIDE DIDN'T FIT ON THE BONE. NO DRILLING WAS DONE, HOWEVER A SECOND SURGERY IS NEEDED. | unstable surgiguide | unstable | during teeth surgery | component defects | additional surgery procedure | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THEY WERE MISSING MEASUREMENTS AND SOME MEASUREMENTS WERE INCORRECT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE (B)(4), COMPLAINT-(B)(4). | loss of measurement | patient data issue | during procedure | software error | inappropriate diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MULTIPLE DIAGNOSTIC MEASUREMENTS WERE POPULATING THE CLINICAL REPORT INCORRECTLY. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE (B)(4), COMPLAINT-(B)(4). | incorrect measurement value | patient data issue | during procedure | software error | inappropriate diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE MEASUREMENT FOR STRAIN IS NOT CORRECT (GLS, 2D). DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE-(B)(4), COMPLAINT-(B)(4). | incorrect measurement value | patient data issue | during procedure | software error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS THAT WERE NEVER TAKEN ON THE MODALITY WERE POPULATING THE CARDIO CLINICAL REPORT AND DIAGNOSTIC MEASUREMENT MITRAL VALVE PEAK E WAS POPULATING WITH THE INCORRECT VALUE. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | incorrect measurement value | patient data issue | during procedure | software error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: INITIALLY PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT WHEN RECONSTRUCTING DATA BY THE INTELLISPACE PORTAL AUTOSPECT_PRO APPLICATION SOME OF THE RESULTS MAY BE DIFFERENT THAN EXPECTED. THE FEEDBACK INDICATE THAT THE AFFECTED DATA IS CREATED ON LEGACY AXIS/IRIX GAMMA CAMERAS MANUFACTURED BY MARCONI MEDICAL SYSTEMS USING 102 DEGREE ACQUISITION PROTOCOL. FURTHER INFORMATION FROM CUSTOMER RECEIVED ON (B)(6) 2015, LEADING OUR ORGANIZATION TO REPORT IT. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. | inconsistent data | patient data issue | during procedure | inappropriate protocol | diagnostic delayed | [5 empty] | FDA2
Event Description: INTELERAD HAS BEEN MADE AWARE OF A PROBLEM WHICH COULD BE A POSSIBLE PATIENT SAFETY: THE ISSUE IS TRIGGERED BY THE PRESENCE OF THE LESS THAN CHARACTER (<) IN THE REPORT. WHEN VIEWING A REPORT, THE LESS THAN CHARACTER IS NOT PRESENT WHILE IT WAS THERE IN THE ORIGINAL REPORT RECEIVED BY INTELEPACS. IN OTHER CASES, WHEN THE CHARACTER IS PRESENT IN A COMBINATION WITH OTHER CHARACTERS, PARTS OF THE REPORT TEXT FOLLOWING THE LESS THAN CHARACTER ARE NOT DISPLAYED DEPENDING ON THE INTELERAD APPLICATION USED BY THE END-USER TO REVIEW THE REPORTS. NO PATIENT WAS REPORTED IMPACTED BY THE PROBLEM. THIS REPORT IS SUBMITTED PREVENTIVELY. TICKET CASE ASSOCIATED: (B)(4). INVESTIGATION DETAILS IS DOCUMENTED IN (B)(4): IN INTELEVIEWER, WHEN VIEWING A REPORT CONTAINING THE LESS THAN (<) CHARACTER WHICH IS NOT PROPERLY FORMATTED FOR HTML, IF YOU ARE AFFECTED BY THIS CATEGORY, YOU SHOULD PROCEED WITH A REVIEW OF CONFIGURATIONS OF THE EXTERNAL SOURCE. PLAIN TEXT REPORTS FROM EXTERNAL SOURCES THAT HAVE BEEN MODIFIED TO BE COMPATIBLE WITH HTML - THE REPORTS ASSOCIATED WITH THIS PROBLEM ARE FROM THIS CATEGORY. RISK ASSESSMENT: THERE IS A RISK OF MISDIAGNOSIS. IS THIS PROBLEM CURRENTLY DOCUMENTED AS A KNOWN RISK? ú YES. IS THERE CURRENTLY A MITIGATION IN PLACE TO REDUCE THIS RISK? YES/NO ú YES. IF YES, HAS THE MITIGATION WORKED TO PREVENT RISK TO PATIENT? YES/NO ú NO. THE MITIGATION APPLIED TO A DIFFERENT CATEGORY OF REPORT TRANSFORMATION. CAN THIS RISK BE FURTHER REDUCED? DESIGN FACTORS THAT MIGHT MITIGATE RISK: ú YES. ENSURE APPLICATIONS THAT RETRIEVE REPORTS FOR RENDERING AUTOMATICALLY DETECT AND CORRECT UNESCAPED HTML. PROBABILITY OF RISK? BEFORE/AFTER MITIGATION ú BEFORE MITIGATION: REMOTE PROBABILITY BECAUSE IT REQUIRES MISSING CONFIGURATION FOR ESCAPING HTML, THE MIS-RENDERING HAVING CLINICAL SIGNIFICANCE AND THE ABSENCE OF THE AFFECTED TEXT NOT BEING EVIDENT TO EXHIBIT THE BEHAVIOR. | missing report character | patient data issue | during review report | missing configuration for html format | inappropriate treatmetn or diagnostic | [5 empty] | FDA2
Event Description: INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM IN ITS (B)(6) SOFTWARE. THE ISSUE IS TRIGGERED BY THE PRESENCE OF THE PIPE CHARACTER ( ¦ ) IN THE REPORT. IF THE TEXT OF AN INTELEPACS-GENERATED REPORT CONTAINS AN UNQUOTED PIPE CHARACTER (I.E. ¦) THEN QUERYING FOR IT VIA HL7 WILL RESULT IN A REPORT THAT APPEARS TO BE TRUNCATED AT THAT POINT. ONLY CLIENTS WITH VERSIONS 4-11-1 AND LATER USING MULTI METHOD REPORTING OR INTELEDICTATION AS WELL AS THE HL7GATEWAYSERVER DYNAMIC REPORT RENDERER ARE AFFECTED. THIS ISSUE COULD IMPACT THE PATIENT SAFETY BECAUSE IT CAN RESULT IN MISSING REPORT TEXT. WHILE IN SOME CASES IT MAY BE OBVIOUS, IN OTHER CASES IT COULD RESULT IN LOST ADDENDA, WHICH WOULD NOT BE VISIBLE TO THE HEALTH PROFESSIONAL WHO COULD MAKE A DIAGNOSIS WITH INCOMPLETE OR MISSING DATA. ROOT CAUSE ANALYSIS: ROOT CAUSE ANALYSIS: THE DYNAMIC REPORT RENDERER OF HL7GATEWAYSERVER APPLICATION FAILED TO ESCAPE A SPECIAL CHARACTER IN A REPORT, CAUSING THE HL7 OBSERVATION FIELD TO BE UNEXPECTEDLY TRUNCATED WHEN RETRIEVING REPORTS THROUGH HL7. THIS ONLY AFFECTS REPORTS STORED IN THE REPORT ACCESS API (I.E., MMR OR INTELEDICTATION), WHERE THE CLIENT IS USING A DYNAMIC REPORT RENDERER. AFFECTED SOFTWARE VERSIONS SYSTEM IS POTENTIALLY AFFECTED IF THE VERSION "R" IDENTIFIER HAS A NUMBER BELOW THE FOLLOWING: PACS 4-11-1 R69, PACS 4-12-1 R81, PACS 4-14-1 R65, PACS 4-15-1 R47, PACS 4-16-1 R46, PACS 4-17-1 R21. RISK ASSESSMENT: | missing report text | patient data issue |  | fail to escape a special character in a report | inapppropriate treatment or diagnostic |  |  |  | bug fixed |  | FDA2
Event Description: ON (B)(6) 2020 A CUSTOMER CALLED MERGE CARDIO SUPPORT STATING THAT A PATIENT'S REPORT WAS IN A CONFIRMED STATUS WHEN IT WAS ACTUALLY ONLY IN A QC STATUS OR A STATUS WHERE A CLINICIAN HAS NOT READ THE IMAGES AND REPORTED ON THEIR FINDINGS. QC REPORTS WITH THE STATUS OF CONFIRMED, IF SENT TO THE PATIENT'S PHYSICIAN, COULD LEAD TO DELAY OF TREATMENT AS THE PATIENT'S IMAGES HAD NOT ACTUALLY BEEN REVIEWED AND THE FINDINGS DOCUMENTED IN THE REPORT. THE LACK OF THE REVIEW AND REPORTING MAY NOT BE READILY APPARENT TO A PHYSICIAN. REFERENCE COMPLAINT-(B)(4). | unable to read image | patient data issue | during procedure | inappropriate image status | diagnostic delayed | [5 empty] | FDA2
Event Description: INTERNALLY FOUND ISSUE. DISTANCE MEASUREMENTS MAY NOT BE CALCULATED ACCURATELY BY ZERO FOOTPRINT VERSIONS 5.0 SP2 TO SP7 WHEN APPLIED TO IMAGES ACQUIRED WITH A MAGNIFICATION FACTOR. | inaccuracy distance measurement | incorrect measurement | [8 empty] | FDA2
Event Description: INVIVO OFFERS THE ELBS BREAST COILS WITH A GRID THAT HAS TWO ADDITIONAL GRID ROWS WHICH ARE OFF-SET FROM THE REST OF THE GRID. SURELOC DID NOT INCLUDE A CONFIGURATION FOR THIS GRID AND THE GENERIC GRID OPTION IN SURELOC ASSUMES THE GRID PLACEMENT IS WITH-OUT ANY OFF SETS. IF THE USER IS NOT EXPERIENCED ENOUGH TO REALIZE THAT THE TOP TWO GRID ROWS ARE OFFSET, THEN THE BIOPSY MAY END UP MISSING THE DESIRED AREA. INVIVO DID NOT MAKE MERGE HEALTHCARE AWARE OF THIS NEW GRID, AND THIS IS THE REASON FOR IT NOT BEING A RECOGNIZED GRID FOR SURELOC. THE USERS ARE TAUGHT TO DO A CONFIRMATION SCAN PRIOR TO PERFORMING BIOPSY. | unexpected grid | application program issue | during scan | software new feature | procedure delayed |  |  |  | confirmation scan |  | FDA2
Event Description: IT HAS BEEN REPORTED TO PHILIPS THAT DURING A CEREBRAL ARTERY ANEURYSM TREATMENT, THE CUSTOMER NOTICED A DIFFERENCE BETWEEN THE MEASUREMENT OF THE ANEURYSM SIZE ON THE 3DRA INTERVENTIONAL TOOL AND THE LIVE X-RAY IMAGE. NO HARM TO THE PATIENT OR USER WAS REPORTED. PHILIPS HAS STARTED AN INVESTIGATION OF THIS COMPLAINT. | inconsistent measurement value | patient data issue | druing treatment | measure from two images | treatment delayed | 2 |  |  |  |  | FDA2
Event Description: IT WAS IDENTIFIED AT CUSTOMER'S SITE THAT 80 STUDIES PRESUMED TO BE MISSING FROM THE SYNGO DYNAMICS (SD) PRODUCT DATING FROM JUNE 2011 THROUGH APRIL 2014. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH WOULD RESULT IN THE NEED FOR A PATIENT RESCAN. THE REPORTED EVENT OCCURRED IN (B)(6). | missing studies | patient data issue | during procedure |  | procedure delayed |  |  |  | rescan |  | FDA2
Event Description: IT WAS REPORTED IN MULTIPLE CASES THE ORTHOSIZE SOFTWARE TEMPLATED A SMALLER SIZE STEM THAN APPROPRIATE FOR THE PATIENT. THE TEMPLATED SIZE WAS IMPLANTED AND SUBSIDENCE ISSUES RESULTED. | incorrect implant size | size incorrect for patient | subsidence issue | software issue | additional surgery procedure | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A CLINICIAN PERFORMED A PROCEDURE TO PLACE DENTAL IMPLANTS IN REGIONS #18, #19 AND #31 USING A SURGIGUIDE. THE SURGIGUIDE FIT PERFECTLY AND IT WAS STATED THAT THE CLINICIAN FOLLOWED THE DRILLING PROTOCOL. IMPLANTS 18 AND 19 WENT APPROXIMATELY 7.5MM INTO THE 9MM OSTEOTOMY AND COULD NOT BE SEATED COMPLETELY. IMPLANT #31 WENT 6MM INTO THE 9MM OSTEOTOMY. THE CLINICIAN REMOVED ALL IMPLANTS AND VERIFIED WITH AN X-RAY THAT OSTEOTOMY WAS DRILLED TO DEPTH. AFTER A SECOND TRY THE IMPLANTS WERE REMOVED AND ALL SITES WERE RE-GRAFTED. A SECOND SURGERY IS NEEDED TO CONTINUE WITH THE IMPLANT PLACEMENT. DURING INVESTIGATION IT WAS NOTICED THAT THE FINAL DRILL STEP FOR TWO OF THE IMPLANTS WAS INCORRECT ON THE DRILLING PROTOCOL. THIS LEAD TO UNDER PREPARATION OF THE OSTEOTOMY AND INABILITY TO INSERT IMPLANTS TO THE CORRECT DEPTH. | fail to seat osteotomy | positioning issue | during procedure | incorrect drilling protocol | additional surgery procedure | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A CUSTOMER USED A SURGIGUIDE COMPATIBLE WITH THE ASTRA TECH IMPLANT SYSTEM TX TO PLACE TWO IMPLANTS IN THE MANDIBLE. THE PATIENT IS PARTIALLY EDENTULOUS AND THE GUIDE WAS PLACED ON THE REMAINING TEETH DURING SURGERY TO PERFORM THE OSTEOTOMIES AND PLACE THE IMPLANTS. THE OSTEOTOMY FOR THE FIRST IMPLANT WAS NOT CREATED AT THE DESIRED POSITION IN THE PATIENT MOUTH AS PLANNED. THE ROOT OF THE NEIGHBORING TOOTH WAS DAMAGED DURING SURGERY. | improper osteotomy position | positioning issue | during surgery | incompatible ssurgiguide | Patient injury | 3 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A DENTAL IMPLANT HAD TO BE REMOVED BECAUSE IT WAS 2MM ABOVE CRESTAL BONE AFTER USING A SURGIGUIDE TO PLACE THE IMPLANT. | improper implant position | positioning issue | during surgery | user error | additional surgery procedure | 3 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A DIFFERENT PATIENT PROFILE WILL LAUNCH WHEN SELECTING A PATIENT IN THE DEXIS IMAGING SUITE PATIENT SELECTION WINDOW. THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT'S NAME IN DEXIS ADMINISTRATION SCREEN. NO INJURIES WERE REPORTED. THERE WAS NO IMPACT TO PATIENT CARE. | incorrect patient image displayed | image display error | during procedure | software defects | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A SERVER ROOM AT THE CUSTOMER'S FACILITY HAD A POWER OUTAGE IN OCTOBER 2016, WHICH RESULTED IN THE FAILURE OF THE RAID (HARD DISK UNIT) ATTACHED TO THE SIEMENS SYNGO IMAGING SYSTEM. A PREVIOUS POWER OUTAGE THAT OCCURRED EARLIER (SEPTEMBER 2016) HAD ALREADY DEACTIVATED THE RAID MIRRORING FUNCTION. THEREFORE, THE DATA WAS NOT PROTECTED BY ANY BACK-UP OPTION. AFTER THE ASSESSMENT WAS CONDUCTED, IT WAS DETERMINED THAT 15 POTENTIALLY CLINICALLY RELEVANT STUDIES DATED BETWEEN OCTOBER 13TH AND 15TH WERE LOST. THIS DATA WILL NOT BE AVAILABLE FOR DIAGNOSIS IF REQUIRED BY PHYSICIAN. THE SYNGO IMAGING SYSTEM DID NOT CONTRIBUTE TO THE REPORTED DATA LOSS. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. THE REPORTED EVENT OCCURRED IN (B)(6). | missing data | loss of data | during procedure | power outage | diagnostic delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE BROKE DURING SURGERY. THE PATIENT WAS ANAESTHETIZED, HOWEVER THE CLINICIAN DELAYED DENTAL IMPLANT PLACEMENT. Manufacturer Narrative: MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO OBTAIN THE DEVICE FOR EVALUATION. | part broken or damaged | break | during procedure |  | procedure delayed | 2 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE BROKE WHILE THE DOCTOR WAS TRYING TO PLACE GUIDE, AFTER REFLECTING THE TISSUE. THE SURGERY WAS ABORTED AND A SECOND SURGERY IS NEEDED. | part broken or damaged | break | during procedure |  | additional surgery procedure | 3 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT FIT IN A PATIENT'S MOUTH. THE DENTAL IMPLANT PLACEMENT SURGERY HAS BEEN DELAYED. Manufacturer Narrative: THERE HAS BEEN A PREVIOUS REPORT RECEIVED WHERE THIS MALFUNCTION RESULTED IN A SERIOUS INJURY. THEREFORE, IT MUST BE PRESUMED THAT RECURRENCE OF THIS MALFUNCTION COULD POSSIBLY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR REQUIRE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE SUCH. AS SUCH, THIS EVENT IS REPORTABLE PER 21CFR PART 803. WE COMPARED THE BUILD MODEL WITH THE PLASTER MODEL. THERE IS A DEVIATION AT THE TEETH 17 AND 28 WHICH PROBABLY CAUSES THE NON FIT IN THE PATIENT MOUTH. THE MODEL IN PLANNING WASN'T THE SAME AS SITUATION IN THE PATIENT'S MOUTH. | unable to fit in patient mouth | component incompatible | during procedure | teeth deviation | serious injury | 4 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT FIT IN A PATIENT'S MOUTH. THE DOCTOR STATED THAT HE OPENED A FLAP TO TRY TO OBTAIN A BETTER FIT AND DRILLED, BUT THEN PERFORATED THE LINGUAL PLATE. THE IMPLANT PLACEMENT SURGERY WAS ABORTED AND A SECOND CORRECTIVE SURGERY IS NEEDED. | unable to fit in patient mouth | component incompatible | during procedure |  | additional surgery procedure | 3 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT PROPERLY FIT IN A PATIENT'S MOUTH. THE DENTAL IMPLANT PLACEMENT SURGERY HAS BEEN DELAYED. | unable to fit in patient mouth | component incompatible | during procedure |  | procedure delayed | 2 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE THAT WAS MEANT TO BE USED TO PLACE THREE IMPLANTS, COULD ONLY BE USED TO PLACE TWO IMPLANTS. THE SURGERY WAS ABORTED. THE DOCTOR COULDN'T FOLLOW THROUGH WITH THE DRILL SEQUENCE BECAUSE OF THE INTEROCCLUSAL SPACE. THE DOCTOR STATED THAT THE SLEEVES IN THE GUIDE WERE TOO HIGH AND VERY LONG AND THEY SIT OFF OF THE PATIENT'S RIDGE. | unable to place implant | difficult to advance | during procedure | sleeves in the guide too high | surgery aborted | 2 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED FOR PREPARATION OF THREE IMPLANTS. THE GUIDE SEATED CORRECTLY AND THE DRILLING SEQUENCE WAS UNEVENTFUL AND IDENTICAL FOR ALL THREE OSTEOTOMIES. THE IMPLANTS WERE PLACED WITHOUT THE GUIDE THROUGH A TISSUE PUNCH. POST OP X-RAY REVEALED THAT ONE OF THE IMPLANTS WAS 3 MM APICAL TO PLANNED LOCATION. THE DOCTOR WAS UNABLE TO RETRIEVE THE IMPLANT. | improper implant position | positioning issue | during procedure | user error | additional surgery procedure | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO CREATE THE OSTEOTOMY FOR THREE DENTAL IMPLANTS (TOOTH POSITION 5, 6 AND 7). AFTER THE FIRST DRILLING STEP THE SURGICAL GUIDE WAS REMOVED TO COMPLETE THE ADDITIONAL DRILLING STEPS FREE HANDED ACCORDING TO THE PILOT GUIDE DRILLING PROTOCOL. AFTER PLACING THE IMPLANTS FREE HANDED, THE CUSTOMER DISCOVERED THAT THE THREE IMPLANTS WERE POSITIONED TOO BUCCAL. A SECOND SURGERY IS NEEDED TO CORRECT THE POSITIONING OF THE IMPLANTS. | improper implant position | positioning issue | during procedure | inappropriate drilling protocol | additional surgery procedure | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO PLACE A DENTAL IMPLANT. THE IMPLANT TRANSFER WAS TOO SHALLOW AND THE GUIDE DID NOT FIT. THE CLINICIAN REPORTED THAT ADDITIONAL STEPS DURING THE SURGERY WERE NEEDED TO PLACE THE IMPLANT DEEPER. A SECOND CORRECTIVE SURGERY WAS NOT NEEDED. | improper implant position | positioning issue | during procedure | inappropriate drilling protocol | surgery delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO PLACE A DENTAL IMPLANT. THE SURGIGUIDE SEATED PERFECTLY AND THE DOCTOR FOLLOWED WITH THE DRILLING SEQUENCE. WHEN THE DOCTOR ATTEMPTED TO PLACE THE IMPLANT, IT SEATED ONLY TO THE CREST OF GINGIVA INSTEAD OF THE CREST OF THE BONE. THE DOCTOR STATED THAT HE HAD TO REMOVE THE IMPLANT AND PLACE BONE GRAFTING MATERIAL IN THE IMPLANT SITE AND THAT A CORRECTIVE SURGERY WILL BE NEEDED IN APPROXIMATELY FOUR MONTHS. ADDITIONAL INFORMATION IS PENDING. | improper implant position | positioning issue | during procedure | inappropriate drilling protocol | additional surgery procedure | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT DUE TO A SOFTWARE ERROR IN THE SYNGO.PLAZA WITH SW VERSIONS VB10A (ALL HOTFIX LEVELS) AND VB10B INCOMPLETE SERIES WERE ARCHIVED IN DICOM LONG TERM STORAGE (LTA) ARCHIVES. A REVIEW OF THE SYNGO.PLAZA SYSTEMS WITH DICOM LTA ARCHIVES WAS TRIGGERED BY A CORRECTIVE ACTION "SYNGO.PLAZA VB10B NON-MEDICAL SOFTWARE APPLICATION CHECK TOOL V1.0 (DISTRIBUTED VIA UPDATE INSTRUCTIONS SY024/17/P AND SY028/17/P, REPORTED TO FDA UNDER C&R # 2240869-05/25/17-0015-C) INTENDED TO DISCOVER INCONSISTENCIES BETWEEN THE SYNGO.PLAZA DATABASE AND THE FILES ARCHIVED IN THE DICOM LTA. DURING THIS CHECK IT WAS DISCOVERED THAT 9249 STUDIES WERE MISSING SERIES IN THE ARCHIVE FOR THIS PARTICULAR CUSTOMER. DUE TO THE HIGH AMOUNT OF AFFECTED STUDIES IT WILL BE VERIFIED IF THIS IDENTIFIED SOFTWARE ERROR IS THE ONLY ROOT CAUSE. THE REPORTED ISSUE OCCURRED IN (B)(6). | incomplete data archived | loss of data | during procedure | software error | diagnostic delayed | 2 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT DURING AN INITIAL TOTAL HIP ARTHROPLASTY PROCEDURE ON AN UNKNOWN DATE THE ORTHOSIZE V1.2.6 SOFTWARE DID NOT CORRECTLY TEMPLATE THE STEM. THE SURGEON BROACHED AND ATTEMPTED TO IMPLANT THE STEM, BUT IT DID NOT FIT. THE STEM WAS REMOVED AND THE SURGEON CONTINUOUSLY BROACHED UNTIL THE APPROPRIATE SIZE WAS REACHED.Event Description: IT WAS REPORTED THAT DURING AN INITIAL TOTAL HIP ARTHROPLASTY PROCEDURE ON (B)(6), 2015 THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 8 STEM. THE SURGEON BROACHED TO SIZE 8 AND ATTEMPTED TO IMPLANT THE STEM; HOWEVER IT DID NOT FIT. THE STEM WAS REMOVED AND THE SURGEON CONTINUOUSLY BROACHED UNTIL THE APPROPRIATE SIZE WAS REACHED. A SIZE 12 STEM WAS IMPLANTED. | incorrect implant size | size incorrect for patient | during hip arthroplasty procedure | software error | procedure delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | unintended part subsidence | migration of device or device component | post procedure | inappropriate part size | procedure delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT IN THE COURSE OF SEVERAL YEARS THE CONCERNED CUSTOMER SITE HAD EXPERIENCED SEVERAL INCIDENTS RELATED TO THEIR NETWORK ATTACHED STORAGE (NAS). FIRST REPORTED OCCURRENCE WAS A HARDWARE CRASH AT THE NAS LEVEL WHILE DATA MIGRATION TO THE NAS WAS IN PROGRESS. SECOND REPORTED OCCURRENCE WAS AN INCORRECT MAINTENANCE OF A PROVISIONALLY ESTABLISHED SET OF SMALLER RAIDS (REDUNDANT ARRAY OF INDEPENDENT DISKS ) USED INSTEAD OF THE CRASHED NAS (LEADING TO WRONG FILE SYSTEM REFERENCES IN THE ARCHIVE DATABASE). THERE ARE NO INJURIES RELATED TO THIS EVENT. THIS EVENT OCCURRED IN (B)(6). | hardware crash | device inoperable | while data migration | hardware failure | diagnostic delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT IN THE SYNGO IMAGING XS THERE IS A POTENTIAL FOR MALFUNCTION WHEN USING THE SYNGO IMAGING XS, VERSION VA70A SOFTWARE AND HIGHER. WITH VA70A APPLICATIONS ROI (REGION OF INTEREST) ARE SHOWING INCORRECT VALUES AND DO NOT MATCH TO SYNGO.VIA AND OSIRIX. THE ISSUE IS ISOLATED TO IMAGES OF PIXEL DEPTH GREATER THAN 12 BIT (99% OF ALL STORED IMAGES). ANALYSIS SHOWED THAT THIS BEHAVIOR IS CAUSED BY A SOFTWARE BUG. A MISCALCULATION DONE FOR THE GREY SCALE VALUES IN THE FUNCTIONS "REGION OF INTEREST", "PIXEL LENS", "EDGE ENHANCEMENT", AND "HISTOGRAMS" MAY OCCUR IN SOME SITUATIONS. THE MIN/MAX VALUE OF PIXEL DENSITY IN THE FUNCTIONS "REGION OF INTEREST", "PIXEL LENS", "EDGE ENHANCEMENT", AND "HISTOGRAMS" IS NOT CORRECT FOR IMAGES WITH A PIXEL DEPTH GREATER THAN 12 BITS. IT HAS TO BE NOTED THAT THE MENTIONED MEASUREMENTS SHOWN IN PRIOR IMAGES ARE ALSO NOT CORRECT. ALL VALUES THAT ARE CALCULATED FOR IMAGES WITH A PIXEL DEPTH LESS THAN 12 BITS ARE CORRECT. THIS INCIDENT OCCURRED IN (B)(6). | incorrect pixel depth calculated | image display error | during procedure | software bug | user confused | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT PATIENT DEMOGRAPHIC DATA UPDATES IN MERGE PACS ARE NOT REFLECTING IN ICONNECT ACDESS 4.0. THE IMAGES ARE BEING PULLED WITH STALE DEMOGRAPHIC DATA. THE CORRECT STUDY LAUNCHES, THE STUDY'S TITLE BAR IS CORRECT, BUT THE DEMOGRAPHICS WITHIN THE VIEWPORT MAY CONFLICT. THE USER CAN GO BACK TO MERGE PACS TO CHECK THE INFORMATION. THE MANUFACTURER REVIEWED THE COMPLAINTS DATABASE TO IDENTIFY ANY ADDITIONAL CUSTOMER COMPLAINTS, AND NONE WERE IDENTIFIED. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | conflicted data | data issue | during review | software bug | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE ISSUE: ONE TRIAL ANALYSIS CHART DOES NOT CORRECTLY REFLECT THE CHANGE TO LEG LENGTH WHEN CHANGING BETWEEN PROXIMAL BODY HEIGHTS. | incorrect chart | data issue | [8 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE RADIOLOGY SERVER AT THE MEDICAL FACILITY WAS INFECTED BY THE WANNACRY VIRUS. THE SEVER WAS DOWN FOR THREE WEEKS WITH NO REQUEST FROM THE USER FOR RECOVERY ASSISTANCE. IT WAS DETERMINED THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED. HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. | software virus infected | application security issue | loss of historical images | software virus | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE SYNGO IMAGING XS SERVICES WOULD NOT START. THE STATION HAD TO BE REBOOTED, HOWEVER, THE SAME ERROR MESSAGE "NETWORK SERVICE STOPPED" KEPT BEING DISPLAYED ON THE SCREEN. THE ONLY INFORMATION THAT COULD BE EXTRACTED FROM THE LOGS WAS "E:\ DRIVE HAS NO ENOUGH SPACE" AND "MAGIC VIEW WAS NOT ABLE TO ACCESS FILES ON IT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE INCIDENT WAS REPORTED IN (B)(6). | loss of connection | connection issue | during procedure | insufficient storage | diagnostic delayed |  |  |  | reboot |  | FDA2
Event Description: IT WAS REPORTED THAT THE SYSTEM FAILED TO BOOT UP. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. | fail to boot up | device inoperable |  |  | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE SYSTEM LOCKED UP. THERE WAS NO PT INJURY OR DEATH REPORTED. | system locked up | application security issue |  |  | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE SYSTEM WOULD NOT BOOT UP. THIS CAUSED A TOTAL LOSS OF NAVIGATION FUNCTION. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY WITH THIS CASE. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE COMPUTER WAS EVALUATED AND REPLACED. THE SYSTEM SOFTWARE WAS RELOADED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | fail to boot up | device inoperable | before procedure | software issue | no harm |  |  |  | reload software |  | FDA2
Event Description: IT WAS REPORTED THAT WHILE USING A SURGIGUIDE TO PLACE SIX IMPLANTS, THE GUIDE SHIFTED DURING THE DRILLING IN THE VESTIBULAR AREA (CLOSE TO SITE 32 AND 42) RESULTING IN A FRACTURE IN THE BONE. THE SURGERY WAS NOT COMPLETED. ALSO IMPLANT 36 DOES NOT SEEM TO BE PLACED IN THE PLANNED AREA. | unstable surgiguide | unstable | during surgery | user error | bone fracture (patient serious injury) | 4 |  |  |  |  | FDA2
Event Description: IT WAS REPORTED THAT WHILE USING TOMTEC-ARENA ZERO IF TWO TABS ARE OPEN ON THE BROWSER AT THE SAME TIME, IF THE USER SWITCHES FROM ONE TAB TO THE OTHER, THEN THE PATIENT INFORMATION FROM THE FIRST TAB IS COPIED TO THE PATIENT STUDY ON THE SECOND TAB. TO DATE NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS ISSUE. | incorrect patient information | patient data issue | during procedure | software issue | diagnostic delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THE SYS EXHIBITED AN UNCOMMANDED SHUT DOWN. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. | unexpected shutdown | device inoperable | during procedure |  | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THE SYSTEM DISPLAYED A FIELD DISTORTION MESSAGE AND LOCKED UP. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY ASSOCIATED W/THE COMPLAINT. | distorted image | image display error | during procedure | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED TO SIEMENS BY A USER THAT A SIMPLE FAILURE TO LOAD AN IMAGE STUDY FROM THE LONG TERM ARCHIVE OCCURRED. HOWEVER, SIEMENS IDENTIFIED THAT AN UNKNOWN AMOUNT OF IMAGE STUDIES WAS MISSING IN THE SITE'S LONG TERM ARCHIVE. THE HISTORICAL STUDY IMAGES ARE NO LONGER AVAILABLE AND CONSIDERED TO BE UNRECOVERABLE AT THIS POINT OF TIME. HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. THE REPORTED INCIDENT OCCURRED IN (B)(6). | fail to load image | patient data issue | loss of data | software issue | diagnostic delayed |  |  |  | rescan |  | FDA2
Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X WORKPLACE SYSTEM. DURING A TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR), WHEN LOADING AN OUTSIDE CTA (COMPUTED TOMOGRAPHY ANGIOGRAPHY) WITH BOOKMARKS, THE 4D VIEWER CLOSED BEFORE OVERLAY FOR 2D3D FUSION COULD BE SELECTED. THE RESULT OF THE ERROR WAS A DELAY IN PROCEDURE. THERE IS NO REPORT OF IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. | unexpected closure | application program issue | during procedure | software defects | diagnostic delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X WORKPLACE SYSTEM. THE USER REPORTED THAT THE 3D IMAGE RECONSTRUCTION HAS DEVIATIONS. THERE IS NO REPORT OF IMPACT TO THE STATE OF HEALTH OF ANY PATIENT OR USER INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | inappropriate 3D image reconstructed | application program issue | pre-operative image | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X- WORKPLACE. DURING AN INTERVENTIONAL PROCEDURE, A 3D DYNACT SPIN FAILED TO RECONSTRUCT. THE PROCEDURE WAS CONTINUED AND COMPLETED ON AN ALTERNATE SYSTEM. WE ARE UNAWARE OF ANY IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | fail to reconstruct 3D model | application program issue | pre-operative image | software issue | [6 empty] | FDA2
Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X-WORKPLACE SYSTEM. THE USER REPORTED POOR IMAGE QUALITY DURING A PROCEDURE. WE ARE UNAWARE OF ANY IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | poor quality image | poor quality image | pre-operative image | software issue | [6 empty] | FDA2
Event Description: IVUS MACHINE NOT HOLDING A CHARGE OR WORKING WHILE PLUGGED IN. BIOMED WAS CALLED YESTERDAY TO HELP TROUBLE SHOOT. MACHINE HAS BEEN PLUGGED INTO 3 DIFFERENT OUTLETS ALL DAY; WRITER PUSHED IN ALL CONNECTIONS TO ELECTRICAL PACK. MACHINE POWERS ON, RED BATTERY LIGHT FLASHES, AND WHEN MACHINE TURNS ON, IT IMMEDIATELY COUNTS DOWN AND SHUTS OFF. | power down unexpected | loss of power | [8 empty] | FDA2
Event Description: MCKESSON IDENTIFIED THAT 8 STUDIES WERE MISSING ON THE SHORT TERM STORAGE LOCATION PRIOR TO PERFORMING A DATA MIGRATION. ALL 8 STUDIES WERE PREVIOUSLY REPORTED . 7 OF THESE REPORTED STUDIES WERE DETERMINED TO BE LOST AND ONE STUDY WAS IDENTIFIED AS BEING INDEXED TO AN INCORRECT PATIENT. NO PATIENT HARM WAS REPORTED AS A RESULT OF THIS ISSUE. | index to incorrect patient | patient data issue |  | user error | no harm | [5 empty] | FDA2
Event Description: MEDTRONIC O-ARM WITH STEALTH SOFTWARE NOT CONNECTING TO PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS) PROPERLY AFTER UPGRADE, DELAYING PATIENT CARE. MANUFACTURER RESPONSE FOR O-ARM, MEDTRONIC STEALTH (PER SITE REPORTER). WORKED TO CONNECT SYSTEM. | loss of connection | connection issue | loss of data | software upgrade | patient care delayed | [5 empty] | FDA2
Event Description: MERGE (B)(4) APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON 23 JUNE 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING STATING THAT "FOLLOWING DATABASE INTEGRITY CHECK, EXAMS STILL MISSING." WITH EXAMS NOT BEING AVAILABLE AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM FOR A PATIENT. THERE HAVE BEEN NO REPORTS OF DELAY IN TREATMENT OR DIAGNOSIS TO A PATIENT, HARM OR POTENTIAL HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | missing exam | patient data issue | during internal review |  | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT IMAGES WERE TAKEN OF A PATIENT WITH THE INCORRECT LINEALITY AT THE MODALITY AND WERE APPEARING THAT WAY IN MERGE CADSTREAM. DUE TO A USER ERROR, IMAGES ACQUIRED DURING MRI SCANNING WERE LABELED WITH THE INCORRECT ORIENTATION ON THE MR SCANNER. PER THE CUSTOMER'S REQUEST, MANUAL PROCESSING OF STUDY BY MERGE SUPPORT RESULTED IN ALL IMAGES REFLECTING THE INCORRECT LATERALITY AND ANTERIOR/POSTERIOR POSITIONING EMBEDDED IN EVERY IMAGE. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND HAVE THE POTENTIAL TO IMPACT PATIENT TREATMENT WHICH COULD LEAD TO DEATH OR SERIOUS INJURY. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | image orientation incorrect | image orientation Incorrect | during procedure | user error | death | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE SUPPORT INDICATING THAT THERE WAS A USER THAT LABELED IMAGES ACQUIRED DURING MRI WITH THE INCORRECT ORIENTATION VIA THE MR SCANNER. I.E. PATIENT WAS PUT IN THE SCANNER FEET FIRST WHEN IT WAS EXPECTED THAT THE HEAD WOULD HAVE BEEN PUT IN FIRST. DUE TO THIS ERROR OF PATIENT POSITIONING AT THE MODALITY, THE CUSTOMER REQUESTED THAT MERGE SUPPORT MANUALLY PROCESS THE STUDY AND CHANGE THE LATERALITY TO ATTEMPT TO RECTIFY THE USER ERROR. THIS THEN CAUSES THE IMAGES TO REFLECT A DIFFERENT LATERALITY AND ANTERIOR/POSTERIOR POSITIONING EMBEDDED IN THE IMAGES VERSUS WHAT WAS ORIGINALLY TAKEN BY THE MODALITY. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | image orientation incorrect | image orientation Incorrect | during procedure | user error | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON DECEMBER 1, 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH PROCESSING A STUDY WHICH ENCOUNTERED A PROCESSING ERROR. THROUGH TROUBLESHOOTING, SUPPORT DISCOVERED THE PATIENT ORIENTATION HAD BEEN ENTERED INCORRECTLY AT THE SCANNER THE CUSTOMER REQUESTED ADJUSTMENTS TO THE PROCESSING INSTRUCTIONS TO REACH A READY TO READ STATUS, HOWEVER, THE PATIENT ORIENTATION COULD NOT BE CORRECTED. CADSTREAM IS UNABLE TO EDIT THE ORIGINAL IMAGE DICOM DATA, AS SUCH, THE PROCESSED STUDY WILL REFLECT THE INCORRECT PATIENT POSITIONING ON ALL POST PROCESSING IMAGES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD, AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | image orientation incorrect | image orientation Incorrect | during procedure | user error | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(4) 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT POSITIONING WHICH WAS ENTERED INCORRECTLY AT THE MODALITY. CADSTREAM IS UNABLE TO EDIT THE ORIGINAL IMAGE DICOM DATA, AS SUCH, THE PROCESSED STUDY WILL REFLECT THE INCORRECT PATIENT POSITING ON ALL POST PROCESSING IMAGES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD, AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect patient position | positioning issue | during procedure | user error | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: MERGE CARDIO 10.1 REPORTS DO NOT DISPLAY PATIENT IDENTIFIERS IN THE HEADER OF EACH PAGE, THEY ARE ONLY DISPLAYED ON THE FIRST PAGE. IF REPORTS ARE PRINTED TO HARD COPY, PAGES COULD POTENTIALLY BECOME MIXED UP BETWEEN DIFFERENT PATIENTS. PRINTING REPORTS TO HARD COPY SO THEY CAN BE FAXED IS NOT A COMMON METHOD OF DELIVERY BUT IS USED BY THE CUSTOMER THAT REPORTED THIS ISSUE. USER ERROR IS THE MAIN CONTRIBUTING FACTOR. THE FREQUENT USE OF MASS FAXES AND PRINTOUTS IS MAKING IT MORE LIKELY THAT PAGES GET MIXED UP. THIS WOULD NOT BE AN ISSUE IF ALL REPORTS ARE ELECTRONICALLY DISTRIBUTED AS PDFS. | missing patient indentifier | patient data issue | patient mixed | user error | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT WHEN USING THE MEASUREMENT TOOL TO MEASURE RIGHT VENTRICLE TAPE, IN M-MODE THAT THE DECIMAL PLACE IS SHIFTED CAUSING THE MEASUREMENT TO LOOK INCORRECT IN THE REPORT AND WORKSTATION. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | decimal place shifted | data issue | incorrect measurement value | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST RE-MAPPING OF 2 MEASUREMENTS: MITRAL REGURGITATION MEASUREMENT IS POPULATING THE MITRAL VALVE GRADIENT FIELD, AND WHEN MEASURING BOTH MITRAL REGURGITATION AND MITRAL VALVE GRADIENT, THE INCORRECT MR MEASUREMENT IS OVERWRITTEN WITH THE CORRECT MV GRADIENT MEASUREMENT. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect MR measurement | data issue | incorrect MR measurement overwrites the correct measurement | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT PREND VMAX WAS MAPPED TO PR PEAK VELOCITY AND SHOULD BE MAPPED TO PR PEAK IN END DIASTOLE AND THAT PREND PG WAS MAPPED TO PR PEAK GRADIENT AND SHOULD BE MAPPED TO PR GRADIENT IN END DIASTOLE. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MROE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement mapping | data issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THE LV EF (LEFT VENTRICULAR EJECTION FRACTION), LV ID ES PLAX (LEFT VENTRICULAR INNER DIMENSION AT END SYSTOLE, PARASTERNAL LONG-AXIS VIEW), AND RV PEAK PRESSURE (RIGHT VENTRICULAR PEAK PRESSURE) MEASUREMENTS WERE DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. THESE MAPPING STRINGS WERE UPDATED AND THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement value | incorrect measurement | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON 1/11/2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | missing report | patient data issue | never confirmed | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON 1/15/2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THE LEFT ATRIAL INTERNAL DIAMETER MEASUREMENT WAS DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. THIS MAPPING STRING WAS UPDATED AND THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement value | incorrect measurement | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, AN INTERNAL INVESTIGATION REVEALED THAT DUE TO TWO MEASUREMENTS BEING MAPPED TO THE SAME CODE, ONE MEASUREMENT COULD OVERWRITE THE OTHER IN THE CLINICAL CARDIO REPORT. THE "MR FLOW RATE" MEASUREMENT AND THE "MAX VELOCITY (MR)" MEASUREMENT WERE BOTH MAPPED TO CPCODE (10607). INTERNAL RESOLUTION IS IN PROCESS. A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION IS AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO CLIENT REPORTS OF THIS ISSUE. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | measurement overwritten | patient data issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT STATING THAT A RECENT REPORT WAS UNAVAILABLE IN THEIR EMR AND THAT THIS HAD RESULTED IN A DELAY OF CARE FOR A PATIENT. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT FAILED TO EXPORT TO THE EMR BECAUSE HOSPITAL STAFF ENTERED "/WIFE" IN THE PHONE NUMBER FIELD OF THE REPORT AND THE HL7 OUTBOUND MESSAGE WILL FAIL IF IT ENCOUNTERS AN UNEXPECTED FORMAT, SUCH AS LETTERS AND/OR SPECIAL CHARACTERS IN A FIELD DESIGNED FOR ONLY NUMBERS. | missing report | patient data issue | user enter in an unexpected format | user error | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, DURING INTERNAL TESTING OF AN UPCOMING VERSION, DUPLICATE MEASUREMENT MAPPING STRINGS WERE DISCOVERED. UPON FURTHER INVESTIGATION, MERGE HEALTHCARE INTERNAL STAFF DISCOVERED THAT FOR GE VIVID 7 MODALITIES WITH XLS FILE DROPS, 2 MEASUREMENTS (AAO AP DIAM, S" AND INFERIOR VENA CAVA DIAM) HAVE DUPLICATE MAPPING STRINGS AND THE 2 STRINGS HAVE DIFFERENT UNITS OF MEASURE. ONE STRING HAS MM AND THE OTHER STRING HAS CM. THE CONFLICT OF UNITS CAUSES THE MEASUREMENTS TO IMPORT INCORRECTLY. THE NUMERICAL VALUE OF THE MEASUREMENT ROUNDS UP AND THE DECIMAL POINT MOVES. FOR EXAMPLE 3.7CM ROUNDS UP TO 4 AND THEN DISPLAYS ON THE REPORT AS 0.4CM. TO RESOLVE THIS ISSUE, MERGE HEALTHCARE INTERNAL STAFF DELETED ALL ROWS IN THE MAPPING TABLE, POPULATED NEW MAPPINGS, AND VERIFIED THAT THE CORRECT MAPPINGS WERE RETAINED FOR EACH OF THE FORMERLY DUPLICATED REPORT TEXTS. NO FURTHER ACTION IS REQUIRED. DUE TO INCORRECT MEASUREMENTS DISPLAYING ON A DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR MISDIAGNOSIS AND INCORRECT TREATMENT THAT MAY LEAD TO HARM. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO CUSTOMER REPORTS OF THIS ISSUE. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | duplicated measurement mapping | patient data issue | during procedure | conflict measurement units | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE MITRAL REGURG PISA RADIUS WAS MAPPED INCORRECTLY. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, THE MEASUREMENT WAS REMOVED AND REMAPPED. IT WAS CONFIRMED TO THEM BE MAPPED CORRECTLY IN THE CUSTOMER'S SYSTEM. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement mapping | patient data issue | during procedure | inconsist mapping | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT STRESS DATA WAS LABELED INCORRECTLY. MERGE HEALTHCARE SUPPORT HAS BEEN IN CONTACT WITH THE CUSTOMER MULTIPLE TIMES SINCE THIS REPORT AND IS CURRENTLY WORKING ON GATHERING FURTHER INFORMATION. AT THIS TIME, THERE ARE NO REPORTS OF PATIENT HARM OR INJURY DUE TO THIS ISSUE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED AS SOON AS MORE INFORMATION IS AVAILABLE. REFERENCE COMPLAINT (B)(4). | incorrect labeling | patient data issue | [8 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THERE WERE ISSUES WITH MEASUREMENT MAPPINGS IN THE CARDIO CLINICAL REPORT. THEY STATED THAT THE IVC (INFERIOR VENA CAVA) MEASUREMENT WAS POPULATING THE IVS (INTERVENTRICULAR SEPTUM) FIELD, THAT THE AORTIC REGURGITATION DECELERATION MEASUREMENT HAD TOO MANY DIGITS, AND THAT THE DESCENDING AORTA ID APPEARED TO BE MISMAPPED. AT THIS TIME, INVESTIGATION BETWEEN THE CUSTOMER AND MERGE HEALTHCARE IS ONGOING. DUE TO INCORRECT VALUES DISPLAYING IN THE CARDIO CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement mapping | patient data issue | [8 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT FROM ONE PHILIPS MODALITY, THE LEFT VENTRICULAR OUTFLOW TRACK (LVOT) MEASUREMENT WAS POPULATING IN THE CLINICAL CARDIO REPORT INCORRECTLY AND THAT THE CORRECT LVOT MEASUREMENT WAS POPULATING IN THE TR (TRICUSPID REGURGITATION) VENA CONTRACTA FIELD. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement value | incorrect measurement | during procedure | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE STATING THAT IVC (INFERIOR VENA CAVA) MEASUREMENTS WERE MISSING IN CLINICAL REPORTING AND THAT THE IVC MEASUREMENT WAS POPULATING THE IVS (INTRAVENTRICULAR SEPTUM) FIELD ON THE CLINICAL CARDIO REPORT. DUE TO INCORRECT VALUES DISPLAYING ON THE CLINICAL CARDIO REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | missing measurement value | patient data issue | during procedure | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, MERGE HEALTHCARE WAS PERFORMING INTERNAL TESTING OF A NEW 5.X PEDIATRIC ECHO KNOWLEDGE BASE FILE FOR MICHIGAN Z-SCORES. IT WAS FOUND THAT WHILE TESTING THIS VERSION 5.X FILE AND COMPARING WITH THE PREVIOUSLY RELEASED VERSION 3.X FILE FOR MICHIGAN Z-SCORE CALCULATIONS, THAT THERE WAS A DIFFERENCE BETWEEN THE TWO. BOTH THE INTERNAL MERGE HEALTHCARE INVESTIGATION AND THE PAEDIATRIC ECHO KNOWLEDGE BASE SUPPLIER ASCEND'S INVESTIGATION CONCLUDED THAT THERE WAS AN ERROR AND THAT THE MICHIGAN REFERENCE RANGES AND Z-SCORES ARE NOT CALCULATING CORRECTLY. FOR CUSTOMERS USING A 3.X VERSION OF THE ASCEND PEDIATRIC ECHO KNOWLEDGE BASE TO CALCULATE MICHIGAN Z-SCORES, THERE IS A POTENTIAL FOR A MISDIAGNOSIS WHICH MAY LEAD TO PATIENT HARM. HOWEVER, THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT-(B)(6). | data calculated incorrectly | incorrect software programming calculations | during software test | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, DURING INTERNAL TESTING OF A NEW RELEASE BY MERGE HEALTHCARE, IT WAS FOUND THAT THE LENGTH MEASUREMENTS IMPORT TO THE DEPTH LOCATION IN THE VEIN MEASUREMENT TABLE FOR THE VEIN ANATOMY OF THE LEG. THIS OCCURS IN THE 5.0 VERSION OF VASCULAR REPORTING WHICH IS CURRENTLY IN CARDIO VERSIONS 11.0-12.0.1. DEPTH MEASUREMENTS CURRENTLY DO NOT MAP SO THEY WERE NOT OVERWRITTEN. THE ISSUE IS THAT LENGTH MEASUREMENTS POPULATE THE DEPTH FIELDS. THIS WAS FOUND TO IMPACT ALL OF THE VEINS FOR LOWER EXTREMITIES. THERE HAVE BEEN ZERO REPORTS OF THIS ISSUE BY A CUSTOMER. DUE TO LENGTH MEASUREMENTS POPULATING THE DEPTH FIELD, THERE IS A POTENTIAL FOR MISDIAGNOSIS. HOWEVER, THERE ARE ZERO REPORTS OF IMPACT TO A PATIENT AND NO CUSTOMERS HAVE REPORTED THIS ISSUE. REFERENCE COMPLAINT-(B)(4). | incorrect measurement mapping | patient data issue | during procedure | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE, AND STATED THAT THREE MEASUREMENTS WERE MAPPED INCORRECTLY (MEAN VELOCITY RATIO, AORTIC VALVE AREA MEAN VELOCITY, AND AORTIC VALVE AREA BSA MEAN VELOCITY). MERGE HEALTHCARE IS CURRENTLY WORKING WITH THE CUSTOMER TO RESOLVE THIS ISSUE WITH THE CUSTOMER . THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY, OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | incorrect measurement mapping | patient data issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | missing report | patient data issue | during reviews | wrong report status | treatment or diagnostic delayed |  |  |  | manual reset status |  | FDA2
Event Description: ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, AT THIS TIME THE CUSTOMER HAS NOT ALLEGED ANY HARM OR IMPACT TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect measurement value | incorrect measurement | during reviews | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT FETAL MEASUREMENTS TAKEN BY THEIR 3RD PARTY MODALITIES WERE POPULATING THE MERGE CARDIO CLINICAL REPORTING FETAL KNOWLEDGEBASE WITH INCORRECT Z-SCORE VALUES. WITH MERGE CARDIO INCORRECTLY POPULATING Z-SCORE VALUES THAT ARE USED FOR DIAGNOSIS, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. AT THIS TIME NO INFORMATION HAS BEEN PROVIDED THAT INDICATES THAT A DEATH OR SERIOUS INJURY HAS OCCURRED AS A RESULT OF THIS ISSUE. (B)(4). | incorrect z-score value | patient data issue | during fetal measurement | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING WITH THE INCORRECT UNIT OF MEASURE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT NUMBER (B)(4). | inappropriate measurement unit | patient data issue | incorrect value displaying | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MERGE CARDIO'S PDF REPORTS WERE POSTING UNDER THE WRONG PATIENT IN THEIR HOSPITAL INFORMATION SYSTEM. CLINICAL REPORTS POSTING UNDER THE WRONG PATIENT HAS POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT WHICH MAY LEAD TO PATIENT HARM. THE CUSTOMER WAS CONTACTED AND STATED THAT NO PATIENT HARM WAS CAUSED DUE TO THIS ISSUE AND THAT IT IS READILY APPARENT TO THE USER. REFERENCE COMPLAINT NUMBER (B)(4). | different patient | patient data issue | incorrect patient displaying | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THEY COULD NOT FIND A STUDY. THIS STUDY WAS COMPLETED UNDER A DIFFERENT PATIENT/ORDER IN ERROR AND THEN THE PHYSICIAN REPORTED ON THE IMAGES WITH THE INCORRECT PATIENT INFORMATION AND PATIENT ORDER. THE CLINICAL REPORT WAS CONFIRMED ON THE WRONG PATIENT WITH THE WRONG PATIENT INFORMATION, THIS HAS THE POTENTIAL FOR INCORRECT DIAGNOSIS OR TREATMENT OF THE PATIENT WHICH MAY LEAD TO PATIENT HARM. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect patient information | patient data issue | incorrect patient info displaying | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER ALLEGED TO MERGE HEALTHCARE THAT A SOFTWARE MALFUNCTION OCCURRED WHICH ALLOWED A CLINICAL REPORT TO BE CONFIRMED WITH AN INCORRECT PATIENT IDENTIFIER. THIS MALFUNCTION, IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. (B)(4). | incorrect confirmation | computer system security issue | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MERGE CARDIO'S PDF REPORTS WERE POSTING UNDER THE WRONG PATIENT IN THEIR HOSPITAL INFORMATION SYSTEM (HIS). CLINICAL REPORTS POSTING UNDER THE WRONG PATIENT HAS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT WHICH MAY LEAD TO HARM. THE CUSTOMER REPORTED THAT NO PATIENT HARM WAS CAUSED AS A RESULT OF THIS ISSUE AND THAT THE ISSUE IS READILY APPARENT TO THE USER. (B)(4) | different patient | patient data issue | incorrect patient displaying | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT HEMODYNAMIC DATA WAS INCORRECTLY IMPORTED TO THE WRONG PATIENT'S STUDY. INCORRECT DIAGNOSTIC DATA BEING IMPORTED TO A STUDY HAS THE POTENTIAL TO CAUSE INCORRECT TREATMENT OF THE PATIENT, WHICH COULD RESULT IN HARM. (B)(4). | incorrect data imported | patient data issue | in reporting | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE PULMONARY ARTERY Z-SCORES WERE NOT CALCULATING AS EXPECTED IN THEIR ECHO.PEDS CLINICAL REPORT. WITH AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THE CUSTOMER INDICATED THAT THERE WERE NO PATIENTS THAT WERE TREATED INCORRECTLY BASED ON THE OUTCOME OF THE PERCEIVED INCORRECT VALUE. (B)(4). | incorrect z-score value | patient data issue | in reporting | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, MERGE HEALTHCARE RECEIVED AN ALLEGATION FROM A CUSTOMER INDICATING THAT THE MEASUREMENT TOOL ON THE CARDIO WORKSTATION PROVIDED INACCURATE RESULTS FOR MEAN GRADIENT. MERGE INVESTIGATED THIS ALLEGATION AND DETERMINED THAT THE CUSTOMER WAS QUESTIONING THE NO VALUE MEAN GRADIENT CALCULATION. ALL OTHER MEAN GRADIENT CALCULATIONS WERE CONSISTENT WITH THE CUSTOMER'S EXPECTATIONS. POTENTIALLY INACCURATE CALCULATIONS CAN LEAD TO CONFUSION BY USER(S) AND MAY LEAD TO INCORRECT TREATMENT AND/OR A DELAY IN CARE THAT MAY CAUSE HARM TO A PATIENT. ACCORDING TO THE CUSTOMER, THE PATIENT IN WHICH THIS ISSUE WAS FOUND, WAS NOT INCORRECTLY DIAGNOSED AS THEY RECOGNIZED THE DISCREPANCY AND MODIFIED THE VALUE PRIOR TO COMPLETING THE REPORT. THE PHYSICIAN STATED THAT THE MEASUREMENT TOOL WAS ONLY USED RARELY TO MANUALLY CALCULATE MEAN GRADIENTS. (B)(4). | data calculated incorrectly | incorrect software programming calculations | during procedure | measurement tool incorrect | user confused | 2 |  |  |  |  | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THE INCORRECT VALUE WAS READILY APPARENT TO PHYSICIANS AND THERE IS NO INFORMATION TO SUGGEST THAT A PATIENT WAS HARMED AS A RESULT OF THIS ISSUE. (B)(4). | incorrect measurement value | incorrect measurement | in reporting | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT CLINICAL REPORTS THAT HAD NOT BEEN REVIEWED BY A PHYSICIAN WERE DISPLAYING WITH A "CONFIRMED" STATUS. AN INCORRECT CONFIRMATION STATUS ON A CLINICAL REPORT THAT HAS NOT BEEN REVIEWED BY A PHYSICIAN COULD CAUSE A DELAY IN PATIENT CARE, WHICH COULD POTENTIALLY HARM A PATIENT. HOWEVER, THE CUSTOMER STATED THAT NO PATIENTS WERE ADVERSELY IMPACTED DUE TO THIS ISSUE. (B)(4). | incorrect confirmation status | data issue | in reporting | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT THE HEIGHT MEASUREMENT WAS IMPORTING INCORRECTLY TO THE CLINICAL REPORT CAUSING MISCALCULATION OF THE BSA. DUE TO THE INCORRECT BSA CALCULATION IN THE CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT THAT COULD RESULT IN HARM. THERE IS NO INDICATION OF ANY PATIENT HARM THAT HAS OCCURRED AS A RESULT OF THIS ISSUE. (B)(4). | incorrect data imported | patient data issue | miscalculation | data error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT THE HEADERS ON ECHO AND STRESS ECHO REPORTS CONTAINING UNIQUE IDENTIFYING INFORMATION (PATIENT NAME, PATIENT ID, AND STUDY DATE/TIME) WAS ONLY VIEWABLE ON THE FIRST PAGE OF THE REPORT AND WAS NOT INCLUDED ON ANY SUBSEQUENT PAGES. WHEN MERGE CARDIO DOES NOT PROVIDE THE CORRECT UNIQUE IDENTIFIER IN THE HEADER OF A PATIENT'S REPORT, THERE IS A POTENTIAL FOR PATIENT DATA TO BE COMPROMISED WHICH COULD POTENTIALLY LEAD TO AN INCORRECT TREATMENT OR DIAGNOSIS THAT MAY RESULT IN HARM TO THE PATIENT. INFORMATION OBTAINED REVEALED THAT THE USER WAS AWARE OF THE PROBLEM, NO PATIENT DATA WAS COMPROMISED, AND NO PATIENT HARM OCCURRED. REFERENCE COMPLAINT NUMBER (B)(4). | missing header | data issue | in reporting | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT DATA WAS NOT IMPORTING INTO THEIR STRESS TABLE WITHIN CLINICAL REPORTING, FOR A STUDY PERFORMED ON (B)(6) 2016. (B)(6) 2016 WHICH REVEALED THAT THE INCORRECT TARGET HEART RATE IS BEING USED FOR STRESS ECHO CALCULATIONS. IF THE WRONG TARGET HEART RATE IS USED FOR DIAGNOSIS, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT. (B)(4). | inappropriate heart rate | patient data issue | during procedure | incorrect data imported | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING WITH THE INCORRECT UNIT OF MEASURE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | inappropriate measurement unit | incorrect measurement | during procedure | incorrect unit imported | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHILE READING IMAGES OF STRESS STUDIES THE SPEED INDICATOR DOES NOT MATCH THE ACTUAL SPEED IN FRAMES PER SECOND (FPS) AND ARE INCONSISTENT. DUE TO THE INCONSISTENT FRAMES PER SECOND READING, THERE IS A POTENTIAL FOR A MISDIAGNOSIS OR MIS-TREATMENT FOR A PATIENT. THERE IS NO INDICATION THAT THERE HAS BEEN ANY PATIENT HARM DUE TO THIS INCONSISTENCY. (B)(4). | inconsistent speed | imprecision | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A FETAL PATIENT REPORT WAS AUTOMATICALLY PULLING PRIOR MEASUREMENT DATA FOR A PRIOR FETUS; HOWEVER, THE ATTENDING MEDICAL STAFF IMMEDIATELY NOTICED THE PROBLEM BEFORE THE PATIENTS' DATA WAS COMPROMISED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THAT THE WRONG FETAL MEASUREMENTS GET REFERENCED, WHICH CAN LEAD TO A POTENTIAL FOR OVER OR UNDER DIAGNOSIS AND/OR TREATMENT OF A PATIENT. (B)(4). Manufacturer Narrative: DUE TO THE POTENTIAL FOR MULTIPLE CUSTOMERS TO HAVE THE SAME REPORTED PROBLEM, A RECALL IS BEING INITIATED. | duplicated data | patient data issue | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT MEASUREMENTS WERE NOT BEING POPULATED INTO THE APPLICATION AND THE REPORT. WHILE NOT SPECIFIC IN THIS CASE, THIS MAY IMPACT THE PHYSICIAN'S ABILITY TO COMPLETE THE REPORT AND/OR FINALIZE DIAGNOSIS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | missing data | loss of data | in reporting | data not populated | diagnostic delayed | [5 empty] | FDA2
Event Description: CUSTOMER REPORTED THAT MEASUREMENTS ARE NOT POPULATING IN THE REPORT. THERE WERE 4 IMAGES THAT HAVE MEASUREMENTS ON THEM, HOWEVER THE MEASUREMENTS ARE NOT POPULATING FROM THE IMAGES FROM THE REPORT. TECHNICAL SUPPORT DETERMINED THAT THE SEQUIA MODALITIES SHOULD NOT SEND IN SR FILES. THIS RESULTED IN A LOT OF UNMAPPED MEASUREMENTS THAT WERE PROPRIETORY MEASUREMENTS SENT FROM THE MODALITY. TECHNICAL SUPPORT ASKED FOR A TEST PATIENT IN ORDER TO MAP SOME THE MEASUREMENTS TO MEET THE MODALITY. THE SITE LATER REPORTED THAT THE MODALITY IS NO LONGER BEING USED. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | fail to populate measurement | incorrect measurement | during procedure | unmapped measurement | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT AN UPDATE STATEMENT WAS RUN ON THE REPORTS TABLE WHICH IN TURN UPDATED THE STUDY TAG FOR ALL ROWS TO ONE STUDY TAG. CUSTOMER SUPPORT UPDATED THE DATABASE TO/FROM BACKUP AND CONFIRMED FUNCTIONALITY. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | fail to update report | patient data issue | in reporting | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE DEVICE SOFTWARE IS INCORRECTLY CALCULATING THE MEAN GRADIENT ACROSS THE AORTIC VALVE. WHEN USING THE GE VIVID E9 ECHO CART, THE MEASURED VALUE WAS 38.16. WHEN USING THE DEVICE SOFTWARE, THE MEASURED VALUE WAS 24.93. THIS HAS BEEN VERIFIED WITH OTHER STUDIES. TECHNICAL SUPPORT WORKED WITH THE SITE TO REPRODUCE THE ISSUE. IT WAS FOUND THAT THEY WERE TAKING THE MEASUREMENT ON THE WORKSTATION AND IT WAS CALCULATING CORRECTLY. THE SITE WAS GOING TO FIND OTHER EXAMPLES TO TEST ON BUT NEVER CALLED BACK IN. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | inconsistent calculation | patient data issue | in reporting | calculation from the different workstations | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT REPORTING WOULD CRASH AND RESULT IN INFORMATION FROM REPORTS AND QC TO GO MISSING. THIS IN TURN IMPACTED WORKFLOW AND CAUSED REPEAT DATA ENTRY FOR PHYSICIANS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | report crashed | loss of data | workflow impacted | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE IE33 CARTS ARE SENDING MV VELOCITY MEASUREMENTS WITH A DISCRIMINATOR FOR LATERAL VS MEDIAL, BUT THE DISCRIMINATOR IS DROPPED IN TRANSLATION RESULTING IN DUPLICATE STRINGS FOR DIFFERENT MEASUREMENTS. THIS IN TURN CAUSES THE FOLLOWING: ONLY HALF THE MEASUREMENTS ARE MAPPING; THE VELOCITY MEASUREMENTS ARE NOT MAPPING. WHEN ONLY HALF THE MEASUREMENTS ARE MAPPING, THIS WAS TRACED TO A STREAMING OF THE WORKFLOW, WHICH IS RESOLVED BY SENDING THE STUDY AT THE END. WHEN THE VELOCITY MEASUREMENTS ARE NOT MAPPING, THIS WAS TRACED TO THE WAY THAT THE STRINGS ARE TRANSLATED BY THE APPLICATION, WHICH RESULTS IN DUPLICATE STRINGS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | device dropped | component falling | incorrect measurement value | hardware failure | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT INCORRECT DATA WAS RECEIVED FROM THE INITIAL HEMO DATA EXPORT. THE REPORT WAS CONFIRMED PRIOR TO THE SECOND HEMO EXPORT. IT WAS FURTHER REPORTED THAT THE HEMO DATA WAS LOCKED IN. A RECOMMENDATION TO CREATE A NEW STUDY WAS MADE. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | incorrect data imported | data issue | in reporting | data was locked in | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT AGE, SEX, DOB, AND THE REFERRING MD ARE NOT POPULATING THE REPORT FROM THE ORDER THAT IS BEING SENT TO VERICIS. IT WAS DETERMINED BY CUSTOMER SUPPORT THAT IF TECHNICIANS ARE NOT SENDING IN ADT MESSAGES FROM THE HIS SYSTEM, THEN INFORMATION SUCH AS DOB, SEX, AND AGE WILL NOT BE INCLUDED IN THE REPORT. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | inadequate patient data | patient data issue | in reporting | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE REPORT IS MISSING A PIECE OF INFORMATION THAT IT SAYS SHOULD BE THERE. SPECIFICALLY, THE REPORT SAYS TO LOOK AT 3RD LESION INTERVENTION, BUT IN THAT INFORMATION WAS NOT AVAILABLE TO THE USER. TECHNICAL SUPPORT DISCUSSED THIS WITH THE CUSTOMER AND EXPLAINED WHAT TO DO TO GET THE 3RD INTERVENTION SECTION TO DISPLAY. STEPS (INCLUDING A SCREEN SHOT) WERE PROVIDED ON HOW TO DO THIS (E.G., PROCEDURE NARRATIVE>INTERVENTIONS AND SELECT THE INTERVENTION NUMBER). THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | incomplete report | patient data issue | in reporting | user error | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT THE ECHOIMS DOES NOT SAVE INFORMATION ON A (B)(4) WORKSTATION. THE INFORMATION FROM QUANTIFY AND SUMMARY REPORTING LOCATIONS DO NOT HOLD FROM WORKSTATIONS BUT WILL SAVE FROM THE HOSPITAL'S PC'S WHEN USED IN CONJUNCTION WITH VERICIS WEB. IT WAS FURTHER REPORTED THAT WHEN THE POP-UP BLOCKER WAS TURNED OFF, THE INFORMATION WAS GETTING SAVED IN THE ECHOIMS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | fail to save info on workstation | loss of data | during procedure | blocker | no harm |  |  |  | turn off pop up blocker |  | FDA2
Event Description: IT WAS REPORTED WITH USERS RIGHTS ACTIVATING GROUP NOT PROPAGATING PROPERLY. THIS COULD PREVENT A USER FROM LOGGING INTO THE SYSTEM TO COMPLETE NECESSARY FUNCTIONS WITHIN THE SYSTEM. IT WAS DETERMINED THAT THE CUSTOMER ACCIDENTALLY ADDED TECHNICIANS TO THE REPORT FELLOW GROUP. ONCE REMOVED, IIS AND REPORTING NEEDED TO BE RESTARTED. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | Inadvertently activating button accidentally | self-activation or keying | log into the system | user error | procedure delayed | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT MEASUREMENTS WERE NOT POPULATING THE REPORT. (B)(4) REPORT MONITOR AND (B)(4) EVENT WERE BOTH RESTARTED. THE SR FILE WAS RECEIVED AND WAS REPROCESSED IN (B)(4). THE FILE WAS THEN RESENT FROM B)(4). MEASUREMENTS NOW POPULATED THE REPORT. THIS WAS CONFIRMED WITH THE CUSTOMER. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | fail to populate measurement | loss of data | in reporting | software issue | no harm |  |  |  | reprocessing |  | FDA2
Event Description: IT WAS REPORTED THAT STUDIES DID NOT CONTAIN HEMO DATA. IT WAS FURTHER REPORTED THAT MULTIPLE STUDIES SHOWED THIS PROBLEM. IF THERE IS AN AUTOMATIC ASSOCIATION OF IMAGES AND HEMO DATA, THEN IT SHOWS ON THE WEBVIEWER AND THE WORKSTATION THAT THERE IS HEMO DATA FOR THE STUDY. IF THERE IS A MANUAL ASSOCIATION OF THE TWO, IT SAYS THERE IS NO HEMO DATA FOR THE STUDY. EVEN WHEN THE REPORT IS STARTED, AND EVEN IF IT SAYS THERE IS HEMO DATA, THERE IS NO HEMO DATA IN THE REPORT. IT WAS DETERMINED BY CUSTOMER SUPPORT THAT THERE WEREN'T ANY HEMO TRANSFERS IN PLACE. THE ISSUE WAS RESOLVED AND DATA ARE POPULATING THE REPORT. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | missing data in report | loss of data | in reporting | software issue | no harm | [5 empty] | FDA2
Event Description: IT WAS REPORTED THAT HEMO DATA WAS NOT CROSSING OVER TO SR. CUSTOMER SUPPORT RESOLVED THE ISSUE BY DISASSOCIATING THE HEMO DATA FROM IMAGES FOLLOWED BY REASSOCIATING THEM AGAIN. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | [8 empty] | FDA2
Event Description: IT WAS REPORTED THAT STRUCTURED REPORT IS SHOWING AS QC STATUS IN CARDIO AS CONFIRMED. THIS ISSUE COULD POTENTIALLY RESULT IN A REPORT NOT BEING REVIEWED BY A PHYSICIAN PRIOR TO SENDING TO EMR OR FAXED. CUSTOMER SUPPORT IDENTIFIED CONFLICTING SQL CONFIRMATION STATES AND CORRECTED THE ISSUE. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | incorrect report status | data issue | unavailable report | conflict SQL confirmation state | no harm | [5 empty] | FDA2
Event Description: MERGE RECEIVED COMPLAINT OF REPORTING TIMING OUT AND SENDING USER BACK TO STUDY LIST WHICH CAUSES ALL REPORTING DATA BEING ENTERED TO BE LOST. THIS OCCURS IF THE USER HAS THE REPORTING OPEN FOR MORE THAN AN HOUR, THE SYSTEM WILL CLOSE THE REPORTING AND WILL NOT SAVE THE INFORMATION ENTERED. THE SYSTEM IS DESIGNED TO TIMEOUT AND CLOSE THE REPORTING AFTER AN HOUR OF INACTIVITY. AFTER ONE HOUR, THE SERVICE WILL CLEAN UP ALL RESOURCES OF THE INACTIVE SESSION. THIS RESULTS IN THE REPORT NOT SAVING ANY OF THE INFORMATION ENTERED. THIS WORKFLOW IS MOST COMMON WITH PEDIATRIC CASES AND EP STUDIES. IN PEDIATRIC CASES, THE PHYSICIANS REQUIRE MORE THAN AN HOUR FOR REPORTING DUE TO THE LEVEL OF DETAIL AND RESEARCH REQUIRED FOR SUCH REPORTS. EP STUDIES TYPICALLY LAST MORE THAN 3 HOURS AND REPORTING IS OPEN FOR THE DURATION OF THE STUDIES. (B)(4). | fail to save report | loss of data | in reporting | close report without saving after an hour timeout | procedure delayed | [5 empty] | FDA2
Event Description: WHILE TESTING ON A CUSTOMER'S TEST ENVIRONMENT WITH CARDIO 10.1, MERGE CUSTOM ENGINEERING DEVELOPER FOUND IF THERE ARE TWO PATIENTS WITH THE SAME LAST NAME, SAME FIRST NAME AND MRN BUT OTHER DIFFERENT ATTRIBUTES, SUCH AS IPID, THE PHYSICIAN SEES STUDIES FROM BOTH PATIENTS IN STUDY COMPARISON VIEW. IN THE STUDY LIST IF TWO PATIENTS HAVE THE SAME FIRST NAME, LAST NAME AND MRN BUT DIFFERENT IPIDS THE SYSTEM TREATS ALL STUDIES AS BEING FROM THE SAME PATIENT. NOT KNOWING THAT THE STUDIES AVAILABLE ARE NOT ALL FOR THE SAME PATIENT WHEN COMPARING THEM IN STUDY COMPARISON MODE, THE PHYSICIAN CAN INTERPRET THE WRONG STUDIES WHICH MAY LEAD TO MISDIAGNOSIS. NOTE THAT THIS ONLY OCCURS WHEN TWO PATIENTS HAVE THE SAME FIRST NAME, LAST NAME AND MRN WHICH IS VERY RARE. (B)(4). | same patient | patient data issue | during procedure | software issue | misdiagnosis or mistreatment | [5 empty] | FDA2
Event Description: MERGE RECEIVED COMPLAINT (B)(4) THAT THE REPORTING FUNCTIONALITY TIMES OUT AND SENDS THE USER BACK TO THE STUDY LIST WHICH CAUSES ALL REPORTING DATA THAT WAS ENTERED TO BE LOST. THIS OCCURS IF THE USER HAS THE REPORTING FUNCTION OPEN FOR MORE THAN AN HOUR, THE SYSTEM WILL TIME OUT AND CLOSE THE REPORTING FUNCTION AND NOT SAVE THE INFORMATION ENTERED. THIS SCENARIO CAN POTENTIALLY LEAD TO DELAY IN CARE. | fail to save report | loss of data | in reporting | close report without saving after an hour timeout | procedure delayed | [5 empty] | FDA2
Event Description: ON 5/5/15, A CUSTOMER REPORTED ((B)(4)) THAT TWO PHYSICIANS FROM ONE CENTER WERE WORKING ON ONE CASE AT THE SAME TIME, AND THEY WERE NOT ALERTED THAT THE CASE WAS OPENED BY THE OTHER PHYSICIAN. THIS PROBLEM WAS IDENTIFIED AS A POTENTIAL SAFETY ISSUE IN THE EVENT CHANGES TO THE SYSTEM AFFECT THE DIAGNOSIS OR TREATMENT OF A PATIENT. NO KNOWN PATIENTS WERE HARMED DUE TO THIS COMPLAINT. | multiple accesses at once time | application security issue | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED ((B)(4)) THAT THEY WENT INTO A STUDY TO VIEW THE IMAGES AND ACCIDENTALLY CLICKED ON THE MULTI STUDY REVIEW BUTTON INSTEAD OF THE REPORTING ICON. WHEN THE CUSTOMER CLICKED BACK ON IMAGE REVIEW, THE APPLICATION HAD STOPPED CACHING THE STUDY WHICH MADE IT LOOK LIKE THE STUDY ONLY HAD A FEW IMAGES. THERE IS THE POTENTIAL SAFETY RISK THAT THE USER BEGINS TO REVIEW A STUDY THAT IS NOT DISPLAYING ALL OF THE AVAILABLE IMAGES. | Inadvertently activating button accidentally | use of device issue | during procedure | user error | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: A MERGE CARDIO PRODUCT PROBLEM WAS IDENTIFIED INTERNALLY. THE PROBLEM COULD LEAD TO INCORRECT ACCESSION NUMBERS ON REPORTS. IF IMAGES ARE SENT WITHOUT AN ORDER, THE IMAGES WILL BE MATCHED WITH THE ORDER BASED ON THE CURRENT PATIENT/MODALITY MATCHING. THE REPORT THEN REFLECTS THE INCORRECT ACCESSION NUMBER FOR THE ORDER AND DOES NOT FILE PROPERLY IN THE EMR. AN INCORRECTLY FILED REPORT IN THE EMR COULD LEAD TO DELAY IN CARE. (B)(4). | incorrect accession number | application security issue | in reporting | image sends withou an order | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: A USER AT A SITE USED THIS WORKFLOW, BUT ACCIDENTALLY CHOSE THE WRONG IMAGE-STUDY WHEN MERGING THE HEMO STUDY. USERS SHOULD NOT BE ABLE TO ADD IMAGES TO A CONFIRMED REPORT VIA MERGE. IT WAS FOUND THAT 9.X USERS CAN MERGE DEVICE IMPORT DATA WITH AN IMAGE-STUDY THAT ALREADY HAS A CONFIRMED REPORT. THIS RESULTS IN ADDING DEVICE IMPORT DATA TO AN IMAGE STUDY WITH A CONFIRMED REPORT THEREFORE IT MAY INTRODUCE INFORMATION THAT WAS NOT AVAILABLE DURING THE INTERPRETATION. (B)(4). | wrong image selected | use of device issue | during procedure | user error | procedure delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS (MR VENA CONTRACTA AND TR VENA CONTRACTA) THAT WERE NEVER TAKEN ON THE MODALITY WERE POPULATING THE CARDIO CLINICAL REPORT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | fail to populate measurement | loss of data | during procedure | incorrect value | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2019, A CUSTOMER ALLEGED THAT CLINICAL REPORTING WAS LOGGING OFF WITHOUT SAVING PREVIOUSLY ENTERED REPORT DATA. THIS COULD LEAD TO A DELAY IN PATIENT CARE WHILE REDOING THE CLINICAL REPORT. DELAY IN TREATMENT HAS A POTENTIAL TO LEAD TO HARM. HOWEVER, THERE IS NO ALLEGATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | log off report without saving | loss of data | in reporting | timeout | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2019 A CUSTOMER CALLED MERGE CARDIO SUPPORT STATING THERE WAS A REPORT WAS IN CONFIRMED STATUS WHEN IT WAS ACTUALLY ONLY QC STATUS. QC REPORTS WITH THE STATUS OF CONFIRMED COULD LEAD TO DELAY OF TREATMENT AS THE REPORT HAS NOT ACTUALLY BEEN READ BY A PHYSICIAN. THE INCORRECT STATUS OF THE REPORT MAY NOT BE READILY APPARENT TO A USER. THERE HAS BEEN NO ALLEGATION OF AN INJURY OR IMPACT TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT-(B)(4). | incorrect report status | data issue | in reporting | software issue | no harm | [5 empty] | FDA2
Event Description: ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS INCORRECT AND NEEDED TO BE DISABLED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT (B)(4). | incorrect measurement value | patient data issue | in reporting | modality issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(4) 2018, IT WAS FOUND INTERNALLY WHILE PATCHING THE CUSTOMERS 3.5 REPORTING KNOWLEDGE BASE TO 5.0 REPORTING KNOWLEDGE BASE THAT SOME OF THE MICHIGAN Z-SCORES WERE CALCULATING INCORRECTLY. THIS IS DUE TO THE FORMULA USING 'B' VS 'BSA' FOR BODY SURFACE AREA. THIS VARIABLE HAS BEEN CAUSING Z-SCORES TO CALCULATE INCORRECTLY. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER AT THIS TIME, THERE HAVE BEEN NO ALLEGATIONS OF PATIENT HARM AS A RESULT OF THIS ISSUE. (B)(4). | incorrect z-score value | patient data issue | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2018 A CUSTOMER REPORTED THAT THE IMAGES OF A STUDY DID NOT MATCH THE NAME OF THE REPORT FOR THAT STUDY. MERGE HEALTHCARE SUPPORT IS WORKING WITH THE CUSTOMER FOR MORE EXAMPLES OF THIS HAPPENING AS THEY ARE UNABLE TO INVESTIGATE THIS ISSUE AS THERE WAS NOT ENOUGH DATA LEFT FOR THEM TO KNOW WHAT HAPPENED DURING THIS INCIDENT. REFERENCE CASE (B)(4). | incorrect image name | patient data issue | during procedure | unkown | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, THE CUSTOMER ALLEGED THE REPORT WAS NOT GOING TO THEIR EMR. UPON INVESTIGATION, IT WAS FOUND THAT THE REASON THE REPORT NEVER EXPORTED TO THE CUSTOMER'S EMR, EPIC. THIS WAS DUE TO IT BEING INCORRECTLY IN CONFIRMED STATUS INSTEAD OF BEING IN QC STATUS AND ACTUALLY CONFIRMED. MERGE CARDIO SUPPORT CORRECTED THE ISSUE BY CHANGING THE STATUS IN THE DATABASE QC REPORTS WITH THE STATUS OF CONFIRMED COULD LEAD TO DELAY OF TREATMENT AS IT HAS NOT ACTUALLY BEEN READ BY A PHYSICIAN AND AS A RESULT MAY NOT BE READILY APPARENT TO A CLINICIAN. HOWEVER, THERE HAVE BEEN NO ALLEGATIONS OF HARM AS A RESULT OF THIS ISSUE. (B)(4). | fail to export report | data issue | in reporting | incorrect report status | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT WRONG RA VOLUME WAS BROUGHT FORWARD. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, THERE HAVE BEEN NO REPORTS OF HARM OR INJURY AS A RESULT OF THIS ISSUE. (B)(4). | incorrect volume calculation | parameter calculation error due to software problem | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, IT WAS FOUND THAT SECONDARY CAPTURES IN TOMTEC WERE SAVING UNDER THE WRONG PATIENT IN MERGE CARDIO. IMAGES SAVING UNDER THE WRONG PATIENT TO LEAD TO PATIENT HARM BY INCORRECT OR MISTREATMENT OF THE PATIENT. (B)(4). | image saved incorrectly | patient data issue | during procedure | user error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON 04/26/2018, A MERGE CARDIO CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT LEFT VENTRICULAR OUTFLOW TRACT (LVOT) MEASUREMENTS WERE INCORRECT. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION. DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY WERE NOT APPEARING IN MERGE CARDIO WITH THE SAME MEASUREMENT VALUE. THE CUSTOMER HAS NOT INDICATED THAT ANY PATIENT WAS IMPACTED PR HARMED AS A RESULT OF THIS ISSUE. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | incorrect measurement value | incorrect measurement | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2018 , A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE HEIGHT MEASUREMENT WAS IMPORTING INCORRECTLY TO THE MERGE CARDIO CLINICAL REPORT CAUSING MISCALCULATION OF OTHER MEASUREMENTS. MERGE HEALTHCARE TECHNICAL SUPPORT'S INITIAL INVESTIGATION FINDINGS APPEAR TO BE RELATED TO SOME NEW ECHOCARDIOGRAM IMAGING CARTS AT THE CUSTOMER'S FACILITY. IT IS POSSIBLE THAT THE NEW ECHO CARTS MAY BE SENDING DATA TO MERGE CARDIO PACS IN AN UNEXPECTED FORMAT THEREFORE RESULTING IN INCORRECT HEIGHT MEASUREMENTS. DUE TO THE MISCALCULATION OF MEASUREMENTS IN THE CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT THAT COULD RESULT IN HARM. HOWEVER, THE CUSTOMER HAS NOT ALLEGED ANY IMPACT OR INJURY TO A PATIENT. (B)(4). | incorrect data imported | patient data issue | miscalculation | unexpected measurement format | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MEASUREMENT CALCULATIONS ARE NOT APPEARING PROPERLY IN THE DIAGNOSTIC REPORT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | data calculated incorrectly | incorrect software programming calculations | in reporting | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT ON (B)(6) 2017 THE MERGE CARDIO OUTBOUND HL7 INTERFACE HAD STOPPED SENDING CLINICAL REPORTS ELECTRONICALLY TO THEIR HOSPITAL INTERFACE SYSTEM FOR APPROXIMATELY 18 HOURS. DURING THIS DOWNTIME, A PATIENT WAS DISCHARGED FROM THE FACILITY WITHOUT A FULL REVIEW OF THE PATIENT'S CLINICAL REPORT. THE CUSTOMER STATED THAT IT WAS NECESSARY FOR THE PATIENT TO RETURN TO THE FACILITY FOR FURTHER TESTING. THE MERGE CARDIO SOFTWARE PROVIDES USERS WITH THE ABILITY TO REVIEW AND PRINT CLINICAL REPORTS WHEN THEY ARE UNABLE TO BE DELIVERED TO THE CUSTOMER'S HOSPITAL INTERFACE SYSTEM. ADDITIONALLY, THE SYSTEM IS DESIGNED TO NOTIFY THE USER IF THE DELIVERY OF CLINICAL REPORTS FAILS. THE NOTIFICATION ALLOWS THE USER TO BE MADE AWARE OF REPORT COMMUNICATION ISSUES IN ORDER FOR OTHER METHODS OF COMMUNICATION TO BE USED (PRINT, SCAN, FAX, ETC.). WITH A PATIENT BEING DISCHARGED BEFORE STAFF COMPLETED A FULL REVIEW OF THAT PATIENT'S CLINICAL REPORT, THERE IS A POTENTIAL FOR A DELAY OR INCORRECT TREATMENT THAT COULD RESULT IN HARM TO THE PATIENT. HOWEVER, THE CUSTOMER REPORTED THAT NO HARM HAD OCCURRED. (B)(4). | fail to send report | Communication or Transmission Issue | review patient report delayed | system downtime | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT AN ISSUE OCCURRED THAT COULD LEAD TO INCORRECT ACCESSION NUMBERS ON CLINICAL REPORTS. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT WITH THEIR CLINICAL WORKFLOW, IF IMAGES ARE SENT WITHOUT AN ORDER, THE IMAGES WILL BE MATCHED WITH THE PATIENT'S ORDER BASED ON THE CURRENT PATIENT/MODALITY MATCHING. THE CLINICAL REPORT THEN REFLECTS THE INCORRECT ACCESSION NUMBER FOR THE ORDER AND DOES NOT FILE PROPERLY IN THE CUSTOMER'S EMR. WHILE THE ACCESSION NUMBER MAY NOT BE THE EXPECTED NUMBER THE PATIENT'S STUDY INFORMATION IS CORRECTLY PLACED WITHIN THE PATIENT'S RECORD ON THE EMR. AN INCORRECTLY FILED REPORT IN THE EMR COULD LEAD TO A DELAY IN CARE. HOWEVER, THE CUSTOMER STATED THAT THERE WAS THERE WAS NO HARM TO THE PATIENT. (B)(4). | incorrect accession number | application security issue | incorrect filed report | images are sent without an order | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER CALLED INTO MERGE HEALTHCARE AND ALLEGED THAT THERE WAS INCORRECT PROCEDURAL VERBIAGE IN A MERGE CARDIO REPORT. INACCURATE CATHERIZATION PROCEDURE DETAILS IN A REPORT COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS OR TREATMENT. MERGE HEALTHCARE IS INVESTIGATING THIS CUSTOMER'S ALLEGATION. THERE IS NO INDICATION FROM THE CUSTOMER THAT THERE WAS ANY CONFIRMED MIS-DIAGNOSIS OR MIS-TREATMENT OF A PATIENT DUE TO INACCURATE VERBIAGE IN A PATIENT'S CARDIOLOGY REPORT. REFERENCE (B)(4). | incorrect procedural verbiage in report | data issue | in reporting | doctor error | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A PATIENT REPORT WAS CONFIRMED (SIGNED) WITH THE INCORRECT PATIENT DATA (SEE ORIGINAL REPORT 2183926-2016-00441). BASED ON THIS ALLEGATION FROM THAT CUSTOMER, MERGE HEALTHCARE INVESTIGATED THE ISSUE AND CONFIRMED THAT THERE WAS A MALFUNCTION OF THE MEDICAL DEVICE. THIS MALFUNCTION IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. MERGE HEALTHCARE WAS ABLE TO QUERY POTENTIALLY AFFECTED CUSTOMER'S DATABASES ON (B)(4) 2016 TO IDENTIFY IF THERE WERE ANY SUSPECT REPORTS STORED THERE. THE QUERY DETECTED A SUSPECT REPORT FOR THIS CUSTOMER AND THEY HAVE BEEN MADE AWARE OF THE ISSUE. (B)(4). | incorrect patient data | patient data issue | during procedure | malfunction | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE ALLEGING THAT TWO (2) STUDIES WERE MISSING AND COULD NOT BE PULLED FROM THE ARCHIVE. MERGE HEALTHCARE TECHNICAL SUPPORT IS INVESTIGATING THE CUSTOMER'S ALLEGATION. AN ANALYSIS OF THE POTENTIAL SOURCES AND STORAGE OF THE STUDIES AND ANY OTHER POTENTIALLY RELATED STUDIES IS ONGOING. DUE TO STUDIES NOT BEING ABLE TO BE FOUND AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN PATIENT TREATMENT OR DIAGNOSIS OR PATIENT HARM. HOWEVER, THERE WERE NO REPORTS OF DELAYS IN PATIENT TREATMENT OR DIAGNOSIS. AT THIS TIME THERE ARE NO REPORTS OF PATIENT HARM. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ISSUE TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. DOES NOT APPLY TO SOFTWARE. | -fail to store medical imaging -loss of image -inadequate image manipulation | -use of device issue -therapeutic or diagnostic output failure -pacing inadequately | [8 empty] | FDA2
Event Description: A PROBLEM WAS OBSERVED IN-HOUSE, WHERE ORIENTATION LABELS ARE NOT READ FROM DICOM FILE AND A DEFAULT ORIENTATION IS ASSUMED IN PLUG IN MODE. WHEN ORTHOCASE IS LAUNCHED IN PLUG IN MODE, ORIENTATION LABELS DISPLAYED IN THE VIEWER MAY BE INCORRECT (I.E. SHOW A DEFAULT VALUE). SAVED DATA (E.G. SECONDARY CAPTURE IMAGES) WILL HAVE THE INCORRECT LABELS BURNED IN. THIS ISSUE DOES NOT OCCUR IN STAND-ALONE MODE. THERE IS NO WORK AROUND. THE ORIENTATION LABELS CANNOT BE DISABLED THROUGH CONFIGURATION. THE ISSUE WAS ENCOUNTERED DURING INTEGRATION TESTING WITH ANOTHER PRODUCT. IN MOST CASES, SURGEONS AND PHYSICIANS WOULD USE OF LEAD MARKERS ON THE IMAGES. THE MARKERS WOULD IDENTIFY ORIENTATION OF THE IMAGE AND ANATOMY. | incorrect orientation label | device markings issue | in testing | configuration | no harm | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED INCORRECT MEASUREMENTS FROM THE TEMPLATE IN MERGE ORTHOCASE. THE CUSTOMER USED THE TEMPLATE TO MEASURE AN AREA FOR SURGERY AND WHAT SIZE IMPLANTS WILL BE USED. DURING THE SURGERY THEY FOUND THAT AREA THEY MEASURED ON THE IMAGE WAS OFF 4-6 MM. THIS EVENT WAS HIGHLY DETECTABLE BY THE SURGEON AND HE WAS ABLE TO RE-MEASURE AND USE CORRECT SIZE IMPLANT. THIS ISSUE HAS A POTENTIAL FOR USE OF AN INCORRECTLY SIZED IMPLANT DURING ORTHOPEDIC SURGERY. THE CUSTOMER CONFIRMED THE SURGERY WAS SUCCESSFUL AND THERE WAS NO HARM TO THE PATIENT. (B)(4). Manufacturer Narrative: AFTER FURTHER INVESTIGATION, IT WAS FOUND THE CUSTOMER WAS USING DIFFERENT SIZE BALLS WHICH CAUSED THE TEMPLATE TO NOT PRODUCE CORRECT MEASUREMENTS. NO ADDITIONAL CORRECTIVE ACTIONS ARE NEEDED. | incorrect measurement value | incorrect measurement | incorrect implant size selected | use selects incorrect size balls | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: (B)(6) HAS REPORTED THAT MAMMO IMAGES FROM (B)(4) ARE OF POOR QUALITY. IN (B)(6) IT WAS A DETERMINED THAT (B)(4) NEEDED TO SUPPLY AN UPDATE TO THE FIRMWARE AND THE CASE WAS CLOSED. (B)(4) HAS TRIED FIX THE ISSUE AND HAS EXHAUSTED ALL OPTIONS AND THAT THE RADIOLOGIST IS REFUSING TO READ MAMMO STUDIES STATING THAT THE IMAGE QUALITY IS POOR. THIS IS ISOLATED TO ONE (B)(4) MODALITY AT THE CUSTOMERS FACILITY. MUCH RESEARCH HAS BEEN DONE BY DEVELOPMENT, SUPPORT AND SOLUTIONS AND AS OF 11/1/15 THERE IS NO ONE ROOT CAUSE. (B)(4). | poor quality image | poor quality image | during procedure | firmware issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED (CASE (B)(4)) THAT UNVIEWED IMAGES WERE NOT DISPLAYING AFTER THE LAST HANGING PROTOCOL STEP. WHEN THE USER PREFERENCE "SHOW UNVIEWED IMAGES AFTER THE LAST HANGING PROTOCOL STEP" IS SELECTED, THE CUSTOMER EXPECTS THAT ANY IMAGES NOT MATCHED TO THE HANGING PROTOCOL WILL DISPLAY IN AN AUTO-GENERATED FINAL VIEW. HOWEVER, IF THE LAST DEFINED VIEW OF THE PROTOCOL WAS SKIPPED (BECAUSE NO IMAGES MATCHED), THE UNVIEWED IMAGES ARE NOT DISPLAYED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE OF THE RISK ASSOCIATED WITH NOT VIEWING ALL OF THE IMAGES IN THE STUDY. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | fail to display unviewed image | image display error | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: A CUSTOMER REPORTED (CASE (B)(4)) THE CUT LINES ARE NOT SHOWING CORRECTLY ON A STUDY IN PACS. THE CUT LINES WOULD GO FROM BEING VERTICAL AND THEN SHIFT ALMOST HORIZONTAL. THE DISPLAY IS VERY OBVIOUS TO THE USER AND SHOWS CONTRADICTORY VIEWS WHICH WOULD CAUSE THE USER TO QUESTION WHAT IS HAPPENING. THE IMAGES CAME FROM SIEMENS (B)(4) SCANNER. AN INVESTIGATION FOUND THAT SIEMENS CHANGED THE RESOLUTION OF THE VALUES OF SOME DICOM ATTRIBUTES CAUSING AN EDGE CASE WHICH MERGE PACS DID NOT HANDLE APPROPRIATELY. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE FOR THE PATIENT IN THE EVENT THERE WAS A MISDIAGNOSIS FROM VIEWING THE INCORRECT MEASUREMENTS OF A PATIENT'S IMAGES. | contradictory view | incorrect display | user confused | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO A 6.5.9 SYSTEM VIA TELMED. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. (B)(4). | fail to transfer image | communication or transmission issue | in reporting | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2013 A CUSTOMER REPORTED THAT A STUDY STATUS HAS BEEN CHANGED INCORRECTLY. THE LOGS REPORTED THAT THE STUDY WAS CHANGED BY A STAFF MEMBER; HOWEVER THE STAFF MEMBER DID NOT HAVE SYSTEM PERMISSIONS TO DO SO. MERGE HEALTHCARE IDENTIFIED THE ROOT CAUSE TO BE A DESIGN FLAW WHERE ANY USER WITH THE CHANGE WORKFLOW STATUS BUTTON CAN PDE ANY STUDY, EVEN IF THE USER DOES NOT HAVE PERMISSIONS TO USE THE PDE TOOL. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | incorrect study status change | application security issue | during procedure | design flow | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2015 A CUSTOMER REPORTED THAT SOME PATIENT PRIORS INCLUDED INCORRECT INFORMATION AFTER A DATA MIGRATION. THE CUSTOMER WAS RECEIVING A PRINT OUT FROM THE RIS SYSTEM AND FOUND THAT MRN NUMBERS WERE MIGRATED OVER WITH THE OLD NUMBERS. THIS RESULTED IN EXAMS IN INCORRECT FOLDERS. MERGE PACS QUALIFIES STUDIES BY MRN WHEN PROCESSING A08 (PATIENT UPDATE) AND A34 (PATIENT MERGE) REQUESTS FROM A HIS/RIS. BY DEFAULT, IF AN A08 OR A34 REQUEST IS RECEIVED THAT HAS AN MRN THAT MATCHES A PATIENT BUT HAS THE WRONG PATIENT NAME, THE REQUEST WILL BE REJECTED BY THE MERGE PACS. MERGE PACS CAN BE CONFIGURED, HOWEVER, TO NOT CHECK TO SEE IF THE PATIENT NAME STORED IN ITS DATABASE MATCHES THAT SENT BY THE HIS/RIS AND WILL INSTEAD RETURN ALL STUDIES THAT HAVE AN MRN THAT MATCHES THAT SENT BY THE HIS/RIS (A PROCESS KNOWN AS "GLOBAL UPDATE"). MERGE WAS ABLE TO PROVIDE THE NECESSARY DATA SO THAT THE CUSTOMER COULD DO A COMPARISON AND IDENTIFY DUPLICATES TO MAKE SURE THE EXAMS WERE IN THE CORRECT FOLDER. THERE WAS NO REPORT OF PATIENT INJURY. REFERENCE: (B)(4). | data migrated | data issue | exam in incorrect folder | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO THE MERGE PACS SYSTEM VIA (B)(4) TRANSMITTER. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. IN THE MEANTIME, A WORKAROUND WAS PROVIDED TO THE SITE TO ONLY SEND LOSSLESS AND NOT ALLOW LOSSY, SO THAT THERE IS A VALID FAILURE WHEN LOSSLESS ISN'T PRESENT, TO RETRIEVE FROM ARCHIVE AND RESEND. THE FAILURES WOULD NO LONGER BE SEEN WITH THIS CONFIGURATION. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO A PATIENT. (B)(4). | incorrect transmission | Communication or Transmission Issue | send old studies | configuration issue | procedure delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER CALLED INTO MERGE SUPPORT AND ALLEGED THAT THEIR VIEWER WAS NOT RESPONDING. ON (B)(6) 2016 THE SAME CLIENT CALLED MERGE SUPPORT ALLEGING THAT THEIR VIEWER WAS CRASHING WHEN A WORKLIST WAS OPENED. BASED ON THE CUSTOMER'S ALLEGATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT DUE TO THE VIEWER NOT FUNCTIONING AS EXPECTED. THE CUSTOMER DID NOT INDICATE ANY PATIENT HARM AS A RESULT OF THE ISSUES WITH THE VIEWER. (B)(4). | viewer no responding | application program issue | during procedure | software issue | procedure delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER CONTACTED SUPPORT ALLEGING THAT THERE WERE MISSING IMAGES WITHIN STUDIES FOR A NUMBER OF PATIENTS. MERGE HEALTHCARE BEGAN INVESTIGATING THE CUSTOMER'S ALLEGATION. IT WAS DETERMINED THAT ALL REPORTS THAT WERE REPORTED TO HAVE MISSING IMAGES HAD BEEN PREVIOUSLY READ AND REPORTED ON. ADDITIONALLY IT WAS DETERMINE THAT THE CUSTOMER HAD A PREVIOUS PRODUCTION DATABASE ISSUE AND HAD ELECTED TO MOVE DATA TO A CONTINGENCY SERVER. THE CLIENT HAD NOT BEEN PERFORMING RECOMMENDED DATABASE AND SERVER MAINTENANCE INCLUDING BACK UPS OF THE SYSTEM. INCONSISTENT AND IRREGULAR BACK-UPS OF CLINICAL DATA INCREASES THE LIKELIHOOD OF DATA LOSS. IT WAS DETERMINED THAT THE MISSING IMAGES WERE DUE TO THE CONTINGENCY SERVER BEING SET TO DO PERIODIC PURGES AS IT IS USED AS A TEMPORARY SOLUTION AND NOT INTENDED TO SUSTAIN LONG TERM USE. ONCE IT WAS DETERMINED THAT CONTINGENCY SERVER PURGING STUDIES, PURGE WAS DISABLED. DUE TO THE LOSS OF IMPAGES THAT CAN BE USED WHEN COMPARING PRIOR STUDIES TO RECENT STUDIES, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER REPORTED THAT THEY WERE NOT AWARE OF ANY INSTANCE WHERE A PATIENT WAS HARMED OR TREATMENT DELAYED. (B)(4). | missing image series | loss of data | during procedure | database and server maintenance issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2017 A CLIENT REPORTED THAT WHEN A RADIOLOGIST BEGINS VIEWING STUDY A IN THE PACS PRIMARY VIEWER AND THEN BEGINS VIEWING STUDY B IN THE SECONDARY VIEWER, IF STUDY A IS DICTATED AND MARKED AS READ, STUDY B IS SET TO A READ STATUS AS WELL. THE CUSTOMER IS EXPECTING THAT ONLY THE STUDY A IS MARKED AS READ, NOT BOTH OF THE STUDIES THAT THE CUSTOMER IS VIEWING (STUDY A AND STUDY B). WHEN A STUDY THAT HAS NOT BEEN DICTATED AND READ IS AUTOMATICALLY CHANGED TO A READ STATUS BY THE SYSTEM, THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS. THIS IS DUE TO THE CUSTOMER POTENTIALLY MISSING THE STUDY AS IT WILL NO LONGER APPEAR ON THEIR PENDING REVIEW WORKLIST. AT THIS TIME, THERE IS NO INDICATION OF DIRECT IMPACT OR HARM TO A PATIENT. (B)(4). | incorrect read status | device operates differently than expected | during procedure | software issue | user inconvience | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, A CUSTOMER REPORTED THEY WERE HAVING ISSUES WITH THE UNVERIFIED-NEW IMAGE STATUS FOR PATIENT REPORTS. MERGE HEALTHCARE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORT STATUS WAS NOT CHANGING TO "UNVERIFIED NEW IMAGES" TO ALERT THE USER THAT ADDITIONAL IMAGES HAD BEEN RECEIVED AND ADDED/APPENDED TO THE PATIENT'S REPORT. THIS ISSUE WAS DETECTABLE TO THE USER AS A STUDY MARKED AS FINAL DID NOT CHANGE STATUS WHEN NEW IMAGES WERE RECEIVED. CURRENTLY, THERE IS NO KNOWN IMPACT TO SPECIFIC OR OVERALL PATIENTS; HOWEVER, THERE IS POTENTIAL FOR A SITUATION TO OCCUR WHERE NEW IMAGES ARE NOT READ RESULTING IN DELAYED OR INCORRECT TREATMENT. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect image status | patient data issue | new images are not read | configuration issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2017 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A STAT EXAM HAD NOT BEEN READ IN A TIMELY MANNER. THE CUSTOMER REQUESTED THAT MERGE HEALTHCARE TECHNICAL SUPPORT PERFORM AN ANALYSIS OF THE ACTIVITIES FOR THE PATIENT'S STUDY. AN ANALYSIS OF THE FILE'S LOGS REVEALED MULTIPLE USERS ACCESSING THE FILE WITHOUT VALIDATING THE STUDY. ADDITIONAL ANALYSIS REVEALED THAT AFTER THE MULTIPLE USERS HAD ACCESSED THE NON-VALIDATED STUDY, A USER DID MANUALLY VALIDATE THE RECORD DURING A SCHEDULED DOWNTIME WHEN THE CUSTOMER'S IT PERSONNEL WERE PERFORMING MAINTENANCE ON THE SYSTEM. PERFORMING A STUDY VALIDATION DURING SYSTEM/SERVER DOWNTIME PREVENTS THE VALIDATION COMMAND FROM GOING THROUGH AND THE FILE IS PUT INTO A "FAILED" STATUS. THE LOGS INDICATED THAT THE NEXT ACTION WAS A USER MARKING THE STUDY AS "FINAL" WITHOUT VALIDATION BEING PERFORMED. FINALLY ON (B)(6) 2017, THE LOGS SHOWED A USER MANUALLY RE-OPENING THE STUDY, READING AND VALIDATING THE REPORT. MERGE HEATLHCARE PERSONNEL REVIEWED THE INVESTIGATION INFORMATION WITH THE CUSTOMER. IT IS THE RESPONSIBILITY OF THE SITE TO REVIEW AND MONITOR STATUS QUEUES FOR ANY FILES THAT HAVE FAILED TO PROCESS. THIS IS AN ACTIVITY THAT IS COMMON TO PACS DEVICES. WHILE THE SITE IS FAIRLY NEW TO MERGE PACS (OCTOBER 2016), IN PERSON TRAINING SESSIONS AND MERGE PACS DEVICE LABELING PROVIDE INFORMATION INSTRUCTING USERS TO MANAGE FILE QUEUES. THE CUSTOMER HAS CONFIRMED THAT NO SERIOUS INJURY OR DEATH RESULTED DUE TO THE DELAY IN VALIDATING THE STAT EXAM HOWEVER, THERE IS A POTENTIAL FOR INJURY SHOULD THIS ISSUE REOCCUR. MERGE HEALTHCARE WORKED WITH THE CUSTOMER TO ENHANCE THE SITE'S ABILITY TO MONITOR REPORT QUEUES. (B)(4). | fail to validate study | patient data issue | during procedure | perform a study validation during system/server downtime or maintenance | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHEN A USER EXITS OUT OF STUDY IN FLUENCY, MERGE PACS CHANGES THE STATUS OF THE STUDY TO "READ." WHEN A STUDY UNEXPECTEDLY GOES TO A READ STATUS WHEN EXITING OUT OF THAT STUDY, THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS DUE TO A CUSTOMER MISSING READING THE STUDY. HOWEVER, THIS ISSUE WAS DETECTED BY THE CUSTOMER AND THERE WAS NO REPORTED IMPACT TO PATIENT CARE. (B)(4). | incorrect report status | data issue | change status to read | unexpected exit out of study | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 THE CUSTOMER REPORTED THAT THEY WERE EXPERIENCING AN ISSUE WITH MERGING PATIENT DEMOGRAPHICS. IT WAS FOUND THAT THE GLOBAL PATIENT UPDATE WAS TURNED ON DURING IMPLEMENTATION AT THE CUSTOMER SITE AND WOULD UPDATE OR MERGE ALL PATIENTS ASSOCIATED WITH A SINGLE MRN/IPID. WHICH RESULTED IN PATIENT INFORMATION TO BE UPDATED TO THE INCORRECT PATIENT IN THEIR SYSTEM WHEN DUPLICATE MRN NUMBERS OCCUR. THIS ISSUE HAS A POTENTIAL TO LEAD TO PATIENT INFORMATION BEING COMPROMISED OR OVERWRITTEN BY A NEW STUDY SUBMISSION. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED AS A RESULT OF THIS ISSUE. ADDITIONALLY, IF GLOBAL PATIENT UPDATE WAS SET TO OFF, PATIENTS WITH THE SAME MRN NUMBERS WOULD NOT BE MERGED. REFERENCE COMPLAINT NUMBER (B)(4). | undesired patient data changes | patient data issue | patient information overwritten | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2016, A CUSTOMER REPORTED STUDIES MARKED AS "READ" OR "FINAL" WERE NOT CHANGING BACK TO "UNREAD" OR "UNVERIFIED" WHEN A NEW DOCUMENT IS ADDED TO THE STUDY. THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS OR TREATMENT DUE TO THE CUSTOMER BEING UNAWARE OF DOCUMENT(S) ADDED TO A STUDY THAT HAS ALREADY BEEN READ. INFORMATION RECEIVED FROM THE CUSTOMER CONFIRMED THERE WAS NO IMPACT TO PATIENT CARE DUE TO THIS ISSUE. (B)(4). | incorrect study status | application security issue | customer unware of document added | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT AN INCIDENT HAD OCCURRED WHERE THE THIRD IMAGE ON A REPORT WAS NOT REVIEWED BY THE RADIOLOGIST. THE THIRD IMAGE SHOWED A FRACTURE THAT INITIALLY WENT UNDETECTED DUE TO NOT BEING REVIEWED. THIS RESULTED IN A DELAY IN DIAGNOSIS AND/OR TREATMENT FOR THE PATIENT. CURRENTLY THERE IS NO INDICATION OF AN ADVERSE EVENT TO THE PATIENT. (B)(4). | undetected image | data issue | not reviewed by radiologist | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 A CUSTOMER REPORTED A PATIENT WAS SENT HOME WITH A CERVICAL FRACTURE BECAUSE IMAGES WERE NOT READ. THE CUSTOMER FOUND THE IMAGES MISSED REVIEW AND CONTACTED THE PATIENT. THE PATIENT WAS ASKED TO GO TO THE CUSTOMER'S MAIN HOSPITAL FOR AN MRI. AFTER THE MRI WAS PERFORMED, THE CUSTOMER HAD SUCCESSFUL SURGERY. THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH CAN LEAD TO PATIENT HARM AS A RESULT OF CLINICAL STAFF NOT REVIEWING ALL IMAGES. IT WAS DETERMINED THAT THE PATIENT RECEIVED AN MRI SCAN AND SURGERY PRIOR TO ANY ADDITIONAL HARM BEYOND THE ORIGINAL CERVICAL FRACTURE. (B)(4). | image missed review | data issue | image not read | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON 04/15/2016 ORTHO (B)(4) REPORTED THAT EXISTING INFORMATION WAS GETTING OVERWRITTEN SINCE EMR INTERFACE WAS ACTIVATED. ORTHO (B)(4) WAS TO CONDUCT A CONVERSION OF THEIR MRN NUMBERS BEFORE EMR WENT LIVE SO THAT MRN NUMBERS WOULD MATCH ATHENA'S MRN NUMBERS. THIS CONVERSION WAS NOT COMPLETED BEFORE EMR WAS ACTIVATED CAUSING EXISTING PATIENT RECORDS TO BE UPDATED INCORRECTLY DUE TO OVERWRITING. OVERWRITTEN INFORMATION COULD POTENTIALLY LEAD TO AN INCORRECT TREATMENT OR DIAGNOSIS. CURRENTLY THERE IS NO INDICATION OF ACTUAL PATIENT HARM. (B)(4). | patient data updated incorrectly | patient data issue | data overwritten | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, THE CLIENT REPORTED A STUDY WAS NOT AVAILABLE ON THE WORKLIST DUE TO THE STUDY STATUS CHANGING TO PRELIMINARY. THE CUSTOMER CONFIRMED THERE WAS A THIRTY MINUTE DELAY IN DIAGNOSIS FOR THE STUDY. INFORMATION RECEIVED FROM THE CUSTOMER CONFIRMED THERE WAS NO DIRECT IMPACT TO PATIENT CARE DUE TO THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect study status change | patient data issue | study status to preliminary | study unavailable | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2018, A CUSTOMER STATED MERGE PACS CHANGED THE STUDY STATUS TO READ, WHICH MADE IT UNAVAILABLE ON THE WORKLIST. WHEN A STUDY AUTOMATICALLY GOES TO READ, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS BECAUSE THE CUSTOMER COULD MISS READING THE STUDY IN A TIMELY MANNER. HOWEVER, FOR THIS CUSTOMER, THE EVENT WAS DETECTED BY THE PHYSICIAN AND THERE WASN'T ANY IMPACT TO PATIENT CARE. OTHER CLINICAL DATA WAS AVAILABLE TO SUPPORT THE PROPER TREATMENT OF THE PATIENT. REFERENCE (B)(4). | incorrect study status change | data issue | unavailable on the worklist | software issue | procedure delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, A CUSTOMER REPORTED A STUDY STATUS CHANGED FROM "UNREAD" TO "READ" WITHIN A FEW SECONDS OF IMPORTING INTO THE PACS SYSTEM. THIS CAUSED THE STUDY TO FALL OFF THE DOCTOR'S WORK LIST AND NOT BE READ WITHIN CUSTOMER'S TIMELINE. THE CUSTOMER CONFIRMED THE STUDY WAS READ WITHIN APPROXIMATELY ONE HOUR AND 20 MINUTES OF ENTERING THE PACS SYSTEM.ÿ WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW STUDY IS ADDED AND THEN REMOVED FROM THE WORKLIST, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS FOUND BY THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. | incorrect study status change | data issue | unavailable on the worklist | software issue | procedure delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, DURING INTERNAL TESTING, A MERGE EMPLOYEE REPORTED AN ISSUE WITH VOICE CLIPS SAVING TO THE WRONG PATIENT. VOICE CLIPS ARE AUDIO ANNOTATIONS AND ARE USED DIFFERENTLY DEPENDING ON USER'S WORK FLOW. VOICE CLIPS COULD CONTAIN ESSENTIAL INFORMATION RELATED TO TREATMENT/DIAGNOSIS. THERE IS A POTENTIAL THAT THE WRONG VOICE CLIP MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITH THE WRONG VOICE CLIP, THERE IS A POTENTIAL FOR AN INCORRECT TREATMENT AND/OR DIAGNOSIS TO THE PATIENT. HOWEVER, THE UNDERLYING IMAGE IS PRESENT FOR THE USER TO READ. THERE WAS NO DIRECT IMPACT TO PATIENT CARE REPORTED. REFERENCE COMPLAINT-(B)(4). | incorrect voice clip | device markings issue | data issue | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6)2017, A CUSTOMER IDENTIFIED AN ISSUE WITH IMAGE LOAD ERRORS WHILE TESTING MERGE PACS VERSION 7.2.1 IN THEIR TEST ENVIRONMENT. IT WAS DETERMINED THAT IMAGES THAT FAILED TO LOAD PROPERLY WERE CORRUPTED. THE CORRUPTED IMAGES WERE SKIPPED DURING MANUAL SCROLLING THROUGH A STUDY. THERE IS NO WARNING OR ERROR THAT INDICATES IMAGES HAVE BEEN SKIPPED. THERE IS A POTENTIAL THAT A SKIPPED IMAGE MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITHOUT NECESSARY INFORMATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS &/OR TREATMENT OR AN INACCURATE DIAGNOSIS &/OR TREATMENT OF A PATIENT. AT THIS TIME, THERE IS NO EVIDENCE OF PATIENT HARM OR DELAY DUE TO THIS ISSUE. (B)(4). | image corrupted | patient data issue | skip image | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: ON 07/10/2017, MERGE HEALTHCARE SUPPORT IDENTIFIED AN ISSUE DURING TROUBLESHOOTING ACTIVITIES WHERE CORRUPT IMAGES ARE SKIPPED DURING MANUAL SCROLLING. THERE IS NO WARNING OR ERROR THAT INDICATES IMAGES HAVE BEEN SKIPPED. THERE IS A POTENTIAL THAT A SKIPPED IMAGE MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITHOUT NECESSARY INFORMATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS &/OR TREATMENT OR AN INACCURATE DIAGNOSIS &/OR TREATMENT OF A PATIENT. AT THIS TIME, THERE IS NO EVIDENCE OF PATIENT HARM OR DELAY DUE TO THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | image corrupted | patient data issue | skip image | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: A CUSTOMER CONTACTED MERGE PACS TECHNICAL SUPPORT ON (B)(6) 2017 AND ALLEGED THAT STUDIES WERE NOT GOING BACK TO AN UNVERIFIED STATE IF MORE IMAGES WERE SENT INTO THE SYSTEM AND APPENDED TO THE PATIENT'S STUDY AFTER THE STUDY WAS EVALUATED. THIS ISSUE IS DUE TO AN OVERWRITE STATUS FIELD BEING SET TO NO. WHEN THE OVERWRITE STATUS IS SET TO YES, IT CHANGES THE STUDY BACK TO UNVERIFIED TO ALERT THE USER OF ADDITIONAL IMAGES. WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW IMAGE IS ADDED TO A STUDY, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS APPARENT TO THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | incorrect study status change | patient data issue | skip image | overwrite status field | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017, A CUSTOMER REPORTED AN ISSUE WITH ANNOTATIONS NOT SAVING WHEN THE STUDY IS DICTATED AND CLOSED. THE PHYSICIAN BECAME AWARE OF THE ISSUE WHEN HE WAS COMPARING A PRIOR IMAGE WITHOUT ANNOTATIONS TO A CURRENT, UNREAD IMAGE. AT THIS TIME, THE MEASUREMENTS WERE STILL AVAILABLE IN THE REPORT AND DID NOT CAUSE AN IMPACT TO THE CARE OF THE PATIENT. Manufacturer Narrative: AFTER FURTHER INVESTIGATION, MERGE HEALTHCARE DEVELOPMENT DETERMINED THIS WAS A DEFECT IN THE CODE AND RESOLVED THIS ISSUE IN V 7.2. TO PREVENT THIS ISSUE, THE USER UNDERSTANDS THEY NEED TO DRAG THE SERIES IN FROM THE THUMBNAIL INSTEAD OF SERIES TOOLBAR. MERGE HEALTHCARE SUPPORT ALSO CREATED A HANGING PROTOCOL TO ASSIST WITH WORKFLOW CHANGES. THE CUSTOMER HAS NOT REPORTED ANY ADDITIONAL ISSUES SINCE CHANGING WORKFLOW. NO FURTHER CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. | fail to save annotation | data issue | inaccuracy image measurement | software defects | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017, A CUSTOMER ALLEGED AN ISSUE WITH A BREAST OUTLINE NOT BEING CORRECT, WHICH RESULTED THE SITE TO TAKING ADDITIONAL VIEWS OF THE PATIENT'S BREAST. THE CUSTOMER ALSO ALLEGED THAT THE CAD MARKINGS FOR MEASUREMENTS AND CALCIFICATION ARE INCORRECT ON 3D IMAGES. THE ALLEGED ISSUE WITH THE BREAST OUTLINE IS READILY APPARENT TO THE USER ON MAMMOGRAPHY IMAGES. HOWEVER, THE ISSUE COULD POTENTIALLY CAUSE A DELAY IN DIAGNOSIS OR INCORRECT TREATMENT DUE TO EITHER ADDITIONAL IMAGES BEING TAKEN AND/OR INCORRECT MEASUREMENTS. IT SHOULD BE NOTED THAT THE CUSTOMER HAS ACCESS TO CORRECT CAD MARKINGS IN 2D IMAGES AND IN SECONDARY CAPTURE. AT THIS TIME THERE IS NO INDICATION OF DIRECT IMPACT TO PATIENT CARE OR HARM AS A RESULT OF THIS ISSUE. | incorrect measurement marking | device markings issue | incorrect measurement value | software issue | procedure delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THE NUMBER OF IMAGES ON THEIR LOCAL MAMMO SERVERS DO NOT MATCH THE NUMBER OF IMAGES ON THEIR PACS SERVER. THERE IS A POTENTIAL RISK OF THE CUSTOMER REMOVING IMAGES FROM THE MAMMO SERVER UNDER THE ASSUMPTION THAT THE IMAGES HAVE BEEN ARCHIVED PROPERLY TO THE PACS SERVER. AT THIS TIME, ALL IMAGES ARE BEING TRANSFERRED TO THE SERVER AND THERE IS NO EVIDENCE OF LOSS OF DATA/IMAGES. HOWEVER, THE ISSUE COULD POTENTIALLY CAUSE A DELAY OR INCORRECT TREATMENT DUE TO A PHYSICIAN NOT BEING ABLE TO VIEW A PATIENT'S PRIOR STUDIES FOR COMPARISON WITH THE CURRENT STUDY. (B)(4). | incorrect image amount | data issue | loss of data | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THAT TWO DIFFERENT PATIENT'S STUDIES WERE MERGED TOGETHER INTO A SINGLE PATIENT'S RECORD BY THE SYSTEM AUTOMATICALLY. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT TWO DIFFERENT PATIENTS HAD BEEN MERGED TOGETHER DUE TO THE PATIENT'S HAVING MATCHING MRN AND IPID IDENTIFIERS. EACH PATIENT'S DEMOGRAPHIC INFORMATION IS STORED AND VISIBLE ON THE IMAGES IN THE REPORT AND WOULD BE DETECTABLE TO THE USER WHEN REVIEWING DEMOGRAPHIC INFORMATION. THERE WAS NO INDICATION FROM THE CUSTOMER REGARDING ANY DIRECT IMPACT TO PATIENT CARE AT THIS TIME. HOWEVER, THIS ISSUE COULD POTENTIALLY CAUSE A DELAY IN PATIENT CARE DUE TO THE INABILITY TO FIND A PATIENT'S INFORMATION OR A MISDIAGNOSIS IF THE PATIENT'S DEMOGRAPHIC INFORMATION IS NOT READ AND A TREATMENT IS APPLIED TO THE INCORRECT PATIENT. (B)(4). | patient data merged | patient data issue | inable to find patient info | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER CONTACTED AFTER HOURS MERGE HEALTHCARE SUPPORT ON (B)(6) 2016 AND REPORTED THAT A PATIENT'S ORDER WAS MERGED WITH ANOTHER PATIENT'S ORDER. MERGE HEALTHCARE SUPPORT BEGAN INVESTIGATING THE ISSUE AND WORKED WITH THE CUSTOMER OVER MULTIPLE WEEKS. THE INVESTIGATION CONCLUDED ON 3/25/2016 AND DETERMINED THAT THE SITE HAD BEEN IMPORTING PATIENT INFORMATION FROM 2 DIFFERENT RIS SYSTEMS. THIS WAS DETERMINED TO BE A USER ERROR THAT CONTRIBUTED TO THE MERGING OF 2 DIFFERENT PATIENT'S INFORMATION. THE SITE HAS DISCONTINUED SENDING PATIENT INFORMATION FROM THE 2ND RIS SYSTEM. MERGE PACS WAS FUNCTIONING AS DESIGNED FOR PATIENT DATA INTERACTIONS WITH ONE RIS SYSTEM; NOT MULTIPLE RIS SYSTEMS. THIS MDR IS BEING SUBMITTED DUE TO THE POTENTIAL FOR A POSSIBLE MISDIAGNOSIS OR MISTREATMENT OF A PATIENT DUE TO THE MERGING OF DIFFERENT PATIENT'S DATA. THE CUSTOMER DID NOT INDICATE ANY HARM OR IMPACT TO A PATIENT(S) DUE TO THE MERGING OF PATIENT INFORMATION BEING IMPORTED FROM THE CUSTOMER'S MULTIPLE RIS SYSTEMS INTO MERGE PACS. (B)(4). | patient order merged | patient data issue | during procedure | user error | inappropriate treatment or diagnostic |  |  | one system designed |  |  | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED ON (B)(6) 2016 THAT THEY WERE EXPERIENCING SLOWNESS ON IMPORTING OF STUDIES AND QUERYING. MERGE HEALTHCARE SUPPORT BEGAN INVESTIGATING THE ISSUE AND DETERMINED THAT THE SITE HAD STUDIES THAT WERE STUCK WHILE IN IMPORT. THE ISSUE WAS FOUND TO BE DUE TO THE ORGANIZATION OF THE DICOM TABLES AND AS A RESULT THESE WERE REORGANIZED. THE ISSUES WITH SLOWNESS AND IMPORTING STUDIES COULD POTENTIALLY LEAD TO A DELAY IN DIAGNOSIS OR TREATMENT. SUPPORT DETERMINED NO ISSUE FOUND WITH THE MERGE PACS SOFTWARE. DATABASE RE-ORGANIZATIONS ARE PART OF NORMAL SERVER UPKEEP. (B)(4). | import study slow | data issue | during procedure | DICOM table not organized | treatment or diagnostic delayed |  |  | database design |  |  | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THE COMMUNICATION NOTES WERE NOT SAVING. WITH MERGE PACS NOT SAVING COMMUNICATION NOTES AS EXPECTED THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT LEADS TO HARM. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. (B)(4). | fail to save note | loss of data | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER REPORTED THAT WHEN THEY MERGED STUDY A AND STUDY B THE COMMENTS FROM STUDY A WERE LOST. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. HOWEVER, THERE IS A POTENTIAL THAT A USER MAY ADD CLINICAL INFORMATION INTO THE COMMENTS SECTION AND THEREFORE, THERE IS A POTENTIAL THAT INFORMATION THAT MAY IMPACT A TREATMENT OR DIAGNOSIS DECISION MAY NOT BE PRESENT ON A MERGED STUDY. (B)(4). Manufacturer Narrative: THE INVESTIGATION FOUND THAT MERGE PACS WAS FUNCTIONING AS DESIGNED. MERGE PACS VERSION 7.1 RELEASE NOTES WILL ADDRESS THE DELETION OF COMMENTS IN STUDY A WHEN MERGING WITH STUDY B. | comment deleted | loss of data | during merging studies | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER ALLEGED TO MERGE HEALTHCARE THAT STUDIES WERE INTERMITTENTLY GOING TO A "READ" STATUS AFTER IMPORTING. WITH NON-READ OR NON-REVIEWED STUDIES INDICATING A READ STATUS COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THE INVESTIGATION FOUND THAT DURING IMPLEMENTATION OF THE UPGRADED VERSION OF MERGE PACS (7.0.2) THE OVERWRITE FUTURE STATUS SETTING WAS SET TO "NO." THIS SETTING WAS INTERMITTENTLY CAUSING NEWLY IMPORTED IMAGES TO NOT CHANGE THE STUDY STATUS TO "UNREAD." MERGE HEALTHCARE SUPPORT RESOLVED THE ISSUE BY CHANGING THE OVERWRITE FUTURE STATUS SETTING TO "NO." | incorrect report status | data issue | read status after importing | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THEY WERE UNABLE TO RETRIEVE SOME IMAGES TAKEN IN 2012 FROM THEIR ARCHIVE. DUE TO THE LOSS OF IMAGES THAT CAN BE USED WHEN COMPARING PRIOR STUDIES TO RECENT STUDIES, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. (B)(4). | fail to retrieve image | loss of data | fail to compare with historical image | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT WHILE VIEWING IMAGES OF CLAVICLE THEY OBSERVED THAT THE DICOM TAGS WERE IDENTIFYING THE IMAGES AS BEING THE PATIENT'S LEFT SIDE REGARDLESS OF WHETHER IT WAS AN IMAGE OF THE PATIENT'S LEFT OR RIGHT CLAVICLE. AS THE CUSTOMER ALSO PLACES MARKERS TO IDENTIFY THE PATIENT'S SIDE THE ISSUE IS EASILY NOTICED. HOWEVER, IF THE CUSTOMER DID NOT PLACE THE MARKER THE DICOM TAGS INCORRECTLY TAGGING THE IMAGES HAS THE POTENTIAL TO LEAD TO INCORRECT TREATMENT OR MISDIAGNOSIS. THERE WAS NO INDICATION FROM THE CUSTOMER THAT ANY PATIENTS WERE HARMED DUE TO THIS COMPLAINT. (B)(4). | incorrect dicom tag | device markings issue | identify image as left side | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT WHILE REVIEWING IMAGES WITH THE USE OF THE "OXFORD PARTIAL KNEE TWIN-PEG" TEMPLATE, FOUND THAT THE TEMPLATE SHOWED TEXT THAT STATED THE INCORRECT VIEW OF THE KNEE IMAGE. THE TEXT SHOULD HAVE READ RIGHT RATHER THEN LEFT. THIS ISSUE HAS A POTENTIAL TO LEAD TO AN INACCURATE DIAGNOSIS OR TREATMENT OF THE WRONG KNEE. THE CUSTOMER REPORTED THAT THE PHYSICIAN NOTICED THE ISSUE BEFORE MISDIAGNOSING THE INCORRECT KNEE. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED DUE TO THIS COMPLAINT. ADDITIONALLY, THE PHYSICIAN WOULD HAVE REFERENCED THE LEAD LATERALITY MARKER WITHIN THE IMAGE TO NOTE WHICH KNEE WILL BE DIAGNOSED. REFERENCE COMPLAINT NUMBER (B)(4). | incorrect image view | image display error | incorrect image text direction | software issue | inaccurate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. THE CUSTOMER CALLED INTO MERGE SUPPORT IN (B)(6) 2016 BECAUSE THEY FOUND THAT SOME OUTSIDE STUDIES THAT THEY WERE BRINGING INTO THEIR SYSTEM WERE CONFLICTING AND MERGING TO SOME CURRENT STUDIES ON THEIR SYSTEM. MERGE SUPPORT RECOGNIZED THE ISSUE AND INFORMED CLIENT THEY NEEDED TO CREATE A DICOM HEADER MODIFICATION RULE TO PREFIX THE MRNS THAT ARE COMING FROM THAT FACILITY. THIS WOULD HAVE FIXED THE ISSUE. CLIENT STATED THAT THEY WOULD WORK WITH MERGE IMPLEMENTATION STAFF TO RESOLVE THE ISSUE. THE CUSTOMER REPORTED THAT THERE HAD BEEN CASES WHERE RAD HAD READ IMAGES OF PATIENT A THINKING THEY WERE PATIENT B AND THE WRONG REPORT WAS ATTACHED TO THE WRONG PATIENT. THE SITE WAS ABLE TO CATCH THESE VERY QUICKLY AND CORRECT THE ISSUE. THERE IS NO INFORMATION SUGGESTING THAT A PATIENT HAS BEEN ADVERSELY AFFECTED BY THIS. TO MAKE SURE THE INCORRECT MERGING DOES NOT HAPPEN, SITE IS MANUALLY QCING EVERYTHING THAT COMES INTO THE SYSTEM. HOWEVER, THERE IS A RISK THAT ONE COULD BE MISSED OR THAT THE STUDY WAS READ AFTER HOURS AND WAS NOT QC'ED. AGAIN, THERE IS NO EVIDENCE OF A SPECIFIC PATIENT BEING ADVERSELY EFFECTED BY THIS. (B)(4). | conflicted study | data issue | import outside study | software issue | treatment or diagnostic delayed |  |  |  | CREATE A DICOM HEADER MODIFICATION RULE |  | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT AN AUTOMATED DAILY REPORT INDICATED THAT THREE OF THEIR ARCHIVED TAPES WERE SHOWING AS UNAVAILABLE. SHOULD A TAPE WITH ARCHIVED PATIENT DATA BE REQUESTED FOR VIEWING, THERE IS A POTENTIAL THAT THE INFORMATION MAY NOT BE AVAILABLE DUE TO THE PACS SYSTEM NOT BEING ABLE TO RETRIEVE/READ THE REQUESTED INFORMATION. THIS COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT OR DIAGNOSIS DUE TO UNAVAILABLE DATA. CURRENTLY, THERE IS NO KNOWN OR ALLEGED IMPACT TO SPECIFIC OR OVERALL PATIENTS. MERGE HEALTHCARE'S INITIAL INVESTIGATION RESULTS INDICATE THAT THE CUSTOMER IS RUNNING ON SERVER SOFTWARE/FIRMWARE THAT IS ANTIQUATED AND IS NO LONGER SUPPORTED BY THE OEM. (B)(4). | unavailable tape | component missing | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. MERGE HEALTHCARE RECEIVED AN ALLEGATION FROM A CUSTOMER ON (B)(6) 2016, CLAIMING THAT LINE MEASUREMENTS WERE NOT MATCHING THE OLDER VERSION OF PACS VERSUS THE NEWER VERSION. WHILE THE DIFFERENCE ALLEGED BY THE CUSTOMER WAS MINIMAL THERE IS A POTENTIAL THAT COMPARISONS MADE BETWEEN OLD IMAGES (~2011) AND NEWER IMAGES COULD RESULT IN AN INACCURATE DIAGNOSIS OR MISTREATMENT OF A PATIENT. MERGE HEALTHCARE IS INVESTIGATING THE ALLEGATION. (B)(4). | inconsistent measurement value | patient data issue | during procedure | dfferent version | inaccurate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT WHEN A NEW STUDY WAS VIEWED, THE PATIENT'S PRIOR STUDIES WERE NOT BEING RETRIEVED FROM THE DATABASE AS EXPECTED. THE CUSTOMER DID NOT ALLEGE ANY DEATH, SERIOUS INJURY OR HARM TO ANY SPECIFIC OR OVERALL PATIENTS. THERE IS HOWEVER, A POTENTIAL FOR A DELAY OR INCORRECT TREATMENT DUE TO A PHYSICIAN IS NOT BEING ABLE TO VIEW A PATIENT'S PRIOR STUDIES FOR COMPARISON WITH THE CURRENT STUDY. | fail to retrieve study | loss of data | compasion studies | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(4) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT WHEN ADDITIONAL IMAGES ARE SENT TO A STUDY THAT HAS ALREADY BEEN READ, THERE WAS NO NOTIFICATION THAT ADDITIONAL STUDIES HAD BEEN RECEIVED. THE STATUS REMAINS AS "READ" OR "FINAL STATUS" RATHER THAN INDICATING "UNVERIFIED NEW IMAGE" HAS BEEN RECEIVED VIA AN ALERT NOTIFICATION. IT IS UNKNOWN TO THE CUSTOMER WHEN THE PROBLEM FOR THE ALERT NOT BEING RECEIVED FIRST OCCURRED. CURRENTLY, THERE IS NO KNOWN IMPACT TO SPECIFIC OR OVERALL PATIENTS; HOWEVER, THERE IS POTENTIAL FOR A SITUATION TO OCCUR WHERE NEW IMAGES ARE NOT READ RESULTING IN DELAYED OR INCORRECT TREATMENT. THE CUSTOMER WILL BE MODIFYING A CONFIGURABLE SOFTWARE SETTING WITHIN THE MERGE PACS DEVICE TO EVALUATE IF THE ALERT IS RECEIVED FOR STUDIES THAT HAVE BEEN READ BUT ADDITIONAL IMAGES ARE SENT FOR THE STUDY. | incorrect image status | patient data issue | new images are not read | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(4) 2017 IT WAS REPORTED, "RADIOLOGIST NOTICED THAT SUV VALUES ON A FUSED PET CT ARE DIFFERENT COMPARED TO WHEN THE SUV IS MEASURED ON SOURCE IMAGES ALONE AND NOT FUSED IMAGES." THERE WAS NO ALLEGATION OF PATIENT INJURY, ADVERSE CONSEQUENCES, OR SIGNIFICANT DELAY IN PATIENT CARE. (B)(4). | inconsistent measurement value | patient data issue | inaccuracy image measurement | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017 INTERNAL TESTING FOUND THAT THERE WERE INCONSISTENCIES BETWEEN THE SUV MEASUREMENTS FOR 2D IMAGES AND SUV MEASUREMENTS FOR SPINNING MIP (MAXIMUM INTENSITY PROJECTION). THIS ISSUE WAS FOUND BY MERGE HEATLHCARE ON (B)(6)2017 AND WAS NOT REPORTED BY A CUSTOMER. THE ISSUE OF POTENTIALLY INACCURATE SUV VALUES MAY HAVE AN IMPACT ON PATIENT TREATMENTS FOR CANCER AND THEREFORE MIGHT RESULT IN HARM. HOWEVER, THERE IS NO INDICATION THAT ANY PATIENTS HAVE BEEN NEGATIVELY IMPACTED BY THIS ISSUE AT THIS TIME (B)(4). | inconsistent measurement value | patient data issue | inaccuracy image measurement | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017 A CUSTOMER REPORTED THAT THEY WERE UNABLE TO CONNECT TO THE SERVER FROM THE VIEWER. IT WAS FOUND THAT THE CUSTOMER HAD A BAD NETWORK INTERFACE CARD ON THEIR SERVER WHICH NEEDED TO BE REPLACED BY (B)(4). IT WAS REPORTED THAT SYSTEM DOWNTIME OCCURRED FROM (B)(6) 2017. THIS MAY POTENTIALLY CAUSE A DELAY IN TREATMENT BUT IS UNLIKELY TO CAUSE HARM AS THERE IS OTHER CLINICAL INFORMATION THAT WILL BE UTILIZED TO SUPPORT THE TREATMENT OF THE PATIENT. THE CUSTOMER DID NOT INDICATE ANY DELAY IN PATIENT CARE DURING ALLEGED DOWNTIME. (B)(4). | fail to connect to server | connection issue | system downtime | bad network interface card | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT A STUDY AND ITS IMAGES WERE DELETED AND ARE NO LONGER ABLE TO BE VIEWED. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. HOWEVER, WITH A PATIENT'S IMAGES OR STUDY BEING UNAVAILABLE FOR VIEWING THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. (B)(4). | image deleted | loss of data | image unavailable to review | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED THEY USED THE QC EDITOR TOOL TO SPLIT A STUDY AND FOUND THE CHANGES WERE NOT SAVED AFTER SAVING AND REOPENING THE STUDY. THE QC EDITOR TOOL DID NOT GIVE THE CUSTOMER A WARNING THAT CHANGES WERE NOT SAVED. THE QC EDITOR TOOL IS USED TO SPLIT, MERGE AND/OR DELETE SERIES OF IMAGES IN A STUDY. THIS MAY POTENTIALLY CAUSE A DELAY IN TREATMENT BUT IS UNLIKELY TO CAUSE HARM AS THERE IS OTHER CLINICAL INFORMATION THAT WILL BE UTILIZED TO SUPPORT THE TREATMENT OF THE PATIENT. HOWEVER THERE IS NO REPORTED SERIOUS INJURY, DEATH OR HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | fail to save change | loss of data | rework on study | software application issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED THEY WERE UNABLE TO LOGIN INTO MERGE PACS FOR APPROXIMATELY 8 HOURS. RATHER THAN USING MERGE PACS, PHYSICIANS AND RADIOLOGISTS WERE ABLE TO VIEW IMAGES AND PATIENT STUDIES AT THE IMAGING MODALITY. WITH CLINICIANS NOT BEING ABLE TO LOG INTO MERGE PACS THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. HOWEVER, THE CUSTOMER CONFIRMED THERE WAS NO DIRECT IMPACT TO PATIENT CARE. (B)(4). | fail to log in | application security issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED IN THEIR PREVIOUS MERGE PACS VERSION THAT IF A DOCUMENT OR IMAGE WAS ADDED TO A STUDY AFTER IT WAS MARKED "READ", IT WOULD CHANGE THE STATUS BACK TO "UNVERIFIED". AFTER THEIR UPGRADE, THEY NOTICED THE STUDY STATUS WAS NOT CHANGING WHEN ADDITIONAL DOCUMENTATION WAS ADDED. WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW IMAGE IS ADDED TO A STUDY, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS APPARENT TO THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT. (B)(4). | incorrect study status | application security issue | unaware new image added | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER REPORTED THAT THE HL7 INTERFACE IS OVERWRITING PATIENT INFORMATION. IT WAS FOUND THAT THE GLOBAL PATIENT UPDATE WAS TURNED ON BY AN EMPLOYEE ON SITE. THIS SETTING AUTOMATICALLY DEFAULTS TO OFF UNLESS THE SETTINGS ARE MANUALLY SWITCHED TO ON. THIS RESULTED IN PATIENT INFORMATION BEING UPDATED TO THE INCORRECT PATIENT IN THEIR SYSTEM WHEN A DUPLICATE MRN NUMBER OCCURRED. THIS ISSUE HAS A POTENTIAL TO LEAD TO PATIENT INFORMATION BEING COMPROMISED OR OVERWRITTEN BY A NEW STUDY SUBMISSION. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED DUE TO THIS COMPLAINT. (B)(4). | patient information overwritten | loss of data | patient information upates to wrong patient | configuration issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THAT WHEN MOVING OR CORRECTING STUDIES USING THE QC EDITOR IN MERGE PACS, THE STUDY COMMENTS MADE ON THE SOURCE STUDY WERE NOT RETAINED AND MOVED TO THE TARGET STUDY. THIS MAY CAUSE A DELAY IN PATIENT CARE OR MISDIAGNOSIS IF IT'S NOT READILY APPARENT TO THE USER AND THEY OVERLOOK IMPORTANT INFORMATION WITHIN THE COMMENTS. THE PATIENT'S IMAGES ARE AVAILABLE FOR DICTATION AND COMMENTS ARE SUPPLEMENTAL INFORMATION. (B)(4). | fail to transfer comment | loss of data | comments are not on the target study | software application issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED ON (B)(6) 2016 THAT THEY WERE EXPERIENCING ENTERPRISE WIDE SLOWNESS. THEY WERE HAVING ISSUES PULLING ORDERS FROM THE WORKLIST. MERGE PACS WAS SLOW TO SCAN, SEND AND RETRIEVE AND THE CUSTOMER STATED THEY WERE NOT ABLE TO BURN A CD DUE TO THESE ISSUES. MERGE SUPPORT BEGAN INVESTIGATING THE ISSUE AND FOUND A RELATED SUPPORT CASE (B)(6). THAT CASE HAD BEEN OPENED ON (B)(6) AND ANOTHER MEMBER OF THE SUPPORT TEAM WAS ALREADY WORKING WITH SITE ON SLOWNESS ISSUES. THIS CASE WAS CLOSED AS A DUPLICATE AND INFORMATION WAS TRANSFERRED TO ORIGINAL CASE. THIS REPORT IS BEING ISSUED AS A RESULT OF THE CUSTOMER REPORTING AGAIN ISSUES WITH SLOWNESS PRIOR TO THE DICOM TABLES WERE BEING REORGANIZED. THE SLOWNESS HAS A POTENTIAL OF LEADING TO A DELAY IN A TREATMENT OF DIAGNOSIS OF A PATIENT. (B)(4). | scan slow | device operates differently than expected | during procedure | enterprise wide slowness | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A SCALABLE SOLUTION THAT SUPPORTS A DISTRIBUTED WORKFORCE AND DELIVERS OPERATIONAL EFFICIENCIES THROUGH WEB-BASED ACCESS, COMMON WORKLISTS, AND COMPREHENSIVE WORKFLOWS. THE CUSTOMER REPORTED THAT A NEW PATIENT CAME IN FOR A PROCEDURES. THE TECHNOLOGIST CLICKED ON PRIOR STUDIES AND NOTICED THERE WAS A BIG LIST AND SINCE IT WAS A NEW PATIENT, THERE SHOULD NOT HAVE BEEN PRIORS AT THAT POINT. SPECIFICALLY, THEY NOTICED A TEST PATIENT'S STUDY WAS ASSOCIATED WITH AN ACTUAL PATIENT'S STUDY. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE DUE TO THE POSSIBILITY OF IN ADVERTENTLY COMBINING STUDIES FROM DIFFERENT PATIENTS. ON 12/4/15, THE OBSERVATION TABLE WAS UPDATED WITH THE CORRECT PATIENT. SISU WILL SUPPRESS THE URN FROM COMING THROUGH ON FUTURE HL7 MESSAGES TO PACS TO PREVENT THIS ISSUE FROM HAPPENING AGAIN IN THE FUTURE. | incorrect patient information | patient data issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). MERGE PACS CAN BE ACCESSED FROM WITHIN THE HOSPITAL OR ENTERPRISE, OR FROM REMOTE LOCATIONS VIA WEB-BASED ACCESS. IMAGES VIEWED ON MOBILE DEVICES MUST NOT BE USED FOR DIAGNOSTIC REVIEW. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THERE WERE DISCREPANT ROI (REGION OF INTEREST) VALUES BETWEEN 2D AND MPR IMAGES. MERGE HEALTHCARE SUPPORT WORKED WITH THE CUSTOMER TO REPRODUCE THE ISSUE. THE ISSUE IS CONFINED TO THE 3D FLAIR SERIES OF IMAGES ONLY. THIS ISSUE HAS BEEN CORRECTED IN THE CURRENT RELEASE OF MERGE PACS AND THE CUSTOMER WAS INFORMED OF THIS. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | inconsistent value | patient data issue | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS), SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT SOME IMAGES IN A MAMMOGRAPHY SERIES WERE NOT DISPLAYING AND WERE MISSED OR SKIPPED. MERGE HEALTHCARE SUPPORT IS CURRENTLY IN THE PROCESS OF ENGAGING THE CUSTOMER TO DETERMINE THE ISSUE AND NEXT STEPS. WHEN FURTHER INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | missing image series | loss of data | during procedure | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, MERGE HEALTHCARE RECEIVED A REPORT FROM A CUSTOMER THAT HOUNSFIELD UNIT VALUES REPORTED IN THE 3D VIEW WERE INCONSISTENT WITH THE HOUNSFIELD UNIT VALUES REPORTED IN THE 2D VIEW OF MERGE PACS. UPON INVESTIGATION, IT WAS DISCOVERED THAT PIXEL VALUE MEASUREMENTS (INCLUDING HOUNSFIELD UNIT MEASUREMENTS AND SUV MEASUREMENTS) FROM THE PROBE, ROI, AND REGION ANALYSIS AREA TOOLS ARE INCORRECT ON MULTI-PLANAR REFORMATS (MPRS) GENERATED BY MERGE PACS. THE ISSUE IS NOTICEABLE WHEN A MODALITY ACQUIRES A SERIES OF IMAGES AND SENDS THE IMAGES TO MERGE PACS. THE RADIOLOGIST THEN GENERATES ADDITIONAL PLANES SUCH AS SAGITTAL OR CORONAL USING MPR. MEASUREMENTS DONE ON THE MERGE PACS GENERATED MPR'S MAY HAVE INCORRECT MEASUREMENTS. INCORRECT PIXEL MEASUREMENTS MAY RESULT IN THE IMPLEMENTATION OF AN INCORRECT OR UNNECESSARY TREATMENT PLAN THAT COULD RESULT IN THE PATIENT RECEIVING UNNECESSARY MEDICATION, CHANGE IN DOSAGE OF EXISTING MEDICATION, OR UNNECESSARY MEDICAL PROCEDURES. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. (B)(4). | inconsistent measurement value | patient data issue | pixal value measurement | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STAND ALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS I CONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, THE CUSTOMER REPORTED THAT A STUDY CAME IN AND THE CT TECH MANUALLY CHANGED FROM UNVERIFIED TO UNREAD AND SET THE CUSTOM STATUS TO ER. IT UPDATED EVERYTHING AS EXPECTED AND SHOWED ON THE RADIOLOGST'S ER WORKLIST. A LITTLE LATER, THE CT TECH SENT IN SCANNED PAPERWORK, WHICH REVERTED THE STATUS BACK TO UNVERIFIED, THEREBY, REMOVING THE STUDY FROM THE RADIOLOGIST'S ER WORKLIST (WHICH IS EXPECTED BEHAVIOR). HOWEVER, THE STATUS OF THE STUDY DID NOT UPDATE TO UNVERIFIED ON THE CT TECH'S WORKLIST AS EXPECTED. THE CT TECH CALLED THE RADIOLOGIST INQUIRING ABOUT THE STUDY AND THE RADIOLOGIST DID NOT HAVE THAT STUDY ON THEIR WORKLIST. THE CT TECH THEN FOUND THAT THE STUDY NEEDED TO BE RE-VERIFIED, THEREBY PLACING ONTO THE RADIOLOGIST'S ER WORKLIST. THE CT TECH WAS NOT ALERTED THAT THE STUDY HAD BEEN PLACED BACK IN THE UNVERIFIED STATUS AND NEEDED TO BE RE-VERIFIED TO UNREAD. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. | incorrect status change | data issue | study needs to reverify | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITIAL MODALITIES. CUSTOMER REPORTED ((B)(4)) THAT THE INCORRECT PATIENT'S IMAGES WERE DISPLAYED IN THE HANGING PROTOCOL COMPARISON WINDOW WHEN USING MERGE PACS. THE USER CLOSED THE VIEWER AND RELAUNCHED THE STUDY. ON THE SECOND LOAD, THE CORRECT COMPARISON WAS DISPLAYED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE INCORRECT THE PATIENT'S IMAGES WERE BEING DISPLAYED FOR COMPARISON. | incomplete image | image display error | fail to compare with historical image | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE PACS PROVIDES A CENTRAL HUB FOR MANAGING AND VIEWING STUDIES, WHILE PROVIDING ACCESS TO REMOTE USERS VIA THE WEB. ITS COMMON WORKLISTS ENABLE EFFICIENT SUB-SPECIALTY READING FROM REMOTE LOCATIONS. ON (B)(6) 2015, A CUSTOMER REPORTED THAT THE STUDY WAS MISSING IMAGES WHEN VIEWED IN MERGE PACS. HOWEVER, THE IMAGES WERE STILL AVAILABLE ON THE MODALITY DEVICE. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE PATIENT IMAGES WERE NOT ACCESSIBLE TO MAKE A COMPREHENSIVE DIAGNOSIS. MERGE REMOVED THE STUDY FROM MERGE PACS AND HAD IT RESENT FROM THE MODALITY TO MAKE THE PATIENT IMAGES AVAILABLE FOR DIAGNOSTIC REVIEW. THE CUSTOMER VERIFIED THE ISSUE WAS RESOLVED. | fail to load image | application program issue | no available for diagnostic review | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016 A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT INDICATING THAT THEY WERE UNABLE TO FIND A STUDY. MERGE INVESTIGATED THE ISSUE AND DETERMINED THAT THE ALLEGED MISSING STUDY WAS STORED IN 2007 AND WATERMARKED OFF IN 2008. MERGE HEALTHCARE HAS NOT BEEN ABLE TO LOCATE THE IMAGE DATA. WITH HISTORIC IMAGES NOT BEING AVAILABLE FOR VIEWING, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS AND/OR TREATMENT. HOWEVER, THERE WERE NO REPORTS OF DELAY OR HARM TO PATIENTS AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THE ISSUE TO DETERMINE WHAT MAY HAVE HAPPENED TO THE ARCHIVED IMAGES. | unable to locate study | data issue | unable study review | unable to locate study | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT X-RAY IMAGES WERE BLURRY. DUE TO IMAGES NOT PRESENTING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM. THERE IS NO INDICATION OF A DELAY IN TREATMENT OR DIAGNOSIS OR HARM THAT HAS OCCURRED AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THE CUSTOMER'S ALLEGATION. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT SHALL BE SUBMITTED. | unclear image | Image Resolution Poor | image not present as expected | poor image quality | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT THEY "CANT INVERT IMAGE WHEN USING WINDOW LEVEL TOOL ON ONE STUDY". WITH IMAGES NOT BEING PRESENTED AS EXPECTED TO A USER, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM. THERE HAVE BEEN NO REPORTS OF DELAY IN TREATMENT OR DIAGNOSIS TO A PATIENT OR OF HARM AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THIS ALLEGATION BY THE CUSTOMER. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | unable to reversal image | image reversal | image not present as expected | software application issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT X-RAY IMAGES WERE DARK AND THAT THESE IMAGES ARE UNABLE TO BE READ. FURTHER DISCUSSIONS WITH THE CUSTOMER REVEALED THAT A RADIOLOGIST OVERLOOKED AN IMAGE DUE TO THE IMAGE BEING BLACK WITH NO STUDY INFORMATION ON THE VIEWPORT. WITH RADSUITE NOT PRESENTING IMAGES IN A MANNER THAT ALLOWS FOR VIEWING, THERE IS A POTENTIAL FOR A DELAY IN A PATIENT'S TREATMENT OR DIAGNOSIS IF AN IMAGE IS OVERLOOKED. THE CUSTOMER HAS INDICATED THAT NO HARM HAS RESULTED TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: THE CUSTOMER'S ALLEGATION IS UNDER FURTHER INVESTIGATION BY MERGE HEALTHCARE. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE FILED. | image dark | poor quality image | lost of study data | software application issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE RADSUITE IS AN APPLICATION THAT PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. THE APPLICATION PROVIDES WORKFLOW INTEGRATION CAPABILITIES FOR HEALTH CARE ENTERPRISES, WHEREIN: RADIOLOGISTS CAN VIEW, ANNOTATE, AND TAG STUDIES AS DIAGNOSTICALLY READ; REFERRING PHYSICIANS CAN VIEW PATIENT IMAGES AND RADIOLOGISTS' ANNOTATIONS AND; TERTIARY CARE PHYSICIANS, MEDICAL TECHNOLOGISTS, AND INFORMATION TECHNOLOGY PROFESSIONALS CAN RECEIVE PATIENT RECORDS. ON (B)(6) 2015, THE CUSTOMER REPORTED ((B)(6)) THAT CASES BEING SENT FROM RADSUITE TO AN OFFSITE DESTINATION WAS SLOW. THIS WAS A THIRD PARTY PACS VENDOR THAT DOES OFFSITE READING FOR THE CUSTOMER. THE CUSTOMER SENDS FROM RADSUITE AT NIGHT FOR THEIR STROKE PATIENTS. THE SITE REQUIRES 3 MINUTE AVERAGE TRANSFER IN ORDER TO MEET 15 MINUTES COMPLETE TO FINAL TURNAROUND TIME FOR STROKE/PERFUSION PATIENTS. EXPORTING AT NIGHT COULD TAKE 20 - 40 MINUTES. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW, LEADING TO A DELAY IN TREATMENT FOR STROKE VICTIMS. THE ROOT CAUSE OF THIS ISSUE WAS NOT LINKED TO RADSUITE OR ANY OTHER MERGE PRODUCT. IT WAS TIED TO A 3RD PARTY VENDOR THAT ST. JOHN'S IS USING TO READ THEIR STUDIES AT NIGHT. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE ROOT CAUSE WAS NOT RELATED TO RADSUITE BUT TO AN OUTSIDE SERVICE USED BY THE CUSTOMER. | reading slow | device operates differently than expected | unavailable for review | communication issue | treatment or diagnostic for stroke delayed | 4 |  |  |  |  | FDA2
Event Description: MERGE RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER REPORTED THAT THE PATIENT LOCATION OF ER WAS NOT COMING ACROSS THE SYSTEM. THE ER WORKLISTS ARE FILTERED BASED ON PATIENT LOCATION. A PRIOR STUDY WAS RETRIEVED AND TRIGGERED AN UPDATE TO THE LOCATION. AS SUCH, THE ER STUDY DROPPED OFF THE ER EXAMLIST. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT IMAGES ARE NOT AVAILABLE FOR A DIAGNOSIS. IT WAS DISCOVERED THAT THE CUSTOMER HAD AN INCORRECT UNDERSTANDING OF THE MERGE SOFTWARE. THE CUSTOMER HAD ASSUMED THAT THE PATIENT LOCATION WAS TRACKED AT AN ENCOUNTER LEVEL, BUT THE PRODUCT DOES NOT HAVE ENCOUNTER DRIVEN WORKFLOW. NO CHANGES BEING MADE TO THE PRODUCT. THIS ISSUE IS NOT A RESULT OF A DEFECT OR PRODUCT LABELING PROBLEMS. (B)(4). Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | incorrect patient information | patient data issue | drop off the exam | user incorrect understanding software | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY (B)(4) A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. INFORMATION WAS RECEIVED FROM A (B)(4) ADMINISTRATOR ON (B)(6) 2016 REPORTING A MISSING REPORT. THE UNITY PACS AUDIT TRAIL STATES THAT THE REPORT WAS SAVED (B)(6) 2016. MERGE SUPPORT INVESTIGATED AND CONFIRMED THAT A REPORT WAS SAVED ON THE STATION. HOWEVER, THERE WAS NO DOCUMENT IN THE REPORT DIRECTORY. A BLANK WORD DOCUMENT WAS PLACED IN THE DIRECTORY AND THE PHYSICIAN RE-READ THE EXAM. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | missing report | loss of data | fail to save report to directory | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS A CUSTOMER CONTACTED MERGE HEALTHCARE ON (B)(6) 2016 AND REPORTED SEEING A MESSAGE THAT A REPORT DID NOT CROSS TO THEIR RADIOLOGY INFORMATION SYSTEM (RIS) SYSTEM (EPIC). AFTER NOTICING THE VISIT NUMBER DID NOT MATCH, THE CUSTOMER CORRECTED IT IN MERGE UNITY AND TRIED TO RESEND. WHEN THE REPORT OPENED, IT WAS BLANK. THE CUSTOMER WAS UNABLE TO VIEW THE REPORT. THE MERGE HEALTHCARE UNITY SUPPORT TECHNOLOGIST PLACED A BLANK WORD DOCUMENT IN THE REPORT DIRECTORY ON THE IMAGE SERVER. THIS ALLOWED THE EXAM TO BE RE-READ BY THE PHYSICIAN AND THE REPORT SAVED AND WAS ACCESSIBLE. (B)(4). Manufacturer Narrative: EXAM WAS AVAILABLE FOR REVIEW. THE ISSUE WAS IDENTIFIED BY THE PACS ADMIN DURING INTERNAL QC WORK VIEWING FAILURES OF HL7 INTERFACES. LOGS FOR THIS STATION/INCIDENT WERE UNAVAILABLE FOR REVIEW. THERE WAS A "REPORT SAVE" EVENT IN THE AUDIT TRAIL, HOWEVER, NO EXAM REPORT UPLOADED TO THE IMAGE SERVER. IT WAS UNABLE TO BE DETERMINED IF THIS INCIDENT WAS RELATED TO SOFTWARE OR TO SITE NETWORK FAILURE. THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | blank report | loss of data | unavailable for review | site network failure, or software failure | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER ALLEGED ON (B)(6) 2016 THAT HE WAS UNABLE TO OPEN AN APPROVED REPORT FROM THE ONLINE TAB. MERGE HEALTHCARE SUPPORT WAS UNABLE TO LOCATE AN ASSOCIATED DOCUMENT IN THE REPORT DIRECTORY ON THE CUSTOMER'S IMAGE SERVER. A BLANK WORD DOCUMENT WAS PLACED IN THE DIRECTORY WHICH ALLOWED THE PHYSICIAN TO RE-READ THE EXAM. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NOT ASSOCIATING TO AN EXAM AND NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: THIS INCIDENT WAS A RESULT OF THE UNIVERSAL MANAGER CRASHING AT THE POINT THE RADIOLOGIST WAS CLOSING THE CASE. SITE SOFTWARE WAS UPDATED ON (B)(6) 2017 FROM (B)(6). Manufacturer Narrative: MERGE HEALTHCARE INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THERE WAS A "REPORTSAVE" EVENT IN THE AUDIT TRAIL BUT THERE WAS NO REPORT AVAILABLE FOR THE PATIENT. MERGE HEALTHCARE'S UNITY PACS SUPPORT WAS UNABLE TO REVIEW THE STATION SPECIFIC LOGS BECAUSE THEY HAD ALREADY BEEN TRUNCATED. MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | unable to open report | patient data issue | unavailable for review | report directory under user image serer | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT ON (B)(6) 2016 AND REPORTED A MISSING REPORT. AN AUDIT OF THE EXAM SHOWED THE CASE DID HAVE A 'REPORT SAVE' ENTRY IN THE TRAIL. THERE IS NO REPORT LISTED FOR THE EXAM. THE STATION SPECIFIC LOGS WERE TRUNCATED PRIOR TO UNITY TECHS REVIEW. THE REPORT WAS NOT SALVAGEABLE AND NEEDED TO BE RE-DICTATED BY THE PHYSICIAN. WITH A PATIENT'S STUDY REQUIRING RE-DICTATION THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT. HOWEVER, THERE HAS BEEN NO INDICATION OF ANY HARM AS A RESULT OF THIS ISSUE. (B)(6). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | missing report | loss of data | unavailable for review | log file trucated | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT ON (B)(6) 2017. THEY ALLEGED THAT THE REPORT FOR AN EXAM READ ON (B)(6) 2017 WAS NOT FOUND. THIS WAS NOTICED BY THE MANAGER WHILE RUNNING A REPORT ON (B)(6) 2017 IN THE MORNING. THE REPORT IS RUN TO MAKE SURE THE EXAMS WERE COMPLETE AND THEIR RESULTS CROSSED FROM UNITY PACS TO THEIR EMR SUCCESSFULLY. A MERGE UNITY TECH SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION, REVIEWED THE LOGS AND DETERMINED THAT THE REPORT WAS NOT SAVED. THE PHYSICIAN RE-READ THE EXAM ON (B)(6) 2017 FOR RESOLUTION. WHILE PATIENT IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER 100615 Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ALLEGATION TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | missing exam | loss of data | unavailable for review | report not saved | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A MERGE UNITY PACS ADMINISTRATOR REPORTS THAT AN EXAM REPORT DONE ON (B)(6) 2016 SHOWS AS NOT AVAILABLE WHEN THEY ATTEMPT TO VIEW THE REPORT. THERE IS A 'REPORT SAVE' EVENT IN THE AUDIT TRAIL. THE EXAM IN QUESTION WAS PART OF AN EXAM THAT WAS SPLIT AFTER ACQUISITION. THERE WAS NO EVIDENCE IN THE LOGS OF THE READING STATION THAT SHOW THE REPORT WAS ACTUALLY SAVED. WITH MERGE UNITY PACS NOT OPERATING AS EXPECTED THERE IS A POTENTIAL FOR A DELAY IN TREATMENT IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM HOWEVER, THE CUSTOMER HAS NOT ALLEGED ANY HARM. (B)(4). Manufacturer Narrative: EXAM WAS OPENED ON THE DOMINATOR READING STATION AND NO EVIDENCE EXISTED THAT A REPORT WAS SAVED. THE EXAM WAS SPLIT OUT FROM ANOTHER, THERE IS A POSSIBILITY THIS WAS A RESULT OF WORK FLOW FROM THE READING PHYSICIAN. PACS ADMIN DID GO OVER THE WORK FLOW WITH THE PHYSICIAN AS A PRECAUTION WHEN READING EXAMS THAT ARE SPLIT INTO TWO SEPARATE EXAMS. SITE SOFTWARE WAS UPGRADED FROM (B)(4) ON 1/21/2017. THE EXAM WAS AVAILABLE FOR REVIEW AT ALL TIMES, AND THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. | missing report | loss of data | unavailable report | software issue | treatment and diagnostic delayed |  |  |  | software upgrade |  | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE WAS NOTIFIED THAT THERE WAS A DISCREPANCY BETWEEN THE OBSTETRIC RISK ANALYSIS PERCENTILES FROM THE CUSTOMER'S EXAMFORM AND MERGE'S STRUCTURED REPORT. THE CUSTOMER REPORTED THAT BECAUSE OF THE INCONSISTENCY IN THE NUMBERS, THE PATIENT WAS CONCERNED ABOUT THE WELLBEING OF THE BABY. THE PATIENT WAS REASSURED THAT THE BABY WAS WITHIN THE NORMAL RANGE. HOWEVER, AN INCONSISTENCY IN THE OBSTETRIC RISK ANALYSIS PERCENTILES CAN CAUSE OR CONTRIBUTE TO POTENTIAL HARM. | inconsistent number | data issue | during analyzing | software issue | user confused | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER ON (B)(6) 2017 THAT ALL USERS WERE NOT ABLE TO CONNECT TO THE SERVER. TESTING AND INVESTIGATION BY MERGE HEALTHCARE TECHNICAL SUPPORT DETERMINED THAT THE SITE HAD BEEN EXPERIENCING ONGOING STORAGE ISSUES WITH THEIR DRIVES. IN ORDER TO RESOLVE THIS HARDWARE STORAGE ISSUE, MERGE HEALHCARE SENT THE CUSTOMER SIX (6) NEW ONE TERABYTE (TB) DRIVES. THE NEW CONFIGURATION WOULD ALLOW FOR DISTRIBUTING THE PACS STORAGE ACROSS MULTIPLE DRIVES ENSURING THAT FREE SPACE AND REGULAR MAINTENANCE (DEFRAGGING) WOULD OCCUR. HOWEVER, DURING THE ON-SITE REPLACEMENT INSTALLATION, THE ARRAY DRIVE IN THE IMAGE SERVER FAILED. THE FAILED ARRAY WAS REPLACED. IN THE EFFORTS OF RESTORING THE DATA, FOUR EXAMS WERE UNABLE TO BE RECOVERED. THEREFORE, EACH PATIENT HAD TO BE RE-SCHEDULED FOR ANOTHER MAMMOGRAM EXAM AT A LATER DATE. THERE IS A POTENTIAL FOR PATIENTS THAT REQUIRE RE-IMAGING DUE TO DATA LOSS, TO EXPERIENCE A DELAY IN TREATMENT AND/OR DIAGNOSIS THAT COULD LEAD TO HARM. AT THIS TIME, THERE IS NO INDICATION OF AN ADVERSE EVENT FOR THE AFFECTED PATIENTS. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ISSUE REPORTED BY THE CUSTOMER. A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING DATABASE/DISK ARRAY GUARDIAN ISSUES. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE ARE THREE SIMILAR COMPLAINTS FOR THIS ISSUE. TRENDING WILL CONTINUE TO BE MONITORED. | fail to connect to server | connection issue | loss of data | insufficient storage | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A REPORT WAS PROVIDED WITH A REQUEST FOR HELP WITH STRUCTURED REPORTS (SR) BY A CUSTOMER USING MERGE UNITY PACS WHO WAS ATTEMPTING TO CREATE A TEMPLATE FOR MULTIPLE FETUS EXAMS. THE MERGE UNITY STRUCTURED REPORT (SR) IS NOT SEPARATING FIELDS FOR ULTRASOUND (US) OBSTETRIC (OB) EXAMS WITH TWINS(AKA MULTIPLE FETUS EXAMS). WITH MERGE UNITY PACS NOT REPORTING PATIENT INFORMATION AS EXPECTED, THERE IS A POTENTIAL FOR MISTREATMENT OR MISDIAGNOSIS. THE ISSUE WAS IDENTIFIED DURING CREATION OF THE SR TEMPLATE AND NO PATIENT INVOLVEMENT WAS REPORTED. (B)(4). | fail to create template | application program issue | incorrect patient information | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING PAGES. THE EXAM WAS CONDUCTED ON (B)(6) 2016; REPORT WAS APPROVED AND DICTATED ON (B)(6) 2016. THE ISSUE WAS REPORTED TO MERGE ON (B)(6) 2016 THAT ONLY 3 OF THE 4 PAGES WERE DISPLAYING. THEREFORE, THE REPORT NEEDED TO BE RE-DICTATED BY THE DOCTOR. THE REPORT WAS RE-READ ON (B)(6) 2016 AND APPROVED ON (B)(6) 2016. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: TRENDING OF MISSING REPORT IS CONTINUED TO BE MONITORED AND ADDITIONAL INVESTIGATION ACTIVITIES ARE BEING PERFORMED TO DETERMINE IF CORRECTIVE ACTIONS ARE REQUIRED. | incomplete report | loss of data | during procedure | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT AN EXAM WAS MARKED READ BY THE DOCTOR, HOWEVER NO REPORT WAS FOUND. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT A REPORT ID WAS CREATED BUT NO WORD DOCUMENTS COULD BE FOUND IN THE AUDIT TRAIL PROPERTIES FOR THE AUDIT TRAIL SAVE EVENT. THE REPORT COULD NOT BE RECOVERED AND THE EXAM NEEDED TO BE RE-READ. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING MISSING REPORTS OR REPORTS NOT SAVING. | missing report | loss of data | exams are readed and no report is foiund | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, WHEN THE PRIMARY MONTAGE WAS TOGGLED OFF, THE USER EXPECTED THE MONTAGE TO TOGGLE OFF AND THE PRIMARY IMAGES WOULD BE RETAINED IN THE CURRENT LOCATION. HOWEVER, THE PRIMARY IMAGES SHIFTED ENTIRELY FROM THE SECOND MONITOR TO THE FIRST MONITOR. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE PRIMARY EXAM SHIFTING TO ANOTHER MONITOR WITHOUT ANY INDICATION OF DOING SO HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT IF THE USER IS UNAWARE THAT THE IMAGES HAVE MOVED. (B)(4). Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE FILED. | image shifted | image display error | the primary exam shifts to another monitor | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A 3D MAMMOGRAM EXAM WAS STUCK ON INCORPORATING AND WAS NOT ABLE TO BE MOVED ONLINE TO BE READ. AN INVESTIGATION DONE BY MERGE DETERMINED THAT THE PATIENT'S MAMMOGRAM EXAM HAD 10 VIEWS, WHICH WAS TOO LARGE TO CROSS OVER TO THE PACS SYSTEM. THE EXAM WAS DELETED AND RESENT, THE RECOMMENDED COMPRESSION WAS APPLIED TO THE IMAGES ON THE MODALITY, THE IMAGES WERE SENT ONE AT A TIME, AND IN ADDITION, THE THIRD PARTY MODALITY VENDOR ASSISTED IN ATTEMPTING TO RESEND THE IMAGES ONE AT A TIME WITH NO SUCCESS. MERGE DETERMINED THAT THE IMAGES WERE TOO LARGE, EVEN AFTER THE COMPRESSION AND SENDING THEM ONE AT A TIME. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE RADIOLOGIST DECIDED TO RE-SCHEDULE THE PATIENT FOR A DIFFERENT EXAM, ONE VIEW, 2D MEDIOLATERAL OBLIQUE (MLO) VIEW MAMMOGRAM WITHOUT TOMOGRAPHY. THEREFORE, IMAGES THAT ARE TOO LARGE TO BE SENT FROM THE MODALITY TO THE PACS SYSTEM HAS THE POTENTIAL TO DELAY TREATMENT AND/OR DIAGNOSIS. (B)(4). | fail to upload report | data issue | no report to be viewed | image too large | treatment and diagnostic delayed |  |  |  | compress file to send |  | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT A REPORT WAS CREATED AND ATTEMPTED TO BE SAVED ON (B)(6) 2016 AT 23:31, HOWEVER, THE REPORT DID NOT UPLOAD SUCCESSFULLY TO THE SERVER. IT WAS NOTED THAT THIS REPORT WAS FOR A STAT CT CHEST/ABDOMEN/PELVIS. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO RECOVER THE REPORT OR FIND THE WORD DOCUMENT FILE FOR THE REPORT THAT WAS READ ON (B)(6) 2016. FURTHER INVESTIGATION DETERMINED THAT THE EXAM WAS RE-READ AND THE REPORT WAS REVISED ON (B)(6) 2016. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING MISSING REPORTS OR REPORTS NOT SAVING. MERGE HEALTHCARE IS EVALUATING THE COMPLAINT TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. | fail to upload report | data issue | no report to be viewed | incorrect report status | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING. AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT A REPORT WAS CREATED & SAVED BY THE RADIOLOGIST. HOWEVER THE REPORT WAS MISSING. THE SYSTEM CRASHED BEFORE THE REPORT WAS SAVED TO THE SERVER. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO RECOVER THE REPORT OR FIND THE WORD DOCUMENT FILE FOR THE REPORT. THEREFORE, THE REPORT WAS LOST AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). | missing report | loss of data | no report to be viewed | system crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER ALLEGED THAT A REPORT FOR A STAT EXAM WAS MISSING. AFTER FURTHER INVESTIGATION BY MERGE HEALTHCARE SUPPORT, A REPORT TAB AND A REPORT ID WAS GENERATED WITHIN THE EXAM PROPERTIES. THIS INDICATED THAT A REPORT EXISTED. AFTER REVIEWING THE EVENT LOGS, IT WAS DETERMINED THAT THE REPORT WAS CORRUPTED. HOWEVER, THE MOST PROBABLE CAUSE FOR THE CORRUPTION COULD NOT BE DETERMINED. MERGE SUPPORT WAS UNABLE TO RECOVER THE REPORT. THEREFORE, THE REPORT NEEDED TO BE RE-DICTATED BY THE CUSTOMER. WHEN A CUSTOMER NEEDS TO RE-DICTATE A STAT EXAM IN ORDER TO OBTAIN A REPORT, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT. HOWEVER, THERE WAS NO REPORTED ADVERSE EVENT OR HARM TO A PATIENT. (B)(4). | file corrupted | loss of data | no report to be viewed | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 A CUSTOMER REPORTED THAT AN EXAM READ BY A RADIOLOGIST IN THE MORNING WAS MISSING THE REPORT HOWEVER, THE EXAM SHOWS THAT IT IS IN AN "APPROVED STATUS." DUE TO A USER NOT BEING ABLE TO LOCATE A REPORT FOR A PATIENT, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. WHILE THE EXAM/IMAGES WERE STILL AVAILABLE, THE PHYSICIAN HAD TO RE-READ THE EXAM TO CREATE A REPORT. THERE WAS NO INDICATION OF ANY HARM OR ADVERSE EVENT TO A PATIENT DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: SITE HAS A QC POLICY IN PLACE CHECKING EXAM REPORTS EACH DAY TO MAKE SURE ALL FAXES WENT OUT, REPORTS PRINTED WHEN NECESSARY. THIS INCIDENT WAS A RESULT OF THE UNIVERSAL MANAGER CRASHING AT THE POINT THE RADIOLOGIST WAS CLOSING THE CASE. SITE SOFTWARE WAS UPDATED ON 2/19/2017 FROM R11.0.1 TO R11.0.4. | missing report | loss of data | fail to locate report | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 AT APPROXIMATELY 1:44AM PST, AFTER HOURS SUPPORT WAS CONTACTED BY MERGE FIELD SERVICE ENGINEER. THE ENGINEER WAS ONSITE PERFORMING A PLANNED AFTER HOURS SERVICE AND UPDATING THE FIRMWARE ON THE IMAGE SERVER. AFTER FIRMWARE UPDATE WAS COMPLETED AND THE IMAGE SERVER WAS RESTARTED, THE NETWORK DRIVE WAS NO LONGER AVAILABLE. ON CALL SPECIALIST WORKED THROUGH THE NIGHT WITH THE FIELD SERVICE ENGINEER AND THEY WERE ABLE TO HAVE THE IMAGE SERVER BACK ONLINE AND AVAILABLE THE NEXT MORNING PRIOR TO THE SITE OPENING. WHILE THE IMAGE SERVER WAS ONLINE AND THEY WERE ABLE TO ACQUIRE EXAMS, THERE WERE SEVERAL MOUNT POINTS (STORAGE LOCATIONS) THAT WERE UNAVAILABLE TO ACCESS. THESE MOUNT POINT WHICH CONTAINED EXAMS ACQUIRED PREVIOUSLY. WITH A COMBINATION OF EXAM RESTORING FROM LONG TERM ARCHIVE AND REPAIR, ALL EXAMS THAT ORIGINATED AT THIS SITE WERE RECOVERED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. (B)(4). | loss of connection | connection issue | unable to access | update firmwark fo image server | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED BY A PACS ADMINISTRATOR THAT SERVER MAINTENANCE HAD BEEN PERFORMED AND THE SITE WAS NOT ABLE TO RESTART THE DATABASE. MERGE HEALTHCARE WORKED WITH THE CUSTOMER TO FURTHER INVESTIGATE THE ISSUE. IT HAS BEEN DETERMINED THAT PRIOR TO THE SCHEDULED SERVER MAINTENANCE, THE DATABASE WAS UNEXPECTEDLY POWERED DOWN WHILE TASKS WERE STILL PENDING. THIS UNEXPECTED SHUTDOWN RESULTED SOME LOST DATA. MERGE HEALTHCARE ASSISTED THE SITE AND UTILIZING BACKUPS AND OTHER RESTORATION PROCEDURES THE DATABASE WAS ABLE TO BE RESTORED AND WAS FUNCTIONAL PRIOR TO THE NEXT BUSINESS DAY. THE PRIMARY IMPACT TO THE CLINICAL WORKFLOW WAS A RADIOLOGIST NEEDING TO RE-READ ABOUT 70 EXAMS. THERE ARE 4 REMAINING REPORTS THAT MERGE IS ASSISTING THE SITE WITH RECONSTRUCTING. THE INABILITY TO VIEW REPORTS AND IMAGES MAY RESULT IN A POTENTIAL TO DELAY IN TREATMENT AND/OR DIAGNOSIS A PATIENT. THE CUSTOMER DID NOT INDICATE ANY HARM TO PATIENTS WHOSE DATA MAY HAVE BEEN IMPACTED BY THE UNEXPECTED SHUTDOWN OF THE DATABASE. (B)(4). Manufacturer Narrative: BASED ON THE INVESTIGATION INTO THE DATABASE AND SCHEDULED SYSTEM MAINTENANCE AT THE SITE, THE DATABASE WAS SHUT DOWN A FEW HOURS PRIOR TO THE SCHEDULED SHUT DOWN. THIS WAS NOT SOMETHING THAT MERGE HEALTHCARE HAD ANY CONTROL OVER. THE DEVICE WAS OPERATING AS EXPECTED AND AS INTENDED UNTIL THE POWER WAS UNEXPECTEDLY SHUT DOWN. THIS IS LIKELY THE RESULT OF A USER ERROR. | unable to restart database | device inoperable | during procedure | database shutdown | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT WAS MISSING. THE CUSTOMER STATED THAT THE PHYSICIAN READ THE REPORT AND SAVED IT ON (B)(6). HOWEVER, WHILE DOING AN AUDIT, IT WAS NOTED THAT THIS REPORT WAS MISSING. IT WAS CONFIRMED THAT THE REPORT DID NOT UPLOAD AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A MISSING REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE REPORT IS MISSING. THIS INVESTIGATION IS IN PROGRESS AND NOT YET COMPLETE. | missing report | loss of data | during procedure | incorrect report status | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED BY A PHYSICIAN THAT AN IMAGE WAS MISSING FROM AN EXAM. THE TECHNOLOGIST WAS ABLE TO DISPLAY THE EXAM FROM THE ONLINE LIST AND ALL IMAGES WERE AVAILABLE. MERGE UNITY PACS IS A SERVER BASED SYSTEM THEREFORE DATA ACCESSED FROM ANY TERMINAL OR WORKSTATION IS EXPECTED TO BE EXACTLY THE SAME, AS IT IS COMING FROM ONE SOURCE. HOWEVER, THE INABILITY TO VIEW ALL IMAGES IN AN EXAM HAS THE POTENTIAL TO DELAY OR INADEQUATELY TREAT AND/OR DIAGNOSIS A PATIENT. AN INVESTIGATION INTO THIS ALLEGATION IS IN PROGRESS. (B)(4). | missing image series | image display error | image was missing from an exam | data access issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE WHEN THE PHYSICIAN OPENED THE COMPARISON STUDY, THE STUDY INITIALLY WAS DISPLAYED ON THE CORRECT MONITOR. HOWEVER, WITHIN A SECOND, THE STUDY INDEPENDENTLY MOVED TO A DIFFERENT MONITOR. TO GET THE MONITOR TO DISPLAY THE CORRECT STUDIES, THE PHYSICIAN CREATED A WORK AROUND. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, STUDIES BEING MOVED TO A DIFFERENT MONITOR WITHOUT THE PHYSICIAN BEING AWARE HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. | display to different monitor | incorrect display | physician unware | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT WHEN A USER DID NOT LOG OUT OF A STATION AND ANOTHER USER LOGGED IN, THE ORIGINAL USER'S NAME WAS STILL SHOWN AS LOGGED IN AT THE LOWER RIGHT HAND CORNER OF THE UNIVERSAL MANAGEMENT. THE CUSTOMER DID NOT KNOW WHAT THEIR AUTO LOG OUT FEATURE WAS SET TO BUT REPORTED THAT WHEN A REPORT IS READ, THE REPORT INDICATED THAT THE ORIGINAL USER READ THE REPORT INSTEAD OF THE INTENDED CURRENT USER. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THIS IS A POTENTIAL SAFETY ISSUE SINCE UNAUTHORIZED USERS CAN ACCESS SENSITIVE PATIENT INFORMATION AND MAKE UNNECESSARY CHANGES OR UPDATES. | incorrect login name displayed | application security issue | display previous user name | configuration issue | inappropriate treatment or diagnostic |  |  |  | auto logout setup |  | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, WHEN THE RADIOLOGIST TOGGLED THE MONTAGE FUNCTION ON THE FAR RIGHT OF THE MONITOR, THE POSITIONS OF THE PRIMARY AND PRIOR EXAMS WERE SWITCHED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, SWITCHING VIEWS FROM THE PRIMARY EXAM AND THE PRIOR EXAM WITHOUT ANY INDICATION OF DOING SO HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. Manufacturer Narrative: (B)(4). CONCLUSIONS CODE: - CONCLUSION NOT YET AVAILABLE-EVALUATION IN PROGRESS. | switch view | incorrect display | primary and prior exams were switched | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, AFTER RESENDING A REPORT WITH THE SAME ORDER NUMBER, THE REPORT UPDATED EVERYWHERE EXCEPT FOR THE ZERO-DOWNLOAD AMBASSADOR (ZDA). THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A REPORT NOT UPDATING WITH THE CORRECT INFORMATION HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. | update report with incorrect info | patient data issue | [8 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS USER INFORMED MERGE UNITY SUPPORT THAT A RADIOLOGIST WAS READING AN EXAM AND WENT TO PLAY THE REPORT. THE MERGE UNITY SYSTEM REPORTED THERE WAS NO EXISTING VOICE CLIP. MERGE UNITY TECHNICAL SUPPORT INVESTIGATED AND DETERMINED THAT THERE WAS A "REPORT SAVE" EVENT IN THE AUDIT TRAIL, BUT NO REPORT EXISTED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ AND SAVED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, A MISSING REPORT HAS THE POTENTIAL TO LEAD TO A DELAY PATIENT CARE THAT MAY LEAD TO HARM. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. WHEN PROMPTED TO LISTEN TO THE VOICE REPORT, THE PHYSICIAN RECEIVED A MESSAGE THAT NO VOICE CLIP EXISTED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | missing voice report | loss of data | missing voice clip | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(4) 2017, MERGE HEALTHCARE WAS NOTIFIED FROM THE CUSTOMER THAT AN INCOMPLETE REPORT WAS SENT OUT TO THE REFERRING PHYSICIAN WITHOUT THE MENTION OF A LUNG NODULE ON (B)(6) 2017. A RADIOLOGIST FROM THE CUSTOMER SITE READ A PATIENT'S CASE ON (B)(6) 2017 AND FOUND AN ABNORMALITY IN THE LUNG. THE RADIOLOGIST CALLED THE REFERRING DOCTOR TO INFORM THEM OF THE ABNORMALITY AND RECOMMENDED THAT THE PATIENT HAVE A FOLLOW UP CT. HOWEVER, THE REPORT THAT WAS SENT OUT TO REFERRING DOCTOR HAD NO MENTION OF THE LUNG NODULE IN THE REPORT BODY. ADDITIONALLY, THERE WAS NO INFORMATION REGARDING THE NODULE IN THE CONCLUSION SECTION OF THE REPORT. THE REFERRING DOCTOR'S OFFICE CALLED BACK FOR CLARIFICATION BECAUSE THEY WERE CONFUSED BY THE REPORT WHICH HAD NO MENTION OF THE LUNG NODULE. THE RADIOLOGIST AT THE SITE THEN SPOKE TO THE REFERRING DOCTOR AND ISSUED AN ADDENDUM REPORT WITH COMPLETE INFORMATION. THERE WERE NO KNOWN ADVERSE EVENTS TO THE PATIENT. HOWEVER, SENDING AN INCOMPLETE REPORT TO A REFERRING PHYSICIAN WITH MISSING INFORMATION COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT AND/OR DIAGNOSIS THAT MAY LEAD TO HARM. (B)(4). | incomplete report | loss of data | without mention of a lung module | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT THE TRANSCRIPTIONIST WAS UNABLE TO OPEN THE DICTATION TO TYPE UP THE REPORT. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE UNITY EXAM AUDIT HAS INFORMATION RELATED TO THE REPORT'S SAVE EVENT. AS A RESULT OF NOT BEING ABLE TO FIND THE DICTATION, THE READING PHYSICIAN RE-READ THE EXAM TO RESOLVE THE ISSUE. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. | unable to open report | patient data issue | unable to type up the report | software issue | treatment and diagnostic delayed |  |  |  | resend the exam |  | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A UNITY PACS ADMINISTRATOR CONTACTED MERGE TECHNICAL SUPPORT ALLEGING THAT DURING THE PROCESS OF ADDING IMAGES TO AN EXAM, THE STATUS OF THAT EXAM DID NOT FUNCTION AS EXPECTED. THE EXAM WAS REVIEWED BY A RADIOLOGIST, AND ADDITIONAL IMAGES WERE ORDERED. THE TECHNOLOGIST REVERTED THE EXAM FROM THE VIEWING/REVIEWING "ONLINE" LOCATION BACK TO THE DICOM GATEWAY (ACQUISITION STATION) WHERE ADDITIONAL IMAGES WERE ACQUIRED AND SUBSEQUENTLY ADDED TO THE EXAM. WITH THE ADDITIONAL IMAGES ADDED, THE EXAM WAS THEN MOVED BACK ONLINE FOR VIEWING/REVIEWING. HOWEVER, THE EXAM WAS IN A READ STATUS (R) INSTEAD OF BEING IN THE EXPECTED SCANNED STATUS (S). A SCANNED STATUS INDICATES TO USERS THAT AN EXAM EITHER HAS NOT BEEN READ/REVIEWED OR THAT ADDITIONAL INFORMATION HAS BEEN ADDED TO THE EXAM AND REQUIRES RE-READING. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. AN EXAM WITH ADDITIONAL IMAGES ADDED, AND MOVED TO THE ONLINE LIST AS A "READ" STUDY HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THERE ARE ADDITIONAL IMAGES TO REVIEW. | incorrect study status | data issue | read status but expected status is scanned | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT WHEN ATTEMPTING TO OPEN THE EXAM A MESSAGE IS DISPLAYED ON THE SCREEN THAT SHOWS "NO REPORT". MERGE SUPPORT INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT THERE WAS A SAVE EVENT FOR THE CUSTOMERS' ALLEGED REPORT. HOWEVER, MERGE UNITY TECHNICAL SUPPORT WAS UNABLE TO FIND THE SAVED REPORT FOR THE CUSTOMER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THIS CASE/SITE WAS LINKED TO (B)(4). MISSING REPORTS ISSUE: REPORTS ARE FAILING TO UPLOAD AFTER A POP-UP INDICATES AN ISSUE COMMUNICATING WITH THE SERVER. RETRYING ALSO FAILS. THE JIRA WAS CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION (B)(4). THE SITE WAS UPGRADED TO (B)(4) ON 08/22/2017. | unable to open report | loss of data | show no report | fail to upload | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A MERGE UNITY CUSTOMER (PACS ADMINISTRATOR) ALLEGED THAT A PATIENT'S EXAM, WHEN VIEWED ON THE ZERO DOWNLOAD AMBASSADOR (ZDA), SHOWS A DIFFERENT EXAM AND PATIENT'S REPORT. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THIS CUSTOMER'S ALLEGATION AND DETERMINED THAT THIS ISSUE WAS SEEN WITH TWO DIFFERENT EXAMS FROM THE SAME PATIENT. ON UNITY PACS SYSTEM, THE USER WILL GET A "NO REPORTS FOUND" MESSAGE, BUT ON THE ZDA A REPORT FROM A DIFFERENT PATIENT IS DISPLAYED. WITH THE ZDA NOT FUNCTIONING AS EXPECTED THERE IS A POTENTIAL THAT A CLINICIAN MAY NOT RECOGNIZE THAT THE EXAM AND REPORT THEY REQUESTED TO BE DISPLAYED VIA THE ZDA IS FOR A DIFFERENT PATIENT. THIS MAY LEAD TO A DELAY AND/OR INCORRECT DIAGNOSIS OR TREATMENT FOR A PATIENT WHICH COULD LEAD TO HARM. HOWEVER, THE CUSTOMER HAS NOT INDICATED THAT ANY HARM HAS RESULTED DUE TO THIS ISSUE. (B)(4). | display to different exam | incorrect display | no report to be viewed | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT REPORTS HAVE NOT BEEN GENERATED. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORTS WERE BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. (B)(4). | incorrect study status | data issue | report held in a paused or pending status | approval process issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT REPORTS HAVE NOT BEEN GENERATED. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORTS WERE BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. REFERENCE COMPLAINT NUMBER 83703. | fail to generate report | data issue | report held in a paused or pending status | approval process issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED ISSUES WITH REPORTS NOT APPROVING. THE REPORT STATUS IS STUCK AFTER SIGNING OFF ON A REPORT. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORT WAS BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. REFERENCE COMPLAINT NUMBER (B)(4). | fail to approve report | data issue | report status is stuck after signing off on a report | approval process issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 08/18/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, UPLOADING A REPORT WOULD CAUSE THE DOMINATOR STATIONS TO CRASH. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE STATION CRASHING WHILE UPLOADING A REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, IF THE REPORT DOES NOT FULLY UPLOAD TO THE SERVER BEFORE THE STATION CRASHES. THE REPORT MAY NEED TO BE RE-DICTATED IF THE REPORT IS NOT RECOVERABLE. (B)(4). Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. | fail to upload report | data issue | station crash | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 06/23/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, SOME PATIENT EXAMS WERE NOT VISIBLE FROM THE ONLINE TAB. THROUGH AN INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT EXAMS WERE PROCESSED THROUGH THE DICOM GATEWAY BUT THE DATE OF BIRTH (DOB) WAS NOT FILLED IN ON AN EXAM. ONCE THE DOB WAS ADDED, THE EXAM WAS THEN PULLED FORWARD AND AVAILABLE ON THE ONLINE LIST OF EXAMS. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, PATIENT EXAMS THAT ARE NOT VISIBLE FROM THE ONLINE TAB HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: THIS ISSUE WAS INVESTIGATED BY UNITY SUPPORT AND DEVELOPMENT UNDER (B)(4). ACCORDING TO THE INVESTIGATION, EXAMS CAN BE VIEWED BY GOING TO THE PATIENTS TAB AND LOOKING AT PRIORS FOR THE ASSOCIATED PATIENT. IF AN EDIT IS MADE TO THE EXAM USING THIS PROCESS, THE EXAM SUBSEQUENTLY SHOWS UP WHEN SEARCHING FOR IT IN THE ONLINE TAB. PER THE INVESTIGATION, THIS ISSUE WAS FIXED IN A LATER VERSION 11.1.2. | fail to display report online | no display or display failure | patient exams invisible | missing DOB (day of birth) | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON JUNE 1, 2016 CUSTOMER REPORTED THE DISTANCE INDICATOR BETWEEN IMAGE SLICES ARE NOT REFLECTING WHAT IS DISPLAYED ON THE MODALITY. IT IS INTERMITTENT AND IS NOT PRESENTLY KNOWN IF IT IS LIMITED TO ONE MODALITY OR VENDOR. MERGE HEALTHCARE WAS NOTIFIED BY THE CUSTOMER THAT THEY WERE INFORMED BY AMERICAN COLLEGE OF RADIOLOGY (ACR) THAT THEIR FACILITY DID NOT PASS ACCREDITATION DUE TO SPACING LISTING ON IMAGES NOT CONSISTENT THROUGHOUT THE SERIES. INITIAL EVALUATION OF IMAGES SHOWS THE SAGITTAL IMAGES WITHIN A SPECIFIC SERIES SHOULD BE SPACED 3.9MM APART. THE MODALITY SHOWS THE CONSISTENT SPACING, HOWEVER, THE UNITY PACS VIEWER SHOWS A 4.0MM SPACING BETWEEN 2 OF THE IMAGES. DUE TO THE POTENTIAL FOR THE INCONSISTENCY IN SPACING MEASUREMENTS THAT IS DISPLAYED ON SOME IMAGES, THERE MAY BE CONFUSION WHICH COULD LEAD TO A POTENTIAL DELAY IN DIAGNOSIS OR TREATMENT FOR A PATIENT. THERE IS NO INDICATION THAT ANY PATIENTS WERE HARMED AS A RESULT OF THE DISPLAYED INCONSISTENCY IN SLICE MEASUREMENTS. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | inconsistent measurement value | patient data issue | customer confused | inconsistency in spacing measurement | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. (B)(6) HOSPITAL REPORTED THAT IF 2 IMAGES ARE CAPTURED FOR THE SAME VIEW AND EACH VIEW IS ASSOCIATED TO A TOMO IMAGE, BOTH TOMOS ARE ASSOCIATED TO THE FIRST OF THOSE TWO IMAGES. THE USER MAY MISS SEEING THE ADDITIONAL VIEW WHICH WOULD HELP WITH DIAGNOSIS. WHAT IS EXPECTED IS THE ABILITY TO VIEW TOMO ON FIRST RCC IMAGE AND ABILITY TO VIEW TOMO ON THE SECOND RCC IMAGE. WHAT IS HAPPENING IS THAT BOTH TOMOS ARE ASSOCIATED TO THE FIRST RCC IMAGE. THE SECOND RCC IMAGE HAS NO TOMO OPTION AVAILABLE. RELEASE 11.0 OF UNITY HAS A DEFECT WHICH MISSED THE ASSOCIATE OF THE SECOND TOMO. DESIGN AND TESTING FAILED TO DETECT THERE COULD BE MORE THAN 2 TOMO IMAGES FOR SAME VIEW. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | unable to display multiple images | incorrect display | missing seeing the additional view | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A MERGE UNITY CUSTOMER CONTACTED MERGE (CASE (B)(4)) AND REPORTED A PHYSICIAN WAS READING TOMO AND THE REFERENCE LINES FOR LATERAL TO MEDIAL ARE NOT CORRECT. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE THE INCORRECT DISPLAY OF THE REFERENCE LINES COULD CONTRIBUTE TO DIAGNOSIS ERRORS. AN INVESTIGATION DETERMINED THAT THE ISSUE WAS NOT CAUSED BY MERGE UNITY BUT WAS AN ERROR BY THE IMAGE MODALITY VENDOR, (B)(4) HAD INCORRECT IMAGE POSITION INFORMATION FOR ALL ANGULATED VIEWS IN THE DICOM MESSAGE. | incorrect reference line | image display error | incorrect image position info for all angulated views | image modality vendor | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. CUSTOMER REPORTED THAT A GOOD NUMBERS OF REPORTS HAVE BEEN MISSING HEADERS. REPORTS ARE FOR DISTRIBUTION TO REFERRING PHYSICIANS. IF THE PATIENT DEMOGRAPHICS IS MISSING AND THE REPORT IS DISTRIBUTED VIA FAX OR PRINTER, THE REPORT COULD NOT BE ASSOCIATED TO ANY PATIENT SINCE NO PATIENT DEMOGRAPHICS ARE AVAILABLE. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO THE PATIENT. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | missing header | loss of data | fail to assocate to any patient | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. DIAGNOSTIC IMAGES SERVICES REPORTED THAT MONTAGE WAS DISPLAYING AN INCORRECT TOMO SLICE. THE FIRST TOMO IMAGE ADDED TO THE MONTAGE IS CHANGED WHEN ADDING A SECOND SLICE. (B)(6) RADIOLOGY ALSO REPORTED WHEN ADDING MULTIPLE TOMO IMAGES TO MONTAGE, THE IMAGES SELECTED ARE NOT THE ONES THAT INSERT TO THE MONTAGE. A PHYSICIAN ENCOUNTERED THIS ISSUE WHEN PRESENTING TO OTHER PHYSICIANS. THE MONTAGE THAT WAS CREATED WITH TOMO IMAGES DID NOT HAVE THE INTENDED TOMO IMAGES. DEFECT WAS FOUND IN UNITY PACS 11.0 CODE. CODE WAS NOT HANDLING MONTAGES CORRECTLY WHEN CLOSING EXAMS. (B)(4) Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | incorrect image slice displayed | image display error | during procedure | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER ALLEGED THAT STRUCTURED REPORTS WERE DISPLAYING MEASUREMENT UNITS IN CM VERSUS MM AS SENT FROM THE MODALITY. DUE TO THE ALLEGED LENGTH UNIT MEASUREMENT ISSUE, THIS ISSUE IS BEING REPORTED AS THERE IS A POTENTIAL FOR A MISDIAGNOSIS OR MISTREATMENT OF A PATIENT. THE CUSTOMER HAS NOT REPORTED ANY KNOWN IMPACTS TO A PATIENT. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | inappropriate measurement unit | patient data issue | length unit measurement issue | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, IMAGES WERE SKIPPED WHILE SCROLLING WITH THE MOUSE WHEEL. THE CUSTOMER REPORTED THAT INITIALLY, A FRACTURE WAS MISSED, HOWEVER WAS CAUGHT AFTER FURTHER INVESTIGATION. IMAGES BEING SKIPPED OVER HAS THE POTENTIAL TO LEAD TO AN INCORRECT AND INCOMPLETE DIAGNOSIS AND/OR TREATMENT, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE IMAGE ARE BEING SKIPPED. A FIELD SERVICE ENGINEER CONDUCTED AN INVESTIGATION ON SITE AND NOTED THAT THERE WAS A BROKEN WIRE IN THE CORD, CAUSING THE MOUSE TO CONNECT AND DISCONNECT FROM WINDOWS. THE MOUSE WAS REPLACED. (B)(4). | image skipped | data issue | image was missing from an exam | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 10/10/2015, A CUSTOMER REPORTED ((B)(4)) THAT READING PREFERENCES UNEXPECTEDLY CHANGED, RESULTING WITH "DISPLAY IMAGES ON PRIMARY MONITOR" AND "DISPLAY PRIMARY MONTAGE" ENABLED. THIS CAUSED THE PRIMARY EXAM ON THE 1+3 MONITOR SETUP TO PLACE THE PRIMARY STUDY ON THE COLOR MONITOR WHICH IS HIDDEN FROM VIEW BECAUSE THE ELECTRONIC REQUISITION WAS MAXIMIZED AND COVERED THE PRIMARY EXAM FROM VIEW. THE PHYSICIAN STOPPED WHEN HE NOTICED THAT THE EXAM HE WAS VIEWING AS THE PRIMARY WAS REALLY THE COMPARISON. HAD HE CONTINUED, THE SYSTEM WOULD HAVE ISSUED A POPUP WARNING MESSAGE WHICH LETS THEM KNOW THAT NOT ALL IMAGES HAVE BEEN VIEWED. ALSO, THE NON-PRIMARY IMAGES ARE HIGHLIGHTED "OTHER" BY THE SYSTEM. THIS PROBLEM WAS IDENTIFIED AS A POTENTIAL SAFETY COMPLAINT BECAUSE IT CAUSED USER CONFUSION. NO KNOWN PATIENTS WERE HARMED DUE TO THIS COMPLAINT. THE PHYSICAN NOTICED THE ISSUE BEFORE MAKING THE MISTAKE OF READING THE COMPARISON EXAM THINKING IT WAS THE PRIMARY. Manufacturer Narrative: THIS IS BEING REPORTED AFTER CONDUCTING A RETROSPECTIVE REVIEW. NO DEFECTS WERE FOUND WITH SOFTWARE. THE DIFFERENCE BETWEEN GLOBAL SETTINGS VS USER PREFERENCE SETTINGS CAUSED CONFUSION AND IMPACTED THE PROPER SET-UP OF THE PREFERENCES. SYSTEM WORKS AS DESIGNED. ADDTIONAL TRAINING WAS PROVIDED TO THE REPORTING USER. | display to different monitor | incorrect display | impact the setup of the preferences | difference between global settings vs. user preference settings | user confused | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A MERGE EMPLOYEE ALLEGED THAT THE 3D CALCIUM SCORING REPORT MAY BE INCORRECT OR INACCURATE FOR PATIENTS. SPECIFICALLY, WHEN A GENDER OF OTHER (O) IS SELECTED, THEY HAVE A CALCIUM SCORE GREATER THAN 75 AND/OR NO AGE IS PROVIDED IN THE STUDY. THIS COULD LEAD TO AN INCORRECT CALCULATION OR DELAY IN PATIENT CARE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION TO DETERMINE IF ANY CORRECTION OR CORRECTIVE ACTIONS ARE NECESSARY. | incorrect score calculation | data issue | incorrect 3D calcium scoring | algorithm issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2017, A CUSTOMER REPORTED THAT A PHYSICIAN DICTATED AN EXAM THE PREVIOUS NIGHT BUT WAS NOT ABLE TO FIND THE REPORT ON THE EXAM. MERGE UNITY TECHNICAL SUPPORT'S INVESTIGATION SHOWED THE CUSTOMER'S DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE PHYSICIAN'S WORK STATION LOGS SHOWED THE REPORT SAVE EVENT, WHICH WAS INTERRUPTED BY A WORK STATION CRASH, WHICH CAUSED THE ISSUE. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. | missing report | loss of data | report did not upload to the server | work station crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2017, A PACS ADMIN REPORTED TO MERGE HEALTHCARE TECHNICAL SUPPORT THAT ONE EXAM REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. THE MERGE UNITY SUPPORT INVESTIGATION SHOWED THAT THE CUSTOMER'S DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE REPORT'S AUDIT TRAIL SHOWED THAT THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4) Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | fail to upload to server | data issue | unavailable report | approval process issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A UNITY PACS ADMINISTRATOR REPORTED TO MERGE TECHNICAL SUPPORT ON (B)(6) 2018 THAT AN INVALID IMAGE FORMAT WAS FOUND ON AN EXAM THAT WAS PREVIOUSLY READ BY A RADIOLOGIST. UPON INITIAL REVIEW ON APRIL 27, 2018 IT WAS BELIEVED THIS ISSUE WAS RELATED TO THE SITE'S SUPPLIED STORAGE VENDOR. THE COMPLAINT REMAINED OPEN FOR ADDITIONAL INFORMATION BUT WAS MARKED AS "NOT REPORTABLE". ON MAY 18, 2018 NEW INFORMATION BECAME AVAILABLE. THE ISSUE WITH THE INVALID IMAGE FORMAT WAS DETERMINED TO BE UNITY SOFTWARE RELATED. WHILE THE EXAM REPORT IS AVAILABLE, THE INABILITY TO RESTORE AN EXAM FROM LONG TERM STORAGE FOR VIEWING COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THE INVESTIGATION INTO THIS ISSUE IS ACTIVE AND ONGOING. UNTY PACS HAS DISCOVERED HD ARCHIVE AND BURN TO CD ARE CORRUPTING IMAGES WITH LARGE LOOK UP TABLES (LUTS). IN THE CASE OF HD ARCHIVE, THE CORRUPTION GOES UNNOTICED UNTIL THE IMAGES ARE RESTORED. SOME IMAGES CAN BE RESTORED WHILE OTHER IMAGES CANNOT BE VIEWED (INVALID IMAGE). THERE ARE EXAM REPORTS AVAILABLE FOR REVIEW FOR ALL EXAMS. MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. ON MAY 23, 2018 THE LEAD UNITY SUPPORT TECHNICIAN CONTACTED THE DESIGNATED COMPLAINT HANDLING UNIT (DCHU) WITH ADDITIONAL INFORMATION LEADING DCHU TO SUBMIT THIS ISSUE TO FDA. | invalid image format | data issue | unable to open image | supplied storage vendor | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER REPORTED THAT NON-ARCHIVED EXAMS HAD ACCIDENTALLY BEEN DELETED. THE CUSTOMER INDICATED THAT THE SYSTEM DID NOT GIVE ANY WARNING WHILE EXAMS THAT WERE NOT ARCHIVED WERE DELETED. WITH MERGE UNITY PACS NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR PATIENT DATA TO BE UNINTENTIONALLY DELETED WHICH MAY RESULT IN DATA USED FOR DIAGNOSIS OR TREATMENT DECISIONS BEING UNAVAILABLE. HOWEVER, THERE WAS NO SPECIFIC REPORTED PATIENT INVOLVEMENT OR HARM AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | data deleted | loss of data | unavailable exams | user error | treatment and diagnostic delayed |  |  |  | delete control | delete alert | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THERE ARE MISSING STANDARD UPTAKE VALUES (SUVS) IN THE DICOM HEADER FROM A PET MODALITY. THE MISSING VALUES ARE CAUSING PROBLEMS WHEN THE RADIOLOGIST REVIEWS THE SUV ON PET EXAMS. DUE TO MERGE UNITY PACS NOT PRESENTING VALUES AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS THAT MAY LEAD TO HARM. HOWEVER, THERE WAS NO REPORTED ADVERSE EVENT OR HARM TO A PATIENT DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: DCHU SPOKE WITH PACS ADMINISTRATOR ON (B)(6) 2018. THE JIRA FOR THIS ISSUE HAS BEEN CLOSED AND VERIFIED AS BEGING RESOLVED IN R11. PACS ADMINISTRATOR REPORTS THIS SITE IS UPGRADING TO R11 ON (B)(6) 2018. UPGRADE TAB IN SALES FORCE CONFIRMS THIS SITE IS SCHEDULED TO GO DOWN AT 8:00 AM ON (B)(6) 2018 FOR UPGRADE TO R11.1.2.6. | incomplete dicom header | data issue | missing standard uptake values | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEIR MERGE UNITY PACS SYSTEM WAS COMPLETELY DOWN. TROUBLESHOOTING AND TECHNICAL SUPPORT IS ONGOING BETWEEN MERGE HEALTHCARE AND THE CUSTOMER. WHEN MORE INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | unintended system shut down | application program issue | during procedure | power issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED VIA APPLICATION LOGS THAT THE REPORT NEVER UPLOADED BECAUSE THE C DRIVE ON THE LOCAL STATION WAS FULL, WHICH PREVENTED THE REPORT FROM SAVING. SUPPORT ASSISTED THE CUSTOMER IN RE-GENERATING THE REPORT AND RECOMMENDED REVIEWING THE C DRIVE OF THE LOCAL STATION. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | fail to read report | data issue | report never uploaded | C drive on the local station was full | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, THE CUSTOMER REPORTED RECEIVING A "NO OBJECT FOUND" ERROR WHEN TRYING TO OPEN A REPORT FOR DICTATION TO IBM WATSON HEALTH IMAGING ON (B)(6) 2020. TROUBLESHOOTING FROM IBM WATSON HEALTH IMAGING SUPPORT FOUND THAT THE REPORT FOLDER WAS EMPTY, CAUSING THE "NO OBJECT FOUND" ERROR, AND THAT THERE WAS A REPORT SAVE EVENT PRESENT IN THE DEVICE AUDIT TRAIL. IBM WATSON HEALTH IMAGING SUPPORT REGENERATED THE REPORT TEMPLATE AND SAVED IT IN DICTATED (DID) STATUS. THE REPORT WAS SAVED WITH THE NEW TEMPLATE. THE REPORT WAS THEN READABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | unable to open report | loss of data | no object found | report folder was empty | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM REPORT IS CORRUPT. THE MERGE HEALTHCARE UNITY TECH'S INVESTIGATION SHOWED THAT THE REPORT DID SAVE SUCCESSFULLY, HOWEVER, THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. THE EXAM REPORT WAS SUCCESSFULLY SAVED IN UNITY PACS. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | file corrupted | data issue | word document corrupted | fail to upload to the image server | treatment and diagnostic delayed |  |  |  |  | display error message | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM WAS COMPLETED BUT THERE IS NO EXAM REPORT. UNITY SUPPORT INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISSING REPORT DOES HAVE THE POTENTIAL TO DELAY PATIENT CARE. REFERENCE SALES FORCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. PENDING FURTHER REVIEW. A SUPPLEMENTAL REPORT WILL BE PROVIDED. | missing report | loss of data | exam was completed but there is no exam report | approval process issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE CUSTOMER REPORTED THAT THE EXAM WAS STUCK AND THE REPORT WAS NOT AVAILABLE. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY REVIEWING THE MERGE UNITY PACS WORKSTATION. THE AUDIT TRAIL OF THE EXAM SHOWED THE ORIGINAL EXAM WAS READ ON THE SAME DAY AS THE EXAM DATE (B)(6) 2018 HOWEVER THE USER DID NOT SAVE THE REPORT AND THAT THERE WAS NO SAVE REPORT EVENT ON (B)(6) 2018 IN THE LOG. THERE WAS NO DATA UPLOADED AND THEREFORE THE REPORT WAS NOT AVAILABLE. THE EXAM WAS REQUIRED TO BE RE-READ BY THE PHYSICIAN. THE EXAM WAS RE-READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER: (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. THE EXAM AUDIT TRAIL DID NOT SHOW A SAVED REPORT EVENT. IT WAS NOT POSSIBLE TO DETERMINE IF THE ERROR MESSAGE DISPLAYED ON THE SCREEN FOR THE DOCTOR WHEN THE REPORT WAS FIRST READ AND SAVED NOTIFIED THE FAILURE TO UPLOAD. THERE WAS NO EVIDENCE THAT REPORT WAS SAVED. THE UNITY SUPPORT TECHNICIAN'S INVESTIGATION SHOWED THAT THE REPORT WAS NOT UPLOADED AND SAVED. AS A RESULT OF FINALIZED REPORT BEING UNAVAILABLE, THE EXAM WAS RE-READ BY A PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | file stuck | data issue | report was not available | user did not save the report | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 4/13/2018, THE CUSTOMER REPORTED THAT THE REPORT WAS MISSING FOR AN EXAM (DATED (B)(6) 2018) THAT WAS DICTATED BY THE DOCTOR ON (B)(6) 2018. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY REVIEWING THE MERGE UNITY PACS WORKSTATION. TECHNICAL SUPPORT DID IDENTIFY A STATION CRASH AND THERE WAS NO DATA UPLOADED AND THEREFORE THE REPORT WAS NOT AVAILABLE. THE EXAM WAS REQUIRED TO BE RE-READ BY THE PHYSICIAN. THE EXAM WAS RE-READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. STATION LOGS WERE UNAVAILABLE FOR REVIEW. IT WAS NOT POSSIBLE TO DETERMINE IF THE ERROR MESSAGE DISPLAYED ON THE SCREEN FOR THE DOCTOR WHEN THE REPORT WAS FIRST READ AND SAVED NOTIFIED THE FAILURE TO UPLOAD. THERE WAS NO EVIDENCE THAT REPORT WAS SAVED. THE UNITY SUPPORT TECHNICIAN'S INVESTIGATION SHOWED THAT THE UNITY PACS STATION CRASHED BEFORE THE REPORT WAS UPLOADED AND SAVED. AS A RESULT OF FINALIZED REPORT BEING UNAVAILABLE, THE EXAM WAS RE-READ BY A PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | missing report | loss of data | no data available | workstation crash | treatment and diagnostic delayed |  |  |  |  | display error message | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE UNITY CUSTOMER REPORTED THEY ARE UNABLE TO VIEW AN EXAM FROM 2013 FOR USE AS A COMPARISON STUDY. THEY ARE MET WITH A MESSAGE WHEN TRYING TO OPEN THIS PARTICULAR STUDY. ERROR READING EXAM DATA FIELD, EXAM CANNOT BE VIEWED. UNITY TECH INVESTIGATION SHOWED THERE WAS AN ONLINE RECORD FOR THIS EXAM AND ALTHOUGH THE DATABASE SHOWED AN OWNING LOCATION, THERE WAS NO IMAGE FOLDER THERE NOR WAS THERE ANY PATIENT FILE. AUDIT TRAIL DOES NOT EXPLAIN WHAT COULD HAVE HAPPENED. EXAM WAS NEVER ARCHIVED AND NOW APPEARS THAT THE IMAGE DATA WAS LOST ALONG THE WAY. THE READING PHYSICIAN DOES HAVE A PHYSICAL REPORT FROM THIS 2013 EXAM FOR REFERENCE. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISSING EXAM DOES HAVE THE POTENTIAL TO DELAY PATIENT CARE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | unable to display report | incorrect display | loss of data | no image folder at database | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS ADMIN REPORTED THEY HAVE 4 EXAMS MARKED AS "READ" BY THE RADIOLOGIST BUT NO REPORT WAS SAVED. THE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE RADIOLOGIST. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | incorrect file status | data issue | report was not available | report did not unpload to server | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS ADMINISTRATOR REPORTED THAT THEY HAVE AN EXAM ONLINE. THE EXAM IS MARKED AS READ BUT DOESN'T LIST A REPORT STATUS OR WHO DICTATED THE REPORT. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THAT THE CUSTOMER'S THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AN AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE READING STATION CRASHED WHILE UPLOADING THE REPORT. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT COMPLETE THE UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. LOGS REVEALED THE STATION CRASHED WHILE UPLOADING THE EXAM REPORT TO THE IMAGE SERVER. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | incorrect file status | loss of data | report did not upload to the server | station crashed while uploading report | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, MERGE HEALTHCARE INTERNALLY FOUND THAT WHEN ADDING AN ADDITIONAL BIOPSY WITHOUT ATYPICAL HYPERPLASIA TO A PATIENT THAT PREVIOUSLY HAD A BIOPSY WITH ATYPICAL HYPERPLASIA, IT WILL OVERWRITE THE CONDITION AND REMOVE THE ATYPICAL HYPERPLASIA COMPLETELY. THIS ISSURE COULD RESULT IN AN INCORRECT GAIL RISK CALCULATION, POTENTIALLY RESULTING IN A DELAY IN TREATMENT AND/OR INCORRECT TREATMENT FOR A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ISSUE FOUND INTERNALLY AND RECENTLY SUBMITTED A CORRECTION TO THE FDA ON 25JUL2017. (2183926-07/25/2017-011-C). ONLY CUSTOMERS WITH A LICENSE FOR MAMMO TRACKING ARE IMPACTED IN RELEASES 10.0.6 AND 10.0.7. | data overwritten | loss of data | incorrect gail risk calculation | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE TECHNICAL SUPPORT WITH QUESTIONS ABOUT HOW THE MERGE UNITY BREAST RISK ASSESSMENT CALCULATES THE RISK. THE CUSTOMER REPORTED THAT A PATIENT HAD TWO PREVIOUS NEGATIVE EXAMS THAT AND THE PATIENT'S FAMILY HAS NO HISTORY OF BREAST CANCER. HOWEVER, WHEN PERFORMING THE RISK ASSESSMENT IN MERGE UNITY FOR THIS PARTICULAR PATIENT, THE LIFETIME RISK LEVEL WAS NOT WHAT THE CUSTOMER EXPECTED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISCALCULATION REGARDING THE POTENTIAL RISK OF BREAST CANCER COULD RESULT IN A PATIENT NOT RECEIVING THE RECOMMENDED TREATMENT AND/OR RECEIVING TREATMENT THAT IS NOT NEEDED. (B)(4). Manufacturer Narrative: UPON RESEARCHING THE CUSTOMER'S QUESTION REGARDING THE CALCULATIONS OF RISK, MERGE HEALTHCARE IDENTIFIED AN ISSUE WHERE ATYPICAL HYPERPLASIA MAY NOT BE FACTORED INTO THE BREAST RISK CALCULATION. A MISCALCULATION REGARDING POTENTIAL RISK COULD RESULT IN A PATIENT NOT RECEIVING THE RECOMMENDED TREATMENT AND/OR RECEIVING TREATMENT THAT IS NOT NEEDED. MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ALLEGATION FROM THE CUSTOMER AND IS PLANNING A CORRECTION (2182926-07/21/2017-010-C). . | incorrect calculation | incorrect software programming calculations | incorrect breast risk assessment calculation | HYPERPLASIA MAY NOT BE FACTORED INTO THE BREAST RISK CALCULATION | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS USER INFORMED MERGE UNITY SUPPORT THAT A RADIOLOGIST WAS COMPLETING A REPORT IN MERGE REPORTING AND RECEIVED A SERIES OF ERROR MESSAGES. THE ERRORS WERE OBSERVED BY THE PHYSICIAN. THE SITE ATTEMPTED TO RESOLVE THIS ISSUE INTERNALLY. AFTER THE CUSTOMER CONTACTED MERGE SUPPORT AN INVESTIGATION SHOWED THAT THE REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | error message | device alarm system issue | fail to upload to image server | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER CONTACT REPORTED ON (B)(6) 2019 THERE WERE FIVE EXAMS WITH CORRUPTED REPORTS FROM (B)(6) 2019 FROM TWO DIFFERENT READING STATIONS. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES WERE AVAILABLE, A CORRUPTED REPORT COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY LOGS FOUND NO ERRORS INDICATING THE FAILURE OF THE REPORTS TO UPLOAD, HOWEVER IT WAS NOTED THAT UNIVERSAL MANAGER (UM) CRASHED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THE EXAMS IN THE FAILED RESULTS OUT QUEUE. THE EXAMS WERE ALL READ AND APPROVED IN APPROVED (APD) STATUS WITH REPORT ID'S. THE REPORTS FOLDER WAS MANUALLY CHECKED, HOWEVER, NONE OF THE REPORTS HAD REPORT FOLDERS. THE REPORT STATUSES WERE UPDATED TO DICTATED (DID) AND RESENT TO THE USER RE-DICTATION. UNITY SUPPORT CONFIRMED THAT ALL FIVE REPORTS WERE RE-READ SUCCESSFULLY. NO FURTHER INVESTIGATION WILL BE PERFORMED. | file corrupted | loss of data | unavailable for review | system crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER CONTACT REPORTED ON (B)(6) 2018 THAT A "NO OBJECT FOUND" ERROR WAS NOTED WHEN TRYING TO VIEW A REPORT FOR A SPECIFIC MAMMOGRAM EXAM, SUBSEQUENTLY, THE STATION KEPT HAVING TO BE RESTARTED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I:\ DRIVE. A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED. THE EXAM WAS RE-READ AND CONFIRMED WITH THE CUSTOMER THAT THE REPORT WAS PRESENT. NO FURTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | error message | device alarm system issue | unavailable report | fail to upload to the image server | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: MERGE UNITY PACS, A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(4) 2017, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE OVER THE WEEKEND OF (B)(6), AN ERROR WAS RECEIVED WHEN ATTEMPTING TO BURN IMAGES ONTO A COMPACT DISC (CD). THE CUSTOMER WAS NOT ABLE TO BURN IMAGES OF A COMPUTED TOMOGRAPHY (CT) OF THE HEAD/NECK ONTO A CD FROM THEIR RIMAGE BURNER, WHICH WAS NEEDED FOR A PATIENT BEING AIR-LIFTED TO ANOTHER FACILITY. A PATIENT WAS BEING EVALUATED IN THE EMERGENCY DEPARTMENT FOR AN ATTEMPTED SUICIDE. IT WAS REPORTED THAT THE PATIENT WAS UNRESPONSIVE AND INTUBATED. THE CUSTOMER REPORTED THAT THE PATIENT WAS AIR LIFTED TO A PEDIATRIC TRAUMA CENTER IN LUBBOCK, TX IN GRAVE CONDITION. AS A RESULT OF THE CUSTOMER NOT BEING ABLE TO BURN IMAGES ONTO THE CD, THE CUSTOMER REPORTED THAT ANOTHER CT HAD TO BE PERFORMED AT THE HOSPITAL RECEIVING THE PATIENT. THERE WERE NO KNOWN ADVERSE EVENTS TO THE PATIENT. HOWEVER, HAVING TO RE-DO A CT OF THE HEAD/NECK CAN DELAY PATIENT CARE AND/OR DIAGNOSIS. IN ADDITION, THIS CAN EXPOSE THE PATIENT TO ADDITIONAL UNNECESSARY RADIATION. (B)(4). Manufacturer Narrative: TESTING AND INVESTIGATION BY MERGE SUPPORT DETERMINED THAT THE ERROR "DEVICE: AUTOLOADER 1, HID-0 ERROR CODE: 363 FAILED TO START LABEL RE-TRANSFER ON EVEREST PRINTER; PRINTER SENSE CODE = 254 (UNDEFINED SENSE CODE)" WAS LIKELY RELATED TO THE LABEL/RIBBON ASSEMBLY. THE CUSTOMER FOUND THAT THE COLOR RIBBON (A PART OF THE LABEL/RIBBON ASSEMBLY) WAS OUT. ONCE THE COLOR RIBBON WAS CHANGED, THE ISSUE WAS RESOLVED. A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING BURNING IMAGES ONTO A CD. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE IS NO SIGNIFICANT TREND. TRENDING WILL CONTINUE TO BE MONITORED. | unable to burn image on CD | data back-up problem | during procedure | hardware issue: color ribbon out | additional unnecessary radiation | [5 empty] | FDA2
Event Description: NO KNOWN EVENT OCCURRED. PROBLEM: THE IMPORT OF ENHANCED ULTRASOUND FORMATTED DATA (NEMA DICOM SUPPLEMENT 43: STORAGE OF 3D ULTRASOUND DATA, VERSION: FINAL TEXT - (B)(6) 2009), CAN RESULT IN INCORRECT DISPLAY OF IMAGE ORIENTATION. THE REASON WAS AN INCORRECT UNDERSTANDING OF THE DATA FORMAT. THIS FAULT DOES NOT RESIDE IN THE TRUE 3D RENDER ENGINE, BUT RATHER IS LOCALIZED TO THE DICOM INPUT TASK OF THE APPLICATION PROBLEM: THE IMPORT OF PHILIPS 3DDCM FORMATTED ULTRASOUND IMAGES CAN RESULT IN INCORRECT DISPLAY OF IMAGE ORIENTATION. THE REASON WAS AN INCORRECT UNDERSTANDING OF THE DATA FORMAT. THIS FAULT DOES NOT RESIDE IN THE TRUE 3D RENDER ENGINE, BUT RATHER IS LOCALIZED TO THE DICOM INPUT TASK OF THE APPLICATION. ISSUE WAS IDENTIFIED ON RELEASE VERSIONS: 1.6.1 AND 1.6.2. | image formatted | incorrect display | unavailable images | enhanced ultrasound formated image | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. CONTEXT: INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM WITH THE MAGNIFICATION GLASS TOOL WHICH, IN SOME INSTANCES, DISPLAYS A DIFFERENT IMAGE THAN THE ONE CURRENTLY BEING SHOWN IN THE SELECTED VIEWPORT. IT APPEARS THAT THE TOOL IS DISPLAYING A MAGNIFIED IMAGE FOR A DIFFERENT SLICE OF THE SERIES. THIS COULD POTENTIALLY DELAY A DIAGNOSIS. THIS PROBLEM WAS REPORTED BY AN INTELERAD EMPLOYEE WHILE AT A CLIENT SITE. INVESTIGATION: (B)(4). THE SORTING ORDER OF AN IMAGE STACK IS NOT TAKEN INTO ACCOUNT BY THE MAGNIFICATION GLASS TOOL. IF THE USER IS USING REVERSE ORDERING (EITHER MANUALLY SELECTED OR THROUGH A PREFERENCE IN THE CASE OF TOMOSYNTHESIS) AND THEN USES THE MAGNIFICATION GLASS TOOL, THE IMAGE SHOWN IN THE TOOL WILL NOT BE THE ONE IN THE CURRENT VIEWPORT BUT RATHER THE CORRESPONDING POSITION FROM THE OPPOSITE END OF THE STACK. THIS DEFECT APPLIES TO ANY MODALITY WITH A STACK OF IMAGES, EX. TOMOSYNTHESIS, CT. THE DISCREPANCY IS HIGHLY DETECTABLE BY USER AND IS EXTREMELY UNLIKELY TO LEAD TO HARM TO THE PATIENT. DESCRIPTION OF HAZARDOUS SITUATION IN USE OF THE DEVICE: INCORRECT IMAGE BEING SHOWN WHILE USING THE MAGNIFICATION TOOL, WHICH COULD CAUSE A LOW POSSIBILITY OF DELAY IN TREATMENT. THE INCORRECT IMAGES DISPLAYED ARE HIGHLY EVIDENT TO RADIOLOGISTS. THIS PROBLEM IS NOT CURRENTLY DOCUMENTED AS A KNOWN RISK. A MITIGATION IS CURRENTLY IN PLACE TO REDUCE THIS RISK AND TO PREVENT RISK TO PATIENT. THE RISK CAN BE FURTHER REDUCED THROUGH ADDITIONAL INTERNAL SOFTWARE CHECKS TO PREVENT INCORRECT IMAGE DISPLAY WHEN INCONSISTENCIES ARE DETECTED. PROBABILITY OF RISK: REMOTE (2) SEVERITY OF RISK: IMPROBABLE (1) DETECTABILITY (USER AWARENESS. IF DEVICE FAILURE OCCURS, IS IT EASILY RECOGNIZED BY USER): HIGHLY DETECTABLE BY USER. RISK VS. BENEFIT COMPARISON: BENEFIT OF AN INCREASE IN AVERAGE TIME SAVINGS, DIAGNOSTIC ACCURACY, DATA INTEGRITY AND IMAGE ACCESSIBILITY OUTWEIGHS THE RISK OF LOW PROBABILITY EVENTS LEADING TO DELAYED DIAGNOSIS. THE PROBLEM IS ONLY CAUSED IN MULTI IMAGE SITUATIONS WHERE THE STACK ORDER IS REVERSED AND THE MAGNIFICATION GLASS TOOL IS USED. THE SOFTWARE FAILURE IS VERY EVIDENT AND EXTREMELY UNLIKELY TO LEAD TO HARM TO PATIENT. THE FEATURE HAS BEEN IN USE FOR EIGHT YEARS WITH ONLY A SINGLE REPORTED PROBLEM, WITH NO INDICATION OF REAL HARM. CORRECTIVE ACTIONS: INTELERAD WILL: 1) IDENTIFY ALL CLIENTS AFFECTED (19 SEPTEMBER 2017). 2) INFORM ALL CONCERNED CLIENTS OF THE ISSUE (SEPTEMBER 2017). 3) PROVIDE ALL CLIENTS WITH A SOFTWARE UPGRADE (INSTALLATION DATE AS AGREED WITH CLIENTS). ESTIMATED COMPLETION DATE OF THE PROPOSED RECALL ACTION IS JUNE 2018. | display to different image | incorrect display | display magnified image for a different slice | did not consider about the sort order of an image | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. FIVE RELATED CUSTOMER COMPLAINTS ((B)(4)) WERE REGISTERED AND INVESTIGATED FROM OCT 2014 TO JUNE 2015 FOR RELATED ISSUES. THE INVESTIGATION OF THESE COMPLAINTS LED TO THE DISCOVERY OF ISSUES WITH IMAGE INGESTION. BUGS IDENTIFIED: (B)(4) LDS: LDS: FAILED TO INGEST IMAGE WITH OVERLAYS BELONGING TO 60XX GROUP HAVING ALPHANUMERIC ID (IE, 600A) (B)(4) LDS: ROWS STUCK IN INCOMING_FILES DUE TO FILES MISSING IN THE INCOMING FILES FOLDER. INVESTIGATION: INVESTIGATION OF THE LOGS IN (B)(4) REVEALED THAT SOME OF THE EXPECTED OVERLAY FILES WERE NOT FOUND ON THE FILE SYSTEM; THIS LED TO DISCOVERY OF THE MISSING FILES LEFT BEHIND AT AN EARLIER STAGE OF IMAGE INGESTION. ROOT CAUSE OF THE PROBLEM: DUE TO THE ASYNCHRONOUS NATURE OF IMAGE INGESTION PROCESSING, A FAILURE AT THE EARLY STAGES OF THE INGESTION PROCESS CAN LEAD TO IMAGES NOT BEING VISIBLE TO THE END USER. THERE IS NO INDICATION TO THE USER THAT THIS HAS OCCURRED. REASON FOR POSSIBLE PATIENT SAFETY: THE RADIOLOGIST MAY PROVIDE A DIAGNOSIS WITHOUT BEING AWARE THAT IMAGES ARE MISSING. DESIGN MITIGATION TO REDUCE THE RISK: PROPOSED NEW MITIGATION: THE PACS WILL REPORT PROBLEMS INGESTING IMAGES AND INDICATE SUCH ISSUES TO THE USER ((B)(4)). THIS MITIGATION WILL BE INCLUDED IN A COMING SOFTWARE RELEASE. (B)(4): LDS: EXPOSE BROKEN SERIES TO USER (WARN ABOUT ADDITIONAL IMAGES PENDING INGESTION WHEN VIEWING IMAGES). DETECTABILITY OF PROBLEM: UNTIL THE SOFTWARE UPDATE THAT WILL INCLUDE THE PROPOSED MITIGATION IS INSTALLED ON THE CLIENT'S SYSTEM, THE USER MAY NOT BE AWARE THAT SOME IMAGES MAY BE MISSING. HOWEVER, IF THE CLIENT'S SYSTEM HAS OTHER MECHANISMS TO VERIFY IMAGE COUNT BETWEEN SOURCE AND DESTINATION, THEN THE USER MIGHT BE ABLE TO KNOW ABOUT THE DISCREPANCY. Manufacturer Narrative: IN ORDER TO CORRECT THIS ISSUE, INTELERAD WILL: PROVIDE THE AFFECTED CLIENTS WITH A LIST OF STUDIES THAT WERE AFFECTED AND POSSIBLY RECOVER THE MISSING IMAGES. DEPLOY A SOFTWARE UPDATE THAT INCLUDES A FIX AND MITIGATION TO ALL CLIENTS WHERE THE ISSUE COULD POTENTIALLY OCCUR. THIS WILL BE DONE TENTATIVELY. PERIODICALLY AUDIT ALL AFFECTED CLIENT, ANY NEW OCCURRENCES OF THIS ISSUE UNTIL THE MITIGATION (SOFTWARE UPDATE) IS IN PLACE. THE SOFTWARE UPDATE WILL INCLUDE ONLY BUG FIXES - YOUR WORKFLOW WILL NOT BE AFFECTED. THE UPDATE WILL RESULT IN ABOUT 5 MINUTES OF DOWNTIME PER SERVER AS WE PERFORM A FULL SERVICE RESTART. IMPACT ON USERS WILL BE MINIMAL, IF ANY. CORRECTIVE ACTION CALENDAR: FOR ALL CLIENTS WITH PACS VERSION 4-5-1 OR HIGHER, THE SOFTWARE UPDATE WILL BE DEPLOYED WITHIN A TIMEFRAME TO BE AGREED WITH CLIENTS (COMPLETION ESTIMATED BY END OF MARCH 2016). | missing image series | loss of data | user not aware that some images be missing | software bug: image ingestion processing | inappropriate treatment or diagnostic |  |  |  |  | provide indication for the failure | FDA2
Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. ONE CUSTOMER (B)(4) WAS REGISTERED AND INVESTIGATED. BUGS IDENTIFIED: (B)(4): (REVERT OF (B)(4)) IV: RADIOGRAPHIC SCALING FACTOR DERIVED FROM DISTANCE TO PATIENT AND DETECTOR IS BEING APPLIED TO NON-PROJECTION RADIOGRAPHY MODALITIES (B)(4): IV (MEASUREMENT): MEASUREMENTS ARE NOT CALIBRATED WHEN DICOM IS MISSING MAG FACTOR, BUT HAS DISTANCE TO PATIENT AND DETECTOR PLATE. INVESTIGATION: THERE IS INDEED A PROBLEM WITH THE MEASUREMENTS AFTER THE FIX INTRODUCED IN (B)(4). IV NOW USES THE RATIO OF THE DISTANCE FROM THE SOURCE TO PATIENT OVER DISTANCE FROM SOURCE TO DETECTOR TO COMPUTE A SCALING FACTOR TO APPLY TO MEASUREMENTS, BUT THAT SCALING SHOULD ONLY BE APPLIED WHEN DEALING WITH A PROJECTION RADIOGRAPHY MODALITY (CR, DX, MG, PX, RF, OR XA). UNFORTUNATELY, IT IS NOW BEING APPLIED TO ALL MODALITIES. THERE IS A GUARD IN THE CODE THAT PREVENTS COMMENTS ABOUT THE SCALING FACTOR TO BE DISPLAYED ON MODALITIES THAT ARE NOT PROJECTION RADIOGRAPHY, BUT THE SCALING FACTOR ITSELF IS APPLIED IN ALL CASES, WHICH IS WRONG. ROOT CAUSE OF THE PROBLEM: SINCE THE INTRODUCTION OF A FIX TO THE MEASUREMENTS ALGORITHM ((B)(4)), INTELEVIEWER NOW USES THE RATIO OF THE DISTANCE FROM THE SOURCE TO PATIENT OVER DISTANCE FROM SOURCE TO DETECTOR TO COMPUTE A SCALING FACTOR TO APPLY TO MEASUREMENTS WHEN THE ESTIMATED RADIOGRAPHIC MAGNIFICATION FACTOR (DICOM (0018,1114)) IS NOT PRESENT. THAT SCALING SHOULD ONLY BE APPLIED WHEN DEALING WITH A PROJECTION RADIOGRAPHY MODALITY (CR, DX, MG, PX, RF, OR XA), HOWEVER, IT ENDED UP BEING APPLIED TO OTHER MODALITIES, SUCH AS CT, GENERATING ERRONEOUS MEASUREMENTS. REASON FOR POSSIBLE PATIENT SAFETY: THE RADIOLOGIST COULD BASE THEIR DIAGNOSIS ON INCORRECT ANATOMICAL MEASUREMENTS. RISK ANALYSIS: THIS PROBLEM IS DOCUMENTED AS A KNOWN RISK AND IS MITIGATED BY DESIGN: CONSISTENT USE OF MAPPING FUNCTIONS TO PATIENT ANATOMICAL SPACE AND CALIBRATED UNITS BEFORE PERFORMING COORDINATE AND INTENSITY CALCULATIONS. FOR CLIENTS USING THIS SPECIFIC SOFTWARE VERSION, THIS PROBLEM COULD WILL OCCUR WITH APPROXIMATELY 50% OF ALL CT STUDIES. RISK TO THE PATIENT COULD OCCUR ONLY IF THE MODALITY IS CT AND IF THE DISCREPANCY OF THE MEASUREMENT IS SMALL. IN ORDER TO CORRECT THIS ISSUE, INTELERAD WILL: IMMEDIATELY: NOTIFY TO CLIENTS TO STOP USING THE AFFECTED VERSION AND TO REVERT TO AN ANTERIOR OR ULTERIOR SOFTWARE VERSION. | out of calibration | calibration error | dicom is missing mag factor | software bug: inappropriate measurement algorithm | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT, THIS ISSUE IS DOCUMENTED AS A PRECAUTION. THE ISSUE WAS REPORTED BY OUR CLIENT. THEY PROVIDE TELERADIOLOGY SERVICES TO VARIOUS HOSPITALS. THEIR PROTOCOL FOR ACCEPTING STAT CASES FROM THESE HOSPITALS REQUIRES THAT THE SENDING FACILITY TRANSMIT A SEPARATE TEXT MESSAGE THAT IDENTIFIES EACH STAT CASE BY ACCESSION NO, TOGETHER WITH THE STUDY DESCRIPTION AND THE EXPECTED NUMBER OF IMAGES TO BE TRANSMITTED. FURTHER, BASED ON THIS PROTOCOL, THE CLIENT'S ADMINISTRATIVE STAFF WHO CONTROL THE WORKFLOW, CONFIRMS THE IMAGE COUNT TRANSMITTED IS THE SAME AS THE IMAGE COUNT AT THE READING CTR (IN INTELEPACS) BEFORE ACCEPTING THE CASE AND ASSIGNING TO A RADIOLOGIST FOR READING. DESCRIPTION OF PROBLEM: IMAGES SUCCESSFULLY TRANSMITTED TO INTELEPACS ARE NOT AVAILABLE FOR READING IN ALL STUDIES, AND THE READER IS UNAWARE OF THE POSSIBILITY OF MISSING IMAGES. THE PROBLEM MANIFESTS UNDER VERY SPECIFIC CIRCUMSTANCES, LIMITED TO A FEW CLIENT SITES AS A POSSIBLE PATIENT SAFETY ISSUE. THE POSSIBLE PATIENT SAFETY ARISES WHEN RECEIVING A SERIES AND CAN MANIFEST IN TWO SCENARIOS: FIRST: FOR THE 4 CLIENTS THAT EMPLOY INGEST SERVERS TO PROVIDE TEMPORARY STORE-AND-FORWARD CAPABILITIES IN THEIR ARCHITECTURE, IN SOME CONDITIONS, THESE SERVERS FAIL TO DELIVER ALL RECEIVED IMAGES TO THE PERMANENT STORAGE SERVERS AND NO RECORD OF THESE IMAGES PERSIST IN THE MAIN ARCHIVE. AS THESE INGEST SERVERS ARE NOT CONFIGURED AS VALID IMAGE LOCATIONS FOR ACCESS BY DIAGNOSTIC VIEWERS, THE STUDY CAN BE DISPLAYED IN INTELEPACS WITH LESS THAN FULL SET OF TRANSMITTED IMAGES. SECOND: FOR THE REMAINING CLIENT ARCHITECTURES THAT DO NOT EMPLOY TEMPORARY INGEST SERVERS, THERE IS STILL A VERY LOW PROBABILITY OF FAILURE (FAR LESS LIKELY THAN THE FIRST SCENARIO) FOR THE DEVICE TO PROPERLY REPLICATE AND REGISTER IMAGES WHEN TRANSFERRING THEM TO ANOTHER SERVER. IF THE SECOND COPY IS A VALID IMAGE LOCATION, IT IS POSSIBLE FOR A SOME IMAGES TO BE MISSING WHEN VIEWING. WE WILL: IMMEDIATELY VERIFY AND APPLY A BUG FIX TO ALL THE CLIENTS WHERE THIS ISSUE CAN OCCUR; APPLY AUDIT SCRIPTS TO IDENTIFY ANY OCCURRENCE OF THE ISSUE AND ALSO VALIDATE THE BUG FIX. CORRECTIVE ACTION CALENDAR: FOR THE THREE CLIENTS USING A CONFIGURATION WITH INGEST SERVERS AS TEMPORARY STORAGE, AN IMMEDIATE UPDATE PROCESS WAS UNDERTAKEN VIA THE CLIENT SPECIFIC RELEASE PROCESS (COMPLETED AS OF 04/23/2015) TO AVOID THE POSSIBILITY OF DATA LOSS. FOR THE REMAINING CLIENTS, THE NORMAL FULL RELEASE UPDATE (FRU) WILL BE EXPEDITED BUT UNDERGO THE USUAL TESTING AND RELEASE PROCESS. Manufacturer Narrative: AFFECTED SOFTWARE VERSIONS: MODELS: 4.7.1 R27 AND UP, 4.8.1 R20 AND UP, 4.9.1 R17 AND UP, 4.10.1 R9 AND UP, 4.11.1 R3 AND UP. | missing image series | loss of data | missing when viewing | software bug | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OCTOBER 5, 2016 SIEMENS BECAME AWARE OF DATA LOSS AT THE CUSTOMER'S SITE IN (B)(6). THE SHORT TERM STORAGE (STS) ATTACHED TO THE SIEMENS SYNGO.PLAZA SYSTEM GOT CORRUPTED DUE TO A HANDLING ERROR DURING ROUTINE MAINTENANCE OF THE HARD DISC UNIT (RAID) ON THE STS. DURING EXCHANGE OF THE HARD DISC, THE CONFIGURATION WAS LOST; THEREFORE DATA STORED ON THE RAID BECAME NO LONGER AVAILABLE ON THE STS. HOWEVER, NOT ALL OF THE DATA WAS COMPLETELY LOST DUE TO REDUNDANCY IN THE SETUP OF THE SYNGO.PLAZA ENVIRONMENT AT THIS SITE: ú A LARGE AMOUNT OF DATA LOST ON THE STS HAS BEEN ARCHIVED ON THE LONG TERM STORAGE (LTS) ATTACHED TO THE SYNGO.PLAZA UNIT. THUS THIS DATA IS STILL AVAILABLE TO THE USER. ú A SEPARATE SUBSET OF THE DATA, WHICH WAS PREVIOUSLY STORED ON THE STS PRIOR TO THE INCIDENT, WAS FOUND ON THE SCANNERS WHERE IT HAD BEEN INITIALLY CREATED. THIS DATA WAS RE-SENT TO THE STS AND IS AVAILABLE TO THE USER AS WELL. ú IN TOTAL 64 SERIES CONTAINING 43 STUDIES WERE PERMANENTLY LOST. THIS DATA WAS OBTAINED BETWEEN APRIL AND AUGUST 2016. SIEMENS SYNGO.PLAZA SYSTEM DID NOT CONTRIBUTE TO THE DATA LOSS. THERE ARE NO INJURIES AND NO NEGATIVE IMPACT HAS BEEN REPORTED BY THE CUSTOMER. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. NO PATIENT RESCANS WERE REQUIRED AS THE LOST STUDIES WERE OBTAINED PRIOR TO AUGUST 31, 2016, THEREFORE THE DATA HAD BEEN ALREADY ASSESSED AND USED FOR CLINICAL DIAGNOSIS PRIOR TO THE DATE OF THE INCIDENT ((B)(6) 2016). HOWEVER, SIEMENS CONSIDERS THIS INCIDENT AN ADVERSE EVENT AS THE DATA OLDER THAN 4 WEEKS MAY STILL HAVE CLINICAL RELEVANCE AND BE USED BY RADIOLOGISTS FOR COMPARISON TO IMAGES RECEIVED DURING NEWER PATIENT EXAMINATIONS. THE CLINICAL IMPACT OF THIS INHIBITED COMPARISON IS ASSESSED AS LOW. (B)(6). | missing data | loss of data | during maintenance | user error: maintenance handling error | no harm | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THAT A MONITOR IMAGE BECAME PURPLISH DURING THE LAPAROSCOPIC CHOLECYSTECTOMY WITH THE SUBJECT DEVICE AND THE LAPARO-THORACO VIDEOSCOPE LTF-VH. THE FACILITY CONTINUED TO USE THE SUBJECT DEVICE AND THE LTF-VH, AND THE PROCEDURE WAS COMPLETED. AFTER THE PROCEDURE, BILE SPILLAGE OCCURRED. THE FACILITY PERFORMED AN OPEN SURGERY BECAUSE THERE WAS NO ALTERNATE DEVICE AT THAT TIME. THE FACILITY COMMENTED THE IMAGE DIFFICULTY MIGHT CAUSE THE BILE SPILLAGE. THE LTF-VH WAS SUSPECTED TO HAVE MALFUNCTION AT FIRST, SO OLYMPUS CHECKED THE LTF-VH ON (B)(6) 2016. BUT THERE WAS NO ABNORMALITY OF THE LTF-VH. AFTER THAT, THE PHENOMENON OCCURRED AGAIN WHEN THE FACILITY USED THE SUBJECT DEVICE CONNECTED WITH ANOTHER SCOPE. THEREFORE, THE SUBJECT DEVICE MIGHT BE ATTRIBUTED TO THE CAUSE OF THIS PHENOMENON. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2016-00237 TO PROVIDE DEVICE EVALUATION RESULTS. THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CHECKED THE SUBJECT DEVICE BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED. OLYMPUS WILL CONTINUE TO INVESTIGATE THE CAUSE OF THIS PHENOMENON. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00106. | image turning color | image display error | bile spillage | image difficulty | additional unnecessary radiation | 3 |  |  |  |  | FDA2
Event Description: OLYMPUS WAS INFORMED THAT AN ENDOSCOPE IMAGE IN THE MONITOR BECAME REDDISH SUDDENLY DURING LAPAROSCOPIC CHOLECYSTECTOMY. THE FACILITY COMPLETED THE PROCEDURE BY EXCHANGING THE CH-S190-XZ-E TO ANOTHER DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT EXCEPT EXCHANGING THE DEVICE. Manufacturer Narrative: OLYMPUS CONFIRMED THAT THE PHENOMENON WAS NOT REPRODUCED IN THE FACILITY AND THE DEVICE WORKED PROPERLY FOR 6 HOURS SINCE THE DEVICE WAS TURNED ON. OLYMPUS COULD NOT DETERMINE THE CAUSE OF THIS PHENOMENON BECAUSE THE PHENOMENON WAS NOT REPRODUCED. THIS PHENOMENON MIGHT BE ATTRIBUTED TO INADEQUATE CONNECTION TO THE VIDEO PROCESSOR DEVICE OR ADHESION OF FOREIGN SUBSTANCES ON ELECTRICAL CONTACTS OF THE DEVICE CAUSED CONTACT FAILURE TEMPORARILY. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | image turning color | image display error | endoscope image became redish | inadequate connection to device; adhesive or foreign substances on device | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THAT DURING CHOLECYSTECTOMY, A MONITOR IMAGE DISAPPEARED WHEN THE FACILITY OUTPUTTED THE ELECTRIC KNIFE. THE FACILITY COMPLETED THE PROCEDURE TO REPLACE A DEVICE TO THEIR SPARE DEVICE BUT THERE WAS NO INFORMATION WHAT DEVICE THE FACILITY REPLACED. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE IS NOT RETURNED TO OLYMPUS YET. OLYMPUS WILL INVESTIGATE THE SUBJECT DEVICE TO DETERMINE THE CAUSE OF THIS PHENOMENON AFTER OLYMPUS RECEIVES IT. OLYMPUS IS UNDER INVESTIGATION FOR THE CAUSE OF THIS PHENOMENON. OLYMPUS WILL SUBMIT A SUPPLEMENTAL MDR REPORT AFTER THE CAUSE OF THIS PHENOMENON IS DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | fail to display image | no display or display failure | when outputted electrical knife | hardware issue: monitor failure | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THAT DURING LAPAROSCOPIC CHOLECYSTECTOMY, A MONITOR IMAGE DISAPPEARED SUDDENLY IN THE COMBINATION WITH THE SUBJECT DEVICE AND THE LTF-VH. THE FACILITY CONSIDERED THE CAUSE OF THIS PHENOMENON AS THE MALFUNCTION OF THE SUBJECT DEVICE OR THE LTF-VH. THE FACILITY TURNED OFF THE SUBJECT DEVICE IMMEDIATELY, AND THEN, THEY REMOVED AND INSERTED THE POWER SUPPLY CABLE. AFTER THAT, THE FACILITY TURNED ON THE SUBJECT DEVICE AGAIN AND THE DEVICES WORKED PROPERLY. THE FACILITY COMPLETED THE PROCEDURE WITHOUT REPLACING THE DEVICES. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2015-00925 TO PROVIDE DEVICE EVALUATION RESULTS. THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR INVESTIGATION. OLYMPUS CHECKED THE SUBJECT DEVICE WITH THE COMBINATION SCOPE WHICH THE FACILITY USED DURING THE PROCEDURE, AND THE PHENOMENON WAS SOMETIMES REPRODUCED WHEN THE VIDEO CONNECTER OF THE SCOPE WAS PULLED. CONTACT FAILURE BETWEEN THE SCOPE AND THE SUBJECT DEVICE MIGHT OCCUR TEMPORARILY BECAUSE THE FACILITY ADDED FORCE TO THE VIDEO CONNECTER OF THE SCOPE TOWARD PULLING DIRECTION FOLLOWING THE CABLE WAS PULLED WHILE USING THE SCOPE. SO THE SUBJECT DEVICE MIGHT NOT BE ABLE TO SUPPLY POWER TO THE SCOPE OR TO PROCESS IMAGE SIGNAL PROPERLY AT THAT TIME. CONSEQUENTLY, A MONITOR IMAGE MIGHT DISAPPEAR. IT WAS ALREADY PASSED OVER 7 YEARS SINCE THE FACILITY BEGAN TO USE THE SUBJECT DEVICE. THE FACILITY HAD INSERTED AND REMOVED A SCOPE FROM THE VIDEO CONNECTER SOCKET REPEATEDLY. SO CONTACT FAILURE BETWEEN THE SCOPE CONNECTER AND THE VIDEO CONNECTER SOCKET MIGHT OCCUR TEMPORARILY BECAUSE THE SOCKET WAS EASILY RATTLED BY WEAR. THERE WAS ALSO A POSSIBILITY THAT DETERIORATING CONTACT PINS OF THE VIDEO CONNECTER SOCKET MIGHT AFFECT THIS PHENOMENON. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS CAUTIONS FOR THE DEVICE HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2015-00928. | fail to display image | no display or display failure | when outputted electrical knife | hardware issue: malfunction | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THAT DURING THE TUR (TRANSURETHRAL RESECTION) PROCEDURE THE HALATION HAD OCCURRED ON THE ENDOSCOPIC IMAGE REPEATEDLY. THE FACILITY HAD COMPLETED THE PROCEDURE WITH USE OF THE OTHER DEVICE. THERE WAS NO REPORT OF THE PT'S INJURY REGARDING THIS EVENT. Manufacturer Narrative: THE REFERENCED OTV-S7V AND CONCOMITANTLY-USED LIGHT SOURCE AND CAMERA HEAD WERE RETURNED TO THE OLYMPUS MEDICAL SYSTEMS CORP (OMSC) FOR EVAL. OMSC EVALUATED THE DEVICES AND FOUND THAT THEY HAD NO ABNORMALITY AND OPERATED CORRECTLY. THE EXACT CAUSE OF THIS PHENOMENON CANNOT BE CONCLUSIVELY DETERMINED, HOWEVER THERE IS THE POSSIBILITY OF THIS PHENOMENON IS ATTRIBUTED TO THE USER HANDLING OF THE DEVICES. OMSC STATED THE APPROPRIATE HANDLING OF THE OTV-S7V IN THE INSTRUCTION MANUAL WHEN THE OTV-S7V HAD ABNORMALITIES. ALSO, OMSC CHECKED THE MFR HISTORY OF THE SUBJECT DEVICE, THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADD'L INFO IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | halation occurred | incorrect display | endoscope image unclear | user handling of device | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THAT NOISE ON THE MONITOR SCREEN OCCURRED DURING LAPAROSCOPIC HERNIA RESTORATION. THE FACILITY REPLACED THE SUBJECT OTV-S190 TO ANOTHER SYSTEM TO COMPLETE THE PROCEDURE. THERE WAS NO INJURY REPORT OF THIS EVENT EXCEPT EXCHANGING THE DEVICE. Manufacturer Narrative: THE SUBJECT OTV-S190 WAS RETURNED TO OLYMPUS FOR EVAL. OLYMPUS COULD NOT CONFIRM THE PHENOMENON BUT THERE WERE FOREIGN SUBSTANCES NEAR THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET. OLYMPUS COULD REPRODUCE THIS PHENOMENON BY CONNECTING A SCOPE TO THE OTV-S190, PLACING FOREIGN SUBSTANCES ON THE SCOPE CONNECTOR PINS FOR INSULATION. THIS PHENOMENON MIGHT BE ATTRIBUTED TO CONDUCTION FAILURE BETWEEN THE SCOPE AND THE OTV-S190 BECAUSE ADHESION OF FOREIGN SUBSTANCES INSIDE THE VIDEO CONNECTOR SOCKET OR WATER ON THE SCOPE CONNECTOR MIGHT INTERRUPT CONNECTION BETWEEN THE TWO DEVICES TEMPORARILY. OLYMPUS ALSO CHECKED THE MFG HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTIONS MANUAL ALREADY STATED NOTIFICATIONS OF ELECTRICAL CONTACT HANDLING. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | noise emitted | audible noise | noise on the monitor screen | foreign substances near the electrical contact | treatment and diagnostic delayed |  |  |  |  | IFU: contact electrical handling | FDA2
Event Description: OLYMPUS WAS INFORMED THAT THE MONITOR SHOWED THE ERROR E315 (UNSUPPORTED SCOPE) DURING UNSPECIFIED PROCEDURE WITH THE SUBJECT DEVICE AND THE LAPARO-THORACO VIDEOSCOPE LTF-VH. THE FACILITY REPLACED THE LTF-VH TO ANOTHER LAPARO-THORACO VIDEOSCOPE LTF-V3, BUT THE SAME PHENOMENON STILL OCCURRED. AFTER THAT, THE LTF-VH WHICH THE FACILITY USED AT FIRST WAS CONNECTED TO THE SUBJECT DEVICE AND THERE WAS NO ERROR. THE FACILITY COMPLETED THE PROCEDURE WITH THE SUBJECT DEVICE AND THE LTF-VH. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION AND OLYMPUS COULD REPRODUCE THE PHENOMENON. OLYMPUS CONFIRMED THERE WAS FOREIGN SUBSTANCE IN THE ELECTRICAL CONTACT OF THE VIDEO CONNECTER SOCKET OF THE SUBJECT DEVICE. OLYMPUS ALSO CONFIRMED AT THE FACILITY THAT THERE WERE FOREIGN SUBSTANCES IN THE ELECTRICAL CONTACT OF THE VIDEO PLUG OF BOTH SCOPES WHICH THE FACILITY USED AT THAT TIME. THERE WAS A POSSIBILITY THAT THE SCOPES WHICH HAD FOREIGN SUBSTANCES IN THE ELECTRICAL CONTACT OF THE VIDEO PLUG WERE CONNECTED TO THE SUBJECT DEVICE, RESULT IN CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE SCOPES. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF OTV-S190, LTF-VH AND LTF-V3 ALREADY MENTIONS CAUTIONS FOR VIDEO CONNECTOR HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT #: 8010047-2016-00245, 8010047-2016-00246. | error message | device displays error message | unsupported videoscope | foreign substances near the electrical contact | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OLYMPUS WAS INFORMED THE FOLLOWING. DURING THE LAPAROSCOPIC CHOLECYSTECTOMY, THE OPERATOR PERFORMED THE CUT PREPARATION. AT THAT TIME, THE ENDOSCOPIC IMAGE WAS DISAPPEARED. THE FACILITY REPLACED THIS OTV-S190 WITH ANOTHER SIMILAR DEVICE MADE BY STRYKER AND THE FACILITY COMPLETED THE PROCEDURE. THERE WAS NO REPORT OF THE PATIENT'S INJURY REGARDING THIS EVENT. ALTHOUGH A CAUSAL RELATIONSHIP WITH THIS EVENT IS UNKNOWN, THE HOSPITALIZATION WAS PROLONGED. Manufacturer Narrative: THE SUBJECTED DEVICE WAS NOT RETURNED TO OLYMPUS MEDICAL SYSTEMS CORP.(OMSC). OMSC CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE OTV-S190 INSTRUCTION MANUALS STATE THE CORRESPONDING METHOD WHEN THERE IS AN ABNORMALITY IN THE ENDOSCOPIC IMAGE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. | loss of visual | no display or display failure | perform cut preparation | hardware issue | prolong surgery | [5 empty] | FDA2
Event Description: ON (B)(6) 2012 A MERGE CARDIO CUSTOMER REPORTED REPORTS DONE UNDER WRONG PATIENT. REPORT FROM 2ND PATIENT IS ALSO SHOWING UP AS A 2ND REPORT THREAD FOR PATIENT #1. THE PDF GENERATION FUNCTIONALITY WAS NOT DESIGNED TO BE EXCLUSIVE PER PATIENT. IF TWO OR MORE USERS WERE REPORTING ON DIFFERENT PATIENTS AND REQUESTED THE GENERATION OF A PDF CONCURRENT, THE SYSTEM WOULD INCORRECTLY ATTACH THE FIRST PDF REPORT FILE GENERATED TO THE OTHER PATIENT(S). Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | incorrect file applied | patient data issue | incorrect report attachment | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2015, VITAL CUSTOMER SUPPORT RECEIVED A CALL FROM (B)(6) REGARDING DISPLAY OF INCORRECTLY INVERTED BRAIN PERFUSION COLOR MAPS (CBF AND MTT) ON OUR VITREA CT BRAIN PERFUSION 4D SOFTWARE APPLICATION. THE REPORTER INDICATED THAT A DELAY IN PATIENT TREATMENT HAD OCCURRED BECAUSE OF THE ISSUE. AGAIN ON (B)(6) 2015, THE SAME REPORTER INFORMED OUR CUSTOMER SUPPORT TEAM REGARDING ANOTHER PATIENT ISSUE WITH THE SOFTWARE IN WHICH THE CT BRAIN PERFUSION RESULT FROM OUR SOFTWARE WAS NOT IN AGREEMENT WITH THE PERFORMED MR RESULTS. THE MRI SHOWED AN OBVIOUS STROKE (INFARCT), WHILE THE CT PERFUSION RESULTS ON OUR SOFTWARE DID NOT. THE PATIENT WAS TREATED BASED ON MRI RESULTS. | incorrect color map | image display error | incorrectly inverted brain perfusion color maps | CT software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2016 PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT ORIENTATIONS MARKED IN SOME IMAGES LOOKS WRONG, THE ISSUE IS STILL UNDER INVESTIGATION. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. Manufacturer Narrative: AN INVESTIGATION WAS PERFORMED BY THE ENGINEERING TEAM AND CLINICAL APPLICATION EXPERTS. FURTHER INVESTIGATION FOUND THAT THIS ISSUE IS SPECIFIC TO SAVING SAGITTAL VOLUME FILES. THIS ISSUE WILL ONLY OCCUR IF THE USER SAVES THE RECONSTRUCTED DATA IN THE SAGITTAL ORIENTATION AND THEN LOADS THE SAGITTAL VOLUME INTO THE VIEWING APPLICATION. CLINICAL APPLICATIONS TRAINS USERS TO ONLY SAVE AND USE THE TRANSVERSE IMAGE VOLUME FILES AFTER EXECUTING THE RECONSTRUCTION APPLICATION (AUTOSPECT PRO). THE ENGINEERING TEAM DETERMINED THAT THE RISK IS ACCEPTABLE REGARDING THIS ISSUE. THE CUSTOMER IS CURRENTLY USING THE SYSTEM AND THEY HAVE BEEN INSTRUCTED BY CLINICAL APPLICATIONS ON HOW TO AVOID THE ISSUE. | image orientation incorrect | image orientation Incorrect | orientation marked in some images looks wrong | medical procedure issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2017 A CUSTOMER SUBMITTED A COMPLAINT INDICATING THAT MAMMOGRAM IMAGES APPEARED DARK WHEN LOADED IN CLEARCANVAS RIS PACS. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: THE CUSTOMER DEMONSTRATED THAT THE ANATOMY WAS NOT VISIBLE THROUGH REMOTE ACCESS. THE PROBE TOOL WAS USED TO CHECK THE BRIGHTNESS VALUES. BRIGHTEST AREAS HAD VALUES AROUND 190-200. DICOM HEADERS DISPLAYED A WINDOW CENTER (0028,1050) VALUE OF 2,700. CENTER VALUE SHOULD BE ROUGHLY BETWEEN THE BRIGHTEST AND DARKEST PIXEL VALUES. IT WAS DETERMINED THAT THE MODALITY IS WRITING THE CENTER VALUE DIFFERING FROM THE BRIGHTEST SPOT. IT WAS DETERMINED AT THIS TIME THAT THE PRODUCT FUNCTIONED AS INTENDED AND THIS ISSUE WAS DUE TO INCORRECT DICOM DATA. A WORKAROUND WAS PROVIDED TO THE CUSTOMER IF THE ISSUE IS ENCOUNTERED AGAIN. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME (B)(4) CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. THIS DEFECT IS ONLY ENCOUNTERED IN SELECT SCENARIOS, WHERE A MODALITY PRODUCES IMAGES THAT ARE COMPRESSED USING (B)(4), AND THE IMAGE PIXEL DATA IS LESS THAN 16-BIT, AND THE "BITNESS" OF THE COMPRESSED DATA STREAM DOES NOT MATCH THAT OF THE IMAGE PIXEL DATA. THE OUTCOME OF THE ABOVE IS A LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA, WHICH CAUSES THE SYMPTOMS THAT WERE REPORTED BY THE CUSTOMER. THE ROOT CAUSE IS THAT A SOFTWARE DEFECT FOUND IN THE DEVICE THAT IS ENCOUNTERED WHEN IT IS USED WITH NON-DICOM COMPLIANT (B)(4) COMPRESSED IMAGES. THE SOFTWARE IMPROPERLY USES THE DICOM BIT DEPTH (I.E. THE BITS STORED TAG) TO DECOMPRESS THE COMPRESSED PIXEL DATA STREAM FOR DISPLAY, INSTEAD OF THE BIT DEPTH THAT IS ENCODED IN THE COMPRESSED PIXEL DATA STREAM ITSELF. IN NON-DICOM COMPLIANT IMAGES WHERE THE DICOM AND COMPRESSED PIXEL DATA STREAM BIT DEPTHS DO NOT MATCH, THE SOFTWARE OUTPUTS AN IMAGE WITH SOME LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA. IN THESE SCENARIOS, THE MODALITY PRODUCING THE IMAGE IS ITSELF ALSO NOT COMPLIANT TO DICOM (B)(4). MDR REPORTS: 3012075008-2020-00002, 3012075008-2020-00003 AND 3012075008-2020-00004 HAVE THE SAME ROOT CAUSE. A CAPA HAS BEEN OPENED AND RECALL INITIATED TO ADDRESS THE ISSUE IN AFFECTED DEVICES. | invisible anatomy | optical issue | invisible through remote access | image is not compliant to dicom | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2017 AGFA BECAME AWARE OF A SECOND OCCURRENCE REPORTED BY THE CUSTOMER IN WHICH REPORTS CROSSING FROM CARDIO SYSTEM (CV12.1SU1) INTO EPIC ARE SHOWING UP WITH INCORRECT DEMOGRAPHICS. THE FIRST OCCURRENCE HAS BEEN REPORTED IN FDA MDR 1225058-2015-14119. PREVIOUSLY THE ROOT CAUSE HAS BEEN IDENTIFIED AS PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. AGFA REPORTED TO FDA VIA 21 CFR 806 REPORTING (Z-0283-2016) A CORRECTION FOR THIS ISSUE. THE CUSTOMER DECLINED THE CORRECTION, HOWEVER THE CUSTOMER DID ACKNOWLEDGE AND ACCEPT THE RISK ASSOCIATED WITH NOT ACCEPTING THE CORRECTED VERSION OF THE SOFTWARE AND NOT REMOVING THE AFFECTED VERSION OF THE SOFTWARE. AGFA IS WORKING WITH THE CUSTOMER TO RESOLVE THIS ISSUE. THERE HAS BEEN NO REPORTED HARM TO PATIENT OR USER DURING THIS EVENT. A SUPPLEMENTAL REPORT WILL BE PROVIDED.Event Description: THIS SUPPLEMENT REPORT #1 IS BEING SUBMITTED TO PROVIDE CORRECTIVE ACTIONS TAKEN. | incorrect demographic | data issue | crossing from cardio system | software defects | no harm | [5 empty] | FDA2
Event Description: ON (B)(6) 2017, PHILIPS RECEIVED FEEDBACK THAT INTELLISPACE PORTAL 7.0.2.20700 MR FIBERTRAK APPLICATION IS NO LONGER ABLE TO POST PROCESS CERTAIN DATA. THE ISSUE IS STILL UNDER INVESTIGATION. Manufacturer Narrative: WE HAVE RECEIVED NEW INFORMATION AND THEREFORE MODEL # BRAND NAME: AND SERIAL # WERE CORRECTED IN THIS REPORT THE RISK ASSESSMENT CONCLUDED THAT THE OVERALL RISK IS ACCEPTABLE. A PHILIPS APPLICATIONS SPECIALIST WENT ONSITE AND TRAINED THE CUSTOMER ON HOW TO PROPERLY CREATE STUDIES TO ANALYZE ON THE SYSTEM. Manufacturer Narrative: WE HAVE NOT COMPLETED OUR INVESTIGATION OF THIS EVENT. WE WILL FILE A FOLLOW-UP EMDR AT THE COMPLETION OF THE INVESTIGATION. (B)(4). | fail to post process data | data issue | during data analysis | software issue | no harm | [5 empty] | FDA2
Event Description: ON (B)(6) 2019 A CUSTOMER REPORTED THAT IMAGES RETRIEVED FROM VRAD (PACS) IS DISPLAYED VERY DARK WHEN LOADED INTO CLEARCANVAS WORKSTATION PERSONAL EDITION 13.1. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: THE CUSTOMER ALSO NOTED THE IMAGES APPEAR FINE WHEN VIEWED IN THE VRAD WEB VIEWER. HE ALSO SAID THAT THEY APPEARED NORMAL IN AN OLDER VERSION OF CC WORKSTATION (2.0 SP1 - OPEN-SOURCE). THE SYNAPTIVE REP INITIATED SCREEN SHARE TO OBSERVE. SAMPLE WAS A DIGITAL X-RAY FROM A KONICA MINOLTA PORTABLE SCANNER. W/L WAS CHANGED TO AROUND 491/177 BEFORE IT WAS VIEWABLE. INITIAL W/L (AUTO) = 4096/2047. OPENED A FEW MORE IMAGES THAT WERE VERY SIMILAR. THE CUSTOMER EXPLAINED THAT THE PREVIOUS VERSION WAS CC WS 2.0 SP1 (OPEN SOURCE), BUT WAS UNABLE TO CONFIRM WHETHER THESE PARTICULAR IMAGES OPENED CORRECTLY WITH THAT VERSION. THE SYNAPTIVE REP GAVE THE CUSTOMER A WORKAROUND VIA THE WINDOW/LEVEL HOTKEY CONFIGURATION. ON 18-MAR-2019 AN INVESTIGATION FOUND THE ISSUE TO BE IDENTICAL TO FEED-3544 (REPORTED AS 3012075008-2020-00001): THE BITS STORED (0028,0101) VALUE IS SET TO 12-BIT IN THE DICOM HEADER HOWEVER, IT APPEARS THE JPEG 2000 COMPRESSED PIXEL DATA IS ACTUALLY 16-BIT. WHEN THE TAG VALUE IS EDITED TO 16, THE IMAGE APPEARS CORRECT (I.E. NOT DARK) ON LOAD. THEREFORE, THE DICOM DATA IS INCONGRUENT. | dark image | image display error | during retrieving image | dicom data is incongruent | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. THE POWERJACKET WINDOW CAN BECOME OUT OF SYNC AND DISPLAY THE DETAILS FOR THE PREVIOUSLY LOADED PATIENT. ON DECEMBER 16, 2020 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER COMP- (B)(4). | display to different patient image | patient data issue | incorrect patient image viewed | software configuration issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) WAS NOTIFIED THAT THE SYNAPSE PACS SHOWS INCORRECT PATIENT STUDY/IMAGE IN THE POWERJACKET. THE ISSUE WAS REVIEWED AND ESCALATED TO THE ENGINEERING DEPARTMENT. ON (B)(6) 2020 A RISK ASSESSMENT WAS PERFORMED BY THE ENGINEERING DEPARTMENT TO INVESTIGATE RISK TO PATIENT SAFETY AND FMSU QUALITY AND REGULATORY AFFAIRS DEPARTMENTS WERE INFORMED OF THE SEVERITY. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A JAVA SCRIPT ERROR. THE ISSUE HAS BEEN RESOLVED; THE CUSTOMER'S SOFTWARE WAS UPGRADED. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER COMP- (B)(4). | display to different patient image | patient data issue | incorrect patient image viewed | java script error | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020 THE FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) ENGINEERING DEPARTMENT INITIATED A COMPLAINT FOR AN INTERNAL ISSUE WHERE THE SYNAPSE PACS WAS SHOWING INCORRECT 3D SPHERE MAX VALUES FOR SPECIFIC CT STUDIES. THE 3D SPHERE MAX SHOULD GIVE A MAX DENSITY VALUE WITHIN THE SPHERE'S RANGE; HOWEVER THE VALUE SHOWN WAS MUCH HIGHER THAN EXPECTED. THE ISSUE IS NOT PRESENT WITH ANY OTHER DENSITY TOOLS. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THIS ISSUE WAS DISCOVERED INTERNALLY; THIS SOFTWARE VERSION HAS NOT BEEN RELEASED TO ANY CUSTOMERS. THE CAUSE WAS TRACED TO INSUFFICIENT TESTING DURING DEVELOPMENT OF SOFTWARE VERSION. THE INCORRECT VALUE IS HIGHLY DETECTIBLE AND WOULD BE NOTICED BY THE CLINICIAN. THE ISSUE WILL BE RESOLVED PRIOR TO RELEASE TO CUSTOMERS. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | incorrect display value | data issue | incorrect 3D sphere max values | software defects | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON (B)(6) 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR AN ISSUE WITH THE EJECTION FRACTION (EF) AND LEFT VENTRICULAR (LV) VOLUME METHOD OF DISK BI-PLANE (BP MOD) CALCULATION. THE EF OR LV VOLUME BP MOD DID NOT UPDATE IN THE CLINICAL REPORTING APPLICATION (CRA) REPORT WHEN MEASUREMENTS WERE CHANGED BY A DIFFERENT USER. AFTER THE SONOGRAPHER PERFORMS LV VOLUME BP MOD AND SAVES THE REPORT, AND ANOTHER USER TRIES TO EDIT THE "TRACE", THE UPDATED/NEW VOLUME AND EJECTION FRACTION VALUES DO NOT UPDATE IN THE CRA. IF THE SONOGRAPHER WHO INITIALLY PERFORMED THE BP MOD EDITS THE "TRACE", THE VALUES UPDATE CORRECTLY. THE ISSUE WAS REVIEWED, AND ESCALATED TO THE ENGINEERING DEPARTMENT. ON OCTOBER 30, 2020, A RISK ASSESSMENT WAS PERFORMED BY THE ENGINEERING DEPARTMENT TO INVESTIGATE RISK TO PATIENT SAFETY, FMSU QUALITY, AND REGULATORY AFFAIRS DEPARTMENTS WERE INFORMED OF THE SEVERITY. THERE WAS NO SERIOUS INJURY, OR DEATH ASSOCIATED WITH THIS EVENT. THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO INSUFFICIENT TESTING DURING DEVELOPMENT. THE ISSUE HAS BEEN RESOLVED IN THE NEXT VERSION, AND THIS CUSTOMER WAS ALREADY PROVIDED WITH A FIX. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF.#: INTERNAL COMPLAINT NUMBER: (B)(4). | fail to update accordingly | data issue | measurements were changed but the trace did not update | software defects: insufficient testing during development | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON (B)(6) 2021 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER NOTIFICATION REGARDING AN ADVERSE EVENT RELATED TO SYNAPSE PACS. TWO REPORTS CHANGED FROM A QUALITATIVE EJECTION FRACTION (EF) VALUE OF 55-60% TO ONE REPORTED AS <20%. THE ISSUE LED TO ONE PATIENT RECEIVING AN UNNECESSARY INVASIVE CORANARY ANGIO PROCEDURE. A SECOND PATIENT WAS UNAFFECTED BECAUSE THE ERROR WAS CAUGHT PRIOR TO ASSESSMENT. THE PATIENT HAS BEEN REPORTED TO US AS BEING HEALTHY WITH NO ISSUES AND WAS DISCHARGED. Manufacturer Narrative: UNKNOWN. THE EF VALUE IS SELECTED BY THE USER VIA A DROP-DOWN MENU. ONCE THE DROP-DOWN MENU HAS BEEN SELECTED, THE EF VALUE CAN BE CHANGED BY THE UP AND DOWN ARROW KEYS. THE USER DID NOT CHANGE FOCUS FROM THE DROP-DOWN MENU IN THE ADVANCED REPORTING WINDOW TO THE VIEWER WINDOW BEFORE USING THE ARROW KEYS IN AN ATTEMPT TO MOVE TO THE NEXT IMAGE. THIS CAUSED THE SELECTED EF VALUE TO UNKNOWINGLY CHANGED. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | value changed incorrectly | data issue | user select EF value via a drop-down menu | user issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON 09-MAY-2016 PHILIPS RECEIVED SERVICE CALL FROM A CUSTOMER REPORTING THAT ORIENTATIONS MARKED IN SOME IMAGES LOOKS WRONG IN TWO EBW-NM STATIONS. THIS ISSUE IS STILL UNDER INVESTIGATION. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENTS, AND NO REPORT OF PATIENT HARM. ON THE 03-MAY-2016 WE ALREADY GOT A FEEDBACK FROM THIS CUSTOMER ABOUT THE SAME ISSUE IN ONE OF THESE SYSTEMS AND WE ALREADY REPORTED THE INCIDENT AS MFR. REPORT #:1525965-2016-00027. THIS REPORT AIMED TO COVER THE SECOND SYSTEM. Manufacturer Narrative: AN INVESTIGATION WAS PERFORMED BY THE ENGINEERING TEAM AND CLINICAL APPLICATION EXPERTS. FURTHER INVESTIGATION FOUND THAT THIS ISSUE IS SPECIFIC TO SAVING SAGITTAL VOLUME FILES. THIS ISSUE WILL ONLY OCCUR IF THE USER SAVES THE RECONSTRUCTED DATA IN THE SAGITTAL ORIENTATION AND THEN LOADS THE SAGITTAL VOLUME INTO THE VIEWING APPLICATION. CLINICAL APPLICATIONS TRAINS USERS TO ONLY SAVE AND USE THE TRANSVERSE IMAGE VOLUME FILES AFTER EXECUTING THE RECONSTRUCTION APPLICATION (AUTOSPECT PRO). THE ENGINEERING TEAM DETERMINED THAT THE RISK IS ACCEPTABLE REGARDING THIS ISSUE. THE CUSTOMER IS CURRENTLY USING THE SYSTEM AND THEY HAVE BEEN INSTRUCTED BY CLINICAL APPLICATIONS ON HOW TO AVOID THE ISSUE. Manufacturer Narrative: (B)(4). | orientation marked incorrectly | image orientation incorrect | save the reconstructed data in the sagittal orientation | software defects | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ON 13-JUN-2017 A CUSTOMER REPORTED THAT IMAGES THAT APPEARED VERY BRIGHT (WASHED OUT) WHEN LOADED IN CLEARCANVAS WORKSTATION PERSONAL EDITION 13.1. SHE RECEIVED THE IMAGES VIA A DVD. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: A SYNAPTIVE REP USED "GOTOASSIST" TO VIEW IMAGES DIRECTLY. THE REP FOUND THE WINDOW CENTER AND WINDOW WIDTH DICOM TAG VALUES WERE SET TO 2047 AND 4096 RESPECTIVELY. THE PROBE TOOL SHOWED THAT THE BRIGHTEST WHITE HAD VALUE OF 60-70, BLACKS WERE 256. MANUAL WINDOW/LEVEL APPLIED TO MAKE IMAGE VISIBLE WAS IN 250/150 RANGE. THE REP FOUND: BITS STORED (0028,0101) VALUE SET TO 12. AFTER CORRECTING THE VALUE TO 16 AND SAVING IMAGE IT DISPLAYED PROPERLY UPON RELOADING. THE ISSUE WAS FOUND TO BE SIMILAR TO FEED-3544 (REPORTED AS 3012075008-2020-00001). IMAGE IS ACTUALLY 16-BIT, BUT DICOM HEADER TAG (0028,0101) INCORRECTLY STATES IT IS 12-BIT. LEFT CUSTOMER TO CONTACT THE SITE THAT PRODUCED THE DVD TO SEE WHY THE DICOM TAG VALUE DOES NOT MATCH WHAT IS IN THE PIXEL DATA. THIS WAS LOGGED AS AND ENHANCEMENT REQUEST FOR IMPROVED INTEROPERABILITY WITH OTHER DICOM VENDORS. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME JPEG 2000 CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. | image brightness out of specification | optical issue | dicom tag value does not match what is in the pixel data | software defects | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON DECEMBER 02, 2020 FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE CARDIOVASCULAR. STUDY WAS SENT IN FOR A NEW PATIENT WITH THE CORRECT MRN. INCORRECT MRN IS SHOWING IN CV CLIENT. ON DECEMBER 16, 2020, A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY, OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. THE NUMERIC ID ASSIGNED TO A PATIENT RECORD COULD BE REUSED ONCE A NEWER PATIENT RECORD WAS MERGED TO AN OLDER PATIENT RECORD. WHEN THIS ACTION IS DONE, THE NEWER ID IS MADE AVAILABLE AGAIN. THIS WAS NOT AN ISSUE PRIOR TO ADVANCED REPORTING. WITH THE INTRODUCTION OF ADVANCED REPORTING THE NUMERIC ID IS NOT RELEASED WITHIN THE ADVANCED REPORTING DATABASE SCHEMA AND THAT IS HOW THE MISMATCH BETWEEN THE PATIENT DEMOGRAPHIC DATA CAN OCCUR. REF: INTERNAL COMPLAINT NUMBER COMP (B)(4). | incorrect patient ID | patient data issue | incorrect MRN is showing in CV client | software configuration issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON FEBRUARY 7, 2017, AGFA BECAME AWARE OF AN EVENT IN WHICH A CUSTOMER HAS LOGGED AN FDA MEDWATCH FDA 3500 FORM STATING THE AGFA IMPAX SOFTWARE IS AT FAULT FOR AN IMPROPERLY PERFORMED MEDICAL PROCEDURE. AT THIS TIME, THE ONLY INFORMATION WE WERE PROVIDED WITH WAS THAT THE FDA MEDWATCH 3500 FORM WAS SUBMITTED TO THE FDA AND THAT AN UNNECESSARY MEDICAL PROCEDURE HAD BEEN PERFORMED. ON FEBRUARY 24, 2017, THE CUSTOMER PROVIDED AGFA WITH THE MEDWATCH FDA 3500 FORM WHICH PROVIDED THE DETAILS OF THE SITUATION. THE CUSTOMER REPORTED A SINGLE IMAGE WITHIN A CT SERIES WAS FLIPPED AND SAVED BY A RADIOLOGIST DURING INTERPRETATION. THE CUSTOMER SUBSEQUENTLY INDICATED THERE WAS NOT A SIGNIFICANT VISUAL INDICATION THAT THE IMAGE HAD BEEN SAVED IN THE OPPOSITE DIRECTION OF ALL OTHER IMAGES IN THE SERIES WHICH THEN LED TO A PATIENT HAVING SURGERY ON THE WRONG SIDE. THE PRELIMINARY INVESTIGATION FROM AGFA SHOWS THIS TO BE A USER ERROR AND THAT THE PRODUCT IS WORKING AS DESIGNED. INVESTIGATION IS UNDER WAY TO DETERMINE THE ROOT CAUSE. A SUPPLEMENTAL REPORT WILL BE PROVIDED.Event Description: THIS SUPPLEMENTAL REPORT # 1 IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION RELATED TO THE DESCRIPTION OF THE OVERALL EVENT AND TO PROVIDE THE ROOT CAUSE. | image flipped | image reversal | surgery on the wrong side | user error | inappropriate surgery | 4 |  |  |  |  | FDA2
Event Description: ON JUNE 26, 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. WHEN OPENING THE POWERJACKET AND SELECTING A PATIENT OFF THE WORKLIST, THE POWERJACKET DISPLAYED INCORRECT PATIENT INFORMATION. THE DOCTOR INDICATED THAT THE ISSUE OCCURRED AGAIN WHEN SELECTING ANOTHER PATIENT AND THE POWERJACKET AGAIN DISPLAYED A DIFFERENT PATIENT. THE FUJIFILM CUSTOMER SERVICE REPRESENTATIVE DID NOT IDENTIFY ANY ISSUES WITHIN THE IIS LOGS OF THE CUSTOMER AND THE ISSUE COULD NOT BE REPRODUCED. ON JULY 08, 2020 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. A BUG FIX WAS DEVELOPED AND PROVIDED TO FIX THE ISSUE. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: (B)(4). | display to different patient image | incorrect display | inappropriate image viewed | software configuration issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON MARCH 15, 2021 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) WAS MADE AWARE OF AN INCIDENT WITH SYNAPSE PACS. WHEN A USER OPENS A STUDY WITH AN URL-CALL FROM THEIR HIS (SAP), IMAGES ARE SHOWN FROM PATIENT A BUT THE INFORMATION PANE SHOWS STUDIES FROM PATIENT B. ON (B)(6) 2021 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. FUJIFILM INITIATED A RECALL ON 3/2/2021 TO ALERT CUSTOMERS OF SEVERAL PATIENT MISMATCH ISSUES EXISTING IN ALL SYNAPSE PACS 5 VERSIONS UP TO 5.7.210. FUJIFILM SUBMITTED C&R REPORT (1000513161-03/11/2021-001-C) TO FDA, WHICH HAS BEEN CLASSIFIED AS CLASS II AND ASSIGNED RECALL NUMBER Z-1348-2021. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | display to different patient image | incorrect display | inappropriate image viewed | software configuration issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON SEPTEMBER 27, 2019, FUJIFILM MEDICAL SYSTEMS USA, INC. SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS TOMO MARKERS AND WAS DETERMINED NOT TO BE A COMPLAINT BASED ON THE SERVICE TICKET. ON JANUARY 24, 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. ENGINEERING DEPARTMENT WAS NOTIFIED OF CUSTOMER CONCERNS REGARDING SYNAPSE PACS TOMO ISSUES. INITIAL INVESTIGATION INDICATED THAT THE ORIENTATION MARKER WAS BEING DISPLAYED ON THE WRONG SIDE OF THE TOMO IMAGE. THE RESOLUTION WAS PROVIDED THROUGH A BUG FIX. BASED ON RISK AND SEVERITY OF THE ISSUE AND FIX, THE COMPLAINT WAS DETERMINED NOT TO BE REPORTABLE. ON APRIL 1, 2020, FUJIFILM MEDICAL SYSTEMS USA, INC QUALITY AND REGULATORY DEPARTMENT WAS NOTIFIED OF CUSTOMER CONCERNS REGARDING ADDITIONAL SYNAPSE PACS TOMO ISSUES. UPON FURTHER INVESTIGATION, ON APRIL 28, 2020 IT WAS DETERMINED THAT THERE WERE THREE (3) ISSUES WITH THE LATERALITY. ISSUE 1: INDICATORS WERE MISSING ON LXCCL, RXCCL, LLMO, AND RLMO. THIS ISSUE WAS FIXED IN VERSION (B)(4), HOWEVER, HAVING THE REFERENCE POINTS ON THE SLIDE BAR BROUGHT ANOTHER ISSUE TO LIGHT. ISSUE 2: THE LLMO AND RLMO ARE DISPLAYING LATERALITY INCORRECTLY BECAUSE THE LATERAL AND MEDIAL REFERENCE POINTS WERE SWITCHED, THE SLIDE BAR WAS IN THE INCORRECT DIRECTION, ISSUE 3: THE IMAGE COUNT WAS DISPLAYING IN REVERSE ORDER. THERE WAS NO PATIENT SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS REPLICATED IN THE FUJIFILM MEDICAL SYSTEMS USA LAB. THE CAUSE WAS TRACED TO THE SOFTWARE NOT READING THE CORRECT DICOM TAB FOR BREAST TOMO IMAGES. A BUG FIX HAS BEEN DEVELOPED AND ALL AFFECTED CUSTOMERS WILL BE NOTIFIED OF THE ISSUE AND PROVIDED THE FIX. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | orientation marked incorrectly | image orientation incorrect | orientation marked was displayed on the wrong side of image | software bug: software not reading the correct dicom tab | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: ON THE (B)(6) 2015 PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT WHEN RECONSTRUCTING DATA BY THE INTELLISPACE PORTAL AUTOSPECT_PRO APPLICATION SOME OF THE RESULTS MAY BE DIFFERENT THAN EXPECTED. THE FEEDBACK INDICATE THAT THE AFFECTED DATA IS CREATED ON LEGACY AXIS/IRIX GAMMA CAMERAS MANUFACTURED BY MARCONI MEDICAL SYSTEMS USING 102 DEGREE ACQUISITION PROTOCOL. THERE WAS NO REPORT OF PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. Manufacturer Narrative: DATA RECEIVED FROM A PREVIOUS SIMILAR COMPLAINT WAS INVESTIGATED BY ENGINEERING AND REVEALED THAT THE AUTOSPECT PRO APPLICATION WAS ONLY DESIGNED TO RECONSTRUCT CARDIAC SPECT DATA OBTAINED WITH DETECTORS POSITIONED AT 90ø OR 180ø RELATIVE TO ONE ANOTHER. HOWEVER, CERTAIN GAMMA CAMERAS, E.G., THE MARCONI AXIS AND LRIX CAMERAS, PERMIT ACQUISITIONS AT OTHER RELATIVE DETECTOR ANGLES. AFFECTED CUSTOMERS WERE NOTIFIED BY FIELD SAFETY NOTICE (FSN 88100036) SENT ON 09-MAR-2016. FIELD CHANGE ORDER ((B)(4)) WAS RELEASED ON (B)(6) 2016 TO IMPLEMENT A SOFTWARE UPDATE TO ELIMINATE THE PROBLEM. Manufacturer Narrative: (B)(4). | fail to reconstruct data | data issue | reconstruct cardiac spect data | inappropriate relative detector angles | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ONE SITE HAS REPORTED MISSING 400-500 STUDIES THAT NEVER MADE IT TO THE ARCHIVE FOR LONG TERM STORAGE BEFORE THE WATERMARKING PROCESS DELETED THEM. THE CODE PURGES SERIES FROM THE RADSUITE (AV) IMAGE CACHE PRIOR TO THE SERIES BEING STORED IN ENTERPRISE CONTENT MANAGER (ECM) LONG TERM IMAGE ARCHIVE, AND STUDIES ARE NOT STORED IN THE LONG TERM ARCHIVE (LTA) OR THE AV IMAGE CACHE FOR FUTURE REFERENCE. Manufacturer Narrative: SUBMITTED THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS TO SECTION. | missing study | loss of data | unavailable studies | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: ORIGINAL ISSUE WAS ESCALATED AS A SIMPLE REQUEST FOR SERVICE SUPPORT IN DETERMINING WHY THE SYNGO DYNAMICS PRODUCT COULD NOT RETRIEVE STUDY IMAGES FROM THE SITE'S ARCHIVE SYSTEM. SERVICE INVESTIGATION FOUND THAT THE SITE'S LONG TERM ARCHIVE SYSTEM (NON-SIEMENS PRODUCT) HAD FAILED. THE SITE HAS TURNED OVER THE ARCHIVE DISKS TO A DATA RECOVERY COMPANY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S HISTORICAL DATA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES WILL STILL BE AVAILABLE IN THE SD SYSTEM UNLESS DELETED FROM THE SD SYSTEM BY SITE IT POLICY OR USER ACTION. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. | fail to retrieve image | loss of data | fail to compare with historical image | site archive system crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OUR FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITHIN THE ACCEPTABLE RISK THRESHOLD FOR THE PRODUCT, HOWEVER, RELEVANT RISK FOR THE SITE WAS IDENTIFIED. THE SYNGO.PLAZA HAS BEEN CONFIGURED TO USE A PRIMARY (NAS1) AND BACKUP (NAS2) NETWORK ATTACHED STORAGE (NAS) MOUNT POINT FOR ARCHIVING PURPOSES. DUE TO AN UNRECOVERABLE HARDWARE FAILURE ON THE NAS1 A RESTORATION OF ALL DATA WAS PERFORMED FROM THE NAS2. HOWEVER, THE NAS2 DID NOT CONTAIN ALL ARCHIVED DATA AS ON NAS1. 8523 SERIES DATED WITHIN THE TIMEFRAME 2011-2016 WERE NOT IN THE ARCHIVE AND NO LONGER AVAILABLE. THIS IS CONSIDERED A DATA LOSS OF CLINICAL RELEVANT DATA. THE ROOT CAUSE ANALYSIS WAS ABLE TO LIMIT DOWN THE ISSUE TO A MANUAL INTERVENTION ON THE NAS2'S FILESYSTEM, WHICH IS NOT NO LONGER TRACEABLE. THE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE MISSING DATA IN THE NAS2 ARCHIVE. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURED IN THE (B)(6). Manufacturer Narrative: SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITHIN THE ACCEPTABLE RISK THRESHOLD FOR THE PRODUCT, HOWEVER, RELEVANT RISK FOR THE SITE DUE TO DATA LOSS OF CLINICAL RELEVANT DATA WAS IDENTIFIED. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. | storage hardware failure | device inoperable | loss of data | hardware issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: OUR HOSPITAL WAS REVIEWING A PATIENT CONCERN. A PATIENT REPORTED THAT WHEN THEY SQUEEZED THE BALL, THERE WAS NO RESPONSE FROM THE CONTROL ROOM. A PNEUMATIC CALL-BALL ATTACHED TO AN AIR FILLED CORD IS A CONNECTION TO AN EXTERNAL AUDIO SYSTEM AND LIGHT SYSTEM. WHEN THE BALL IS SQUEEZED, AN AUDIBLE SOUND/ ALERT IS SOUNDED, HEARD AND VISUALIZED AS A LIGHT IN THE MRI CONTROL ROOM. THE CALL-BALL IS AN AIR FILLED BULB-BALL THAT IS CONNECTED TO A FLEXIBLE TUBING WHICH UTILIZES A MALE CONNECTOR TO A FEMALE PORT. THIS AIR FILLED TUBING IS CONNECTED TO THE FOOT OF THE MRI PATIENT TABLE. THE PATIENT ACTIVATES THE SYSTEM BY SQUEEZING THE CALL-BALL WHICH ALERTS THE TECHNOLOGIST IN THE CONTROL AREA WITH BOTH VISUAL AND AUDIBLE ALARM AND THEN THIS HAS TO BE DEACTIVATED BY THE MRI TECHNOLOGIST. WE FOUND THAT IF THE CORD IS CRIMPED IN ANYWAY OR CLOSURE OF THE AIR FILLED TUBING OR FAILURE TO SQUEEZE THE BALL WITH SUFFICIENT FORCE DEACTIVATES THE FEATURE AND NO SOUND OR LIGHT IS SEEN. WE ARE REPORTING THIS AS A SAFETY CONCERN. | fail to activate alarm | not audible alarm | during MRI procedure | insufficient force | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: PACS APPLICATION FAILS TO IDENTIFY THAT AN EXTERNAL RADIOLOGY REPORT HAS BEEN RECEIVED AND PERMITS A RADIOLOGIST TO CONTINUE DICTATING A CASE THAT HAS BEEN ALREADY FINALIZED. THIS ALLOWS THE DISCREPANT REPORTS TO EXIST (1 INTERNAL TO OUR PACS AND 1 FROM THE LOCAL RADIOLOGIST WHO HAS THEIR REPORT ENTERED INTO THE PERMANENT MEDICAL RECORD). THE LOCALLY GENERATED REPORT IS NOT ARCHIVED IN PACS AND ANY SUBSEQUENT REVIEW OF THE CASE AS A PRIOR OR FOR QA PURPOSES, WILL HAVE CONCLUSIONS BASED ON THE INTERNALLY GENERATED REPORT WHICH IS NOT THE REPORT OF RECORD. | fail to identify report received | data issue | unable to continue a case | software issue | process delayed | [5 empty] | FDA2
Event Description: PATIENT HAD A PROCEDURE DONE IN CATHETER LABORATORY. UNABLE TO FIND THE FILMS ON THE MCKESSON ARCHIVAL SYSTEM. THERE WAS AN ERROR MESSAGE ON THE COMPUTER DURING THIS TIME. THE FILMS COULD NOT BE TRANSFERRED OVER UNTIL THE ERROR MESSAGE WAS CLEARED. ONCE THE ERROR MESSAGE WAS CORRECTED THE FILMS WERE AVAILABLE. ALL FILMS FOR THIS PATIENT ARE PRESENT. THERE IS A PROCESS IN PLACE, SO THAT THIS WILL NOT HAPPEN. FILMS ARE TO BE CHECKED DAILY TO BE SURE THEY HAVE TRANSFERRED OVER TO MCKESSON. ANY MISSING FILMS DUE TO ERROR MESSAGES SHOULD BE ADDRESSED IN REAL TIME. CATH LAB ASSISTANT NURSE MANAGER HUDDLED WITH TEAM MEMBERS REGARDING THIS PROCESS AND TO BE SURE THAT THEY ARE FOLLOWING IT DAILY AND WITH ALL CASES. | missing film | loss of data | transfer film over to lab | error message | process delayed | [5 empty] | FDA2
Event Description: PATIENT PRESENTED FOR A MAMMOGRAM AND ULTRASOUND. HE INDICATED THAT HE COULDN'T STAND FOR THE EXAM, SO HE WAS POSITIONED SITTING DOWN. WHILE SETTING UP FOR THE IMAGE THE PATIENT WAS HEARD YELLING AND IT WAS NOTED THE C-ARM OF THE UNIT HAD COME DOWN ON THE PATIENT'S LEFT FEMUR. THE UNIT WAS RAISED AND PATIENT COMPLAINED OF LEFT HEEL PAIN. HE WAS SUBSEQUENTLY TAKEN TO THE EMERGENCY ROOM WHERE IMAGING STUDIES CONFIRMED A COMMINUTED CALCANEAL FRACTURE. | c-arm dropped | component falling | during mammogram and ultrasound | hardware issue | user calcaneal fracture | 4 |  |  |  |  | FDA2
Event Description: PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. DUE TO AN IDENTIFIED HARDWARE FAILURE IN NETWORK ATTACHED STORAGE (NAS), THERE ARE MISSING FOLDERS CONTAINING PATIENT IMAGES IN THE CONFIGURED NAS (NAS1). A BACKUP NAS IS CONFIGURED, HOWEVER, IT IS MANAGED BY A 3RD PARTY APPLICATION AND NOT CONFIGURED IN THE SYNGO.PLAZA SYSTEM. CONTENTS OF NAS1 AND NAS2 ARE IDENTICAL. BOTH NAS SYSTEMS HAVE A LARGE AMOUNT OF EMPTY FOLDERS FOR DATES RANGING BETWEEN JANUARY 6, 2017 AND APRIL 19, 2017. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED DATA; HOWEVER, THE AFFECTED DATA IS CONSIDERED AS CLINICALLY RELEVANT. THE PRODUCT SYNGO.PLAZA DID NOT CONTRIBUTE TO THE DATA LOSS, AS THE ISSUE IS RELATED TO A HARDWARE FAILURE AND THE SYNGO.PLAZA IS A SOFTWARE SOLUTION ONLY (B)(6). | storage hardware failure | device inoperable | unavailable report | hardware failure | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO MIX-UP OCCURRED, HENCE, NO PATIENT RESCAN WAS REQUIRED. CUSTOMER REPORTED THAT CERTAIN STUDIES CANNOT BE FOUND IN SYNGO.PLAZA PATIENT BROWSER. THE SYSTEM WAS MIGRATED AND UPGRADED FROM AN OLD SYNGO.PLAZA VERSION TO A NEW SERVER AND NEW SW VERSION VB20A IN (B)(6) 2017. DURING THE ANALYSIS OF THIS ISSUE IT HAS BEEN IDENTIFIED THAT AN INCORRECT LICENSE FILE IS PRESENT ON THE SYSTEM. THUS STUDIES WERE NOT ARCHIVED TO SYNGO.PLAZA'S LONG TERM ARCHIVE. THE AFFECTED STUDIES WERE PRESENT IN THE ONLINE STORAGE FOR REPORTING OF CURRENT EXAMINATIONS. AFTER A CERTAIN FILL LEVEL WAS REACHED IN THE ONLINE STORAGE, THE STUDIES WERE REMOVED BY AUTODELETION FUNCTION. THE AFFECTED STUDIES RANGE FROM 01/2017 TO 07/2017 AND ARE NOT AVAILABLE FOR PRIOR COMPARISON. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED STUDIES AND THE POSSIBILITY OF RESTORING THEM. NO ISSUE FROM SYNGO.PLAZA SOFTWARE SINCE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE DATA LOSS AT THE SITE. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. (B)(6). | fail to display report online | no display or display failure | the studies were not archived | incorrect license file | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: PERMANENT LOSS OF 15 STUDIES WAS DISCOVERED IN THE SYNGO.PLAZA'S DATABASE. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE DATA LOSS AS THE ISSUE IS RELATED TO A HARDWARE FAILURE AND THE SYNGO.PLAZA SYSTEM IS A SOFTWARE ONLY SOLUTION. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. THE LOST DATA IS CONSIDERED CLINICALLY RELEVANT AS THE STUDIES WERE BETWEEN 1 AND 6 MONTHS OLD AND MAY BE STILL REQUESTED FOR COMPARISON. THE REPORTED EVENT OCCURRED IN (B)(6. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. ON 2017-10-26 CUSTOMER REPORTED MISSING OR CORRUPTED IMAGES APPEARING IN SYNGO.PLAZA. DURING THE ANALYSIS IT WAS IDENTIFIED THAT ON THE (B)(6) 2017, ONE OF THE HARD DRIVES USED BY THE SYNGO.PLAZA'S SHORT TERM STORAGE HAD FILE SYSTEM PROBLEMS CAUSED BY HARDWARE FAILURES. (B)(6. | missing image series | loss of data | fail to compare with historical image | hard drive failure | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: POTENTIAL HARM WITH SOFTWARE; I WORK FOR THE DEPARTMENT OF (B)(6). OUR RADIOLOGY PACS IS SUPPLIED BY INTELERAD MEDICAL SYSTEMS. IN THE PAST, A CODE DEFECT WAS IDENTIFIED THAT ALLOWED OUR RADIOLOGISTS TO DICTATE A RADIOLOGY REPORT ON A STUDY THAT HAD ALREADY BEEN READ BY A LOCAL RADIOLOGIST. THE PROBLEM HAS RETURNED AND HAS RESULTED IN CONFLICTING REPORTS AND CALLBACKS TO ORDERING PROVIDERS FOR CRITICAL RESULTS, WITH RESULTS THAT DIFFERED FROM THE LOCAL RADIOLOGIST REPORT, LEADING TO CONFUSION ON THE PART OF THE ORDERING PROVIDER. THIS SOFTWARE DEFICIENCY COULD RESULT IN A DIFFERING DIAGNOSIS AND/OR CHANGE IN TREATMENT PLANS IF THE LOCAL RADIOLOGIST REPORT DIFFERS SIGNIFICANTLY FROM THE (B)(6) RADIOLOGY REPORT. SINCE THIS IS NOW THE SECOND TIME THAT THIS PROBLEM HAS OCCURRED AND IS A PATIENT SAFETY ISSUE, I AM ELEVATING THIS TO THE FDA. FDA SAFETY REPORT ID# (B)(4). | conflicted report | data issue | user confused | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: PROBLEM DESCRIPTION: FOLLOWING A CLIENT COMPLAINT, INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM IN THE IMAGE FUSION MODULE WHICH MAY DISPLAY AN INCORRECT CALCULATION OF STANDARD UPTAKE VALUE (SUV) WHEN LOADING THE PRESENTATION STATE OF IMAGE FOR A PREVIOUSLY CALCULATED (AND SAVED) REGION OF INTEREST (ROI). THIS PROBLEM COULD LEAD RADIOLOGISTS TO BASE THEIR DIAGNOSIS/REPORT ON INCORRECT ANATOMICAL MEASUREMENTS LEADING TO INCORRECT MEDICAL TREATMENT. NO PATIENT INCIDENT HAS BEEN REPORTED AT THIS MOMENT - THIS PROBLEM IS DOCUMENTED AS A PRECAUTION. INVESTIGATION DETAILS: ONE CUSTOMER COMPLAINT ((B)(4)) BUG IDENTIFIED: (B)(4). | incorrect display value | incorrect display | calculation of standard uptake value | software bug | inappropriate treatmetn or diagnostic | [5 empty] | FDA2
Event Description: PROBLEM: A FEW OF OUR CARDIOLOGISTS HAVE HAD ISSUES WITH FUJI CHANGING EJECTION FRACTION (EF) VALUES IN THE PAST TWO MONTHS. ONE WEEK AGO, THIS HAPPENED ON ONE OUTPATIENT STRESS ECHOCARDIOGRAMS AND ONE OF THE INPATIENT CASES. THIS INPATIENT UNDERWENT A CORONARY ANGIOGRAM AND MRI DUE TO THE ABERRANCY ON HIS ECHO REPORT. BOTH REPORTS CHANGED FROM A QUALITATIVE EF 55-60% TO ONE REPORTED AS <20%. PER THE MRI, EJECTION FRACTION = NORMAL. IMPACT ON THE INPATIENT: ADDITIONAL CARDIOLOGY FOLLOW UP (CONSULT, PROGRESS NOTES, ECHOCARDIOGRAM, CARDIAC MRI, ANGIOGRAM, AND EXTRA HOSPITAL DAY'S STAY). FOLLOW UP: THE EVENT RESULTING IN THE ERROR WAS IDENTIFIED (FOUND ONE OF OUR MDS AND TECHNICIANS) AND IS TIED TO THE PROVIDER WORKFLOW. WHEN THE READING MD SIGNS THE REPORT (I.E. THE READER CHOOSES A NORMAL EF AND SIGNS THE REPORT, BUT THE FINAL REPORT STATES EF 20%). THIS OCCURS WHEN YOU SELECT THE EF STATEMENT AND YOU SELECT ONE OF THE ARROW KEYS ON THE KEYBOARD INSTEAD OF CHANGING TO THE NEXT IMAGE, THE EF STATEMENT WILL CHANGE TO ANOTHER OPTION ON THE LIST. INTERIM ACTION: FACILITY CARDIOLOGY DEPARTMENT LEADS HAVE SPOKEN WITH THE MANUFACTURER AND INDICATED THAT THIS IS A TOP PRIORITY FOR COMPANY TO FIND A RESOLUTION. IN THE INTERIM, ALL USERS WERE NOTIFIED TO BE AWARE WHEN USING THE ARROWS ON THE KEYBOARD TO MOVE THROUGH THE REPORTS. ALSO INSTRUCTED "IF YOU SEE AN ECHO THAT STATES EF < 20%, (ESPECIALLY IF THIS IS AN UNEXPECTED RESULT OR IS ASSOCIATED WITH OTHER DISCREPANT STATEMENTS SUCH AS ¨NORMAL WALL MOTION¨ OR ¨NO CHANGE FROM PRIOR¨ WHEN THE PRIOR EF WAS NORMAL), PLEASE PULL UP THE ECHO IMAGES TO VERIFY THE LVEF. IF YOU FEEL THE <20% WAS AN ERROR PLEASE REACH OUT TO THE READING CARDIOLOGIST SO THE REPORT CAN BE CORRECTED..." | value changed incorrectly | data issue | error was identified | user error: use the arrows on the keyboard to move through the report | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: PROBLEM: NO PATIENT INVOLVEMENT, PROBLEM IDENTIFIED DURING INTERNAL TESTING. POTENTIAL HAZARD(S): WHEN USING ENHANCED CT FORMAT DICOM DATA, THERE MAY BE: INCORRECT ORIENTATION OF DISPLAYED IMAGE DATA INDICATING INCORRECT PATIENT SIDE (LEFT, RIGHT, SUPERIOR, INFERIOR, ANTERIOR OR POSTERIOR) (SIDEDNESS). INACCURATE MEASUREMENT. Manufacturer Narrative: NO SPECIFIC MEDICAL DEVICE FAILURE. NO PATIENTS INVOLVED. ISSUE IDENTIFIED DURING ECHOPIXEL INTERNAL TESTING. HEALTH HAZARD EVALUATION COMPLETED BY QUALITY MANAGEMENT.. ALL CURRENT USERS OF PRODUCT HAVE BEEN CONTACTED, PROVIDED INFORMATION ON THE ISSUE. ALL CURRENT USERS HAVE INDICATED THAT NONE USE THE STATED ENHANCED MULTI-FRAME CT FORMATTED DICOM DATA, THEREFORE NO RISK TO PATIENTS CORRECTIVE ACTION PROCESS IN PLACE. NO SPECIFIC DEVICE FAILURE INVOLVED. | image orientation marked incorrectly | image orientation incorrect | during internal testing | enhanced CT format dicom data | no harm | [5 empty] | FDA2
Event Description: Q-STATION MIXED INFORMATION.Event Description: THE CUSTOMER REPORTED THAT WHEN USING Q-STATION, DATA FROM A PREVIOUS PATIENT APPEARS ON THE CURRENT PATIENT'S DATA REPORT. THERE WAS NO REPORT OF INJURY TO PATIENT OR NEGATIVE IMPACT ON CLINICAL JUDGMENT. Manufacturer Narrative: THIS ISSUE IS CURRENTLY UNDER INVESTIGATION. FINAL RESULTS OF THE EVALUATION WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS RESULTING FROM THIS ISSUE. SOFTWARE. Manufacturer Narrative: IT WAS DETERMINED THAT THE ISSUE WAS A KNOWN SOFTWARE DEFECT. MANDATORY FIELD CHANGE ORDER,(B)(4), WAS DEVELOPED AND RELEASED TO CORRECT THIS ISSUE. THE FIX IS INCLUDED IN Q-STATION VERSION 3.3.2 AND HIGHER. THE AFFECTED CUSTOMERS WILL BE PROVIDED WITH THE NEWEST VERSION OF Q-STATION (3.3.2) ONCE THE NEW ORDERS HAVE BEEN CREATED AND SHIPPED. | display to different patient image | patient data issue | during procedure | software defect: Q-station | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE IS A MEDICAL IMAGE VIEWING SOFTWARE APPLICATION THAT ALLOWS IMAGE DATA TO BE VIEWED FROM A FACILITY IMAGE CACHE (FIC) OR FROM A LONG TERM ARCHIVE (LTA). A CUSTOMER NOTICED THAT PATIENT A'S STUDY WAS STORED UNDER THE CORRECT PATIENT IN THE FIC BUT WAS STORED AS A DIFFERENT, PATIENT B, IN THE LTA. WHEN PATIENT A'S CURRENT STUDY WAS OPENED IN THE RADSUITE VIEWER ON THE LTA, PATIENT B'S STUDY WAS OPENED AS A PRIOR INSTEAD OF PATIENT A'S STUDY, AND WAS LABELED WITH PATIENT A'S INFORMATION. | save image to different patient | patient data issue | during procedure | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER FILED A COMPLAINT (CASE (B)(4)) AGAINST RADSUITE STATING THAT THE PIXEL VALUE TOOL'S BODY WEIGHT AND IDEAL BODY WEIGHT VALUES ARE OFF IN RADSUITE COMPARED TO THE SCANNER. THE CUSTOMER CONTACT STATED THAT THE RADIOLOGIST'S USE THIS INFORMATION TO DETERMINE IF THE TUMOR IS MALIGNANT OR BENIGN AND NEEDS IT TO BE ACCURATE. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. NOTE THAT FDA TERMINATED THIS RECALL ON FEBRUARY 13, 2017. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE INVESTIGATION BY MERGE HEALTHCARE SHOWED THAT THE VALUE IS OFF BY APPROXIMATELY A FACTOR OF 20. THIS SHOULD BE EASILY RECOGNIZED BY THE USER BECAUSE OF THE DRASTIC DIFFERENCE BETWEEN THE VALUE AND THE VISUAL (QUALITATIVE) APPEARANCE OF THE IMAGE. THE DEFECT WAS INTRODUCED INTO THE RADSUITE V8.30.7.8 RELEASE. (B)(6) WAS THE FIRST AND ONLY CUSTOMER TO RECEIVE THE V8.30.7.8 RELEASE THAT CONTAINED THE DEFECT. A FIX HAS BEEN APPLIED AT (B)(6) TO CORRECT THE CALCULATION ERROR. | body weight value off | patient data issue | use body weight value to determine if the tumor is malignant or benign | configuration issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED ((B)(4)) THAT AN INCORRECT STUDY WAS ATTACHED TO A PATIENT AND THEY WERE TRYING TO DELETE THE IMAGES. THEY WERE UNABLE TO DELETE THE IMAGES FROM THE STUDY. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION COULD NOT BE CORRECTED ON THE STUDY, LEADING TO A POTENTIAL MISDIAGNOSIS. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | unable to delete image | patient data issue | incorrect study was attached to a patient study | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED (CASE (B)(4)) THAT STUDIES ARE STORING BUT ARE ALSO LOCKING UP. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION COULD NOT BE AVAILABLE FOR REVIEW. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | report locked | data issue | fail to open report | unavailable for reviewing | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED (CASE (B)(4)) THAT THEY WERE UNABLE TO RETRIEVE A STUDY BACK FROM THE ARCHIVE TO THE RADSUITE VIEWER. THE CUSTOMER WAS TRYING TO BRING THE STUDY BACK FROM THE ARCHIVE TO CORRECT MRN INFORMATION. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION CANNOT BE CORRECTED ON THE STUDY. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | fail to retrieve study | loss of data | can not correct study information | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER COMPLAINT (B)(4) FROM (B)(6) OPENED (B)(6) 2015 STATES PREFETCHER IS NOT WORKING AS EXPECTED (APPEARS DELAYED). CUSTOMER REPORTED THAT PREFETCHER IS NOT WORKING AND IS AFFECTED , POTENTIALLY, PATIENT CARE. THE PREFETCHER IS INTENDED TO PULL PATIENT PRIORS FOR RADIOLOGIST TO REVIEW DURING DIAGNOSIS. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | low display speed | erratic display | fail to pull patient info for review | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER REPORTED ISSUES WITH THEIR PRIMARY PRODUCTION RADSUITE CAUSING "DO NOT ROUTE" EXCEPTIONS. ALL CASES INVESTIGATED BY CUSTOMER ROUTED AND THEN THE EXCEPTIONS POPPED UP AFTERWARDS. ALSO, AN ICON NEXT TO THE STUDY WILL SHOW A RED "X" NOTIFYING USERS OF AN EXCEPTION. PER THE COMPLAINT CLAIM, THE APPEARANCE OF THOSE ICONS CONFUSED A RADIOLOGIST AND DELAYED READING OF A STAT STUDY BY 14 HOURS ON A CRITICAL FINDING. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | fail to route | device alarm system issue | user confused | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. ON (B)(6) 2013 A CUSTOMER REPORTED A PROBLEM WHEN THE MOUSE CURSOR HOVERED OVER THE ROI IN THE SECONDARY SERIES IT DISPLAYED THE STATISTICAL DATA BELONGING TO THE FIRST IMAGE. THIS IS A POTENTIAL SAFETY ISSUE SINCE THE ROI DISPLAYS AND MEASURES HOUNSFIELD UNITS (HUS). HOUNSFIELD MEASUREMENTS ARE USED TO MEASURE TISSUE DENSITY AND MAY BE USED AS AN INPUT TO CLINICAL DECISION MAKING. THERE WAS NO REPORT OF ADVERSE PATIENT IMPACT. (B)(4). Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | display to different patient image | patient data issue | measure tissue density | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. STUDIES ARE WATERMARKING OFF IN AV WITHIN 5-6 DAYS. THIS IS CAUSING THE RADIOLOGISTS TO HAVE TO PRE-FETCH STUDIES, CAUSING DISSATISFACTION. PREVIOUSLY THERE WERE 27 DAYS BEFORE STUDIES WOULD WATERMARK OFF AV. STUDY TODAY (B)(6) DID NOT PRE-FETCH PRIOR STUDY FROM (B)(6) 2015. RADIOLOGIST FEELS THIS IS A PATIENT SAFETY ISSUE. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | report watermarked off early | data issue | pre-fetch study | software configuration issue | user dissatisfaction | 1 |  |  |  |  | FDA2
Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. THE APPLICATION PROVIDES WORKFLOW INTEGRATION CAPABILITIES FOR HEALTH CARE ENTERPRISES, WHEREIN: RADIOLOGISTS CAN VIEW, ANNOTATE, AND TAG STUDIES AS DIAGNOSTICALLY READ. REFERRING PHYSICIANS CAN VIEW PATIENT IMAGES AND RADIOLOGISTS' ANNOTATIONS. TERTIARY CARE PHYSICIANS, MEDICAL TECHNOLOGISTS, AND INFORMATION TECHNOLOGY PROFESSIONALS CAN RECEIVE PATIENT RECORDS. ON 5/25/2017, MERGE HEALTHCARE TECHNICAL SUPPORT WAS NOTIFIED BY A RADSUITE CUSTOMER THAT PROCEDURES WERE NOT ABLE TO BE PERFORMED CAUSING A DELAY IN TREATMENT FOR 2 PATIENTS. PER THE CUSTOMER SITE, EMERGENCY DEPARTMENT STUDIES ARE TYPICALLY READ WITHIN 30 MINUTES HOWEVER, THERE WERE TWO STUDIES THAT WERE NOT READ UNTIL THE NEXT MORNING. WITH MERGE RADSUITE NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. THERE WAS NO REPORTED HARM TO EITHER PATIENT A OR PATIENT B AS A RESULT OF THE DELAY. (B)(4). | read report delayed | data issue | procedures were not able to be performed | user error | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: REPORTEDLY, AN ATTENDING PHYSICIAN WAS INTERPRETING AN ECHOCARDIOGRAM IN SYNAPSE CARDIOVASCULAR AND WAS MEASURING DOPPLER. THE PATIENT'S HEART RHYTHM WAS IRREGULAR, SO THE PHYSICIAN HAD TO AVERAGE DOPPLER SIGNAL MEASUREMENTS FOR EACH VALVE. WHILE DOING DOPPLER SIGNAL TRACINGS OF THE MITRAL VALVE, HE NOTICED THAT VALUES THAT SHOULD HAVE BEEN CONSTANTS IN THE CALCULATIONS FROM THE INITIAL TRACINGS WOULD SPONTANEOUSLY CHANGE AS MORE SIGNALS WERE AVERAGED (SPECIFICALLY MEAN PRESSURE GRADIENTS). REPORTEDLY, THIS HAS BEEN AN ONGOING ISSUE THAT IS CURRENTLY BEING INVESTIGATED BY THE MANUFACTURER. UNKNOWN HOW LONG THIS HAS BEEN AN ISSUE AND HOW MANY PATIENTS HAVE BEEN AFFECTED. | fail to update accordingly | data issue | average doppler signal measurements for each value shall spontaneosusly change | algorithm issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: SICAT IMPLANT IS A STANDALONE SOFTWARE APPLICATION, I.E. A RADIOLOGICAL VISUALIZATION SOFTWARE FOR DIAGNOSIS AND DENTAL IMPLANT PLANNING. A DENTIST FOUND THAT IMPLANT PLANNING DATA EXPORTED FROM SICAT IMPLANT ARE NOT CORRECT FOR ONE SPECIFIC EXPORT FORMAT. IF IT WERE TO RECUR, THE ERROR IN THE EXPORT FORMAT MIGHT CAUSE DENTISTS TO MANUFACTURE (OR HAVE MANUFACTURED) SURGICAL GUIDES WHICH ARE INCORRECT. USING THOSE GUIDES MIGHT CAUSE DRILLING INTO THE JAW TO BE EITHER NOT DEEP ENOUGH OR TO DEEP. A DENTIST MIGHT HAVE TO STOP A SURGICAL PROCEDURE. OR THE DRILLING WHEN TO DEEP MIGHT CAUSE DAMAGE TO STRUCTURES WITHIN THE JAW WHICH NEED INTERVENTION. | incorrect report format | data issue | manufacture incorrect surgical guides | software issue | damage to jaw structure | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF A CATASTROPHIC LOSS OF THE CUSTOMER'S STORAGE AREA NETWORK (SAN) AT THIS SITE. THE SAN IS PROVIDED AND MAINTAINED BY THE USER; IT IS THEIR CHOSEN LOCATION FOR THE DEPLOYMENT AND MANAGEMENT OF MULTIPLE MEDICAL SOFTWARE PRODUCTS, INCLUDING SIEMENS HEALTHCARE PRODUCTS. THE SPECIFIC CAUSE OF THE FAILURE OF THE SAN WAS NOT SHARED BY THE CUSTOMER WITH SIEMENS. THERE IS NO INDICATION THAT THE DEPLOYED SIEMENS PRODUCTS CAUSED THE FAILURE OF THE SAN, EITHER BY THE CUSTOMER OR AS A TECHNICAL POSSIBILITY. IT WAS SUBSEQUENTLY DETERMINED THAT THERE WERE NO RECENT DATA BACKUPS FROM WHICH DATA RECOVERY WOULD BE POSSIBLE AS THE CUSTOMER ALSO USED THEIR SAN AS THE SOLE BACKUP LOCATION. THE SITE HAS AN ONGOING RECOVERY EFFORT WITH A 3RD PARTY DATA RECOVERY COMPANY. IT HAS BEEN IDENTIFIED THAT ONLY SOME DRIVES FROM THE NETWORK HAVE A POTENTIAL FOR DATA RECOVERY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S DATA. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE LOSS OF POTENTIALLY CLINICALLY RELEVANT DATA MAY CONTRIBUTE TO A DELAY IN TREATMENT OR TO A PATH FOR INCORRECT TREATMENT. THE EXTENT OF THE PROBABILITY OF HARM OCCURRENCE FOR THE SITE IS INCALCULABLE AS THE EXTENT OF THE TOTAL RECOVERY AND INTERIM CLINICAL OPERATIONS AT THE CONCERNED SITE IS UNKNOWN TO SIEMENS. THE REPORTED EVENT OCCURED IN (B)(6). | loss of patient data | loss of data | loss of the customer storage area network at site | failure of storage area network | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF A RANSOMWARE VIRUS ATTACK ON THE SYNGO.PLAZA SERVER. THE NAME OF THE VIRUS IS UNKNOWN; HOWEVER, IT COMES FROM THE FAMILY OF RANSOMWARE VIRUSES. THE SYNGO.PLAZA SERVER WAS NOT INFECTED, ONLY THE FILES ACCESSIBLE THROUGH SAMBA (SMB) SHARES WERE ENCRYPTED AS A RESULT OF THE ATTACK. THE VIRUS ATTACKED THE SYNGO.PLAZA THROUGH THE CLIENTS OR OTHER COMPUTERS IN THE NETWORK. AT THE TIME OF THE ATTACK THE ANTIVIRUS WAS NOT INSTALLED ON THE SYNGO.PLAZA SERVER. ACCESS TO SHARE WAS NOT LIMITED AS RECOMMENDED BY THE SYNGO.PLAZA PRODUCT DOCUMENTATION AS IT WAS SET WITH PERMISSION TO "EVERYONE" WITH "READ, WRITE, EXECUTE". THE ATTACK RESULTED IN 310,075 SOPS (SERVICE OBJECT PAIRS) IN THE STS (SHORT TERM STORAGE) TO BE ENCRYPTED. THESE SOPS BELONG TO 50,504 SERIES (FULLY OR PARTIALLY AFFECTED) OUT OF 665,628 STORED SERIES IN THE STS. MULTIPLE SEPARATE LTS (LONG TERM STORAGE) SYSTEMS WERE AFFECTED BY THIS ATTACK AS WELL. THIS INCIDENT OCCURRED IN (B)(6). Manufacturer Narrative: POSSIBILITY TO RECOVER DATA FROM THE ENCRYPTED LTS IS CURRENTLY BEING ASSESSED BY SIEMENS. HOWEVER, WITH CURRENTLY AVAILABLE INFORMATION, IT CAN BE CONFIRMED THAT THE SYNGO.PLAZA SYSTEM DID NOT CONTRIBUTE TO THE DATA LOSS. | ransomeware virus attack | application security issue | attack through clients in the network | lack of anti-virus | security issue |  |  | permission setup (read, write, exceute) | install anti-virus software |  | FDA2
Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. ALL IMAGES IN THE SHORT TERM STORAGE (STS) WERE CORRUPTED AFTER A HARD DISK (RAID) FAILURE OCCURRED. AT THE USER SITE THE SYNGO.PLAZA SYSTEM IS OPERATED IN A STANDALONE MODE, I.E. WITHOUT ARCHIVING FUNCTIONALITY. FURTHERMORE ,CUSTOMER HAS NO OTHER ARCHIVING SOLUTION AVAILABLE. NONE OF THE PATIENT DATA COULD BE RECOVERED DUE TO ALL DATABASE DRIVES AND BACKUP DRIVES BEING CORRUPTED. TOTAL NUMBER AND TIME PERIOD OF LOST DATA ARE AS FOLLOWS: TOTAL NUMBER OF IMAGES LOST: 97868. TOTAL NUMBER OF STUDIES CORRUPTED: 19343. TIMELINE OF THE LOST DATA CANNOT BE DETERMINED AS ALL CORRUPTED STUDY RECORDS HAVE BEEN REMOVED FROM THE DATABASE AND THE DATA WAS PRESENT ONLY UNTIL DECEMBER 6, 2019. NONE OF THE STUDIES INCLUDING THE RECENT STUDIES (DATED BETWEEN DECEMBER 1 AND 6, 2019) WERE AVAILABLE IN THE MODALITY TO RESEND. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATION. THE REPORTED INCIDENT OCCURRED DUE TO A HARDWARE ISSUE WITHOUT CONTRIBUTION FROM THE SYNGO.PLAZA. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. INTERNAL ID # (B)(4). | image corrupted | patient data issue | unavailable images to compare | hardware disk failure: database drive and backup drive | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. THE APPLICATION LOG FILES SHOWED MULTIPLE INSTANCES OF ABNORMAL SHUTDOWN (POWER OUTAGE OR WINDOWS BLUE SCREEN/BSOD ISSUES) OF THE MACHINE. DUE TO THIS INTERRUPTION IN FILE TRANSFER PROCESS, THE IMAGE FILES WERE NOT SUCCESSFULLY STORED TO THE LOCAL FILE SYSTEM, RESULTING IN ARCHIVE FAILURE AND LOSS OF PHYSICAL FILES. TOTAL # OF IMAGES LOST - 2067. TOTAL # OF STUDIES LOST- 52. TOTAL # OF SERIES AFFECTED - 73. AFFECTED STUDIES ARE DATED 2014, 2015, 2016, 2017, 2018 AND 2019. LOADING, ARCHIVING AND AUTO-ROUTING OF SUCH IMAGES WILL FAIL AS SOME OF THE IMAGES ARE CORRUPTED AND UNREADABLE. PRIOR IMAGES ARE NO AVAILABLE FOR COMPARISON EITHER. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATIONS. THE REPORTED ISSUE AT THE CONCERNED SITE IS CAUSED BY OUTDATED BSOD - MICROSOFT SOFTWARE. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID #(B)(4). | fail to save report | loss of data | during file transfer process | power outage | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH THE POTENTIAL OF DATA PATIENT MIX-UP FROM A SITE IN (B)(6). PER THE COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. HOWEVER, SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE AS LOW RISK FOR THE PRODUCT AND RISK FOR THE SITE. IT WAS REPORTED THAT PATIENT DATA GOT MIXED UP FOR ABOUT 3 DAYS DURING 2 OCCURRENCES. THE MIX-UP WAS CORRECTED AND NO IMPACT FOR PATIENTS WAS REPORTED. SOME OF THE DATA MIX-UP WAS DETECTED BY THE USER. THE COMPLETE AMOUNT OF DATA MIX-UP WAS DETERMINED AND ALL DATA WAS CORRECTED. THE SYNGO.PLAZA PRODUCT IS WORKING AS DESIGNED AND INTENDED; THE ESTABLISHED MITIGATION TO ROLL BACK A FAILED STORING ACTIVITY IS CURRENTLY IN PLACE AND EFFECTIVE UNDER NORMAL CIRCUMSTANCES. HOWEVER, THE IMPLEMENTATION DID NOT CONSIDER SUCH AN EXCEPTIONAL SCENARIO AS IN THIS CASE. THEREFORE AN INVESTIGATION HAS STARTED TO DETERMINE WHETHER THE MITIGATION CAN BE ENHANCED. | data mix-up | data issue | data patient mix-up | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH TWO ORACLE DATABASE PARTITIONS ON THE SYNGO IMAGING SYSTEM. THE ORACLE DATABASE PARTITION IS A TABLE THAT CONTAINS DICOM INSTANCE DETAILS FOR THE SYNGO IMAGING ARCHIVE. THESE PARTITIONS ARE NOT ACCESSIBLE DUE TO I/O ERRORS IN THE FILESYSTEM. IMAGES MAINTAINED IN THESE PARTS OF THE DATABASE CANNOT BE OPENED. NO IMPACT ON CLINICAL ROUTINE AT THE SITE HAS BEEN OBSERVED. THE USER CONTINUES TO READ AND SAVE CURRENT EXAMINATIONS. HOWEVER, IF PRIOR IMAGES ARE NEEDED FOR A FOLLOW UP READING AND THEY ARE LOCATED IN THE INOPERABLE PARTS OF THE DATABASE, IT WOULD NOT BE POSSIBLE TO ACCESS THEM LEADING TO A REDUCED BASIS FOR CLINICAL DIAGNOSIS. THE ROOT CAUSE OF THE REPORTED ISSUE IS A HARDWARE ERROR. THE SYNGO IMAGING PRODUCT DID NOT CONTRIBUTE TO THE UNAVAILABILITY OF IMAGES. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/19. ACTIVITIES TO REPAIR THE DEFECTIVE DATABASE AT THE USER SITE AS WELL AS REVIEW FOR POTENTIALLY MISSING DATA AND ITS FURTHER RESTORATION ARE ON-GOING. THE FINAL AMOUNT OF DATA LOSS (IF ANY) CAN BE DETERMINED ONLY AFTER THE RECOVERY ACTIVITIES AT SITE ARE COMPLETED. NO CONSEQUENCES FROM THE USER HAVE BEEN REPORTED SO FAR. CUSTOMER'S ADDRESS: (B)(6). | inaccessible database partition | application program issue | database can not open | I/O error in filesystem | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF DATA LOSS CAUSED BY A HARDWARE MAINTENANCE ISSUE. OUR FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITH NO RISK FOR THE PRODUCT WITH POTENTIAL RISK FOR THE USER. THE DRIVES D:\ (INSTALLATION DRIVE -VIRTUAL), F:\ AND G:\ ON THE SYNGO.PLAZA SERVER WERE EITHER NOT ACCESSIBLE OR EMPTY. FIVE HDD'S WERE FOUND TO BE DEFECTIVE ON THE STORAGE/DATASTORE FOR THE PLAZA VM, WHICH IS CONFIGURED FROM RHDASSRV1. THESE DRIVES HOLD THE PARTITIONS OF THE SYNGO PLAZA SERVER. NAS1 SERVER/LONG TERM STORAGE (LTS) HAS 2 HDD'S NOT ACCESSIBLE/EMPTY AS WELL. TOTAL OF 7 HDD'S ARE EITHER DEFECTIVE/EMPTY OR NOT ACCESSIBLE BY SERVERS. ALL PRIOR INFORMATION IS LOST. CLINICAL DIAGNOSIS MAY BE POTENTIALLY AFFECTED AS OLD ARCHIVED DATA IS NOT AVAILABLE TO PRIOR COMPARISON. THE REPORTED DATA LOSS WAS NOT RELATED TO THE SYNGO.PLAZA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED INCIDENT OCCURRED IN (B)(6). Manufacturer Narrative: THE ADMINISTRATOR MANUAL HAS A CAUTION "PLAZA_RISK_DOCUPS", INFORMING USER TO CONNECT ALL SYSTEM COMPONENTS TO AN UNINTERRUPTIBLE POWER SUPPLY (UPS). ANOTHER CAUTION - "PLAZA_RISK_DOCDATABASENOTAVAILABLE" - EXISTS TO INDICATE THE NEED TO MAINTAIN BACKUP TO AVOID DATA LOSS DURING HARDWARE FAILURES. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. | inaccessible server | data issue | unavailable data to compare | hardware maintenance issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE (B)(6) SYSTEM. HDDS AT THE CLINICAL FACILITY WERE DAMAGED OR BECAME UNAVAILABLE DUE TO POWER LOSS. THIS SCENARIO CORRUPTED USER'S ENTIRE RAID STORAGE VOLUME (HOSTING SYSTEM AND DATA). AFTER SYSTEM RECOVERY WAS PERFORMED AND DATA WAS RESTORED FROM ALL AVAILABLE SOURCES, THE TOTAL EXTENT OF DATA LOSS WAS DETERMINED AS FOLLOWS: (B)(4) STUDIES, (B)(4) SERIES, (B)(4) IMAGES. THESE (B)(4) STUDIES ARE DATED WITHIN THE TIME PERIOD OF (B)(6) 2002 TO (B)(6) 2019. THEY ARE ASSESSED AS PARTIALLY CLINICAL RELEVANT. THE (B)(6) IS WORKING AS DESIGNED AND INTENDED, AND THE REPORTED DATA LOSS WAS UNRELATED TO THE (B)(6) PRODUCT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE (B)(6) IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE (B)(6) PRODUCT. THE (B)(6) PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS USER. (B)(6). | hard drive failure | data back-up problem | unavailable data to compare | hardware damaged | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE SYNGO PLAZA SYSTEM. THE ONLINE CONSISTENCY CHECK (OCC) TOOL REPORTED DATA LOSS IN ONLINE PARTITION (FILES NOT AVAILABLE AND FILES CORRUPTED). THERE WERE MULTIPLE OCCURRENCES OF BLUE SCREEN OF DEATH (BSOD) 0X51 ERRORS CAUSING LOSS OR CORRUPTION OF IMAGE FILES. SOME OF THE MISSING/CORRUPT PATIENT DATA WAS SUCCESSFULLY RECOVERED. HOWEVER, THE TOTAL EXTENT OF REPORTED DATA LOSS IS AS FOLLOWS: 28 PATIENTS. 28 STUDIES: ALL STUDIES ARE DATED BETWEEN (B)(6) 2016 AND (B)(6) 2018. 53 SERIES. 1577 IMAGES. THE AFFECTED DATA WAS STORED ON THE SHORT TERM STORAGE (STS) AND WAS NOT YET ARCHIVED. THE STUDY REPORTS WERE NOT YET GENERATED IN THE SYNGO PLAZA SYSTEM. SINCE THE STUDIES ARE FROM RECENT TIMELINE, THEY ARE STILL ASSESSED AS CLINICALLY RELEVANT FOR FOLLOW-UP DIAGNOSIS. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. STS IS INTENDED ONLY FOR TEMPORARY STORAGE OF PATIENT DATA UNTIL IT IS TRANSFERRED TO LTS (LONG TERM STORAGE). THE SYNGO.PLAZA IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE SYNGO.PLAZA PRODUCT. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS USER. (B)(6). | blue screen of death | device inoperable | loss of data | system crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SIEMENS BECAME AWARE OF FILES BEING CORRUPTED ON THE SYNGO.PLAZA SYSTEM WITH SOFTWARE VERSION (B)(4). FROM THE AVAILABLE LOGS, IT WAS OBSERVED THAT ALL SYNGO.PLAZA SERVICES SHUT DOWN ABNORMALLY AT THE TIME OF FILE CREATION. THIS RESULTED IN CORRUPTED FILES FOR STUDIES, WHICH WERE RECEIVED DURING THE SHUTDOWN. IN TOTAL 125 FILES FROM 10 STUDIES WERE CORRUPTED - 100 FILES FROM 2019 AND 25 FILES FROM 2017. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: WHEN THE USER LOADED THE STUDIES, AN ERROR SHOULD HAVE BEEN DISPLAYED TO THE USER INDICATING AN ISSUE. IT WAS CONFIRMED THAT THE SYNGO.PLAZA HAD SENT THE ERROR STATUS TO MODALITY FOR THE CORRUPTED IMAGES. IN THIS CASE SYNGO.PLAZA APPLICATION WORKED AS SPECIFIED. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID # (B)(4). | file corrupted | loss of data | during file creation | service shut down abnormally at the time of file creation | no harm | [5 empty] | FDA2
Event Description: SIEMENS HEALTHINEERS BECAME AWARE OF A DATA LOSS ON THE SYNGO IMAGING SYSTEM. INITIALLY 100 SERIES WERE REPORTED WITH NOTIFICATION "ARCHIVE FAILED". HOWEVER, THE ACTUAL INVESTIGATION SHOWED THAT ONLY 30 IMAGES OUT OF 2 SERIES WERE REALLY LOST DUE TO A HARDWARE FAILURE. THE LOST DATA WAS CLINICALLY RELEVANT; HOWEVER, NO PATIENT RESCANS WERE REQUIRED. THE INCOMPLETE DATA WAS INDICATED TO THE USER AT THE TIME OF IMAGE LOADING. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. NO CONSEQUENCES FROM THIS INCIDENT HAVE BEEN REPORTED. WITH THIS INFORMATION, A DEATH OR SERIOUS INJURY IS HIGHLY IMPROBABLE. | archive failed | data issue | during file creation | hardware failure | no harm | [5 empty] | FDA2
Event Description: SIEMENS RECEIVED FEEDBACK FROM THE CUSTOMER REPORTING THAT THE DISPLAYED STANDARD UPTAKE VALUE (SUV) WAS INCORRECT DURING THE CASE AS IT WAS NOT RECONSTRUCTED FROM THE VERY BEGINNING. THIS OCCURRED WHEN THE CUSTOMER VIEWED THE LIST MODE RECONSTRUCTION DATA WITH MI APPLICATIONS ON THE SYNGO MULTIMODALITY WORKPLACE, WHICH WAS INSTALLED ON CUSTOMER'S BIOGRAPH 40 UNIT. ACCORDING TO THE CUSTOMER, THE DISPLAYED SUV WAS WRONG IN THE REPORTED CASE AS IT WAS NOT BEING RECONSTRUCTED FROM THE VERY BEGINNING. INCORRECT SUV MAY LEAD TO A POTENTIAL MISDIAGNOSIS AND INCORRECT PATIENT TREATMENT. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: **RESUBMISSION OF INITIAL REPORT AS PER FDA ON (B)(6) 2020** OUR EXPERTS WERE ABLE TO REPRODUCE THE REPORTED ISSUE. THE REPORTED ISSUE IS UNDER INVESTIGATION AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. CUSTOMER'S ADDRESS: (B)(6). | incorrect display value | incorrect display | incorrect standard uptake value | it was not reconstructed from the very beginning | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: SIEMENS RECENTLY BECAME AWARE OF DATA LOSS THAT HAD OCCURRED ON THE SYNGO IMAGING SYSTEM. FEBRUARY 12, 2016, AFTER PERFORMING REPAIR ACTIONS ON THE HARD DISC UNIT (RAID) OF THE ARCHIVE THAT IS ATTACHED TO THE SYNGO IMAGING PRODUCT, IT WAS NOTICED THAT 6 STUDIES WERE MISSING AND 6 STUDIES WERE PARTIALLY MISSING. THESE STUDIES RANGED FROM JANUARY 31, 2016 TO FEBRUARY 5, 2016. THIS DATA LOSS WAS DISCOVERED BY A SIEMENS SERVICE ENGINEER WHO WAS PERFORMING THE REPAIR ACTION. THE ENGINEER COMPARED THE ARCHIVE STATUS BEFORE AND AFTER THE REPAIR ACTION AND DISCOVERED DATA LOSS. HOWEVER, THE CUSTOMER ALLOWED TO DELETE STUDIES CREATED PRIOR TO FEBRUARY 6, 2016 ON AFFECTED IMAGE PARTITION ON THE COURSE OF THE REPAIR ACTIONS. THE DATA LOSS WAS CAUSED BY A HARDWARE ERROR AS THE RAID UNIT PARTIALLY CRASHED. THE DEFECTIVE DISC WAS REPLACED BY A (B)(4) SERVICE ENGINEER. THE LOST DATA WAS ABOUT A WEEK OLD. IT IS UNKNOWN WHETHER IT WAS USED FOR CLINICAL DIAGNOSIS. HOWEVER, THE CUSTOMER CONFIRMED THAT DATA FROM THAT PERIOD OF THE TIME CAN BE DELETED. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE REPORTED EVENT WAS EVALUATED BY OUR FACTORY EXPERTS. THE SYNGO IMAGING SYSTEM IS FUNCTIONING ACCORDING TO SPECIFICATIONS AND DID NOT IN ANY WAY CONTRIBUTE TO THE DATA LOSS. THE REPORTED DATA LOSS WAS CAUSED BY A HARDWARE PROBLEM IN THE HARD DISC UNIT (RAID) AT THE CUSTOMER SITE. NO CONSEQUENCES FOR PATIENTS HAVE REPORTED. (B)(6). | missing partial reports | loss of data | during repair action | hardware error: defective disc | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: SINCE 2010, THE IMAGE ARCHIVE IS DONE IN ONE NAS (NETWORK ATTACHED STORAGE), WHICH IS ATTACHED TO THE SYNGO IMAGING XS SERVER. ON (B)(6), THE NAS SYSTEM CRASHED SUDDENLY AND AFTER A RESTARTED, ONE OF THE LOGICAL DRIVES DISAPPEARED FROM THE OPERATING SYSTEM. THE DRIVE HAS BEEN RE-ENABLED AND DURING THE INVESTIGATIONS IT WAS IDENTIFIED THAT SEVERAL ZIP FILES CONTAINING THE PATIENT STUDIES ARE CORRUPTED. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 7/28/20. THERE ARE NO INJURIES ATTRIBUTED TO THE REPORTED EVENT. HOWEVER, THE INVESTIGATION WHETHER CLINICALLY RELEVANT DATA LOSS OCCURRED IS ON-GOING. THE AMOUNT OF DATA LOSS IS CURRENTLY UNKNOWN. ACTIVITIES ARE ONGOING AT SITE TO EVALUATE THE AMOUNT OF DATA LOST AND AFTER THOSE ACTIVITIES ARE FINISHED THE AMOUNT OF DATA LOST WILL BE KNOWN. THIS ACTIVITY WILL TAKE SEVERAL WEEKS. THIS REPORT WAS SUBMITTED JANUARY 19, 2017. CUSTOMER'S ADDRESS: (B)(6). | file corrupted | patient data issue | unavailable image to compare | system crash | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER BELIEVES THAT THE STANDARDIZED UPTAKE VALUES (SUV) MEASUREMENTS ARE NOT AS ACCURATE WHEN THE MPR IS DONE USING A 2D IMAGE. AT THE TIME OF THIS REPORT, THE DEGREE OF INACCURACY AND ITS CLINICAL SIGNIFICANCE HAS NOT BEEN CONFIRMED. Narrative: INVESTIGATION INTO THE ROOT CAUSE OF THE PROBLEM IS UNDERWAY AT THE TIME OF THIS REPORT. | inaccurate value | data issue | standaardize uptake values measurement | using a 2D image | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE CUSTOMER CALLED IN A COMPLAINT ON (B)(6) 2015, STATING THEY HAD SENT TWO ACCESSION NUMBERS OVER THAT SAME MORNING AND ON THE AGILITY SIDE, THE PATIENT NAME IS INCORRECT. UPON INVESTIGATION, THE STUDY USER ID MATCHED SHOWING THE STUDY WAS IN IMPAX CORRECTLY AND IN AGILITY WITH THE SAME STUDY_USER ID, BUT IT HAD THE WRONG PATIENT DEMOGRAPHICS. AGFA DISCOVERED THERE WERE CASES WHERE SEPARATE PATIENTS EXISTED IN IMPAX WITH THE SAME PATIENT ID. | incorrect patient information | patient data issue | patient name is incorrect | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER COMPLAINT, THAT BASED ON A STITCHING IMAGE PATIENT WAS TAKEN TO SURGERY UNNECESSARILY.Event Description: THE CUSTOMER COMPLAINT, THAT BASED ON A STITCHING IMAGE, WHEREON BROKEN RODS FOR THE SPINE WAS SHOWN, PATIENT WAS TAKEN TO SURGERY. CUSTOMER FOUND THAT SURGERY WAS NOT NEEDED AS RODS IN THE PATIENTS BACK WERE NOT BROKEN. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. (B)(4). Manufacturer Narrative: STITCHING OF X-RAY IMAGES IS OF INTEREST IN CASE OF DISEASE PATTERNS LIKE SCOLIOSIS OR ASYMMETRIES IN THE STRUCTURE OF LEG BONES. THE SYSTEM OPTIONALLY SUPPORTS THE AUTOMATIC AND MANUAL JOINING OF SEVERAL IMAGES TO PRODUCE ONE LARGE IMAGE THIS EVENT OCCURRED SINCE THE OPERATOR FLIPPED ONLY THE FIRST IMAGE MANUALLY OF THE TWO GENERATED IMAGES, THEN BOTH WERE STITCHED TOGETHER, SHOWING A WRONG COMPOSITE (MISALIGNED OR BROKEN RODS IN THE BACK.) BASED ON THIS STITCHING RESULT, A SURGERY WAS DONE, NO CONTROL X-RAY WAS PERFORMED FOR VERIFYING THE ISSUE. (B)(4). | inappropriate stitching image | image display error | stitching of x-ray image to show a wrong composite | user error | unneccessary surgery | 4 |  |  |  |  | FDA2
Event Description: THE CUSTOMER DISCOVERED A MISMATCH IN THE SYNGO.PLAZA PATIENT BROWSER FOR A CERTAIN PATIENT BETWEEN THE STUDY DESCRIPTION AND THE RELATED THUMBNAILS. THIS WAS NOTICED ON TIME; THEREFORE NO ACTUAL PATIENT MIX-UP OCCURRED. THE ROOT CAUSE OF THE PATIENT MISMATCH WAS THE MICROSOFT CLUSTER STARTING APPLICATION PROCESSES ON TWO SYNGO.PLAZA SYSTEMS SIMULTANEOUSLY. AFTER RESENDING THE AFFECTED STUDY, THE ISSUE COULD BE SOLVED BY THE SIEMENS HC CUSTOMER SERVICE. Manufacturer Narrative: THE ISSUE IS A REPORTABLE ADVERSE EVENT IN THE U.S. DUE TO PATIENT DATA MISMATCH, WHICH HAS A POTENTIAL TO BECOME PATIENT MIX-UP, WITHOUT ANY CONTRIBUTION OF THE SIEMENS SYNGO.PLAZA PRODUCT. (B)(6). | inconsistent report | data issue | mismatch between study description and the related thumbnails | microsoft cluster starting application process on two simultaneously | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER HAS CONFIGURED HANGING PROTOCOL (HP) CONFIGURATION OF "FIT TO WINDOW" AS THE ZOOM FACTOR. THE CUSTOMER STATED THAT THERE WAS AN INTERMITTENT ISSUE IN THAT SOME SERIES WERE ZOOMED AND SOME WERE NOT WHEN AXIAL IMAGES ARE LINKED, AND THEN ¨DRAGGED AND DROPPED¨ ONTO THE MONITOR. THE CUSTOMER INDICATED THAT A PARTICULAR REGION WAS NOT INCLUDED IN ONE OF THE IMAGES THAT WERE ZOOMED. THE CUSTOMER INDICATED THAT THE REGION IS CAPTURED IF USING HP WITH A MATRIX OF 1 PRESENTATION. AN ADVERSE EVENT DID NOT OCCURRED, NO MISDIAGNOSIS WAS MADE; NO TREATMENT RENDERED; SAME PHYSICIAN WAS MANUALLY CHANGING THE ZOOM ON THE IMAGES AND NOTICED THAT THE DEFAULT FIT TO WINDOW ZOOM DID NOT SHOW ALL THE PATHOLOGY. Manufacturer Narrative: INVESTIGATION IDENTIFIED A PRODUCT ANOMALY WAS FOUND RESULTING FROM CHANGES IN VALUES OF PIXEL SPACING (0028,0030) DICOM TAG BETWEEN IMAGES. WHENEVER A USER MOVES FROM AN IMAGE WITH ONE TYPE OF PIXEL SPACING TO ANOTHER IMAGE WHICH HAS A DIFFERENT PIXEL SPACING ON THE SAME WINDOW WHERE LINKING IS ENABLED, THEN FIT TO WINDOW ZOOM FACTOR MAY BE INCORRECT. INVESTIGATION OF ISSUE INDICATES THAT IT WOULD BE OBVIOUS TO THE USER THAT IF THE FIT TO WINDOW IMAGES SHOWED THE FULL REGION OF INTEREST AND THE ZOOMED IMAGES DID NOT AS THEY ARE VIEWED AT THE SAME TIME ON THE MONITOR; LINKED IMAGES ARE VIEWED AS A SUPPLEMENT REFERENCE ONLY TO INDIVIDUALLY VIEWED IMAGES WHICH ARE PRIMARILY USED FOR DIAGNOSIS; THERE IS A REFERENCE IMAGE AVAILABLE WHICH WOULD IMMEDIATELY INDICATE TO THE USER THAT SOME OF THE IMAGES ARE ZOOMED. THE ZOOM MAGNIFICATION VALUE IS SHOWN ON THE IMAGES. A ¨LINKED ICON¨ IS ALSO SHOWN ON THE SERIES/IMAGES THAT ARE LINKED TO REMIND THE RADIOLOGIST THEY ARE VIEWING LINKED IMAGES. IN CONCLUSION, THE RISK REPORTED ISSUE HAS BEEN REVIEWED AND IT DOES NOT POSE A RISK THAT WOULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT, IF IT WERE TO RECUR, WOULD NOT BE LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | incorrect zoom factor | use of incorrect control settings | linked images are viewed as supplement reference only | software bug | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED A DELAY IN SURGERY (X 8 DAYS) FOLLOWING A MISSED FRACTURE. PATIENT REQUIRED HOSPITALIZATION (UNKNOWN TIME) POST SURGERY. THE PATIENT PRESENTED IN THE ER ON (B)(6) 2018 WITH ACUTE TRAUMA. AN XR ANKLE RIGHT COMPLETE WAS ORDERED BY EMERGENCY ROOM CLINICIAN. THREE VIEWS WERE OBTAINED ON A LONG CASSETTE, PROVIDING ANATOMY BEYOND ANKLE TO INCLUDE MID SHAFT CALF. THESE VIEWS WERE SUBMITTED FOR INTERPRETATION ON THE RA 1000 AFTER THEY WERE MAGNIFIED BY THE TECHNOLOGIST AND MANUALLY SAVED. THE RADIOLOGIST RECEIVED THE EXAM IN THE MOST RECENT SAVED STATE (MAGNIFIED) AND PROVIDED REPORT BASED ON THIS PRESENTATION. THE RADIOLOGIST HAD THE OPTION TO VIEW IMAGES AS ACQUIRED AT ALL TIMES. THE FINDINGS IDENTIFIED NO RADIOGRAPHIC EVIDENCE OF ACUTE DISPLACED "OSSIOUS" FRACTURE OR DISLOCATION INVOLVING THE RIGHT ANKLE. THE PATIENT SOUGHT REPEATED CARE ON (B)(6) 2018, WITH A HISTORY OF RIGHT LOWER LEG PAIN. AS A RESULT, THE CLINICIAN ORDERED A XR LEG RIGHT LOWER TIBIA/FIBULA EXAM. FOUR VIEWS WERE OBTAINED. THESE VIEWS WERE SUBMITTED FOR INTERPRETATION ON THE RA 1000, REVEALING "A TRANSVERSE FRACTURE THROUGH THE TIBIAL SHAFT." THE INITIAL IMAGES FROM (B)(6) 2018 WERE RE-ASSESSED | incorrrect report | data issue | fail to identify evidence of acute displaced fracture or dislocation | user error: fail to check image | surgery delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED A PATIENT CARDIAC SCAN WAS PROCESSED ON A XELERIS 3.1 WORKSTATION. THE PATIENT WAS DIAGNOSED AS "NORMAL" FOLLOWING AN INTERPRETATION THE PHYSICIAN HAD MADE. THE CARDIAC EXAM WAS PERFORMED AS A PRE-OPERATIVE TEST FOR A GALLBLADDER SURGERY THAT WAS PRESCRIBED FOR THE PATIENT. A WEEK AFTER THE CARDIAC EXAM READING, A GALLBLADDER SURGERY WAS PERFORMED. ON THE SAME DAY OF THE SURGERY, THE PATIENT SUFFERED A HEART ATTACK. AS A RESULT, A CATHETERIZATION WAS PERFORMED THAT DAY, AND A BYPASS SURGERY WAS PERFORMED ON THE PATIENT THE FOLLOWING DAY. THE EVENT IS CURRENTLY UNDER INVESTIGATION. Manufacturer Narrative: PATIENT WEIGHT UNKNOWN AT THIS TIME. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE INVESTIGATION REVEALED THE DEVICE AND INTERFACE PERFORMED AS DESIGNED AND INTENDED WITH A CORRECT RESPONSE. THERE WAS NO MALFUNCTION OF THE DEVICE AND THE DATA WAS ACCURATE. THE CAMERA AND SOFTWARE PERFORMED TECHNICALLY AS INTENDED WITH NO MALFUNCTION OR FAILURE. THE ROOT CAUSE WAS IDENTIFIED AS UNINTENDED MISINTERPRETATION OF SIGNIFICANT CLINICAL INFORMATION BY USER. | incorrect interpretation | use of device issue | the patient was diagnosed as normal | user error | heart attack | 4 |  |  |  |  | FDA2
Event Description: THE CUSTOMER REPORTED CANCER WAS NOT DETECTED IN A TIMELY MANNER DUE TO NOT VIEWING THE ADDITIONAL SERIES THAT WAS OBTAINED. AN MRI STUDY OF THE LEFT KNEE WAS PERFORMED ON (B)(6) 2019. IT WAS ORDERED STAT. THE FULL STUDY, CONTAINING 9 SERIES WAS SENT INTO THE SITES CPACS SYSTEM WITHIN 30 MINUTES OF ACQUISITION. USING UNIVERSAL VIEWER, THE RADIOLOGIST OPENED THE STUDY ON (B)(6) WHERE A LESION IN THE DISTAL FEMUR SHAFT WAS IDENTIFIED. DEDICATED RADIOLOGY OF FEMUR WAS RECOMMENDED. IN (B)(6) OF 2019, THE RADIOLOGIST WHO READ THE REPORT IDENTIFIED NEED FOR MORE URGENT FOLLOW UP IMAGING AFTER A RE-REVIEW OF MRI KNEE STUDY. IT WAS IDENTIFIED THAT THE 9TH SERIES HAD NOT BEEN VIEWED ON (B)(6). FURTHER TESTING WAS PERFORMED ON (B)(6), WHERE A PET SCAN WAS DONE TO CLASSIFY THE LESION AS CANCEROUS. Manufacturer Narrative: (B)(4). LEGAL MANUFACTURER: HCIT (B)(4) - (B)(4). INITIAL REPORTER: (B)(6). GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | fail to read report | patient data issue | the 9th series had not been viewed | not viewing the additional image series | fail to detect cancer | 4 |  |  |  |  | FDA2
Event Description: THE CUSTOMER REPORTED THAT A PATIENT MAY HAVE HAD AN UNNECESSARY SURGERY. THE CUSTOMER ALLEGES THE EVENT OCCURRED DUE TO THE SURGEON NOT FINDING THE RADIOLOGY REPORT IN PACS. THE CUSTOMER STATED THE SURGEON SAW THE IMAGES IN [?]VERIFIED' STATUS IN PACS AND CHOSE TO REVIEW HIMSELF. THIS RESULTED IN THE SURGEON DETERMINING SURGERY WAS APPROPRIATE. AT THE TIME OF THIS EVENT THE RADIOLOGIST'S REPORT WAS COMPLETE AND AVAILABLE IN RISI AND THE CUSTOMER'S HIS. NO FURTHER INFORMATION AVAILABLE AT THIS TIME. | fail to read report | patient data issue | surgeon determined surgery was appropriate | user error | unneccessary surgery | 4 |  |  |  |  | FDA2
Event Description: THE CUSTOMER REPORTED THAT INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER USED A SCENARIO WHICH WAS NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | display to different patient image | patient data issue | click directly on a patient name in the administration screen | software defects: this sscenairo was not indicated in user manual | user confused |  |  |  |  | add the scenairo in the user manual | FDA2
Event Description: THE CUSTOMER REPORTED THAT PANORAMIC X-RAY IMAGES ACQUIRED AT THE END OF THE WORD DAY WERE DUPLICATED AND SAVED IN THE WRONG PATIENT PROFILE. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER CONTACTED GENDEX TECHNICAL SUPPORT ON (B)(6) 2015. GENDEX TECHNICAL SUPPORT UPDATED SOFTWARE DRIVER GXPICTURE FROM VERSION 3.4.1 TO 3.5.3. THE UPDATE TO THE SOFTWARE DRIVER RESOLVED THE ISSUE OF DUPLICATING IMAGES AND THE CUSTOMER REPORTED THAT EVERYTHING IS NOW WORKING CORRECTLY. NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | save image to different patient | patient data issue | save to wrong patient profile | software defect: software driver upgrade | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED (LIST IS JUMPING) WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: INVESTIGATION REVEALED THAT THE CUSTOMER USED THE SOFTWARE IN A MULTI-WORKSTATION ENVIRONMENT WHICH IS A SCENARIO NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. TECHNICAL SUPPORT REBUILT THE SQL DATABASE AND RETESTED WITH SUCCESSFUL RETREIVAL OF CORRECT PATIENT FILE. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. Manufacturer Narrative: INITIAL INVESTIGATION CONFIRMED REPORT OF "INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE (B)(6) (PATIENT) ADMINISTRATION SCREEN". A RESOLUTION WAS PROVIDED TO THE CUSTOMER. HOWEVER, FURTHER INVESTIGATION IS NEEDED TO DETERMINE ROOT CAUSE OF THE REPORTED ISSUE. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON INVESTIGATION COMPLETION. | display to different patient image | incorrect display | when clicking directly on a patient name in the administration screen | software defect: not a normal scenairo | user confused |  |  |  |  | add the scenairo in the user manual | FDA2
Event Description: THE CUSTOMER REPORTED THAT THE MOUSE NAVIGATES THROUGH SYSTEM WAS FROZEN. IF THE MOUSE AND THE KEYBOARD ARE NOT WORKING, THE SYSTEM WILL NOT BOOT. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN OVER THE PHONE INVESTIGATION. THE REPORTED ISSUE COULD NOT BE DUPLICATED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | mouse navigation frozen | device operates differently than expected | system will not boot | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THAT THE STAT PRIORITY ICON WAS NOT DISPLAYED IN THE STUDY LIST WHEN THE STUDY WAS OPENED. NO PATIENT HARM WAS REPORTED DUE TO THIS ISSUE. Manufacturer Narrative: MCKESSON COMPLETED AN INVESTIGATION AND CONCLUDED THAT THE ROOT CAUSE OF THE PROBLEM WAS A SOFTWARE DEFECT. MCKESSON FOUND THAT THE STAT PRIORITY ALERT ICON MAY NOT BE DISPLAYED IN THE LIST, FOR ANY USER, WHEN ALL STAT STUDIES ARE OPENED AT THE SAME TIME AND THERE ARE NO "HIGH" PRIORITY STUDIES IN THE UNREPORTED STUDY LIST. UNDER THESE CONDITIONS, THE STAT PRIORITY BUTTON WILL NOT DISPLAY THE CORRECT NUMBER OF STAT STUDIES. MCKESSON WILL WORK WITH AFFECTED CUSTOMERS TO APPLY A SOFTWARE UPDATE TO PREVENT THE RECURRENCE OF THE PROBLEM. | fail to display icon | incorrect display | stat priority alert icon may not be displayed in the list | software defect | user confused | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM WOULD NOT RESPOND TO KEYBOARD OR MOUSE INPUTS AND THE SYSTEM WOULD NOT BOOT UP. THERE WAS NO PT INJURY OR DEATH REPORTED. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ONSITE INVESTIGATION. THE KEYBOARD WAS REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | unrespond to keyboard | human-device interface issue | fail to respond inputs | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THAT WHEN VIEWING STUDIES IN CHEV, THEY DID NOT RETAIN THE PRESENTATION STATE FROM THE SAME STUDY IN PACS. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. Manufacturer Narrative: THE CUSTOMER REPORTED THAT WHEN VIEWING STUDIES IN CHANGE HEALTHCARE ENTERPRISE VIEWER (CHEV), IT DID NOT RETAIN THE PRESENTATION STATE FOR THE SAME STUDY SAVED IN PACS. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. THE INVESTIGATION CONCLUDED THAT WHILE THIS ISSUE WAS THE RESULT OF A MALFUNCTION, IT WAS NOT LIKELY TO RESULT IN A REPORTABLE ADVERSE EVENT OR SERIOUS INJURY. FURTHER INVESTIGATIONS DETERMINED THIS IS CAUSED BY A SOFTWARE DEFECT THAT WHEN IMAGING STUDIES ARE VIEWED IN CHEV VERSION 2.1, THE STUDIES DO NOT RETAIN THE IMAGE STYLES DEFINED BY PRESENTATION STATES CREATED AND SAVED IN CHANGE HEALTHCARE RADIOLOGY SOLUTIONS (CHRS). IMAGE STYLES INCLUDE THE FOLLOWING ATTRIBUTES OF IMAGE DISPLAY: ORIENTATION, WINDOW/LEVEL, AND ROI (ZOOM/PAN). CHANGE HEALTHCARE WILL BE NOTIFYING AFFECTED CUSTOMERS OF THIS ISSUE AND PROVIDING A SOFTWARE PATCH TO CORRECT THE ISSUE. | fail to retain image styles | data issue | did not retain the presentation state | software defect | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THE INSTATRAK 3500 COMPUTER TOWER WOULD NOT TURN ON AND FAILED TO BOOT UP. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE POWER SUPPLY WAS REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | fail to boot up | device inoperable | computer tower not turn on | power supply issue | no harm | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THE MOUSE WAS NOT WORKING AND NO KEYBOARD WAS ATTACHED. WHEN THE MOUSE AND THE KEYBOARD FAILED, THE SYSTEM WILL NOT BOOT. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE FAILURE COULD NOT BE DUPLICATED. AS A PRECAUTION THE MOUSE WAS REPLACED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | unresponsive mouse click | human-device interface issue | system will not boot | hardware failure: mouse issue | no harm | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THE PATIENT WAS DISCHARGED FROM INPATIENT HOSPITAL STAY FOLLOWING RADIOLOGIST REVIEW OF HISTORIC ABDOMINAL CT EXAM FROM (B)(6) 2008 THAT CLEARED PATIENT OF ANY PATHOLOGY. AN ABDOMINAL CT EXAM WAS ALSO PERFORMED ON PATIENT ON (B)(6) 2017. THE RADIOLOGIST REVIEWED THE (B)(6) 2008 EXAM AS THE CURRENT EXAM AND PROVIDED REPORT OF FINDINGS OF THIS IMAGE REVIEW ON THE (B)(6) 2017 REPORT. THE PATIENT THEN CAME BACK INTO ER (EMERGENCY ROOM) 2 DAYS LATER, ON (B)(6) 2017 WHERE SHE WAS DIAGNOSED WITH A LARGE ABSCESS WHICH REQUIRED IMMEDIATE SURGICAL INTERVENTION AND 2 DAYS HOSPITALIZATION. THE CUSTOMER REPORTS THAT THE PATIENT IS FULLY RECOVERED. THE CUSTOMER REPORTS THAT THE EXAMS DISPLAYED ON MONITORS IN A GE DEFAULT DDP (DEFAULT DISPLAY PROTOCOL) WHICH WAS NOT WHAT USER EXPECTED. Manufacturer Narrative: CORRECTED DATA - REPORT SOURCE IDENTIFIED AS HEALTHCARE PROFESSIONAL. CORRECTED DATA - DATE OF DEVICE MANUFACTURE CHANGED TO DEC. 2, 2014. ADDITIONAL MANUFACTURER NARRATIVE - EVENT PROBLEM EVALUATION CODES ADDED. INVESTIGATION COMPLETED. ROOT CAUSE WAS DETERMINED TO BE A USE ERROR. NO PRODUCT DEFECT IDENTIFIED. | incorrect file applied | patient data issue | inappropriate surgical intervention | user error | unneccessary surgery | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED THE PATIENT WAS HOSPITALIZED FOR A SMALL BOWEL OBSTRUCTION AND PAIN. A CT WAS ORDERED FOR THE PATIENT. THE CT SCAN WAS PERFORMED ON (B)(6) 2016. THIS STUDY WAS VIEWED BY RADIOLOGIST ON (B)(6) 2016 WHEN AN EXAM NOTE WAS PERFORMED BY THE RADIOLOGIST THAT INDICATED: SBO WITH FREE AIR AND FLUID. THE WARD DOCTOR WAS NOTIFIED OF THESE FINDINGS ON (B)(6) 2016. THE WARD DOCTOR DETERMINED THAT THE PATIENT WAS TOO FRAIL FOR SURGERY, SO NO SURGERY WAS PERFORMED. THE PATIENT DIED ON (B)(6) 2016. THE FINAL REPORT (SIGNED (B)(6) 2016) INDICATED THAT THERE WAS A DELAY IN READING BECAUSE THE STUDY WAS INCORRECTLY DATE AND TIME STAMPED Manufacturer Narrative: THE INVESTIGATION REVEALED THAT THE MERGING OF TWO PATIENT STUDIES WAS PERFORMED CORRECTLY AND THAT THE STUDY HAD THE APPROPRIATE TIME AND DATE STAMP AFTER THE MERGE. THE INVESTIGATION ALSO REVEALED THAT THE EXAM WAS PRESENT ON THE WORKLIST AND AVAILABLE FOR USE BY THE RADIOLOGIST. THE INVESTIGATION DID NOT IDENTIFY ANY DEFICIENCY WITH THE CPACS PRODUCT. THE CPACS PRODUCT IS WORKING AS DESIGNED. Manufacturer Narrative: PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS. THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. MANUFACTURE DATE IS UNKNOWN. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | incorrect report date or time stamped | data issue | delay in reading | software issue | treatment and diagnostic delayed | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THAT THE DEMOGRAPHICS MANAGER (VERSION 12.0.1-BUILD-03) CAUSED REPORT DATA MISATTRIBUTION. SEVERAL RESULT PDFS DID NOT CROSS TO THE CUSTOMER'S HIS, MEDITECH. ROOT CAUSE AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV. HOWEVER, THE PATIENT DEMOGRAPHIC DATA MAY CHANGE AFTER THE REPORT IS SIGNED AND THIS NEW INFORMATION SHOULD BE REFLECTED IN THOSE PREVIOUSLY-SIGNED SYSTEM REPORTS. THE DEMOGRAPHICS MANAGER UPDATES PREVIOUSLY SIGNED REPORTS (XML AND PDF FILES) WITH THE REVISED PATIENT DEMOGRAPHICS INFORMATION AND RESENDS THE INFORMATION TO THE PROPER EXTERNAL SYSTEMS. THE DEMOGRAPHICS MANAGER IS ABLE TO REVISE THE REPORTS WITHOUT HAVING PHYSICIANS RE-SIGN THEIR PREVIOUSLY SIGNED REPORTS. THE DEMOGRAPHICS MANAGER SERVICE WILL ALWAYS REGENERATE THE PDFS OF ALL SIGNED REPORTS FOR A PATIENT AND/OR STUDY. THE ARCHIVE VERSIONS OF THE REPORT ARE THE TWO COPIES OF XML FILES LOCATED ON THE REPORT STORAGE SERVER AND THE BACKUP SERVER. A PRODUCT DEFECT IN THE DEMOGRAPHICS MANAGER CODE MAY CAUSE SOME REPORTING OBJECTS TO BE POPULATED WITH ANOTHER REPORT'S DATA CAUSING MISATTRIBUTION. UNDER NORMAL CONFIGURATION/OPERATION, THE DM SERVICE FINISHES THE DEMOGRAPHIC UPDATES WITHIN A 120 SECOND WINDOW. IF HOWEVER A LARGE NUMBER OF JOBS ARE QUEUED UP AND THEY DO NOT FINISH PROCESSING WITHIN THE 120 SECOND WINDOW, MISATTRIBUTION MAY OCCUR. | incorrect file applied | data issue | patient demographic data may change | software defect | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THAT THEY DISCOVERED ON (B)(6) 2015, CARDIO ECHO RESULTS FROM THEIR AGFA SYSTEM PROCESSED UNSUCCESSFULLY INTO THEIR EPIC SYSTEM. REPORTS AT THE TIME OF PHYSICIAN SIGNATURE WERE CORRECT, BUT THE IMPAX CV 12.1 DEMOGRAPHICS MANAGER DID NOT PRODUCE THE CORRECT PDF REPORTS. AGFA CARDIO TECHNICAL SUPPORT CONFIGURED AN AUTO CORRECTION ON AGFA REPORT AND THE SYSTEM PERFORMED BETTER. AT THIS TIME, AGFA'S INITIAL RISK MANAGEMENT PROCESS AND EVALUATION OF SEVERITY/FREQUENCY RATING DID NOT RESULT IN CFR PART 803 REPORTING. AGFA TECHNICAL SUPPORT PROCEEDED AND CHANGED THE STORED PROCEDURE THAT DEMOGRAPHICS MANAGER USES AND RESTARTED DEMOGRAPHICS MANAGER. ALL PDF REPORTS OUTLINED BY THE CUSTOMER WERE REVIEWED AND CORRECTED. ON (B)(6) 2015, AFTER MONITORING THE SYSTEM, FURTHER REVIEW AND REPROCESSING OF REPORTS, AGFA TECHNICAL SUPPORT DETERMINED SOME REPORTS STILL CONTAINED INACCURATE DATA DUE TO HOW DEMOGRAPHICS MANAGER PROCESSES UPDATES. | incorrect report | data issue | unavailable for review | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE HOSPITAL SITE REPORTED THAT THE CARESTREAM VUE PACS ALLEGEDLY CONTRIBUTED TO A PATIENT DEATH DUE TO THE SYSTEM BEING DOWN AND PATIENT IMAGES NOT ACCESSIBLE. Manufacturer Narrative: CARESTREAM RECEIVED A COMPLAINT ON (B)(6) 2020 AND HAS STARTED THE INVESTIGATION PROCESS. WHEN ADDITIONAL INFORMATION AND FACTS BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED WITHIN 30 DAYS. | fail to access image | patient data issue | patient images not accessible | system crash | death | 5 |  |  |  |  | FDA2
Event Description: THE ISSUE OCCURRED DURING A SIEMENS-HEALTHCARE INSTALLATION UPGRADE OF THE SITE'S (B)(6) A (MLA) SYSTEM FROM SW VERSION VA20 TO VA30. AN ERROR OCCURRED DURING THE INSTALLATION IN THE CONFIGURATION BETWEEN THE MLA AND SYNGO DYNAMICS (SD) PRODUCTS. IT RESULTED IN A CROSSED PATIENT DATA SITUATION FOR 350 PATIENTS. THE PATIENT DEMOGRAPHIC DATA DISPLAYED FOR THE STUDIES IN THE SYNGO DYNAMICS PRODUCT WAS AFFECTED BUT THERE WAS NO CHANGE IN THE IMAGING MODALITIES STUDY PATIENT DEMOGRAPHICS DATA. THE ERROR WAS IDENTIFIED AND THE CONFIGURATION WAS CORRECTED. THERE IS NO REASONABLY FORESEEABLE CLINICAL IMPACT TO THE CROSSED DATA AS THE SITE IS AWARE TO CROSS-CHECK THE PATIENT DEMOGRAPHIC INFORMATION WITH OTHER AVAILABLE SOURCES WHILE THE MANUAL CORRECTIONS ARE TAKING PLACE. THE IMAGING MODALITY STUDY AND PATIENT DEMOGRAPHIC DATA IS NOT AFFECTED. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | incorrect patient data | patient data issue | crossed patient data situation | installation in the configuration | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE ISSUE OCCURRED DURING A SIEMENS-HEALTHCARE PROFESSIONAL SERVICES MIGRATION OF TWO HOSPITAL'S ARCHIVE SYSTEMS INTO ONE (THE NEW PARENT ORGANIZATION SYSTEM). UNRECOVERABLE LOSS OF 13 IMAGE STUDIES DURING THE MIGRATION OCCURRED. THE ISSUE APPEARS TO BE RELATED TO A MISUNDERSTANDING OF THE STATE OF REPORTS BETWEEN THE "CHILD" SITE'S SOURCE SYSTEM AND THE "PARENT" SITE'S. PRODUCTION DESTINATION SYSTEM. FOR THIS MIGRATION, THE EXPECTATION WAS THAT THE "CHILD" STUDIES AND REPORTS WOULD BE INSERTED INTO THE "PARENT'S" PRODUCTION SYSTEM. THE WORKFLOW DISCREPANCY HAS BEEN IDENTIFIED BY SIEMENS. THE CLINICAL RELEVANCE OF THE LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES ARE STILL AVAILABLE IN THE SITES HIS. THERE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO REASONABLY FORESEEABLE PATH TO OCCURRENCE OF HARM CAN BE DETERMINED. | missing study | loss of data | migration of two archive systems into one | user error: misunderstanding of the state of report | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE ISSUE WAS DISCOVERED DURING A SIMPLE REQUEST FOR SERVICE SUPPORT IN DETERMINING WHY ONE PATIENT'S INFORMATION WAS NOT ABLE TO BE CALLED UP FROM THE ARCHIVE BY THE SYNGO DYNAMICS PRODUCT. THE ISSUE WAS ESCALATED BY THE CUSTOMER IN (B)(6) 2017. IT WAS FOUND THAT 155 DIRECTORIES (5170 OBJECTS) WERE MISSING FROM THE SITE'S LONG TERM ARCHIVE SYSTEM. THERE WAS NO INDICATION OR INFORMATION FROM THE SITE OF ANY POLICY OR EVENT WHICH WOULD HAVE CAUSED THIS LOSS. SUBSEQUENT SERVICE INVESTIGATION FOUND A CLOSED SERVICE TICKET DOCUMENTING A HARDWARE FAILURE OF THE SITE'S LONG TERM ARCHIVE SYSTEM (NAS), WHICH WAS ADDRESSED BY A DATA RECOVERY COMPANY IN JANUARY 2017. THE SITE'S NAS IS NOT A SIEMEN'S PRODUCT OR RESPONSIBILITY. | missing report | loss of data | unavailable to report compared | hardware failure | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. A CUSTOMER REPORTED TO MERGE HEALTHCARE SUPPORT ON (B)(6) 2017 THAT THERE WERE "BAD IMAGES PINK/GREN/YELLOW/RED" FOR AN ULTRASOUND IMAGE THAT WAS BEING VIEWED FROM THE ARCHIVE. THE CUSTOMER SITE HAS NOT REPORTED ANY PATIENT SAFETY CONCERNS OR DELAYS, HOWEVER IF THE IMAGES ARE BLURRY OR ARE NOT CONSIDERED TO BE DIAGNOSTIC QUALITY, THERE IS A POTENTIAL FOR PATIENT HARM, DELAY IN TREATMENT OR DIAGNOSIS. (B)(4). Manufacturer Narrative: INITIAL INVESTIGATION FINDINGS DETERMINED THAT THE SITE WAS UTILIZING NEW STORAGE AND OPTIMAL IMAGE/DATA COMPRESSION RULES WERE NOT BEING APPLIED WHEN THE STUDY WAS TRANSFERRED TO THE ARCHIVE. THE STORAGE RULES HAVE BEEN RESOLVED. MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ALLEGATION TO DETERMINE IF ANY ADDITIONAL CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. THIS IS NOT HARDWARE. | poor quality image | poor quality image | unavailable to image compared | image compression rules | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER SITE REPORTED THAT THEY COULD NOT GET A REPORT OUT FOR AN URGENTLY ILL PATIENT UNTIL THE IMAGING CAPTURE SITE'S IDS WERE ADDED. IT TOOK APPROXIMATELY 9 HOURS TO RESOLVE THE PROBLEM WITH THE REPORT. NO REPORT OF INJURY WAS RECEIVED. Manufacturer Narrative: AS PART OF THE NORMAL PROCESS, WHEN THE CUSTOMER ADDS A NEW SECONDARY CAPTURE SITE ID FOR A NEW SITE THEY NEED TO REQUEST IN ADVANCE THAT A TOPLOGO BE CREATED SO THAT THEY CAN GET REPORTS OUT. IF NOT, AS LONG AS THE SITE IS ABLE TO READ/DICTATE CASES THEY SHOULD BE ABLE TO GET A REPORT ON THE CASE ALTHOUGH IT MAY NOT HAVE THE CORRECT ORIGINATING SITE LISTED UNTIL THE CAPTURE SITE ID IS ENTERED - THIS COULD BE EXPLAINED IN THE BODY OF THE REPORT. THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. PROBLEM NOT CAUSED BY A DEVICE PROBLEM. | fail to retrieve study | patient data issue | can not get a report for an urgently ill patient | missing site ID | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT A PATIENT ID WAS ORIGINALLY FOR ONE PATIENT, BUT DUE TO A DIGITIZED MAMMO FROM OUTSIDE FACILITY COMING ACROSS WITH SAME PATIENT ID BUT FOR A DIFFERENT PATIENT, IT CAUSED A DISCREPANCY WITH PRIORS COMING UP FOR BOTH PATIENTS AND PATIENTS BEING LINKED. NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | different patient | patient data issue | same patient ID but for a different patient | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT USING THE "UPDATE DICOM HEADER" OPTION IN UNIVERSAL MANAGER RESULTS IN ALL IMAGES BEING UNVIEWABLE. PANELS ALL DISPLAY "INVALID IMAGE SIGNATURE". NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | invalid image signature | patient data issue | all images being unviewable | use the function of "update dicom header" | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT IMAGES FROM EXAM A WERE SHOWING UP IN MONTAGES FOR EXAM B. NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | display to different exam | incorrect display | images from exam a were showing up in montages fro exam b | software issue | user confused | [5 empty] | FDA2
Event Description: THE MONITOR IMAGE BECAME NOISY WHEN THE USER FACILITY CHECKED THE IMAGE FROM THE ENDOEYE FLEX DEFLECTABLE VIDEOSCOPE LTF-S190-10 IN PREPARATION FOR THE PROCEDURE AFTER THE PATIENT WAS ANESTHETIZED. THE USER FACILITY TURNED OFF AND ON THE SUBJECT DEVICE, AND DISCONNECTED AND CONNECTED THE VIDEO CONNECTOR TO THE SUBJECT DEVICE. HOWEVER, THE PHENOMENON STILL OCCURRED. THE FACILITY COMPLETED THE PROCEDURE WITH REPLACING THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE AND CONCOMITANT THE LTF-S190-10 WERE RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE SUBJECT DEVICE WORKED PROPERLY. ALTHOUGH THE SUBJECT DEVICE CONNECTED TO THE LTF-S190-10, THE REPORTED PHENOMENON COULD NOT BE REPRODUCED. THE SUBJECT DEVICE WORKED PROPERLY AFTER THE USER FACILITY REPLACED THE LTF-S190-10 TO ANOTHER VIDEOSCPE. SO THERE WAS A POSSIBILITY THAT CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE LTF-S190-10 MIGHT OCCUR TEMPORARILY DUE TO ADHESION OF FOREIGN SUBSTANCES ON THE ELECTRICAL CONTACT OF THE VIDEO CONNECTER IN THE LTF-S190-10. | noise emitted | audible noise | monitor image became noisy | adhesion or foreign substances on the electrical contact | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE MONITOR IMAGE WAS DISTURBED DURING THE UNKNOWN PROCEDURE. THE FACILITY COMPLETED THE PROCEDURE TO REPLACE THE SUBJECT OTV-S7PRO TO ANOTHER DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. THE FACILITY WAS AWARE THAT THE PHENOMENON WAS CAUSED BY THEIR HANDLING. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS AND THERE WAS CORROSION ON CONTACT OF THE VIDEO CONNECTOR SOCKET. THE CONTACT WAS CORRODED BECAUSE THE FACILITY MIGHT CONNECT A VIDEO CONNECTOR WITH WET CONDITION OF A CAMERA HEAD TO THE SUBJECT DEVICE. SO THE MONITOR IMAGE WAS FINALLY DISTURBED DUE TO CONTACT FAILURE CAUSED FROM THE CORROSION. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS CAUTIONS FOR VIDEO CONNECTOR HANDLING. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | erratic of intermittent display | erratic display | monitor image was disturbed | corrosion on contact | prolong surgery | [5 empty] | FDA2
Event Description: THE MONITOR IMAGE WAS LOST WHEN THE USER FACILITY INSERTED THE ENDOEYE FLEX DEFLECTABLE VIDEOSCOPE LTF-S190-10 TO THE PATIENT'S BODY CAVITY. THE USER FACILITY TURNED OFF AND ON THE SUBJECT DEVICE, AND DISCONNECTED AND CONNECTED THE VIDEO CONNECTOR TO THE SUBJECT DEVICE. HOWEVER, THE PHENOMENON STILL OCCURRED. THE FACILITY COMPLETED THE PROCEDURE WITH REPLACING THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE AND CONCOMITANT THE LTF-S190-10 WERE RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE SUBJECT DEVICE WORKED PROPERLY. OLYMPUS COULD REPRODUCE THE PHENOMENON WHEN THE SUBJECT DEVICE WAS CONNECTED TO THE LTF-S190-10. THE SUBJECT DEVICE DETECTED CCD OVER CURRENT WHEN THE LTF-S190-10 WAS CONNECTED TO THE SUBJECT DEVICE. THEREFORE, THE LTF-S190-10 MIGHT HAVE SOME MALFUNCTIONS. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED AT THIS TIME. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00231. | fail to display image | no display or display failure | when videoscope inserts to patient's body cavity | hardware issue: videscope CCD over current | prolong surgery | [5 empty] | FDA2
Event Description: THE PHYSICIAN LAUNCHED A CASE IN A FUSION WORKSTATION (FW) INTEGRATED WITH MERGE MAMMO AND ONLY PRIORS CAME UP, NOT THE STUDY REQUESTED. THE PHYSICIAN CLICKED TO VIEW THE (B)(6) 2013 STUDY AND THE FIRST CASE TO LOAD WAS THE (B)(6) 2011 STUDY. THEREFORE THE CURRENT STUDY DID NOT LAUNCH IN THE VIEWER. THERE WAS NO MESSAGE TO THE USER THAT THE RECENT CASE WAS NOT LOADED HOWEVER ALL IMAGES MAINTAINED THE DATE AND TIME STAMP. THE REPORTER IDENTIFIED THIS AS A POTENTIAL SAFETY ISSUE IN THE EVENT THE PHYSICIAN MISDIAGNOSED THE PATIENT, BEING UNAWARE OF THE MISSING STUDIES. Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE ISSUE WAS OCCURRING AT THIS INSTITUTION BECAUSE OF THE PARTICULAR WORKFLOW THEY USE TO LOAD STUDIES. THE NORMAL WORKFLOW IS TO DOUBLE-CLICK ON THE WORKLIST AND ALL OF THE STUDIES WILL LOAD. THE PHYSICIAN WAS TRYING TO SELECT PARTICULAR STUDIES TO LOAD. THE CODE WAS FIXED IN (B)(4) VERSION 3.3.4., AND WAS APPLIED TO THE SYSTEM AT THE INSTITUTION ON (B)(4) 2014. | incorrect report displayed | display misread | user unware of the missing studies | particular workflow to load studies | misdiagnosed | [5 empty] | FDA2
Event Description: THE REPORTED FACILITY CONTACTED MCKESSON TO REQUEST ASSISTANCE WITH AN INVESTIGATION INTO INTERMITTENT DOWNTIME OF THE HORIZON MEDICAL IMAGING (HMI) SYSTEM FOR A SPECIFIC WORKSTATION AT THE SITE. THE SITE REPORTED THAT A PATIENT WAS KEPT UNDER ANESTHESIA 30 MINUTES LONGER THAN PLANNED PRIOR TO TH START OF A SURGICAL PROCEDURE. THE WORKSTATION FORM THE OPERATING ROOM WAS NOT ABLE TO CONNECT TO THE HMI PACS SYSTEM LEADING TO IMAGES FOR THE PATIENT BEING UNAVAILABLE. SOMEONE FROM THE RADIOLOGY DEPARTMENT RETRIEVED THE IMAGES ON FILM FROM THE RADIOLOGY DEPARTMENT. NO PATIENT BEING UNAVAILABLE. SOMEONE FROM THE RADIOLOGY DEPARTMENT RETRIEVED THE IMAGES ON FILM FROM THE RADIOLOGY DEPARTMENT. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: THE REPORTING FACILITY IS MULTI-FACILITY SITE. A SATELLITE FACILITY OF THE REPORTING FACILITY NEEDED A NEW WORKSTATION CONFIGURED AND MCKESSON SUPPORT HAD INADVERTENTLY PROVIDED INCORRECT INFORMATION DURING THE CONFIGURATION OF THE NEW WORKSTATION CONTAINED THE INCORRECT USER NAME AND PASSWORD FOR THE NEW WORKSTATION AT THE SATELLITE FACILITY, THE LOG-IN INFORMATION WAS FOR THE REPORTING FACILITY. WHEN USERS LOG-IN TO THE NEWLY CONFIGURED WORKSTATION AND THE LOG-ON WAS AUTHENTICATED, IT BROUGHT DOWN THE HMI SYSTEM AT THE REPORTING FACILITY SINCE THE USER NAME AND PASSWORD WERE USED TO LOG-IN TO ANOTHER SITE. BOTH THE FACILITIES HAVE CONFIRMED TO MCKESSON THAT THEY HAVE RESOLVED THAT THEY HAVE RESOLVED THIS ISSUE BY THEMSELVES BY CORRECTING THE CONFIGURATION. | fail to retrieve image | communication or transmission issue | under anesthesia 30 minutes longer than planned | incorrect configuration | prolong surgery | [5 empty] | FDA2
Event Description: THE REPORTED ISSUE WAS IDENTIFIED INTERNALLY BY SIEMENS R&D DEPARTMENT ON THE SYNGOVIA SYSTEMS WITH VERSION VB50 IN THE LUNGCAD NAVIGATION TOOL. AN INCORRECT NODULE GETS CORRECTED IN THE LUNGCAD NAVIGATION TOOL AFTER ACCEPTING THE CURRENTLY EDITED NODULE FOR THE 2ND TIME. CORRECTION FOCUS DOES NOT SWITCH TO THE NEW NODULE AND AFFECTS PREVIOUS NODULE, ON WHICH CORRECTION PEN WAS INVOKED. AS A RESULT, INCORRECT INFORMATION MAY CONTRIBUTE TO MEDICAL DIAGNOSIS AS A DIFFERENT NODULE QUANTIFICATION IS CHANGED AND IT COULD LEAD TO A WRONG NODULE ASSESSMENT. WRONG NODULE ASSESSMENT MAY CONTRIBUTE TO INADEQUATE CLINICAL PATIENT MANAGEMENT. THE DESCRIBED ISSUE IS DETECTIBLE TO THE USER; HOWEVER, THE USER MAY NOT REALIZE THAT VALUES HAD BEEN CHANGED FOR THE ACCEPTED FINDINGS. Manufacturer Narrative: SIEMENS STOPPED DISTRIBUTION OF THE AFFECTED SOFTWARE AND IS WORKING ON A CORRECTIVE ACTION FOR THE REPORTED ISSUE. INTERNAL ID#: (B)(4). | incorrect nodule information | data issue | wrong nodule assessment | software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE REPORTING FACILITY ALLEGED THAT THE ECHO IMAGES FOR PATIENT A WERE INADVERTENTLY SENT TO THE FOLDER OF PATIENT B, WHOSE FOLDER SUBSEQUENTLY CONTAINED THE WRONG ECHO IMAGES. THE ISSUE OCCURRED AT 4:20 PM. AS A RESULT OF THIS ISSUE, PATIENT B WAS TRANSFERRED TO THE CARDIAC CARE UNIT (CCU) AT APPROXIMATELY 6:30 PM, AFTER THE DEPARTMENT RECEIVED THE WRONG ECHO IMAGES. THE ISSUE WAS DISCOVERED WHEN THE REPORTING FACILITY STAFF REVIEWED PATIENT B'S RECORDS THAT EVENING. PATIENT A'S ECHO IMAGES HAVE NOT BEEN AFFECTED, SO THERE ARE NO CONSEQUENCES TO PATIENT A AS A RESULT OF THIS ISSUE AND THERE HAS BEEN NO REPORT OF INJURY OR HARM TO THE PATIENT B. Manufacturer Narrative: THE ROOT CAUSE OF THE ISSUE IS CURRENTLY UNKNOWN AND IS UNDER INVESTIGATION. (B)(4) WILL PROVIDE A FOLLOW UP REPORT TO THE AGENCY ONCE ADDITIONAL INFORMATION IS AVAILABLE. | save image to different patient | patient data issue | department receive wrong echo images | user error or software issue | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE REPORTING FACILITY CONTACTED CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM TO DETERMINE WHETHER AN IMAGING STUDY PERFORMED FOR AN EMERGENCY PATIENT HAD BEEN IMPORTED INTO MCKESSON RADIOLOGY PACS FROM THE MODALITY, AS THE REPORTING FACILITY HAD BEEN UNABLE TO LOCATE THE STUDY IN THE SYSTEM. THE REPORTING FACILITY INDICATED THE PATIENT HAD SUBSEQUENTLY EXPIRED. Manufacturer Narrative: CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM SUCCESSFULLY LOCATED THE IMAGING STUDY FOR THE PATIENT IN MCKESSON RADIOLOGY PACS. UPON FURTHER INVESTIGATION, CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM DETERMINED THAT THE STUDY HAD BEEN SENT TO MCKESSON RADIOLOGY FROM AN UPSTREAM SYSTEM WITHOUT AN ASSIGNED MEDICAL RECORD NUMBER (MRN). IN THIS SITUATION, MCKESSON RADIOLOGY PACS IS DESIGNED TO ASSIGN A TEMPORARY MRN USING A DESIGNATED PREFIX UPON IMPORT TO CLEARLY IDENTIFY IMAGING STUDIES NEEDING FURTHER DEMOGRAPHIC REVIEW. USERS ARE EXPECTED TO IDENTIFY AND UPDATE THE TEMPORARY MRN FIELD DURING THE REVIEW WORKFLOW FOLLOWING IMPORT OF AN IMAGING STUDY. MCKESSON RADIOLOGY PACS PROVIDES ADDITIONAL FEATURES TO ENABLE USERS TO SEARCH FOR AND LOCATE IMAGING STUDIES IN THE SYSTEM USING DIFFERENT METHODS; FOR EXAMPLE, BY USING THE ACCESSION NUMBER SEARCH OR USING THE "RECENTLY PERFORMED" STUDY LIST FILTER. MCKESSON RADIOLOGY PACS PERFORMED AS DESIGNED. | unable to locate study | patient data issue | unavailable for review | without an assigned medical record number (MRN) | procedure delayed |  |  |  | additional search method |  | FDA2
Event Description: THE REPORTING FACILITY HAD REPORTED A ONE HOUR DELAY IN CARE FOR A (B)(6) YEAR OLD STROKE PATIENT DUE TO A FAILURE TO IMPORT ALL STUDY IMAGES INTO THE PACS SYSTEM. THE REPORTING FACILITY PERFORMED ALTERNATIVE MEASURES TO VIEW THE STUDY IMAGES TO TREAT THE STROKE PATIENT. NO HARM TO THE PATIENT WAS REPORTED AS A RESULT OF THE DELAY. Manufacturer Narrative: MCKESSON HAS IDENTIFIED A SOFTWARE DEFECT THAT IS ENCOUNTERED WHEN USERS LOGIN TO THE PACS FROM AN EMR, OR USING LEGACY WEB URL. AS A RESULT OF THE SOFTWARE DEFECT, THE USER MAY EXPERIENCE MISSING IMAGES IN A NEWLY IMPORTED STUDY, AND/OR STUDY IMPORTS THAT REMAIN IN AN "IN-PROGRESS" STATUS. MCKESSON WILL WORK WITH AFFECTED CUSTOMERS TO APPLY A SOFTWARE UPDATE TO PREVENT THE RECURRENCE OF THE PROBLEM. | fail to retrieve image | communication or transmission issue | fail to import all study images | software defect: login and user may experiencew missing images | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE SITE REPORTED THAT WHEN THEY ASSIGN A STUDY AS "NEEDS-OVER READ" USING THE AUTO-ASSIGN FUNCTIONALITY IN MCKESSON RADIOLOGY PACS, THE STUDY IS NOT AUTOMATICALLY ASSIGNED TO THE EMERGENCY DEPARTMENT PHYSICIAN GROUP. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: MCKESSON HAS INVESTIGATED THIS REPORTED ISSUE AND CONCLUDED THAT THE ROOT CAUSE OF THE PROBLEM IS A SOFTWARE DEFECT. AS PART OF THE SITE SPECIFIC WORKFLOW, FOR PATIENT IN THE EMERGENCY DEPARTMENT, RADIOLOGISTS SAVE A PREFERENCE SETTING TO AUTOMATICALLY ASSIGN STUDIES WITH POSITIVE OR SUSPICIOUS FINDINGS, TO THE EMERGENCY DEPARTMENT PHYSICIAN GROUP FOR FURTHER REVIEW. DUE TO THE SOFTWARE DEFECT THE PREFERENCE SETTING WAS NOT SAVED RESULTING IN NOT AUTOMATICALLY ASSIGNING SELECTED STUDIES AND ALERTING THE EMERGENCY DEPARTMENT PHYSICIANS FOR FURTHER REVIEW. THIS COULD POTENTIALLY CAUSE A DELAY IN CARE FOR PATIENT REQUIRING FURTHER TREATMENT. A SOFTWARE UPDATE HAS BEEN PROVIDED TO THE REPORTING FACILITY TO RESOLVE THE ISSUE. | fail to assign study | data issue | fail to assign a study as "needs over read" | software defect | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THE STEREOTACTIC IMAGING ON THE CDS WAS NOT WORKING. REQUIRED A TOTAL OF 3 DISCS WITH IMAGING TO BE RETRIEVED, WHILE THE PATIENT WAS UNDER ANESTHESIA (APPROX. 14 HOURS). THIS IS WHAT THE PATIENT SAFETY REPORT STATED: PA WAS ATTEMPTING TO BRING UP THE STEREOTACTIC IMAGE ON DISC ON THE STRYKER UNIT PRIOR TO DR BEGINNING SURGERY. THE DISC WOULD NOT FUNCTION OR OPEN THE IMAGE. THE PATIENT WAS UNDER ANESTHESIA AT THIS POINT WITHOUT AVAILABLE IMAGE. PA CALLED THE FILE ROOM. THE CD BURNER WAS NOT WORKING. HE WAS CONNECTED TO MEDICAL RECORD DEPARTMENT WHO BURNT A DISC FOR HIM. ONCE BACK IN THE OPERATING ROOM, IT WAS REALIZED THAT THE SECOND DISC HAD THE INCORRECT IMAGE BURNT ON IT. WHILE PA WAS RETRIEVING THE 3RD DISC FROM MEDICAL RECORDS, THE STRYKER REP MADE AVAILABLE A SECOND STRYKER UNIT. PRIOR TO PA RETURNING TO OPERATING ROOM WITH A THIRD DISC, THE FIRST DISC WAS ABLE TO BE UTILIZED TO OPEN THE IMAGE ON THE SECOND STRYKER UNIT. DR WAS ABLE TO BEGIN THE SURGERY. IT IS REPORTED THAT ONCE SURGERY BEGAN, THE PATIENT HAD BEEN UNDER ANESTHESIA WITHOUT ANY SURGICAL INTERVENTION FOR APPROXIMATELY ONE HOUR. IT WAS LATER LEARNED THAT THE DISC DRIVE IN THE FIRST STRYKER UNIT WAS DYSFUNCTIONAL. IT IS POSSIBLE THAT IF THE DISC DRIVE WERE TO HAVE WORKED IN THE FIRST STRYKER UNIT, THAT THE FIRST DISC MAY HAVE OPENED THE IMAGE. | unable to read image | patient data issue | patient was under asesthesia without available image | hardware failure: CD drive dysfunctional | surgery delayed | [5 empty] | FDA2
Event Description: THE SUV VALUES THAT ARE BEING CALCULATED IN THE PET/CT FUSION TOOL ARE INCORRECT. THE ANOMALY PRESENTED ITSELF WHEN THE SPHERE ANNOTATION WAS DRAWN OVER THE IMAGE, THE VALUES FOR MINIMUM, AVERAGE, AND MAXIMUM WERE ALL EQUAL OR IN SOME CASES VERY CLOSE. THIS IS AN INDICATION THAT THE VALUES ARE INCORRECT. THE VOI (VOLUME OF INTEREST) ONLY REFLECTS THE SMALL REGION WHERE THE USER FIRST BEGAN DRAGGING THE RIGHT-MOUSE BUTTON TO DRAW THE SPHERE. THIS COULD LEAD TO MISDIAGNOSIS IN PET CASES WHERE THE SUV (MAXIMUM) WOULD DISPLAY A LOWER VALUE THAN IT SHOULD. Manufacturer Narrative: CUSTOMERS HAVE BEEN NOTIFIED OF THE MALFUNCTION AND HAVE BEEN GIVEN CLEAR INSTRUCTIONS ON HOW TO PREVENT THE MALFUNCTION FROM OCCURRING AS WELL AS INSTRUCTIONS FOR REQUESTING A FREE UPGRADE TO A CORRECTED VERSION OF THE SOFTWARE. | incorrect calculation | data issue | the SUV maximum display a lower value | software defect | inappropriate treatment or diagnostic | [5 empty] | FDA2
Event Description: THE USER FACILITY INFORMED OLYMPUS THAT AN ABNORMAL MONITOR IMAGE OCCURRED WITH USING THE SUBJECT DEVICE ON (B)(6) 2015. OLYMPUS ASKED THE USER FACILITY ABOUT THE DETAILED SITUATION WHEN THE PHENOMENON OCCURRED, AND OBTAINED FOLLOWING ADDITIONAL INFORMATION ON FEBRUARY 19, 2016. DURING HYSTEROSCOPY, AN ABNORMAL MONITOR IMAGE OCCURRED WHEN THE USER FACILITY TURNED ON THE SUBJECT DEVICE AFTER THE PATIENT WAS ANESTHETIZED. THE USER FACILITY ABORTED THE PROCEDURE BECAUSE OF THIS PHENOMENON. THERE WAS NO REPORT OF PATIENT INJURY. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. THE REPORTED PHENOMENON WAS REPRODUCED WHEN OLYMPUS JIGGLED THE VIDEO CONNECTOR PLUG OF A CAMERA HEAD UP AND DOWN WITH CONNECTING A CAMERA HEAD TO THE SUBJECT DEVICE. THE SUBJECT DEVICE WORKED PROPERLY WITH REPLACING THE VIDEO CONNECTOR SOCKET TO ANOTHER ONE. THERE WAS CORROSION ON THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET. THE USER FACILITY MIGHT CONNECT THE SUBJECT DEVICE WITH A CAMERA HEAD WHOSE ELECTRICAL CONTACT OF THE VIDEO PLUG WAS NOT DRY AFTER THE CAMERA HEAD WAS REPROCESSED. THEREFORE, THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET MIGHT GET CORRODED, RESULTING IN CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE CAMERA HEAD. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS THAT AN OPERATOR SHOULD MAKE SURE THE VIDEO PLUG AND ITS ELECTRICAL CONTACTS OF A CAMERA HEAD ARE COMPLETELY DRY BEFORE CONNECTING THE PLUG TO THIS DEVICE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | abnormal image | image display error | after the patient was anesthetized | corrosion on contact | procedure aborted | [5 empty] | FDA2
Event Description: THE VIDISTAR PACS AND DICOM VIEWER SOFTWARE SYSTEM IS USED FOR THE TRANSFER, RECORDING, STORAGE, DISPLAY, AND POST-PROCESSING OF DIGITAL MEDICAL IMAGES, INCLUDING DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE (DICOM) DATA, MEASUREMENT QUANTIFICATION AND DATA MANIPULATION. VIDISTAR PACS IS SUPPORTING AN INTERFACE THAT PROVIDES THE IMAGES AND MEASUREMENTS TO A REMOTE EHR SYSTEM. SSHFS IS A SECURE FILE SYSTEM CLIENT USED TO MOUNT AND INTERACT WITH DIRECTORIES AND FILES LOCATED ON A REMOTE SERVER. DURING A STUDY TRANSFER TO THE REMOTE EHR SYSTEM, THE PATIENT IMAGES INCLUDED AN IMAGE FROM ANOTHER PATIENT'S STUDY. THE ADDITIONAL IMAGE WAS LABELED CORRECTLY, BUT WAS MISFILED. THE PROBLEM WAS CAUGHT BY THE TECHNICIAN BEFORE THE STUDY WAS INTERPRETED. Manufacturer Narrative: OUR FIRM INVESTIGATED THE CUSTOMER COMPLAINT OF INTERMIXED IMAGES AND WERE ABLE TO DETERMINE THAT THE PIECE OF SOFTWARE RESPONSIBLE FOR TRANSFERRING THE IMAGES FROM ONE SYSTEM TO ANOTHER (SSHFS VERSION 2.4) WAS INTERMIXING IMAGES UNDER SPECIFIC CIRCUMSTANCES. IMPLEMENTING CHECKSUM LOGIC TO THE APPLICATION'S LOGS REVEALED WHEN IN THE TRANSFER PROCESS IMAGES WERE INTERMIXED. THIS INDICATED THAT THE CULPRIT WAS A BUG IN THE TRANSFER SOFTWARE (SSHFS VERSION 2.4) AND NOT THE MODALITY, VIDISTAR PACS, NOR THE EHR SYSTEM. THE TRANSFER SOFTWARE (SSHFS) IS A THIRD-PARTY OPEN SOURCE SOFTWARE THAT WAS ORIGINALLY CHOSEN AND VALIDATED FOR THE INTERFACE BY VIDISTAR AND THE EHR SYSTEM. FURTHER INVESTIGATION SHOWED INSTABILITY OF THE MOUNT CREATED BY THE TRANSFER SOFTWARE (SSHFS VERSION 2.4) AT THE LEVELS OF TRAFFIC CREATED BY THE INTERFACE THUS OFFERING AN INCREASED LIKELIHOOD OF THE SPECIFIC CIRCUMSTANCES NECESSARY FOR THE BUG TO PRESENT. | image intermixed | data issue | during a study transfer to the remote system | software bug in the transfer software | inappropriate treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THIS IS A COMMON AND WIDESPREAD PROBLEM AFFECTING MANY CASES. WHEN THE PATIENT'S EHR IS OPEN AND BEING SEARCHED FOR RADIOLOGY RESULTS, THERE IS A CAMERA ICON ON WHICH TO CLICK WHICH OPENS THE RADIOGRAPH IMAGES TO BE SEEN. THE PATIENT'S CHEST RADIOGRAPH WAS THE SUBJECT OF THIS QUERY. THE ICON WAS CLICKED AND A NORMAL APPEARING CHEST RADIOGRAPH APPEARED AND OTHER RESULTS WERE THEN ACCESSED. THE PROBLEM IS THAT THE RADIOGRAPH THAT APPEARED AND WAS NORMAL, WAS NOT OF THIS PATIENT, BUT OF SOME OTHER PATIENT THAT WAS NOT EVEN ON MY PATIENT LIST. THE PATIENT IN QUESTION HAD A RADIOGRAPH THAT INDICATED PULMONARY EDEMA AND WAS SOUGHT AFTER THE PHYSICAL LUNG EXAM WAS FOUND TO BE INCONSISTENT WITH A NORMAL CHEST RADIOGRAPH. THE DEFECTIVE PRESENTATION OF ANOTHER PATIENT'S RADIOGRAPH IS COMMON WITH THIS INTERFACE OF DEVICES BUT IF THE IMAGE WAS ALLOWED TO STAY ON THE SCREEN, THE CORRECT IMAGE OF THE RIGHT PATIENT WOULD LOAD. THIS DEFECT CAUSES MISDIAGNOSIS AND WRONG TREATMENTS. | display to different patient image | patient data issue | click on cameria icon to opent the radiograph images | software issue | inappropriate treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: DURING A SITE CHECK IN (B)(6) ("ATRIUM (B)(6)", (B)(6)), WHERE THE (B)(6) SERVER CRASHED, (B)(4) MIXED UP IMAGES WERE DETECTED BY SIEMENS SERVICE PERSONNEL (B)(6) 2014. THIS ISSUE IS IDENTICAL TO THE EARLIER CASE TRACKED WITH COMPLAINT (B)(4), REPORTED AS ADVERSE EVENT UNDER REPORT # 2240869-2014-00003625. FOR THE SITE IN (B)(6) THE FOLLOWING WAS DETERMINED : "TODAY WE HAD AN UNEXPECTED CRASH AT (B)(6). SO ONCE THE SYSTEM WAS OK I RUNNED THE QUERY TO CHECK FOR IMAGE FILENAME DUPLICATES AND WE ALSO HAVE HERE SOME DUPLICATES, BUT FROM THE PAST (B)(6)." (B)(4) ADDITIONAL IMAGES WITH DUPLICATE REFERENCE WERE IDENTIFIED FROM AN EARLIER CRASH (B)(6) 2014. . IN CASE OF AN OPERATING SYSTEM CRASH (BLUESCREEN) OR AN UNEXPECTED CRASH OF THE (B)(6) DICOM SERVICES, EXACTLY DURING WORK-IN-PROGRESS OF STORING RECEIVED IMAGES, IT MAY HAPPEN (ARBITRARILY, DEPENDING ON THE CACHE BEHAVIOR OF THE OPERATING SYSTEM) THAT SOME IMAGE FILENAMES WHICH USED LAST BEFORE THE CRASH ARE USED AGAIN FOR THE NEXT IMAGES RECEIVED AFTER THE SYSTEM RESUMED WORKING. IF THIS HAPPENS, A SITUATION MAY ARISE WHEN IN THE IMAGE TABLE WITHIN THE (B)(6) DATABASE TWO REFERENCES WILL BE POINTING TO THE SAME IMAGE FILENAME. THE REFERENCE FROM BEFORE THE CRASH NOW ALSO POINTS TO THE IMAGE WRITTEN AFTER THE CRASH, WHICH HAS CONTENT FROM AN EXAMINATION/PATIENT AFTER CRASH. THESE REFERENCES ARE UNIDIRECTIONAL, THEREFORE NO TRACE BACK FROM THE IMAGE TO THE PATIENT IS POSSIBLE. | image intermixed | data issue | during storing received images | server crash | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THIS MORNING, THE TELEMETRY SYSTEM (SCOTTCARE) FROZE. STAFF WAS IN THE PROCESS OF ADDING PATIENTS TO THE SYSTEM. AFTER THE THIRD WAS ENTERED THE SYSTEM DID NOT RESPOND. SYSTEM REBOOT OCCURRED PER STAFF. BIOMED WAS CALLED AND THEY CAME TO WORK ON THE SYSTEM. AFTER 3 HOURS OF TIME, THE SYSTEM BECAME FUNCTIONAL. STAFF CALLED TO CANCEL PATIENT FOR THE DAY, AS WE DID NOT HAVE AN IDEA TO HOW LONG THE SYSTEM WOULD BE DOWN, WHILE THE FIXES WERE BEING WORKED ON. | system frozen | device inoperable | during the process of adding patients to the system | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: THIS WAS A (B)(6) PATIENT WITH A HISTORY OF SEIZURES AND A ASTROCYTOMA. THE PATIENT WAS HAVING LASER INTERSTITIAL THERMAL THERAPY PROCEDURE ON HER ASTROCYTOMA. SHE WAS TAKEN TO THE SIEMENS REALTIME MIR FOR THE PROCEDURE WHICH WAS PERFORMED WITHOUT COMPLICATION. THE COMPLICATION OCCURRED WHEN THE LASER PROBE WAS BEING REMOVED. DURING THE COMPLETION OF THE ABLATION PROCEDURE, THE LASER TIP WAS WITHDRAWN FROM THE CRANIUM BOLT AND IT WAS NOTICED THAT THE LASER CRYSTAL, OR STYLUS, WAS NOT ATTACHED TO THE WAND. IT WAS NECESSARY FOR THE PATIENT TO BE TAKEN TO THE OPERATING ROOM AND THE BURR HOLE ENLARGED TO REMOVE THE STYLUS. THE PATIENT WENT HOME THE FOLLOWING DAY AND POST OP FOLLOW UP WAS VERY POSITIVE. THERE WERE NO COMPLICATIONS WITH THE SECOND OPERATING ROOM VISIT TO RETRIEVE THE STYLUS. | tip detached | detachment of device component | during ablation procedure | connection failure | additional surgery procedure | [5 empty] | FDA2
Event Description: UNABLE TO LOCATE FILMS ON MCKESSON SYSTEM. PATIENT WAS DONE IN CATHETER LAB. WORK ORDER PUT INTO IS, BUT THERE WAS AN ERROR MESSAGE ON THE COMPUTER DURING THIS TIME. THE FILMS COULD NOT BE TRANSFERRED OVER UNTIL THE ERROR MESSAGE WAS CLEARED. ONCE THE ERROR MESSAGE WAS CORRECTED, THE FILMS WERE AVAILABLE. ALL FILMS FOR THIS PATIENT ARE PRESENT. THERE IS A PROCESS IN PLACE SO THAT THIS WILL NOT HAPPEN. FILMS ARE TO BE CHECKED DAILY TO BE SURE THEY HAVE TRANSFERRED OVER TO MCKESSON. ANY MISSING FILMS DUE TO ERROR MESSAGES SHOULD BE ADDRESSED IN REAL TIME. CATH LAB ASSISTANT NURSE MANAGER HUDDLED WITH TEAM MEMBERS REGARDING THIS PROCESS AND TO BE SURE THAT THEY ARE FOLLOWING IT DAILY AND WITH ALL CASES. | fail to locate image | patient data issue | fail to transfer film | error message | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WE ARE REPORTING THIS MDR REPORTABLE EVENT ON DECEMBER 22, 2016. ON (B)(6) 2016, THE CUSTOMER REPORTED TO AGFA THAT WHEN USING MPAX CV REPORTING FOR NON-INVASIVE VASCULAR ULTRASOUND (NIV), WHEN BUILDING A NIV REPORT, IN THE CONCLUSION SECTION OF THE UPPER ARTERIAL DUPLEX, IF YOU SELECT ANY CHECK BOXES UNDER THE LEFT ANEURYSM IT PUTS IN THE REPORT THAT THERE WAS NO ANEURYSM. THIS REPORT IS FOR THE 3ND OF THREE EVENTS FOR THE SAME ISSUE AT THIS SITE. THE FIRST EVENT IS REPORTED IN FDA MDR 1225058-2016-00001. THE SECOND EVENT IS REPORTED IN FDA MDR 1225058-2016-00002. ON (B)(6) 2016, AGFA TECHNICAL SUPPORT CONTACTED THE CUSTOMER AND UPDATED HIM THAT ONLY THREE (3) REPORTS WERE AFFECTED AND THOSE THREE (3) REPORTS WERE ALL LISTED AS TEST REPORTS. | incorrect report | data issue | missing aneurysm in the report | software issue | inappropriate treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WE ARE SUBMITTING THIS MEDICAL DEVICE REPORT TODAY DECEMBER 13, 2016. ON NOVEMBER 22, 2016, THE CUSTOMER REPORTED THAT WHILE USING AN NX WORKSTATION WITH SOFTWARE VERSION NX 3.0.8950 SOFTWARE AND SELECTING THE AFFECTED PATIENT/EXAM FROM CLOSED EXAMS, INITIALLY THE WRONG IMAGE WAS LINKED TO THE EXAM AND APPEARED. AFTER A SHORT TIME THE WRONG IMAGE WAS REPLACED BY THE CORRECT IMAGE, HOWEVER THE WRONG IMAGE WAS USED FOR TRANSMITTING TO PACS. TO CORRECT THIS ISSUE, THE USER BACKED UP ALL LOGS AND REBOOTED THE SYSTEM. THE CORRECT IMAGE WAS THEN TRANSFERRED TO PACS. DURING THE INVESTIGATION BY AGFA NOVEMBER 24, 2016 - DECEMBER 1, 2016, THE ROOT CAUSE WAS IDENTIFIED AS A MEMORY LEAK WHICH WAS INTRODUCED IN THE IMAGE PROCESSING SOFTWARE IN NX 3.0.8950. | incorrect image linked to report | data issue | transfer incorrect image to system | memory leak | inappropriate treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WHEN AN FCR IMAGE WAS FORWARDED FROM SYNAPSE PACS TO ANOTHER MANUFACTURER'S PACS, THE IMAGE FOR PATIENT A WAS ALSO FOUND IN PATIENT B'S STUDY IN THE DESTINATION PACS. (PATIENT B'S ORIGINAL IMAGE WAS NO LONGER IN THE STUDY). PATIENT A'S STUDY WAS CORRECT IN THE DESTINATION PACS. BOTH PATIENT'S STUDIES REMAINED CORRECT IN SYNAPSE PACS. NO INJURY WAS REPORTED; DIAGNOSIS COMPLETED USING SYNAPSE PACS, PRIOR TO SENDING. Manufacturer Narrative: THIS ISSUE WAS REVIEWED FOR MDR REPORTABILITY WHEN IT WAS FIRST RECEIVED. BECAUSE THE IMAGES RETAINED IN SYNAPSE PACS WERE STILL CORRECT AND DIAGNOSIS WAS PERFORMED IN SYNAPSE, FMSU DETERMINED THAT THIS WAS A MALFUNCTION THAT WOULD NOT BE LIKELY TO RESULT IN SERIOUS INJURY IN DEATH. MOVING FCR IMAGES TO ANOTHER SYSTEM IS TYPICALLY DONE FOR STORAGE OR TO SEND IMAGES AND REPORT TO A REFERRING PHYSICIAN. IN THE EVENT IMAGE PIXEL DATA WAS IMPROPERLY WRITTEN TO A PATIENT, THE REFERRING PHYSICIAN WOULD BE ABLE TO LIKELY IDENTIFY THE ISSUE WHEN REVIEWING THE RADIOLOGIST'S REPORT. (B)(4) LATER IDENTIFIED THAT THERE ARE SITES WHERE IMAGES ARE SENT TO OTHER PACS FOR DIAGNOSIS (FOR EXAMPLE FROM REMOTE LOCATIONS WHERE A RADIOLOGIST IS NOT AVAILABLE). IN THIS CASE A RADIOLOGIST MAY NOT HAVE ANY INDICATION THAT THE PATIENT IMAGE IS AN INCORRECT ONE, E.G., IF PATIENT A AND PATIENT B IMAGES WERE OF THE SAME ANATOMY. | send report to wrong place | data issue | image of patient a was found in patient b study | software defect | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WHEN ENTERING THE CORRECT ID NUMBER OF A PATIENT WITH A SLASH, X-RAY IMAGES ARE ASSIGNED TO A DIFFERENT PATIENT. IF THE LABORATORY TECHNICIAN DOES NOT NOTICE THIS IMMEDIATELY, THE X-RAY IMAGES RETURN TO THE PACS AS A STUDY WITH THE WRONG NAME AND BIRTH-DATE NUMBER. THIS ISSUE EXISTS ONLY WHEN A NEW PATIENT WITH A SLASH IS ENTERED IN THE PATIENT ID IN THE DATABASE AND STORED FOR USE LATER (E.G: XXX/XXX/XXX,XXX.XXX). LATER WHEN THE USER DESIRES TO OPEN A NEW STUDY AND SEARCH BY NAME, THE NAME CAN BE SEEN BUT WHEN IT IS RETRIEVED EITHER NOTHING IS RETRIEVED OR A DIFFERENT PATIENT IS RETRIEVED. THE PATIENT LIST USES A "/" TO SPLIT NAMES AND WHEN THE NAME ITSELF CONTAINS A "/", THE SOFTWARE FAILS TO SPLIT THE PATIENT NAMES. Manufacturer Narrative: CARESTREAM PERFORMED AN INVESTIGATION AND IDENTIFIED THE ROOT CAUSE ANALYSIS IN CAPA (B)(4). TO MITIGATE RISK TO PATIENT FROM USE OF THE DEVICE, A SOFTWARE DESIGN CHANGE WILL BE IMPLEMENTED IN IMAGE SUITE SOFTWARE VMR8 AND RELEASED IN MANUFACTURING. NO FURTHER REPORTING IS REQUIRED. | assign image to different patient | patient data issue | when a new patient with a slash is entered in patient ID | software defect: splash is used to split names but will be error if the name itself contains a slash | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WHEN THE PATIENT UNDERWENT GENERAL ANESTHESIA DURING UNSPECIFIED PROCEDURE, THE USER FACILITY FOUND A VIDEOSCOPE IMAGE ON THE MONITOR BECAME DARK EVEN THOUGH MENU DISPLAY WAS SHOWN ON THE MONITOR PROPERLY. THE USER FACILITY REPLACED THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE, BUT THE PHENOMENON STILL OCCURRED. AFTER THAT, THE USER FACILITY REPLACED THE SUBJECT DEVICE TO ANOTHER DEVICE AND THE PHENOMENON DID NOT OCCUR. THE USER FACILITY COMPLETED THE PROCEDURE USING THE REPLACED DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE VIDEOSCOPE LOCKING MECHANISM OF THE VIDEO CONNECTOR BROKE AND A VIDEOSCOPE WAS CONNECTED TO THE SUBJECT DEVICE UNSTABLY. THE USER FACILITY MIGHT DISCONNECT A VIDEOSCOPE FROM THE SUBJECT DEVICE WITHOUT PRESSING THE LOCKING LEVER. THEREFORE, THE SPRING WHICH COULD KEEP THE VIDEOSCOPE CONNECTION LOCKED BROKE AND THE VIDEOSCOPE COULD NOT CONNECT THE SUBJECT DEVICE FIRMLY. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS THAT AN OPERATOR SHOULD PUSH THE LOCKING LEVER DOWN TO DISCONNECT THE VIDEO PLUG OF THE VIDEOSCOPE FROM THIS DEVICE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | dark image | image display error | patient underwent general anesthesia | user error: videoscope locking mechanism connector broken | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WHILE EVALUATING A REPORTED ISSUE OF A SINGLE INSTANCE OF DATA LOAD FAILURE, SIEMENS FACTORY AND SERVICE EXPERTS FOUND A DATA LOSS OF 1858 CLINICAL IMAGE STUDIES OVER A TIME PERIOD OF 2012 TO 2016 (ENTIRE DIRECTORIES MISSING BETWEEN JUNE 2015 AND DECEMBER 2016 WITH A FEW OUTLIERS MISSING FROM 2012-2014). THE EXACT CAUSE OF THE DATA LOSS COULD NOT BE DETERMINED. THE HISTORICAL STUDY IMAGES WERE NOT FOUND IN THE PERMANENT ARCHIVE, THEREFORE, THEY ARE CONSIDERED TO BE UNRECOVERABLE AT THIS POINT IN TIME. THE CLINICAL RELEVANCE IS THAT THE HISTORICAL IMAGES ARE NOT AVAILABLE FOR COMPARISON WITH NEWER IMAGES HOWEVER THE CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | missing data | loss of data | unavailable image to comparison | software issue | treatment or diagnostic delayed | [5 empty] | FDA2
Event Description: WHILE PERFORMING PROACTIVE MAINTENANCE OF UNARCHIVED STUDIES, MCKESSON IDENTIFIED 3 UNREPORTED STUDIES FROM THE SHORT TERM STORAGE IN THE HMI PACS THAT FAILED TO ARCHIVE TO LONG TERM STORAGE. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: DURING INTERNAL REVIEW, MCKESSON DETERMINED THAT THIS ISSUE IS SIMILAR TO 8022257-2013-00004, AND IS REPORTABLE UNDER PER 21 CFR 803. MCKESSON HAS IDENTIFIED AN ISSUE WHERE APPLICATION LOGIC LEADS TO DISCREPANCIES BETWEEN THE IMAGE INDEX AND THE CONTENT OF THE IMAGE FOLDER WHICH CAN CAUSE IMAGE LOSS. THE IMAGE FOLDER IS A FILE SYSTEM DIRECTORY WHERE IMAGES ARE STORED; IT ALSO CONTAINS THE IMAGE INDEX. THE IMAGE INDEX INDICATES TO THE APPLICATION A LIST OF IMAGE FILES TO EXPECT FOR USE WITHIN THE APPLICATION. IF THE IMAGE INDEX AND THE CONTENTS OF THE IMAGE FOLDER ARE NOT SYNCHRONIZED, THE FOLLOWING MAY OCCUR: IMAGE MAY BE LOST OR UNAVAILABLE, LEADING TO CLINICAL DECISIONS BEING MADE ON INCOMPLETE STUDIES AND/OR THE STUDY CANNOT BE ARCHIVED. IN THE EVENT OF A HARDWARE FAILURE, UNARCHIVED STUDIES MAY BE LOST LEADING TO PRIOR STUDIES NOT BEING AVAILABLE FOR COMPARISON. THE FOLLOWING FACTORS MAY CONTRIBUTE TO THE OCCURRENCE OF THE DISCREPANCY: NETWORK STABILITY ISSUES (FOR EXAMPLE: POOR PERFORMANCE, INSTABILITY AND COMMUNICATION ERRORS (DROPPED CONNECTIONS, PACKET LOSS)); DISK WRITE FAILURES (CANNOT WRITE TO DISK). THE PRODUCT MAY FAIL TO HANDLE THE OUTCOME OF SUCH FACTORS RESULTING IN THE DISCREPANCIES DESCRIBED ABOVE. | archive failed | data issue | during proactive maintenance | discrepancy between image index and the content of the image folder | no harm | [5 empty] | FDA2
Event Description: X-RAY IMAGES WERE ADDITIONALLY LINKED TO WRONG PATIENT RECORD. Manufacturer Narrative: THE INVESTIGATION FROM THE R&D DEPARTMENT HAS RESULTED: SIDEXIS 4 IS USED IN COMBINATION WITH SIDICOM WLS AND A CENTRAL PACS SYSTEM. SIDICOM WLS IS CONFIGURATED TO SEND AUTOMATICALLY THE IMAGES TO THE PACS SYSTEM AFTER THE ACQUISITION. MULTIPLE SIDEXIS 4 USERS AT ONE WINDOWS PC'S WITH INDIVIDUAL USER ACCOUNTS LOGGED-IN, IN PARALLEL. AFTER AN EXAMINATION THE USER DOES NO LOGOUT OF HIS WINDOWS SESSION AND A DIFFERENT USER STARTS WITH HIS ADDITIONAL LOG-IN. ADDITIONALLY, SIDEXIS 4 IS NOT CLOSED FROM THE PREVIOUS USER. IF THE ABOVE LISTED CONDITIONS ARE GIVEN THE ISSUE DESCRIBED WILL OCCUR. THE IMMEDIATE ACTION IS TO ADD AN ADDITIONAL INSTRUCTION INTO THE SIDICOM WLS PACKAGES THAT ARE SHIPPED. FURTHER CAPA ACTION IS ONGOING. | image link to wrong patient | patient data issue | send automatically the images | multiple users are login at the same time | inappropriate treatment or diagnostic | [5 empty] | FDA2
Fetal patient report was automatically pulling prior measurement data for a prior fetus since the fetal study is based on the mother's MRN, resulting in the wrong fetal measurements getting referenced. | -fail to create related measurement -loss of measurement -inaccurate measurement item | -incorrect measurement -imprecision -break | [8 empty] | FDA_unkown
Following update from 3.3 to 3.5 of the DynaCAD software it was noted that the Ktrans map was not rendered correctly on the DynaCad Client. The defect causes the pharmacokinetic (PK) color maps to display incorrectly when viewed from remote DynaCAD client computers and could result in visually underestimating calculated Ktrans, Kep, and iAUGC values. The defect also impacts DynaCAD s on-the-fly calculation of Apparent Diffusion Coefficient (ADC) maps. If the ADC values are computed on-the-fly by DynaCAD, the ADC values and colors will also display incorrectly if viewed on a remote DynaCAD client. | -inappropriate color code -missing color code -fail to set value correctly | -device operates differently than expected -use of device issue -difficult to open or close | [8 empty] | FDA_unkown
For sites using the Issuer of Patient ID (IPID), the system will display the study list and images from different patients with different IPIDs as though they are for the same patient if they all have the first name, last name, and medical record number in common. | -inadequate patient data -incorrect patient information -fail to enter patient data | -Image Resolution Poor -incorrect display -image display error | [8 empty] | FDA_unkown
FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected. | -incorrect 3d model -incomplete 3d model -adjustment out of specification | -therapeutic or diagnostic output failure -use of device issue -improper or incorrect procedure or method | [8 empty] | FDA_unkown
Functionality in the report sections "Findings Information" and "Summary of Measured Findings" not functioning properly. The corrected values modified in the report are neither saved, nor printed, nor sent to the information system. The system will still use original values. | -incorrect measurement -inadequate ifu -incorrect readings | -imprecision -Inaccurate Flowrate -leak | [8 empty] | FDA_unkown
GE Healthcare is conducting a recall due to the potential fall of aged CRT monitors - Prestige 2, Prestilix, Prestige VH, Prestige SI, RFX/SFX, Legacy, Precision 500D, ADVANTX and Innova 2000 systems. | -ageing -loss of function -poor telemetry | -incorrect measurement -user used incorrect product for intended use -use of device issue | [8 empty] | FDA_unkown
If a reader selects Multi-study review prior to the original study completing its loading process and then immediately returns to image review of the original study, it is possible that not all current images will be present for review, which could result in incorrect treatment/diagnosis. | -inappropriate program selected -lack of review function -unable to review loop | -device cleaning issue -unstable -device handling issue | [8 empty] | FDA_unkown
If images are sent without an order in the system, they will be matched with the latest order on the current patient/modality matching potentially resulting in the matching of the report to the incorrect accession number (but still associated to the correct patient). | -fail to display message -fail to generate notification -fail to follow plan | -device markings issue -no display or display failure -application program issue | [8 empty] | FDA_unkown
Image data for a patients image may not be correct | -loss of image -data corrupted -inadequate image manipulation | -data issue -pacing inadequately -application security issue | [8 empty] | FDA_unkown
Images may potentially become corrupt while using the system. | -system abused -software corrupted -data corrupted | -Shelf Life Exceeded -pacing inadequately -malposition of device | [8 empty] | FDA_unkown
Inaccuracy of MR thermometry during MRI-guided laser ablation procedures using the Medtronic Visualase Thermal Therapy System may result in unaccounted spread of thermal energy to the surrounding tissue. | -poor thermal expansion -thermal event -incorrect measurement | -incorrect measurement -electromagnetic interference (emi) -optical issue | [8 empty] | FDA_unkown
Incorrect values for the volume calculation. Software update VB30B via Update Instructions SY018/16/P to resolve software errors. | -incorrect volume calculation -error message -fail to scan model volume | -use of device issue -failure to service -lack of maintenance documentation or guidelines | [8 empty] | FDA_unkown
Interventional Radiology (IR) images are stored as JPEG2k Lossless in Merge Enterprise Archive (EA) and are not displaying correctly in RadSuite. | -inadequate image types -loss of image -inadequate image manipulation | -therapeutic or diagnostic output failure -patient data issue -solder joint failure | [8 empty] | FDA_unkown
Issue for customers that use an EMR login or legacy web URL login or legacy web URL login for McKesson Radiology PACS that may result in missing images in a newly imported study, and/or study imports that remain in an "in-progress" status. | -missing data -missing maintenance -fail to transfer data | -use of device issue -human-device interface issue -device alarm system issue | [8 empty] | FDA_unkown
It was reported by a customer that RadSuite images are not appearing as readable images. This product malfunction has the potential to result in delay in care and a possible re-exposure of the patient to radiation (x-ray). | -loss of image -inadequate image manipulation -unreadable display | -fail-safe design failure -device stops intermittently -break | [8 empty] | FDA_unkown
It was reported from a complaint that study category labels of "Prior" and "Baseline" and hanging position appeared incorrectly swapped when the hanging protocol for three comparison studies is manually applied from the patient folder. | -labeling misaligned -incorrect labeling -inadequate labeling | -human-device interface issue -No Known Device Problem -failure to sense | [8 empty] | FDA_unkown
Lateromedial (LM) and Lateromedial Oblique (LMO) Orientation Markers may be displayed incorrectly on the 3D TOMO slider bar. In addition, the slice location of the image and/or the slice direction could be incorrect as well, and result in misdiagnosis | -image orientation incorrect -incorrect display format -incorrect 3d model | -application program issue -device sensing issue -no display or display failure | [8 empty] | FDA_unkown
Measurements done on the Merge PACS generated MPR s may have incorrect measurements. | -incorrect measurement -fail to measure -inaccurate measurement | -application program issue -imprecision -use of device issue | [8 empty] | FDA_unkown
Measurements in the software are changing after saving a plan with a measurement result, which may result in potential patient injury or delay in diagnosis or treatment. | -loss of measurement -fail to measure -incorrect measurement | -user used incorrect product for intended use -therapeutic or diagnostic output failure -defective item | [8 empty] | FDA_unkown
Merge PACS did not show unviewed images when the last view was skipped. There is a potential risk to health of a physician misdiagnosis because not all images are available for viewing. | -loss of image -fail to interpret images -unable to view display | -unstable -device alarm system issue -device unsafe to use in environment | [8 empty] | FDA_unkown
Multiple studies were missing at a customer location that never made it to the archive before watermarking | -missing maintenance -missing filter -part missing | -component missing -device maintenance issue -device markings issue | [8 empty] | FDA_unkown
Novarad Corporation announces a voluntary field action for the NovaPACS versions 8.4.7, 8.5.3, and 8.5.6 due to a "pop-up screen" that was prematurely released. | -release delayed -fail to pop-up -inappropriate release | -defective component -device operates differently than expected -Application Program Issue | [8 empty] | FDA_unkown
Observed an internal repetitive software glitch in ImmersiveView. | -software corrupted -software corruption -software task fault displayed | -material disintegration -device handling problem -Device displays error message | [8 empty] | FDA_unkown
Ortho Kinematics Inc. sent a Notice of Correction to Released Testing Results, Radiological Read Report for the Vertebral Motion Analyzer (VMA) test because it contained an error. The error occurred due to a software bug that has been corrected. | -incorrect readings -fail to disclose error message -error message | -Communication or Transmission Issue -Device displays error message -Incorrect Measurement | [8 empty] | FDA_unkown
Ortho Kinematics notified customers that errors were contained in Radiological Read Report and VMA Report for the VMA test. The error resulted from an administrative error of VMA user account settings and has been fixed. | -incorrect readings -error message -inappropriate symbol | -Improper or Incorrect Procedure or Method -no display or display failure -Defective Component | [8 empty] | FDA_unkown
Ortho Kinematics, Inc. notified customers of erroneous results of VMA tests due to a software bug that has been corrected. | -software defect -fail to monitor feedback -error message | -output issue -device maintenance issue -operating system version or upgrade problem | [8 empty] | FDA_unkown
Panasonic has issued a Mandatory Product Safety Notice, indicating a possibility of battery ignition (fire) in CF-C2 "Toughbook" laptops. Since Konica Minolta uses this laptop in some of its product configurations, it notified its customers to utilize a BIOS Utility for the laptops to minimize the risk of an incident while Panasonic develops a final solution. | -fail to indicate battery charging status -ignition -inadequate safety information | -market issue -improper or incorrect procedure or method -compliance issue | [8 empty] | FDA_unkown
Patient information in the header is only found on the first page of the report and not on the subsequent pages. | -missing information -loss of patient data -incorrect message display | -failure to read input signal -patient data issue -no display or display failure | [8 empty] | FDA_unkown
Potential data loss occurs as a result of product archiving not working properly. | -loss of data -fail to record data -data corrupted | -patient data issue -environmental control or utility issue -market issue | [8 empty] | FDA_unkown
Potential data loss occurs as a result of the software archiving not working properly. | -loss of data -data corrupted -incorrect data | -patient data issue -Decreased Sensitivity -not audible alarm | [8 empty] | FDA_unkown
Potential exists for an incorrect patient image being displayed which could result in the delay in diagnosis or treatment. | -image orientation incorrect -incorrect display -incorrect message display | -application program issue -no display or display failure -failure to align | [8 empty] | FDA_unkown
Potential for images to be flipped while streaming, which could incorrectly display image orientation markers. | -image orientation incorrect -fail to interpret images -inappropriate image format | -poor quality image -application program issue -chemical issue | [8 empty] | FDA_unkown
Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). | -unable to scroll image -image orientation incorrect -incorrect display format | -use of device issue -dose calculation error due to software problem -not audible alarm | [8 empty] | FDA_unkown
Potential for RadSuite AV viewer to skip image slices when the user presses page up or page down when the viewer displays multiple viewports and the user has mouse focused on any viewport other than the first viewport. | -skip failed -fail to toggle image visibility -fail to read images | -use of device issue -device alarm system issue -not audible alarm | [8 empty] | FDA_unkown
Potential that one or more image series may be missing from an exam without a user warning displayed in the Viewer. | -fail to verify image -missing symbol -loss of warning | -no display or display failure -device alarm system issue -device markings issue | [8 empty] | FDA_unkown
Remote comparison studies displayed may be associated with the wrong patient, without a caution icon present. | -image orientation incorrect -missing symbol -fail to generate indication | -device markings issue -inadequate instructions for healthcare professional -no display or display failure | [8 empty] | FDA_unkown
Reporting feature times out after inactivity for more than an hour sending the user back to the study list, which causes all reporting data being entered to be lost. | -incorrect time configuration -lack of time selection -lack of time indication | -use of device issue -use of incorrect control settings -device difficult to setup or prepare | [8 empty] | FDA_unkown
Reports were missing demographic header information when they were faxed to referring physicians. | -missing information -incorrect patient information -missing data | -component missing -device markings issue -inadequate instructions for healthcare professional | [8 empty] | FDA_unkown
Siemens has identified changes in default carry forward workflow using VA10B_HF03 version of Syngo Dynamics. If the template design has measurements inside a carry forward-configured report section, this will result in carrying forward the entire section of the report including data from previous studies. This may result in the physician making analysis based on the old measurements. | -fail to identify change -lack of change control -fail to measure | -optical issue -material disintegration -compliance issue | [8 empty] | FDA_unkown
Siemens is releasing a Customer Advisory Notice to explain meaning of a suffix o.p. that shows on some DICOM images on area, distance or perimeter values. This suffix stands for on projection and indicates that a shown measurement value may be different from the real value. | -inadequate warning -incorrect display -lack of orientation | -application program issue -break -use of incorrect control settings | [8 empty] | FDA_unkown
Siemens is releasing a letter to inform about potential data loss relevant to diagnosis. | -loss of patient data -loss of data -loss of message | -communication or transmission issue -application security issue -failure to sense | [8 empty] | FDA_unkown
Siemens is releasing a non-medical software application LTA Incomplete Archive Check Tool V1.0. It is intended to eliminate an issue that occurs during syngo.plaza de-archiving from Dicom LTA. It has been defined that the number of de-archived images is less than the count of the archived images for the series. The issues were discovered through Siemens Global Complaint Handling System. | -incomplete software upgrade -fail to schedule check -software version incorrect | -use of device issue -device alarm system issue -deployment issue | [8 empty] | FDA_unkown
Software changes now available to address several issues | -fail to address problem -fail to implement change -inappropriate change | -compliance issue -Mechanical Issue -fail-safe design failure | [8 empty] | FDA_unkown
Software defect which may potentially result in one or more images missing in a study | -component missing -missing symbol -fail to identify defect | -defective item -device markings issue -size incorrect for patient | [8 empty] | FDA_unkown
Software displayed incorrect prior reports in the viewport area, only when more than one prior study (2 or more) was viewed. | -incorrect display format -incorrect message display -incorrect display | -device expiration issue -inappropriate prompts -false alarm | [8 empty] | FDA_unkown
Software error in previous software versions in which two references for the same image may exist in the database. | -software version incorrect -fail to interpret images -error message | -no pacing -failure to pace or properly pace -computer software issue | [8 empty] | FDA_unkown
Software error. Incorrect values for the volume calculation from a freehand VOI at the customer site. In volume calculations of prostate as well as in volume calculations of liver were too high. | -incorrect volume calculation -fail to calculate volume -fail to set value correctly | -use of device issue -output below specifications -parameter calculation error due to software problem | [8 empty] | FDA_unkown
Software issues. Siemens is voluntarily initiating a recall after they became aware of the following system behavior: 1) At times, the view and laterality marker is overlaid by patient demographic information, thus becoming difficult to read. 2) The facility name and address are not shown per default (in the current software versions this information is displayed only when the reader chooses an appropriate image text display mode). | -fail to specify patient population -device malfunction -fail to reattach | -no display or display failure -data back-up problem -human-device interface issue | [8 empty] | FDA_unkown
Software update for improvements and to resolve several issues | -fail to update software -fail to update indicator -fail to update | -use of device issue -device contamination with biological material -device cleaning issue | [8 empty] | FDA_unkown
Software update to correct several issues that include (1)Potential data loss, (2) Study mix-up, (3) Incorrect measurements on multi-frame images, (4) Dearchiving issue, and (5) Unauthorized access of data due to inadequate permissions for shared folders. | -incorrect measurement -unauthorized access -incorrect data | -unauthorized access to computer system -human-device interface issue -User used incorrect product for intended use | [8 empty] | FDA_unkown
Software upgrade to correct format of study dates and issues with Legacy Presentation States (annotations) in order to prevent potential patient misdiagnosis. | -incorrect data format -software version incorrect -incorrect display format | -compliance issue -therapeutic or diagnostic output failure -use of device issue | [8 empty] | FDA_unkown
Studies coming over via telmed were missing patient's DOB, procedure, and referring physician. | -missing review of pre-operative patient images -patient trigger failed -incorrect patient information | -component missing -device contamination with biological material -compliance issue | [8 empty] | FDA_unkown
Studies that are viewed in iConnect Access that originate on Merge PACS / Ortho PACS may not display the correct current patient demographics if there has been a PDE (Patient / Study Demographic Edit) performed after the images were ingested in PACS. | -incorrect patient information -fail to enter patient data -fail to display data | -inadequate user interface -human-device interface issue -use of device issue | [8 empty] | FDA_unkown
Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | -fail to update configuration -fail to update -fail to display data | -disconnection -improper flow or infusion -use of device issue | [8 empty] | FDA_unkown
Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | -fail to update configuration -fail to update -fail to display data | -disconnection -improper flow or infusion -use of device issue | [8 empty] | FDA_unkown
Synapse cannot display image files, DICOM SR files, and/or Annotation files. The "Image Not Loaded" message displays instead | -incorrect message display -inadequate message display -fail to display message | -incorrect display -difficult to interrogate -poor quality image | [8 empty] | FDA_unkown
The bug causes BOLD activation maps to be visualized as overlays without taking the coregistration into account. Any output created from these activation maps as overlays will not be adjusted according to the coregistration. | -fail to adjust image brightness -fail to adjust size -fail to adjust structure opacity | -defective item -device difficult to setup or prepare -device cleaning issue | [8 empty] | FDA_unkown
The error occurs in the following situations: A group of fibers has been selected using the VOI functionality and put into a fiber group. This is done with images series A visualized in the 3D viewer. Then a new images series B is visualized in the 3D viewer. If image series A and B have the same geometrical resolution (pixel sizes and slice distances), the coregistration matrix taking A to B will not be applied to the fiber group, and therefore the fiber group will in general not be positioned correctly on B in the 3D viewer. Furthermore, if the fiber group is exported as a new image series the fibers could be misplaced, and similarly if the fibers are visualized in the MPR on B, the positioning will not be correct. The misalignment will be equal to the rotations/translations necessary to align A with B. | -incorrect geometry -loss of image -image orientation incorrect | -improper or incorrect procedure or method -failure to sense -fitting problem | [8 empty] | FDA_unkown
The firm discovered a software issue associated with the Image-Com module/clinical application package (CAP) interface for 3D application in the ARENA software, where measurements from one study might be stored to another study. | -software installation problem -loss of measurement -inappropriate measurement algorithm | -therapeutic or diagnostic output failure -material disintegration -human-device interface issue | [8 empty] | FDA_unkown
The firm received a report of an atypical dataset being generated. When using the cross-localization feature and images from a modality that generates asymmetrical images, the cross-localization reference line may be inaccurately placed on any of the corresponding images that are open. | -unclear image -inappropriate image format -illegible data | -Human-Device Interface Issue -no display or display failure -filling problem | [8 empty] | FDA_unkown
The Neuroreader has been distributed with a reference normative database that is different from the version reviewed and cleared by FDA. | -data corrupted -unauthorized access -inappropriate certificate scope | -application program version or upgrade problem -operating system version or upgrade problem -problem with software installation | [8 empty] | FDA_unkown
The patient name in the Halo title bar and the thumbnails do not match the name on displayed images. | -incorrect patient information -incorrect message display -incorrect display format | -device markings issue -no display or display failure -incorrect display | [8 empty] | FDA_unkown
The software did not show unviewed images when the last view was skipped. The physician may potentially misdiagnose when not all images are available for viewing. | -loss of image -unable to view display -fail to interpret images | -unstable -no display or display failure -alarm not visible | [8 empty] | FDA_unkown
The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user. | -fail to update indicator -fail to update configuration -software version incorrect | -alarm not visible -patient data issue -not audible alarm | [8 empty] | FDA_unkown
The software fails to associate to the correct MG image if there are two images for the same view. | -fail to interpret images -loss of image -fail to identify use error | -device alarm system issue -connection issue -poor quality image | [8 empty] | FDA_unkown
The software is not adding the correct tomo image. | -image orientation incorrect -fail to support correct use -inappropriate image rendered | -use of device issue -pacing inadequately -reset issue | [8 empty] | FDA_unkown
The software is not identifying the patient as having atypical hyperplasia, resulting in an incorrect Gail Risk calculation. | -incorrect patient information -incorrect ratio calculation -inadequate risk analysis | -improper or incorrect procedure or method -use of device issue -user used incorrect product for intended use | [8 empty] | FDA_unkown
The software is unable to accurately determine the calcium score of scans with a slice thickness not equal to 3 mm. | -wrong scan condition -inaccurate measurement item -incorrect sample size | -unstable -failure to recalibrate -poor quality image | [8 empty] | FDA_unkown
The software was not presenting the PURE filter user interface message on DICOM transfer studies. If a site starts using PURE filter, the user interface message will not display. This may result in incorrect patient follow-up. | -incorrect message display -inadequate message display -fail to display message | -failure to deploy -improper or incorrect procedure or method -reset issue | [8 empty] | FDA_unkown
The STAT priority icon was not displayed in the Study List when the study was opened. | -fail to mark icon -incorrect screen display -unreadable display | -No Display or Display Failure -Air Leak -incorrect display | [8 empty] | FDA_unkown
The SUV values that are being calculated in the PET/CT fusion tool are incorrect. | -incorrect value -incorrect measurement -incorrect data | -therapeutic or diagnostic output failure -failure to read input signal -device alarm system issue | [8 empty] | FDA_unkown
The values provided from the Pixel Value tool do not appear to be correct, which may result in potential patient injury or delay in diagnosis or treatment. | -incorrect tool identification -incorrect value -incorrect display value | -use of device issue -user used incorrect product for intended use -incorrect display | [8 empty] | FDA_unkown
There is a potential for the wrong patient information may be displayed in the viewer or PowerJacket. Additionally, the incorrect max standard uptake values (SUV) for PET ad CT studies may be returned for Siemens Modality. | -incorrect display value -incorrect display format -incorrect message display | -device markings issue -pacing inadequately -incorrect measurement | [8 empty] | FDA_unkown
There is a potential that one or more image series (i.e., all images within an image set) may be missing from an exam without a warning displayed in the viewer. | -fail to verify image -loss of warning -missing symbol | -no display or display failure -device alarm system issue -human-device interface issue | [8 empty] | FDA_unkown
There is a potential that one or more images or image series may be missing from exams without a warning displayed in the viewer. | -component missing -missing symbol -loss of warning | -no display or display failure -device alarm system issue -human-device interface issue | [8 empty] | FDA_unkown
There is the possibility of viewing studies directly from the Enterprise Archive or VNA with incorrect patient images because the updated series or study is not present in the archive. | -incorrect patient information -inappropriate patient type -inadequate patient data | -parameter calculation error due to software problem -therapeutic or diagnostic output failure -electronic property issue | [8 empty] | FDA_unkown
This email is to provide notification that, due to a software bug that has been corrected, the Radiological Read Report for the VMA test(s) contained an error. | -software defect -error message -incorrect readings | -Defective Component -break -application program issue | [8 empty] | FDA_unkown
To inform users about the possible incorrect values for Distance Measurements when using certain modalities in combination with syngo Imaging. | -incorrect distance measurement -incorrect measurement -incorrect data | -imprecision -Communication or Transmission Issue -sticking | [8 empty] | FDA_unkown
Users can merge a device import file with an image study that already has a confirmed report, which may result in including information not present at the time of physician interpretation. | -lack of confirmation -fail to save surgical plan -unable to switch | -device markings issue -failure of device to self-test -component missing | [8 empty] | FDA_unkown
When a prior study is being replaced in the workflow step, in certain scenarios (based on the Display Protocol configuration) the prior study will only be replaced in the active Workflow Step and not show in all other workflow steps. The other workflow steps will continue to show the initially loaded study. | -fail to display previous data -fail to analyze new use error -lack of intended users | -user used incorrect product for intended use -incorrect display -no display or display failure | [8 empty] | FDA_unkown
When an image is flipped horizontally and vertically, the orientation markers appear correctly on the image to which they were applied. However, once new image rendering happens on this image or series (with scrolling; playing cin¨; maximizing or changing viewports/layouts; or synchronizing viewports and resizing the browser), the new image orientation is correctly maintained but the display of orientation markers can become incorrect. | -image orientation incorrect -inappropriate software release -incorrect display format | -device sensing issue -therapeutic or diagnostic output failure -failure to read input signal | [8 empty] | FDA_unkown
When cut lines are shown on an image, the lines may go from being vertical and then shift to horizontal causing confusion to the physician. This may cause a misdiagnosis. | -confused indication -loss of image -confused interface | -improper or incorrect procedure or method -use of device issue -device markings issue | [8 empty] | FDA_unkown
When measuring a lesion on an unmagnified mammography image then performing the same measurement on an image magnified by the Hologic imager (different image), the measurements are not the same. | -incorrect measurement -inaccurate measurement -fail to interpret images | -application program issue -incorrect measurement -device alarm system issue | [8 empty] | FDA_unkown
When RadSuite is used with IPID (Issuer of Patient ID) as a part of the "Patient Identifier," it is possible in some circumstances that the demographics of one patient will be applied to a study or studies for another patient. | -incorrect patient information -inappropriate patient type -incorrect identification | -instruction for use issue -defective item -no display or display failure | [8 empty] | FDA_unkown
When switching back & forth between multiple UV instances in the Windows taskbar, the patient images displayed on the mammo high resolution monitors may not show images of the patient selected on UV, which could result in an incorrect diagnosis. | -switch back and forth -image orientation incorrect -inadequate message display | -no display or display failure -application program issue -Image Resolution Poor | [8 empty] | FDA_unkown
When taking measurements from images on the Cardio workstation or from the US cart, numbers are not crossing to the report in the correct unit of measure. | -inadequate measure unit -incorrect measurement -fail to place measurements | -parameter calculation error due to software problem -incorrect measurement -use of device issue | [8 empty] | FDA_unkown
When the archiving configuration is changed, data received/created after the upgrade may be flagged as "Not to be archived". This is caused by the automatic function for cleaning up temporary data being disabled by the software upgrade. Due to the disabled cleanup function, disc capacity for free space decreases faster than usual. Unless the archiving configuration and the cleanup automatic function are reverted back to the original state (prior to the software upgrade), data that was incorrectly flagged "Not to be archived" must be manually prevented from being deleted by either clinical administrator or service engineer. | -fail to update configuration -fail to extract data -fail to delete data | -failure to deliver -application security issue -patient data issue | [8 empty] | FDA_unkown
While merging exams in a test instance, merges carry forward in the production system, when the middle tier is configured to the wrong IP address of the Centricity Exam Manager. When a current and historical exam are opened at the same time, the system asynchronously tries to access the operation specifying how each image should be grouped, causing random image-series grouping errors. | -fail to setup configuration -incorrect configuration -incorrect test procedure | -use of device issue -device operates differently than expected -failure of device to self-test | [8 empty] | FDA_unkown
After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks | unable to read image | application program issue | Unable to Display Patient Image during pre-operative | User unaware viewing options are available; User cannot find viewing options control; User does not understand parameters for viewing options; Complex or confusing control system. | Pre-operative Delayed | No Harm - (S0) |  | Viewing panel shall be appeared at the top of the toolbar so that it is noticable and uses a quadrant icon; 2D Source Images shall be appeared in the Viewing panel. with a toggle switch. | software shall provide viewing options controls; software shall provide image configuration controls; software shall provide loaded image format controls | software shall indicate the toggle status; software shall indicate the load imge status; software shall alert if the "the image is not dicom format" | u_FMEA
User monitors loading of the 2D images on the "Viewing" panel | inadequate display message | human-device interface issue | Unable to Display Patient Image during intra-operative | User unaware viewing options are available; User cannot find viewing options control; User does not understand parameters for viewing options; Complex or confusing control system. | Surgery Prolonged | Reversible - (S2) |  | Viewing panel shall be appeared at the top of the toolbar so that it is noticable and uses a quadrant icon. 2D Source Images shall be appeared in the Viewing panel. with a toggle switch. | software shall provide monitor information controls | software shall indicate image loading state; software shall alert if the "the image is not dicom format"; IFU shall provide guides on how to load images and error codes | u_FMEA
User selects which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas | wrong image selected | use of device issue | Unable to Display Patient Image during pre-operative | User inadvertently selects incorrect images | Pre-operative Delayed | No Harm - (S0) |  | Icons and a informational tip shall be presented to distinguish between Axial, Sagittal and Coronal which prevents selection error; Anatomical Atlas shall be clearly labeled and provides options for infill and contour. |  | software shall indicate the selected 2D source image as Axial; software shall indicate the selected 2D source image as Sagittal; software shall indicate the selected 2D source image as Coronal; software shall indicate the selected 2D source image as Anatomical Atlas; user manual/IFU shall define about "axial, sagittal, corronal, anatomical atlas" iamges; | u_FMEA
User selects which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas | selecting confused | inadequate user interface | Unable to Display Patient Image during intra-operative | User unaware of which images are to be used; User cannot determine which images are the correct ones; Complex or confusing control system. | Surgery Prolonged | Reversible - (S2) |  | Icons and a informational tip shall be presented to distinguish between Axial, Sagittal and Coronal which prevents selection error; Anatomical Atlas shall be clearly labeled and provides options for infill and contour. |  | software shall indicate the selected 2D source image as Axial; software shall indicate the selected 2D source image as Sagittal; software shall indicate the selected 2D source image as Coronal; software shall indicate the selected 2D source image as Anatomical Atlas; user manual/IFU shall define about "axial, sagittal, corronal, anatomical atlas" iamges; | u_FMEA
User selects whether they would like the segmented anatomic regions to be displayed by infill or contour | incorrect configuration | use of incorrect control settings | Unable to Manipulate Patient Image during pre-operative | User unaware Anatomical Atlas options are available; User cannot find Anatonical Atlas options control; User does not understand parameters for Anatomical Atlas; User fails to properly configure Anatomical Atlas | Pre-operative Delayed | No Harm - (S0) |  | Anatomical Atlas shall be clearly labeled and provides options for infill and contour. |  | software shall indicate infill status during operation; software shall indicate contour status during operation; IFU/user manual shall define the infill and contour of anatomic | u_FMEA
User selects whether they would like the segmented anatomic regions to be displayed by infill or contour | incorrect selecting | use of device issue | Unable to Manipulate Patient Image during intra-operative | User unaware Anatomical Atlas options are available; User cannot find Anatonical Atlas options control; User does not understand parameters for Anatomical Atlas | Surgery Prolonged | Reversible - (S2) |  | Anatomical Atlas shall be clearly labeled and provided options for infill and contour. |  | software shall indicate infill status during operation; software shall indicate contour status during operation; IFU/user manual shall define the infill and contour of anatomic | u_FMEA
User selects stack they would like to manipulate | incorrect selecting | use of device issue | Unable to Manipulate Patient Image during pre-operative | User does not see the stack options controls; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design select icon rather than invisible control | software shall provide select lists to prevent out of scope | software shall indicate the select status | u_FMEA
User adjusts image using pan, rotate, zoom or scroll. | fail to adjust image | application program issue | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the scroll/pan/zoom/rotate options controls; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design pan, rotate, zoom, or scroll icon rather than invisible control | software shall provide zoom/pan/rrotate/scroll specifications to prevent out of scope | software shall indicate the pan status; software shall indicate the rotate status; software shall indicate the zoom status; software shall indicate the scroll status; software shall display the rotation angle when user rotates an image; software shall display the zoom ratio when user zooms an image' | u_FMEA
user undoes a prior action. | fail to undo prior action | use of device issue | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the undo function; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design undo icon rather than invisible control | software shall proivde undo control (e.g. how many backward steps to undo) | software shall indicate the undo status | u_FMEA
user resets the view to default position | fail to reset | failure to reset | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the reset function; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design reset icon rather than invisible control | software shall proivde reset view control (e.g. position, angle) | software shall indicate the reset status | u_FMEA
user adjusts contrast/brightness of image | fail to adjust contrast | use of device issue | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the adjust function; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design adjust icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
user selects a contrast/brightness setting preset | fail to start with preset value | use of incorrect control settings | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the retreat preset function; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design preset icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
user makes further adjustments to contrast/brightness of image | adjustment difficult | device difficult to setup or prepare | Unable to Manipulate Patient Image during pre-operative | User does not understand how to use the adjustment function; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | software shall design adjustment icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
User selects stack they would like to manipulate | incorrect selecting | use of device issue | Unable to Manipulate Patient Image during intra-operative | User does not see the stack options controls; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design select icon rather than invisible control | software shall provide select lists to prevent out of scope | software shall indicate the select status | u_FMEA
User adjusts image using pan, rotate, zoom or scroll. | fail to adjust image | application program issue | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the scroll/pan/zoom/rotate options controls; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design pan, rotate, zoom, or scroll icon rather than invisible control | software shall provide zoom/pan/rrotate/scroll specifications to prevent out of scope | software shall indicate the pan status; software shall indicate the rotate status; software shall indicate the zoom status; software shall indicate the scroll status; software shall display the rotation angle when user rotates an image; software shall display the zoom ratio when user zooms an image' | u_FMEA
user undoes a prior action. | fail to undo prior action | use of device issue | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the undo function; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design undo icon rather than invisible control | software shall proivde undo control (e.g. how many backward steps to undo) | software shall indicate the undo status | u_FMEA
user resets the view to default position | fail to reset | failure to reset | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the reset function; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design reset icon rather than invisible control | software shall proivde reset view control (e.g. position, angle) | software shall indicate the reset state | u_FMEA
user adjusts contrast/brightness of image | fail to adjust contrast | use of device issue | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the adjust function; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design adjust icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall indicate the adjustment state; software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
user selects a contrast/brightness setting preset | fail to start with preset value | use of incorrect control settings | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the retreat preset function; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design preset icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall indicate the preset state; software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
user makes further adjustments to contrast/brightness of image | adjustment difficult | device difficult to setup or prepare | Unable to Manipulate Patient Image during intra-operative | User does not understand how to use the adjustment function; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | software shall design adjustment icon rather than invisible control | software shall proivde image contrast/brightness control (e.g. contrast ratio, brightness specification) | software shall indicate the adjustment state; software shall display the contrast/brightnest value when user adjusts the contrast/brightness of imatge; | u_FMEA
User selects the annotation menu from the sidebar | fail to place annotation on image | human-device interface issue | Delayed Pre-Operative Plan during pre-operative | User unaware annotation options are available; User cannot find annotation options control; User does not understand parameters for annotation options | Pre-operative Delayed | No Harm - (S0) |  | Annotations shall be clearly labeled with an icon in the toolbar. | software shall provide annotation control (e.g. format, text direction, delete permission) | software shall alert users if incorrect annotation direction; software shall alert users if incorrect annotation format | u_FMEA
User selects the annotation menu from the sidebar | fail to place annotation on image | human-device interface issue | Delayed Pre-Operative Plan during intra-operative | User unaware annotation options are available User cannot find annotation options control User does not understand parameters for annotation options | Surgery Prolonged | Reversible - (S2) |  | Annotations shall be clearly labeled with an icon in the toolbar. | software shall provide annotation control (e.g. format, text direction, delete permission) | software shall alert users if incorrect annotation direction; software shall alert users if incorrect annotation format | u_FMEA
User selects "distance" tool to measure the straight-line distance between two points on the image | an erroneous measurement value | incorrect measurement | Inappropriate Pre-Operative Plan during pre-operative | User does not know measurement feature is available; User unsure how to access measurement feature; User does not read measured distance correctly; User unintentionally adds extra measurement segment; User does not know renaming feature is available; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | Annotations shall be clearly labeled with an icon in the toolbar; Distance is clearly labeled with an icon in the toolbar under Annotations. | software shall provide measurement accuracy controls; software shall provide measurement range controls; | software shall indicate the measurement state | u_FMEA
User clicks on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen | fail to pick end point | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User places start/end point in incorrect location; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide pick up points controls | software shall indicate the measurement state; software shall alert if user clicks on wrong location | u_FMEA
User may drag the endpoints to resize the line, or drag the line itself to reposition the line on the image | fail to drag operation | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User unintentionally removes measurement segment; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide drag operation controls | software shall indicate the drag state; software shall alert if user drags improperly | u_FMEA
User may delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu | incorrect deletion | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User unintentionally deletes annotations; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide delete permission controls | software shall indicate the delete state; software shall alert if users operates delete operation | u_FMEA
User may select any of the annotation distances or arrows from the Annotation menu list and rename them | invalid name | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User gives measurement incorrect name; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide name speciffication controls | software shall indicate the rename state; software shall alert if users enter invalid name | u_FMEA
User may reset all annotations using the "Reset" button. | inappropriate reset | reset issue | Inappropriate Pre-Operative Plan during pre-operative | User unintentionally clicks on reset button; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide reset button controls | software shall indicate the reset state; software shall alert if user clicks on the reset during improper times | u_FMEA
User selects "distance" tool to measure the straight-line distance between two points on the image | an erroneous measurement value | incorrect measurement | Inappropriate Pre-Operative Plan during intra-operative | User does not know measurement feature is available; User unsure how to access measurement feature; User does not read measured distance correctly; User unintentionally adds extra measurement segment; User does not know renaming feature is available; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | Annotations shall be clearly labeled with an icon in the toolbar; Distance is clearly labeled with an icon in the toolbar under Annotations. | software shall provide measurement accuracy controls; software shall provide measurement range controls; | software shall indicate the measurement state | u_FMEA
User clicks on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen | fail to pick end point | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User places start/end point in incorrect location; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide pick up points controls | software shall indicate the measurement state; software shall alert if user clicks on wrong location | u_FMEA
User may drag the endpoints to resize the line, or drag the line itself to reposition the line on the image | fail to drag operation | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User unintentionally removes measurement segment; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide drag operation controls | software shall indicate the drag state; software shall alert if user drags improperly | u_FMEA
User may delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu | incorrect deletion | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User unintentionally deletes annotations; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide delete permission controls | software shall indicate the delete state; software shall alert if users operates delete operation | u_FMEA
User may select any of the annotation distances or arrows from the Annotation menu list and rename them | invalid name | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User gives measurement incorrect name; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide name speciffication controls | software shall indicate the rename state; software shall alert if users enter invalid name | u_FMEA
User may reset all annotations using the "Reset" button. | inappropriate reset | reset issue | Inappropriate Pre-Operative Plan during intra-operative | User unintentionally clicks on reset button; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide reset button controls | software shall indicate the reset state; software shall alert if user clicks on the reset during improper times | u_FMEA
User may add and delete arrows in a similar fashion after clicking the "Arrow" menu option. | inappropriate arrow fashion | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User does not know arrow feature is available; User unsure how to access arrow feature; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  | Annotations shall be clearly labeled with an icon in the toolbar; Arrow shall be clearly labeled with an icon in the toolbar under Annotations. | software shall provide arrow delete permission controls | software shall indicate the states of add arrow, delete arrow, select annotation distance, reset annotation | u_FMEA
Arrows added to the screen will also appear in the Annotation sidebar menu list | fail to add arrow | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User draws arrow incorrectly; User places start/end point in incorrect location; User unintentionally removes arrow; User unintentinally adds extra arrow; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide arrow fashion controls | software shall indicate arrow state; | u_FMEA
User may select any of the annotation distances or arrows from the Annotation menu list and rename them | invalid name | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User gives arrow incorrect name; User does not know renaming feature is available; Complex or confusing control system | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide name controls | Software shall alert if user input invalid name such as special characters as: #, %, / | u_FMEA
User may add and delete arrows in a similar fashion after clicking the "Arrow" menu option. | inappropriate arrow fashion | use of device issue | Inappropriate Intra-Operative Plan during intra-operative | User does not know arrow feature is available; User unsure how to access arrow feature; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  | Annotations shall be clearly labeled with an icon in the toolbar; Arrow shall be clearly labeled with an icon in the toolbar under Annotations. | software shall provide arrow delete permission controls | software shall indicate the states of add arrow, delete arrow, select annotation distance, reset annotation | u_FMEA
Arrows added to the screen will also appear in the Annotation sidebar menu list | fail to add arrow | use of device issue | Inappropriate Intra-Operative Plan during intra-operative | User draws arrow incorrectly; User places start/end point in incorrect location; User unintentionally removes arrow; User unintentinally adds extra arrow; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide arrow fashion controls | software shall indicate arrow state; | u_FMEA
User may select any of the annotation distances or arrows from the Annotation menu list and rename them | invalid name | use of device issue | Inappropriate Intra-Operative Plan during intra-operative | User gives arrow incorrect name; User does not know renaming feature is available; Complex or confusing control system | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide name controls | Software shall alert if user input invalid name such as special characters as: #, %, / | u_FMEA
User may add a note by selecting the "note" annotation option and clicking anywhere on the image | fail to add notes at anywhere | data issue | Inappropriate Pre-Operative Plan during pre-operative | User enters incorrect information in the note; User adds note but is unable to add text to the note;User unsure how to access note feature | Pre-operative Delayed | No Harm - (S0) |  | Annotations shall be clearly labeled with an icon in the toolbar; Note is clearly labeled with an icon in the toolbar under Annotations. | software shall provide image click location controls | software shall provide notes to show the operation status during process (e.g. aadd note, delete note, edit, reset annotation) to prevent user confused | u_FMEA
A textbox will appear; user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu | incorrect text direction | data issue | Inappropriate Pre-Operative Plan during pre-operative | User places note in incorrect location; User places note in incorrect direction | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide text direction controls | software shall indicate textbox state | u_FMEA
User may return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox | inappropriate annotation format | data issue | Inappropriate Pre-Operative Plan during pre-operative | User unaware of note feature | Pre-operative Delayed | No Harm - (S0) |  |  | software shall annotation format controls | software shall indicate annotation state | u_FMEA
User may remove any added note using the trash icon in the Annotation menu or on the note itself | unintend to delete note | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User unintentionally moves note | Pre-operative Delayed | No Harm - (S0) |  | Annotations shall be clearly labeled with an icon in the toolbar; Note is clearly labeled with an icon in the toolbar under Annotations. | software shall provide delete permission controls | software shall indicate delete state; software shall alert if user delete note | u_FMEA
User may add a note by selecting the "note" annotation option and clicking anywhere on the image | fail to add notes at anywhere | data issue | Inappropriate Intra-Operative Plan during intra-operative | User enters incorrect information in the note; User adds note but is unable to add text to the note;User unsure how to access note feature | Surgery Prolonged | Reversible - (S2) |  | Annotations shall be clearly labeled with an icon in the toolbar; Note is clearly labeled with an icon in the toolbar under Annotations. | software shall provide image click location controls | software shall provide notes to show the operation status during process (e.g. aadd note, delete note, edit, reset annotation) to prevent user confused | u_FMEA
A textbox will appear; user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu | incorrect text direction | data issue | Inappropriate Intra-Operative Plan during intra-operative | User places note in incorrect location; User places note in incorrect direction | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide text direction controls | software shall indicate textbox state | u_FMEA
User may return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox | inappropriate annotation format | data issue | Inappropriate Intra-Operative Plan during intra-operative | User unaware of note feature | Surgery Prolonged | Reversible - (S2) |  |  | software shall annotation format controls | software shall indicate annotation state | u_FMEA
User may remove any added note using the trash icon in the Annotation menu or on the note itself | unintend to delete note | use of device issue | Inappropriate Intra-Operative Plan during intra-operative | User unintentionally moves note | Surgery Prolonged | Reversible - (S2) |  | Annotations shall be clearly labeled with an icon in the toolbar; Note is clearly labeled with an icon in the toolbar under Annotations. | software shall provide delete permission controls | software shall indicate delete state; software shall alert if user delete note | u_FMEA
5.1 - User will toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu | confused toggle | human-device interface issue | Inaccessible Patient 3D Model during intra-operative | User unaware viewing options are available; User cannot find viewing options control; User does not understand parameters for viewing options | Pre-operative Delayed | No Harm - (S0) |  | Viewing panel shall be appeared at the top of the toolbar so that it is noticable and uses a quadrant icon; 3D Model shall be appeared in the Viewing panel with a toggle switch. |  | software shall indicaate toggle state (on, off) | u_FMEA
User selects "3D Model" menu to see which anatomic structures are available | configuration confused | inadequate user interface | Unable to Manipulate 3D Model during pre-operative | User unaware structure options are available; User cannot find structure options control; User does not understand parameters for structure | Pre-operative Delayed | No Harm - (S0) |  | 3D Model shall be appeared in the Viewing panel. with a toggle switch; structure shall be appeared under 3D Model with a checkbox; Checkbox shall be appeared to turn structures on and off. | software shall provide select controls | software shall indicate the selected anatomic structures | u_FMEA
User selects/deselects the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available | incorrect deselecting | use of device issue | Unable to Manipulate 3D Model during pre-operative | User unaware structure options are available; User cannot find structure options control; User does not understand parameters for structure | Pre-operative Delayed | No Harm - (S0) |  | 3D Model shall be appeared in the Viewing panel. with a toggle switch; structure shall be appeared under 3D Model with a checkbox; Checkbox shall be appeared to turn structures on and off. | software shall provide select controls; software shall provide deselect controls | software shall indicate the deselected anatomic structures | u_FMEA
User selects "3D Model" menu to see which anatomic structures are available | configuration confused | inadequate user interface | Unable to Manipulate 3D Model during intra-operative | User unaware structure options are available; User cannot find structure options control; User does not understand parameters for structure | Surgery Prolonged | Reversible - (S2) |  | 3D Model shall be appeared in the Viewing panel. with a toggle switch; structure shall be appeared under 3D Model with a checkbox; Checkbox shall be appeared to turn structures on and off. | software shall provide select controls | software shall indicate the selected anatomic structures | u_FMEA
User selects/deselects the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available | incorrect deselecting | use of device issue | Unable to Manipulate 3D Model during intra-operative | User unaware structure options are available; User cannot find structure options control; User does not understand parameters for structure | Surgery Prolonged | Reversible - (S2) |  | 3D Model shall be appeared in the Viewing panel. with a toggle switch; structure shall be appeared under 3D Model with a checkbox; Checkbox shall be appeared to turn structures on and off. | software shall provide select controls; software shall provide deselect controls | software shall indicate the deselected anatomic structures | u_FMEA
User scans the volumes of the different models to check against expectations | misread volume | use of device issue | Unable to verify volume during pro-operative | User unaware volumes are available; User cannot find volumes; User misreads volumes | No harm | No Harm - (S0) |  | 3D Model shall be clearly labeled with an icon in the toolbar. Volumns in CC shall be clearly labeled for each structure under 3D Model. | software shall provide volume accuracy controls; software shall provide volume specification controls | software shall indicate the 3D volume state | u_FMEA
User selects "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships | fail to manipulate 3D model | use of device issue | Unable to Manipulate 3D Model during pre-operative | User unaware pan/rotate/zoom/recenter options controls are available; User cannot find pan/rotate/zoom/recenter options controls; User does not understand how to use the pan/rotate/zoom/recenter options controls | Pre-operative Delayed | No Harm - (S0) |  | Pan/Rotate/Zoom shall be clearly labeld in the toolbar with an icon rather than using an invisible control | software shall provide pan parameter controls; software shall provide rotate parameter controls; software shall provide zoom parameter controls | software shall indicate pan state; software shall indicate rotate state; software shall indicate zoom state | u_FMEA
Alternatively, user selects from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view | fail to auto-orient model | image orientation incorrect | Unable to Manipulate 3D Model during pre-operative | User unaware anatomical plane icons or cube options controls are available; User cannot find anatomical plane icons or cube options controls; User does not understand how to use the anatomical plane icons or cube options controls | Pre-operative Delayed | No Harm - (S0) |  | anatomical plane icons shall be clearly labeld in the toolbar with an icon rather than using an invisible control | software shall provide anatomical plane parameter controls; software shall provide cube parameter controls | software shall indicate auto-orient state | u_FMEA
User may enter fullscreen at times for a better view of the model | fail to enter fullscreen | use of device issue | Unable to Manipulate 3D Model during pre-operative | User unaware fullscreen options controls are available; User cannot find fullscreen options controls; User does not understand how to use the fullscreen options controls | Pre-operative Delayed | No Harm - (S0) |  |  | software shall provide fullscreen parameter controls | software shall indicate fullscreen state | u_FMEA
User selects "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships | fail to manipulate 3D model | use of device issue | Unable to Manipulate 3D Model during intra-operative | User unaware pan/rotate/zoom/recenter options controls are available; User cannot find pan/rotate/zoom/recenter options controls; User does not understand how to use the pan/rotate/zoom/recenter options controls | Surgery Prolonged | Reversible - (S2) |  | Pan/Rotate/Zoom shall be clearly labeld in the toolbar with an icon rather than using an invisible control | software shall provide pan parameter controls; software shall provide rotate parameter controls; software shall provide zoom parameter controls | software shall indicate pan state; software shall indicate rotate state; software shall indicate zoom state | u_FMEA
Alternatively, user selects from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view | fail to auto-orient model | image orientation incorrect | Unable to Manipulate 3D Model during intra-operative | User unaware anatomical plane icons or cube options controls are available; User cannot find anatomical plane icons or cube options controls; User does not understand how to use the anatomical plane icons or cube options controls | Surgery Prolonged | Reversible - (S2) |  | anatomical plane icons shall be clearly labeld in the toolbar with an icon rather than using an invisible control | software shall provide anatomical plane parameter controls; software shall provide cube parameter controls | software shall indicate auto-orient state | u_FMEA
User may enter fullscreen at times for a better view of the model | fail to enter fullscreen | use of device issue | Unable to Manipulate 3D Model during intra-operative | User unaware fullscreen options controls are available; User cannot find fullscreen options controls; User does not understand how to use the fullscreen options controls | Surgery Prolonged | Reversible - (S2) |  |  | software shall provide fullscreen parameter controls | software shall indicate fullscreen state | u_FMEA
User selects an individual structure | incorrect selecting | use of device issue | Unable to Manipulate 3D Model during pre-operative | User unable to select intended structure using controls | Pre-operative Delayed | No Harm - (S0) |  |  |  | software shall indicate the selected structure | u_FMEA
User adjusts the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) | fail to adjust opacity level | optical issue | Unable to Manipulate 3D Model during pre-operative | User unaware opacity options controls are available; User cannot find the opacity levels options controls; User does not understand how to use the opacity options controls | Pre-operative Delayed | No Harm - (S0) |  | Individual structure shall be displayed opacity level number value; Slider controls shall allow the user to adjusts the opacity and see the number value change. | software shall provide opacity range controls | software shall indicate opacity percentage | u_FMEA
User selects "Done" button to complete the adjustment. | fail to save adjustments | device difficult to setup or prepare | Unable to Manipulate 3D Model during pre-operative | User unaware done button is available | Pre-operative Delayed | No Harm - (S0) |  |  |  | software shall indicate the adjustment is completed | u_FMEA
User selects an individual structure | incorrect selecting | use of device issue | Unable to Manipulate 3D Model during intra-operative | User unable to select intended structure using controls | Surgery Prolonged | Reversible - (S2) |  |  |  | software shall indicate the selected structure | u_FMEA
User adjusts the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) | fail to adjust opacity level | optical issue | Unable to Manipulate 3D Model during intra-operative | User unaware opacity options controls are available; User cannot find the opacity levels options controls; User does not understand how to use the opacity options controls | Surgery Prolonged | Reversible - (S2) |  | Individual structure shall be displayed opacity level number value; Slider controls shall allow the user to adjusts the opacity and see the number value change. | software shall provide opacity range controls | software shall indicate opacity percentage | u_FMEA
User selects "Done" button to complete the adjustment. | fail to save adjustments | device difficult to setup or prepare | Unable to Manipulate 3D Model during intra-operative | User unaware done button is available | Surgery Prolonged | Reversible - (S2) |  |  |  | software shall indicate the adjustment is completed | u_FMEA
User selects "Clip Apply Model" menu and allows module to load | load incorrect 3D model | device handling issue | Delayed Pre-Operative Plan during pre-operative | User unaware clip apply options are available User cannot find clip apply options control User does not understand parameters for clip apply options | Pre-operative Delayed | No Harm - (S0) |  | Clip Apply Model shall be clearly labeled with an icon in the toolbar; A spinner with the messages "Preparing Clip Apply model…" shall be appeared to indicate that the Clip Apply Model is loading; A tip reading "Right click on the model to apply a clip." shall be appeared when the Clip Apply Model is loaded on the screen. | software shall provide module controls | software shall indicate patient name; software shall indicate model name; software shall indicate patient number; software shall provide alerts when the uploaded models are corrupted, incorrect format | u_FMEA
User selects "Clip Apply Model" menu and allows module to load | fail to upload 3D model | device handling issue | Delayed Pre-Operative Plan during pre-operative | User unaware clip apply options are available User cannot find clip apply options control User does not understand parameters for clip apply options | Need to add to HHL | Need to add to HHL |  | Clip Apply Model shall be clearly labeled with an icon in the toolbar; A spinner with the messages "Preparing Clip Apply model…" shall be appeared to indicate that the Clip Apply Model is loading; A tip reading "Right click on the model to apply a clip." shall be appeared when the Clip Apply Model is loaded on the screen. | software shall provide module controls | software shall indicate patient name; software shall indicate model name; software shall indicate patient number; software shall provide alerts when the uploaded models are corrupted, incorrect format | u_FMEA
User manipulates the model as necessary to establish the correct orientation to apply a clip | inappropriate model orientation | image orientation incorrect | Unable to Manipulate 3D Model during pre-operative | User unaware pan/rotate/zoom/recenter/icons/cube/fullscreen options controls are available; User cannot find pan/rotate/zoom/recenter/icons/cube/fullscreen options controls; User does not understand how to use the pan/rotate/zoom/recenter/icons/cube/fullscreen options controls | Pre-operative Delayed | No Harm - (S0) |  | Pan/Rotate/Zoom and Recenter shall be clearly labeld in the toolbar with an icons rather than using an invisible controls; Icons in the 3D Model areas shall allow the user to snap to Axial, Sagittal and Coronal. | software shall provide pan parameter controls; software shall provide rotate parameter controls; software shall provide zoom parameter controls | software shall indicate pan state; software shall indicate rotate state; software shall indicate zoom state; software shall provide alerts when applying a clip on wrong model, clip disengaged, clip mishandled | u_FMEA
User manipulates the model as necessary to establish the correct orientation to apply a clip | orientation marked incorrectly | image orientation incorrect | Unable to Manipulate 3D Model during intra-operative | User unaware pan/rotate/zoom/recenter/icons/cube/fullscreen options controls are available; User cannot find pan/rotate/zoom/recenter/icons/cube/fullscreen options controls; User does not understand how to use the pan/rotate/zoom/recenter/icons/cube/fullscreen options controls | Surgery Prolonged | Reversible - (S2) |  | Pan/Rotate/Zoom and Recenter shall be clearly labeld in the toolbar with an icons rather than using an invisible controls; Icons in the 3D Model areas shall allow the user to snap to Axial, Sagittal and Coronal. | software shall provide pan parameter controls; software shall provide rotate parameter controls; software shall provide zoom parameter controls | software shall indicate pan state; software shall indicate rotate state; software shall indicate zoom state; software shall provide alerts when applying a clip on wrong model, clip disengaged, clip mishandled | u_FMEA
User right clicks on the model the apply the clip | clip mishandled | device handling issue | Inappropriate Pre-Operative Plan during pre-operative | User unaware of clip apply feature; User unsure how to access clip apply feature; User does not understand how to apply clip; User places clip in incorrect location; User unintentionally place clip; User unintentionally deletes an added clip | Pre-operative Delayed | No Harm - (S0) |  | When in Clip Apply, a tip reading "Right click on the model to apply a clip." shall be appeared when the Clip Apply Model is loaded on the screen; When clips are applied they shall be appeared in a list format with an icon to remove each clip. | software shall provide clip location controls; software shall provide clip delete permission controls | software shall provide clip info; software shall indicate clip operation state; software shall provide alerts when applying a clip on wrong model; software shall provide alerts when clip mishandled; software shall provide alerts when clip incorrect location | u_FMEA
If the user is unhappy with the clip placement they chose, they may either remove a single clip, or remove all clips | fail to deselect clip | device handling issue | Inappropriate Pre-Operative Plan during pre-operative | User unaware of clip apply feature; User unsure how to access clip apply feature; User does not understand how to remove clip | Pre-operative Delayed | No Harm - (S0) |  |  |  | software shall indicate clip removed | u_FMEA
User right clicks on the model the apply the clip | clip mishandled | device handling issue | Inappropriate Pre-Operative Plan during intra-operative | User unaware of clip apply feature; User unsure how to access clip apply feature; User does not understand how to apply clip; User places clip in incorrect location; User unintentionally place clip; User unintentionally deletes an added clip | Surgery Prolonged | Reversible - (S2) |  | When in Clip Apply, a tip reading "Right click on the model to apply a clip." shall be appeared when the Clip Apply Model is loaded on the screen; When clips are applied they shall be appeared in a list format with an icon to remove each clip. | software shall provide clip location controls; software shall provide clip delete permission controls | software shall provide clip info; software shall indicate clip operation state; software shall provide alerts when applying a clip on wrong model; software shall provide alerts when clip mishandled; software shall provide alerts when clip incorrect location | u_FMEA
If the user is unhappy with the clip placement they chose, they may either remove a single clip, or remove all clips | fail to deselect clip | device handling issue | Inappropriate Pre-Operative Plan during intra-operative | User unaware of clip apply feature; User unsure how to access clip apply feature; User does not understand how to remove clip | Surgery Prolonged | Reversible - (S2) |  |  |  | software shall indicate clip removed | u_FMEA
User observes the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip | misread volume | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User unaware volumes are available; User cannot find volumes; User misreads volumes; User mistakes volume simulation of one clip for that of another clip; User misunderstands volume labels | Pre-operative Delayed | No Harm - (S0) |  | After a clip is applied, volumn simulation data shall be clearly labed on the screen for organ volumn and healthy tissue. | software shall provide volume calculation controls; software shall provide volume accuracy controls | software shall provide clip info; software shall display resected organ volume; shoftware shall display health tissue volume; software shall provide alerts when applying a clip on wrong model; software shall provide alerts when clip disengaged; software shall provide alerts when clip mishandled; software shall provide alerts when clip incorrect location | u_FMEA
User observes the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip | fail to verify volume | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User unaware volumes are available; User cannot find volumes; User misreads volumes; User mistakes volume simulation of one clip for that of another clip; User misunderstands volume labels | Surgery Prolonged | Reversible - (S2) |  | After a clip is applied, volumn simulation data shall be clearly labed on the screen for organ volumn and healthy tissue. | software shall provide volume calculation controls; software shall provide volume accuracy controls | software shall provide clip info; software shall display resected organ volume; shoftware shall display health tissue volume; software shall provide alerts when applying a clip on wrong model; software shall provide alerts when clip disengaged; software shall provide alerts when clip mishandled; software shall provide alerts when clip incorrect location | u_FMEA
At any time, the user may select any of the clips in the menu in order to review the volume simulation details for that clip | invisible clip | no display or display failure | Delayed Pre-Operative Plan during pre-operative | User unaware clip review is available; User cannot find clip review options control; User does not understand parameters for clip review | Pre-operative Delayed | No Harm - (S0) |  | When clips are applied they shall be appeared in a list format for the user to select individaul clips and see the volumn simulation data. | software shall provide clip parameter controls; software shall provide volume unit controls; software shall provide volume accuracy controls | software shall indicate clip state; software shall indicate volume unit; software shall provide alerts when volume is simulated incorrectly | u_FMEA
At any time, the user may select any of the clips in the menu in order to review the volume simulation details for that clip | misread volume | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User does not read volume correctly | Surgery Prolonged | Reversible - (S2) |  | When clips are applied they shall be appeared in a list format for the user to select individaul clips and see the volumn simulation data. | software shall provide clip parameter controls; software shall provide volume unit controls; software shall provide volume accuracy controls | software shall indicate clip state; software shall indicate volume unit; software shall provide alerts when volume is simulated incorrectly | u_FMEA
When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons | fail to save surgical plan | loss of data | Delayed Pre-Operative Plan during pre-operative | User unaware plan saving is available; User cannot find plan saving options control; User does not understand parameters for plan saving | Pre-operative Delayed | No Harm - (S0) |  | Save New View button shall be appeared clearly in the header so that it is always prominently displayed on the screen and noticeable to the user;When the the plan is saved, it shall be appeared clearly in the My Views list. | software shall provide save controls | software shall indicate save state(e.g. saving, saving succesful, saving failed); software shall provide alerts when the surgical plan fails to be saved (due to space is full, system crash) | u_FMEA
When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons | save plan to different patient | patient data issue | Delayed Intra-Operative Plan during intra-operative | two patients have the same name | Need to add to HHL | Need to add to HHL |  |  | software shall provide save controls (patient name, patient ID) | software shall indicate full patient info; software shall provide alerts when multiple patients have the same name | u_FMEA
When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons | fail to name surgical plan | loss of data | Loss of Patient Data during pre-operative | User unaware rule to create name | Pre-operative Delayed | No Harm - (S0) |  | When the the plan is saved, it shall be appeared clearly in the My Views list. | software shall provide naming controls | software shall provide alerts when the surgical plan is named incorrectly | u_FMEA
When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons | inappropriate module selected | use of device issue | Loss of Patient Data during intra-operative | User unaware wrong module selected | Surgery Prolonged | Reversible - (S2) |  | Save New View button shall be appeared clearly in the header so that it is always prominently displayed on the screen and noticeable to the user. When the the plan is saved, it shall be appeared clearly in the My Views list. | software shall provide module controls | software shall indicate the selected module state | u_FMEA
User navigates to "My Views" menu | fail to navigate to menu | use of device issue | Delayed Pre-Operative Plan during pre-operative | User cannot navigate properly; Complex or confusing control system. | Pre-operative Delayed | No Harm - (S0) |  | Save New View button shall be appeared clearly in the header so that it is always prominently displayed on the screen and noticeable to the user | software shall provide navigation controls | software shall indicate navigate state | u_FMEA
User selects "SAVE NEW VIEW" and pop-up appears | selecting confused | use of device issue | Delayed Pre-Operative Plan during pre-operative | User unaware pop-up appear; Complex or confusing control system. | Pre-operative Delayed | No Harm - (S0) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide pop-up view controls | software shall indicate selecting state | u_FMEA
User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" | unable to rename view | use of device issue | Delayed Pre-Operative Plan during pre-operative | User unaware difference between save and replace; User unaware save as a new view is available | Pre-operative Delayed | No Harm - (S0) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide renaming controls | software shall indicate save state; software shall indicate replace state; software shall indicate rename state; software shall provide alerts when the view name is invalid, unsuccesssful saving | u_FMEA
User presses "CONFIRM" a second time and the view is saved | fail to save view | data back-up problem | Delayed Pre-Operative Plan during pre-operative | User unaware it needs to press confirm a second time | Pre-operative Delayed | No Harm - (S0) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide confirm controls | software provide notes to show the saving state (e.g. saving, saving succesful, view title) | u_FMEA
User navigates to "My Views" menu | fail to navigate to menu | use of device issue | Delayed Intra-Operative Plan during intra-operative | User cannot navigate properly; Complex or confusing control system. | Surgery Prolonged | Reversible - (S2) |  | Save New View button shall be appeared clearly in the header so that it is always prominently displayed on the screen and noticeable to the user | software shall provide navigation controls | software shall indicate navigate state | u_FMEA
User selects "SAVE NEW VIEW" and pop-up appears | selecting confused | use of device issue | Delayed Intra-Operative Plan during intra-operative | User unaware pop-up appear; Complex or confusing control system. | Surgery Prolonged | Reversible - (S2) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide pop-up view controls | software shall indicate selecting state | u_FMEA
User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" | unable to rename view | use of device issue | Delayed Intra-Operative Plan during intra-operative | User unaware difference between save and replace; User unaware save as a new view is available | Surgery Prolonged | Reversible - (S2) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide renaming controls | software shall indicate save state; software shall indicate replace state; software shall indicate rename state; software shall provide alerts when the view name is invalid, unsuccesssful saving | u_FMEA
User presses "CONFIRM" a second time and the view is saved | fail to save view | data back-up problem | Delayed Intra-Operative Plan during intra-operative | User unaware it needs to press confirm a second time | Surgery Prolonged | Reversible - (S2) |  | When the the plan is saved, it shall be appeared clearly in the My Views list | software shall provide confirm controls | software provide notes to show the saving state (e.g. saving, saving succesful, view title) | u_FMEA
User navigates to "My Views" menu and selects a file to view | difficult to navigate function | use of device issue | Inappropriate Pre-Operative Plan during pre-operative | User unaware saved view is available; User cannot find saved view options control | Pre-operative Delayed | No Harm - (S0) |  | My Views shall be clearly labeld in the toolbar with an icon. When Views are saved, they shall be appeared in a list format under My Views for the user to select and display on the screen. | software shall provide file format controls; software shall provide navigate controls | software shall indicate navigate state; software shall indicate file upload state; software shall provide alerts when the file has invalid format | u_FMEA
User navigates to "My Views" menu and selects a file to view | fail to navigate to menu | use of device issue | Inappropriate Pre-Operative Plan during intra-operative | User unaware saved view is available; User cannot find saved view options control | Surgery Prolonged | Reversible - (S2) |  | My Views shall be clearly labeld in the toolbar with an icon. When Views are saved, they shall be appeared in a list format under My Views for the user to select and display on the screen. | software shall provide file format controls; software shall provide navigate controls | software shall indicate navigate state; software shall indicate file upload state; software shall provide alerts when the file has invalid format | u_FMEA

## Combined_forGPT2
Index | category | Inputs | Failure_Modes | Hazard | Source | Harm
| Market | Pre-op planning is a critical step to ensure safety, efficacy and efficiency of resection procedures. | inadequate pre-op planning | application program issue | Viewer
| Market | Pre-op planning is intended to inform 3 clinical decisions as framed by surgeons: “where to enter”, “how to move” and the “amount of resection”. | incorrect clinical decision | Poor Quality Image | Viewer
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 1. Size, location and vascularity of the tumor(s) | poor tumor image | Poor Quality Image | Viewer
| Market | Larger tumors are more concerning for spread. Anything above 6 or 7 cm might mean a full removal and how it sits on the kidney in relationship to other organs. | tumor removed inappropriately | Poor Quality Image | Viewer
| Market | Where is the cancer? Has it invaded the outside wall? Where has it invaded? What structures have been invaded? | fail to visualize cancer location | poor quality image | Viewer
| Market | My focus is to see the proper anatomy of the mass and its vascular anatomy | poor anatomy image | optical issue | Viewer
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 2. Tumor relationship to surrounding critical structures (i.e. vasculature, ureter, chest wall, etc.) | fail to visualize surrounding critical structures | Poor Quality Image | Viewer
| Market | I need to avoid critical structures like the ureter. Knowing where it is can speed things up for us. | fail to visualize surrounding critical structures | Poor Quality Image | Viewer
| Market | I’m looking for the relationship of tumor to surrounding structures, particularly the vasculature. The hepatic veins and vena cava are the bigger ones for liver | fail to visualize surrounding critical structures | Poor Quality Image | Viewer
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 3. Patient-specific anatomical variations or aberrant anatomy. | fail to visualize anatomical variation | poor quality image | Viewer
| Market | There is a great deal of variety within the anatomy. This is what we're looking for in the CT | fail to visualize anatomy variety | poor quality image | Viewer
| Market | Because of the anatomical variability, we are often operating blindly. We don't know how many bronchi we will come into.” | fail to visualize anatomy variety | poor quality image | Viewer
| Market | By integrating 3D models and patient simulation from Visible Patient into J&J ecosystem, we’ll provide the following value to the clinicians: | inadequate 3D model | application program issue | Viewer
| Market | we’ll provide the following value to the clinicians: Highly detailed, virtual 3D model, using patients’ pre-op images (CT, MRI) | fail to develop detailed 3D model | application program issue | Viewer
| Market | we’ll provide the following value to the clinicians: Color-coded, interactive experience that allows you to explore and visualize spatial relationships between tumors, critical structures and surrounding anatomy of your patient | fail to visualize spatial relationship | image display error | Viewer
| Market | we’ll provide the following value to the clinicians: Easy-to-use interface that is accessible on your team’s PC/Mac, tablet, and mobile devices for reference in the OR, at home, and on-the-go | difficult to use | human-device interface issue | Viewer
| Market | we’ll provide the following value to the clinicians: Allows you to explore anatomy on your schedule, without reliance on interpretation by your radiology team | fail to explore anatomy independently | human-device interface issue | Viewer
| Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to provide virtual clip-applying tool | human-device interface issue | Viewer
| Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to real time assess organ | poor quality image | Viewer
| Market | Surgeons use 2D images to “envision a 3D model” in their mind when planning their approach. More experienced surgeons are more likely to be confident in this ability. However, even experienced surgeons acknowledge inefficiency and communication challenges with the surgical team when developing a surgical plan, e.g. reliance on the radiologist to develop a surgical plan. | fail to develop surgical plan independently | poor quality image | Viewer
| Market | In addition to 3D models, 2D visualization, with color, of a 3D model (Anatomical Atlas) is also a valuable tool for clinicians as it is “what I’m used to and know how to interpret, but can see things better”. | inappropriate color code | use of device issue | Viewer
| Market | The patient simulation (Clip Applying) tool provides significant value when planning painstaking, tissue-sparing procedures. For many, the patient simulation tool is considered a “breakthrough” because of high-quality image and instant volume calculations that allow for simulation of different surgical plans pre-operatively (e.g. segmentectomies vs a lobectomy). | fail to provide real time organ volume calculation | poor quality image | Viewer
| Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | inefficacy pre-op planning | application program issue | Viewer
| Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | unclear simulation | poor quality image | Viewer
| Market | Surgeons see the flexibility in access across devices and locations (at home, at the office, on the go, in the OR) as a critical attribute of the solution | fail to access across devices and locations | human-device interface issue | Viewer
| Market | All surgeons for the target specialties (thoracic, HBP, colorectal and urology) are procedurally required to review pre-operative patient images before the procedure. However, only a subset of surgeons refers back to pre-operative patient images intra-operatively, mostly due to an unexpected finding or uncertainty. Hence, the primary use case for pre-operative planning is before the procedure begins; intraoperative reference of the preoperative planning is secondary. | missing review of pre-operative patient images | use of device issue | Viewer
| Market | The primary clinical goals of this project are to design, develop and deploy a solution that will enable Digital Solution to ● launch surgical decision support, especially medical imaging software application products | incorrect surgical decision | application program issue | Viewer
| Market | enable surgeons to order, manage, and review patient-specific 2D and 3D anatomical models | fail to manage 3D model | application program issue | Viewer
| Market | enable surgeons to view pre-operative images and 2D and 3D anatomical models online | fail to manage 2D images online | application program issue | Viewer
| Market | enable surgeons to perform a simulated resection by applying virtual clips to approximated segments in order to simulate post-operative organ volume | fail to provide real time organ volume calculation | application program issue | Viewer
| Market | enable surgeons to save a viewing and simulation session and continue at a later time | fail to continue the paused session | application program issue | Viewer
| Market | enable surgeons to continue their viewing and simulation session in the OR using the Digital Hub hardware device after the commercial release of the Digital Hub (viewing of data sets on the Digital Hub will be outlined in a future Charter), | insufficient memory size | data back-up problem | Viewer
| Market | enable surgeons to continue viewing their images and simulation in the OR on the web portal | fail to access through web portal | human-device interface issue | Viewer
| Market | Integrated solution should be comprised of the following major components: surgeon portal, digital solutions viewer, digital surgery enablement platform. | fail to access through web portal | application program issue | Viewer
| Market | Surgeon Portal - a web portal for surgeons to log into, select the appropriate language, order the modeling service, review the status of the modeling service, upload DICOM images, start the Digital Solutions Viewer. | insecure login | unauthorized access to computer system | Viewer
| Market | Digital Solutions Viewer - a DICOM viewer that allows for interactions with the anatomical model, allow for simulation of resection by placing a virtual clip. | fail to provide simulation resection | application program issue | Viewer
| Market | Digital Solutions Viewer - a DICOM viewer that allows for viewing a DICOM image and/or the 3D anatomical model selected by the surgeon through the surgeon portal. | fail to view images through web portal | application program issue | Viewer
| Market | Digital Solutions Viewer - a DICOM viewer allows for placement of annotations and measurements, show the anatomical model, allow for interactions with the anatomical model | fail to place measurements | application program issue | Viewer
| Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables user management, data storage | fail to manage users | application program issue | Viewer
| Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables data storage, and provides DICOM studies to the external 3D model creation service as well as to the Digital Solutions Viewer. | fail to provide external 3D model creation | application program issue | Viewer
| Market | Bandit Phase 2 will deliver an integrated experiences for the customer, which can be categorized in three groups: ● 3D Model Submission and Delivery ● Digital Solution Viewer ● Patient Modeling and Outcome Simulation | fail to deliver 3D model | application program issue | Viewer
| Market | AI-enhanced visualization and personalized insights to help surgeons get a clearer roadmap for surgery | incorrect surgical decision | poor quality image | Viewer
| Market | 3D model submission and delivery: 3D model order submission by clinical user; 3D model order management by order admin; 3D model order billing; PACS integration | inappropriate user | application program issue | Viewer
| Market | Image viewer: image viewer; meaurement tools | lack of measurement tool | application program issue | Viewer
| Market | patient modeling and outcome simulation: integrated 3D models; patient specific simulation of resection | fail to provide simulation resection | application program issue | Viewer
| Market | Target for the Digital Solutions Viewer is the EU. The Digital Solutions Viewer will be the first CE marked product and first product under the new EU Medical Device Regulation (MDR) for Auris. | fail to comply with CE Mark | compliance issue | Viewer
| Market | The Digital Solutions Viewer is considered software as a medical device (SaMD) and will need to conform to the requirements of the Medical Device Regulation (EU) 2017/745. The Viewer will need to provide Technical Documentation (TD) covering the design, development, and state of the art with respect to safety and performance and demonstrate compliance to the General safety and Performance requirements in Annex I. | IEC 62304 violated | computer software issue | Viewer
| Market | The proposed classification of the Digital Solutions Viewer is a Class IIa Rule 11 (Software intended to provide information which is used to make decisions with diagnosis or therapeutic purposes). | inappropriate medical device classification | compliance issue | Viewer
| Market | As the Digital Solutions viewer will be the first MDR product for Auris, the Auris QMS must be MDR ready to support certification. The proposed certificate scope will be for image visualization and planning software. | inappropriate certificate scope | compliance issue | Viewer
| Market | The MDR stipulates that the manufacturer must provide sufficient clinical evidence (Article 61) to justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR. | insufficient clinical evidence | data issue | Viewer
| Market | The MDR provides a possibility to use clinical data related to a marketed equivalent device (found in the clinical evaluation required for a device) under conformity assessment. | lack of clinical data from equivalent device | data issue | Viewer
| Market | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes of the software are intended to be similar for the Auris Viewer. | inconsistent algorithm | programming issue | Viewer
| Market | The equivalence requirement to demonstrate sufficient levels of access to comparator data should be met as Visible Patient is collaborating with J&J and the Auris Viewer is intended to interface with the Visible Patient Lab 3D recreated models. | fail to access across devices and locations | human-device interface issue | Viewer
| Market | Primary User Targets: Surgeons in Thoracic, Urology, Colorectal, and Hepato-biliary | inappropriate user | user used incorrect product for intended use | Viewer
| Market | Secondary User Targets (secondary expansion within primary user hospitals): Lower volume surgeons in Primary User Target specialties, Surgeons in GYN | inappropriate user | user used incorrect product for intended use | Viewer
| Market | Builds Loyal Digital Surgery Customer Base: Increases the value of the online Ethicon Digital Solutions video, image & case library, creating network effects & switching costs | lack of case library | inadequate instructions for healthcare professional | Viewer
| Market | Provides Critical Data Pipeline: Adds pre-op image and 3D model data to feed the Ethicon AI Insights platform | incompatible accessory | component or accessory incompatibility | Viewer
| Market | Pre-operative 3D planning software is in-market today, but most focus on orthopaedics and neurosurgery. Bandit has three main competitors in the US (Fujifilm Synapse, TeraRecon iNtuition, and Intuitive Surgical IRIS™), but only Fujifilm and TeraRecon have a significant presence in EMEA, Bandit’s initial regional foothold market. Only Fujifilm is used today in APAC. | inadequate pre-op planning | application program issue | Viewer
| Market | Fujifilm and TeraRecon require radiologist creation of 3D models and involvement with interpretation and dictation, whereas Bandit eliminates this step, giving surgeons a streamlined pre-op planning sequence that puts the control in their hands on their schedule. | fail to develop surgical plan independently | application program issue | Viewer
| Market | Intuitive IRIS™ is currently cleared for use with CT only, whereas the others, including Bandit, are applicable for MRI as well. | fail to use 2D MRI images to envision a 3D model | application program issue | Viewer
| Market | The major project milestones and estimated timing are listed below. This schedule assumes a fully staffed development team and approved program budget by the end of Phase 1. An updated schedule and list of major milestones along with program risks will be provided in the Design and Development Plan (DDP) prior to exiting the Planning and Definition Phase. | fail to identify risk | market issue | Viewer
| Market | The following assumptions were made as basis of this Project Authorization Charter: ● CE Mark is prioritized (EMEA launch) and 510K Approval is performed afterwards | inappropriate authorization lists | market issue | Viewer
| Market | Expedited Review (MDR) has minimal questions and no major revisions needed | prolonged Eu MDR | compliance issue | Viewer
| Market | FDA acceptance of substantial equivalence with minimal questions | prolonged FDR approval | compliance issue | Viewer
| Market | Code the complete the Digital Solutions Viewer is solid with few bugs/issues by mid-March 2021 | major bug existed | computer software issue | Viewer
| Market | Web Portal & DSEP are not part of MDR submission | MDR submission delayed | compliance issue | Viewer
| Market | QMS audit can completed by regulatory body by end of May (earlier than planned) | audit observations | compliance issue | Viewer
| Market | QMS documents can be reviewed and released in Agile within 2 weeks (documents complete to released) | inadequate QMS document lists | compliance issue | Viewer
| Clinical_Workflow | The intended use of the device is an aid in pre-surgical planning for procedures requiring spatial understanding of complex patient anatomy. | fail to provide spatial relationship | application program issue | Viewer
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | inadequate surgical plan | improper or incorrect procedure or method | Viewer
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor defects | application program issue | Viewer
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor abnormalities | application program issue | Viewer
| Clinical_Workflow | Patients undergoing procedures in the general, bariatric, HPN, gynecology, thoracic, and urology specialties, who require medical images to be taken to plan for their surgical procedure. | poor medical image | Poor Quality Image | Viewer
| Clinical_Workflow | If the medical device fails while being used to reference medical images or 3D models during a procedure, the user may need to access the patient images on another PACS viewer to proceed. | lack of backup viewer | fail-safe mechanism issue | Viewer
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can take linear measurements between two points displayed on the DICOM or 3D model | inappropriate measurement algorithm | measurement system incompatibility | Viewer
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can calculate a designated area shown on a 2D DICOM slice | fail to calculate slice area | incorrect software programming calculations | Viewer
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can through the clip apply feature, calculate volume of a designated portion of an organ to be resected | fail to calculate volume of resected organ | incorrect software programming calculations | Viewer
| Clinical_Workflow | Measurements taken on a 2D device should be accurate for images acquired with high resolution. | low resolution image | Image Resolution Poor | Viewer
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | fail to reconstruct 3D model | application program issue | Viewer
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | inappropriate segnmentation alogorithm applied | application program issue | Viewer
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | incorrect volume calculation | parameter calculation error due to software problem | Viewer
| Clinical_Workflow | The medical device may be used with a radiology monitor | fail to compatible with radiology monitor | Incompatibility Problem | Viewer
| Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | software installation problem | problem with software installation | Viewer
| Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | obsolete software used | application program version or upgrade problem | Viewer
| Clinical_Workflow | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | fail to compatible with operating system | Incompatibility Problem | Viewer
| Clinical | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | short software life cycle | computer software issue | Viewer
| Clinical_Workflow | When using a tablet to interface with the device during a procedure, there is a risk that the procedure would coincide with battery depletion. | premature battery depletion | Premature Discharge of Battery | Viewer
| Clinical_Workflow | successful application of the medical device is critically dependent on human factors such as the user interface. The user interface is fundamental to interacting with and displaying measurements of the patient images and anatomic models. | inappropriate interface design | inadequate user interface | Viewer
| Clinical_Workflow | use errors can arise from not understanding or properly navigating the user interface. Symbols for the user interface have been based on those common in other PACS viewer products, but may be arranged differently or applied slightly differently for this medical device. | unclear symbol | device markings issue | Viewer
| Clinical_Workflow | In general, the medical device is used in the office environment for planning prior to a case. When using the medical device in the operating room, the user may be more distracted and could make use errors such as loading the incorrect patient’s data, or ignoring the orientation indicators. | improper or incorrect use of device | improper or incorrect procedure or method | Viewer
| Clinical_Workflow | The medical device displays information. The display relies on text, colors, sizing, etc to relay spatial information and distinguish tissue types. These can be adjusted for accessibility but must take colorblind users into account | fail to consider colorblind user | incorrect display | Viewer
| Clinical_Workflow | the medical device must be used with a display large enough and with high enough resolution so that the user can clearly see details of the patient images and anatomic models. | low resolution | Image Resolution Poor | Viewer
| Clinical_Workflow | The device must render the patient models taking all these factors into account. | inappropriate model rendered | Poor Quality Image | Viewer
| Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to navigate functions | inadequate user interface | Viewer
| Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to access settings | inadequate user interface | Viewer
| Clinical_Workflow | User shall launch program by openning Aurora viewer during the selecting an order phase | fail to launch program | application program issue | Viewer
| Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to navigate tab | human-device interface issue | Viewer
| Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to select order | human-device interface issue | Viewer
| Clinical_Workflow | User shall find patient(s) of interest by searching for patient by name, MRN, surgeon during the selecting an order phase | fail to search patient name | application program issue | Viewer
| Clinical_Workflow | User shall find patient(s) of interest through sorting lists by specialty, patient, surgeon, delivery date, or status to find patient of interest during the selecting an order phase | fail to sort patient interest list | application program issue | Viewer
| Clinical_Workflow | User shall find patient(s) of interest through filters list using "FILTER BY" tool in order to find patient during the selecting an order phase | fail to filter patient interest list | application program issue | Viewer
| Clinical_Workflow | User shall select patient and launch image viewer by touching/clicking on row and row will highlight when selected during the selecting an order phase | unresponsive touch | failure to deploy | Viewer
| Clinical_Workflow | User shall select patient and launch image viewer by clicking "NEXT" to load patient images and 3D model during the selecting an order phase | unresponsive click | failure to deploy | Viewer
| Clinical_Workflow | User shall verify images are correct by double checking patient informaiton during the selecting an order phase | fail to verify image | application program issue | Viewer
| Clinical_Workflow | User shall view 2D image stacks by monitoring loading of the 2D images on the "Viewing" panel during the load and view desired image stacks phase | fail to monitor loading images | application program issue | Viewer
| Clinical_Workflow | After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks. | fail to view image stack | application program issue | Viewer
| Clinical_Workflow | User shall select which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas during setting screen layout phase | inadequate image types | application program issue | Viewer
| Clinical_Workflow | User shall select whether they would like the segmented anatomic regions to be displayed by infill or contour when turning on and off segmentation layers | fail to display anatomic contour | application program issue | Viewer
| Clinical_Workflow | User shall adjust view of desired image stacks by operating pan, rotate, zoom, and scroll to select stack they would like to manipulate during the adjust view of desired image stacks phase | unable to adjust view of image stack | application program issue | Viewer
| Clinical_Workflow | User shall use touchscreen, mouse, or side menu to select "Pan, Rotate, Zoom or Scroll" | insensitive touch screen | decreased sensitivity | Viewer
| Clinical_Workflow | User shall undo a single adjustment with the "Undo" button during the adjust view of desired image stacks phase | fail to undo an adjustment | application program issue | Viewer
| Clinical_Workflow | User shall reset the view altogether with the "Re-Center" button during the adjust view of desired image stacks phase | fail to reset view | application program issue | Viewer
| Clinical_Workflow | User shall navigate to the Windowing menu on the sidebar to adjust the contrast or brightness of the image during the adjust view of desired image stacks phase | fail to adjust image contrast | application program issue | Viewer
| Clinical_Workflow | User shall select from presets to auto-adjust window width and level based on target anatomy the image was acquired to show during the adjust view of desired image stacks phase | inappropriate window width | application program issue | Viewer
| Clinical_Workflow | User shall choose to select "Width/Level" button to make customize adjustments to the brightness and contrast during the adjust view of desired image stacks phase | fail to adjust image brightness | application program issue | Viewer
| Clinical_Workflow | User shall drag slider(s) to desire width and level percentages when adjust windowing during the adjust view of desired image stacks phase | fail to drag slider | human-device interface issue | Viewer
| Clinical_Workflow | User shall reset the view using the "Reset" button when adjust windowing during the adjust view of desired image stacks phase | fail to reset view | application program issue | Viewer
| Clinical_Workflow | User shall select the annotation menu by navigate to tools from the sidebar during the annotation tool phase | inappropriate annotation tool | human-device interface issue | Viewer
| Clinical_Workflow | User shall select "distance" tool to measure the straight-line distance between two points on the image during the annotation tool phase | incorrect distance measurement | incorrect measurement | Viewer
| Clinical_Workflow | User shall click on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen during the annotation tool phase | fail to pick measurement point | human-device interface issue | Viewer
| Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to drag and drop line | human-device interface issue | Viewer
| Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to resize line | human-device interface issue | Viewer
| Clinical_Workflow | User shall delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu by using distance tool during the annotation tool phase | fail to add line | application program issue | Viewer
| Clinical_Workflow | User shall add and delete arrows in a similar fashion after clicking the "Arrow" menu option. Arrows added to the screen will also appear in the Annotation sidebar menu list by arrow tool during the annotation tool phase | fail to delete arrow | human-device interface issue | Viewer
| Clinical_Workflow | User shall add a note by selecting the "note" annotation option and clicking anywhere on the image during the annotation tool phase | fail to add notes at anywhere | application program issue | Viewer
| Clinical_Workflow | when a textbox appears, user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu during the annotation tool phase | fail to add any notes | application program issue | Viewer
| Clinical_Workflow | User shall return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox during the annotation tool phase | fail to edit notes | application program issue | Viewer
| Clinical_Workflow | User shall remove any added note using the trash icon in the Annotation menu or on the note itself during the annotation tool phase | fail to remove any notes | application program issue | Viewer
| Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename arrow | human-device interface issue | Viewer
| Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename distance | human-device interface issue | Viewer
| Clinical_Workflow | User shall reset all annotations using the "Reset" button during the annotation tool phase | fail to reset annotations | application program issue | Viewer
| Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle on 3D model | application program issue | Viewer
| Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle off 2D images | application program issue | Viewer
| Clinical_Workflow | User shall confirm they are viewing the correct 3D model by checking the information shown onscreen. They toggle on "Patient Info" if not already on when openning 3D Model View during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Viewer
| Clinical_Workflow | User shall select "3D Model" menu to see which anatomic structures are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Viewer
| Clinical_Workflow | User shall select/deselect the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to deselect 3D model | application program issue | Viewer
| Clinical_Workflow | User shall scan the volumes of the different models to check against expectations when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to scan model volume | application program issue | Viewer
| Clinical_Workflow | User shall select "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships when setting desired orientation during the visualizing and manipulating 3D model phase | inappropriate model orientation | image orientation incorrect | Viewer
| Clinical_Workflow | Alternatively, user shall select from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view when setting desired orientation during the visualizing and manipulating 3D model phase | fail to auto-orient model | image orientation incorrect | Viewer
| Clinical_Workflow | User shall enter fullscreen at times for a better view of the model when setting desired orientation during the visualizing and manipulating 3D model phase | fail to enter fullscreen | application program issue | Viewer
| Clinical_Workflow | User shall select an individual structure when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure visibility | optical issue | Viewer
| Clinical_Workflow | User shall adjust the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure opacity | optical issue | Viewer
| Clinical_Workflow | User shall select "Done" button to complete the adjustment when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to done an adjustment | failure to deploy | Viewer
| Clinical_Workflow | User shall select "Clip Apply Model" menu and allows module to load when navigating to Clip Apply feature during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | human-device interface issue | Viewer
| Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | Viewer
| Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | inappropriate model orientation | image orientation incorrect | Viewer
| Clinical_Workflow | User shall right click on the model the apply the clip when adding and removing Clips during applying clips to simulate resection volume phase | wrong mouse click | use of device issue | Viewer
| Clinical_Workflow | User shall observe the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to place next clip | application program issue | Viewer
| Clinical_Workflow | If the user is unhappy with the clip placement they chose, they shall either remove a single clip, or remove all clips when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | Viewer
| Clinical_Workflow | At any time, the user shall select any of the clips in the menu in order to review the volume simulation details for that clip when adding and removing Clips when saving the plan during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | application program issue | Viewer
| Clinical_Workflow | When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons when saving the plan during applying clips to simulate resection volume phase | fail to save surgical plan | data back-up problem | Viewer
| Clinical_Workflow | If they choose to save the plan, user names the plans and confirms. (Need to check) when saving the plan during applying clips to simulate resection volume phase | fail to name surgical plan | data back-up problem | Viewer
| Clinical_Workflow | User shall navigate to "My Views" menu when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | Viewer
| Clinical_Workflow | User shall select "SAVE NEW VIEW" and pop-up appears when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | Viewer
| Clinical_Workflow | User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" when saving current view for later access during saving current view phase | fail to replace current view | data back-up problem | Viewer
| Clinical_Workflow | User shall press "CONFIRM" a second time and the view is saved when saving current view for later access during saving current view phase | fail to confirm current view | application program issue | Viewer
| DIR | Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer. | fail to compatible with regular computer | Incompatibility Problem | Viewer
| DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | fail to open images | image display error | Viewer
| DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | inappropriate image format | data issue | Viewer
| DIR | The system shall have the capability to render 2D views of the preoperative patient images. | unable to render 2D views | application program issue | Viewer
| DIR | The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study. | unable to render 3D views | application program issue | Viewer
| DIR | The system shall have the capability to render a view of 3D model segmentations | unable to render 3D views | application program issue | Viewer
| DIR | The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view | unable to render 2D overlay on 3D segmentation | application program issue | Viewer
| DIR | The system shall provide a method to zoom in and out on the views. | unable to zoom in views | application program issue | Viewer
| DIR | The system shall provide a method to pan a viewing panel. | unable to pan viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to rotate a 2D viewing panel. | unable to rotate 2D viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to rotate a 3D viewing panel. | unable to rotate 3D viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel. | fail to toggle image visibility | application program issue | Viewer
| DIR | The system shall provide a method to rotate the cut plane within an MPR view. | unable to rotate cut plane | application program issue | Viewer
| DIR | The system shall provide a method to undo and redo user interactions. | fail to undo user interaction | use of device issue | Viewer
| DIR | The system shall provide a method to undo and redo user interactions. | fail to redo user interaction | use of device issue | Viewer
| DIR | The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size. | unable to resize viewing panel | application program issue | Viewer
| DIR | The system shall provide a method for hiding a viewing panel. | unable to hide viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to adjust the Window Center and Window Width adjustments using presets. | unable to adjust window center | application program issue | Viewer
| DIR | The system shall provide a method for manually adjusting window width and center values for the intensity range | unable to adjust window width | application program issue | Viewer
| DIR | The system shall provide a method for a reset all viewing panels simultaneously | unable to reset viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to select the layout for the viewing panels | unable to select layout for viewing panel | application program issue | Viewer
| DIR | The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume. | unable to advance MPR cut plane | application program issue | Viewer
| DIR | The system shall be able to display metadata contained in the DICOM data. | fail to display DICOM data | image display error | Viewer
| DIR | The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects. | fail to list 3D models | application program issue | Viewer
| DIR | The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model. | fail to display volumetric sizing information | no display or display failure | Viewer
| DIR | The system shall be able to display metadata contained in the 3D model container | fail to display metadata | no display or display failure | Viewer
| DIR | The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels | fail to display semi-transparent overlay | application program issue | Viewer
| DIR | The system shall use SI units and their derivatives | inappropriate measurement unit | incorrect measurement | Viewer
| DIR | The system shall provide a method for performing one or more measurements within one of the viewing panels | unable to perform multiple measurements | application program issue | Viewer
| DIR | The system shall label and list measurements automatically | fail to label measurement automatically | application program issue | Viewer
| DIR | The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object. | fail to place annotation on 3D model | application program issue | Viewer
| DIR | The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel. | fail to edit notes | application program issue | Viewer
| DIR | The system shall allow for simulation of resection by applying a clip to approximated segments. | fail to simulate resection | application program issue | Viewer
| DIR | The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip. | fail to display volumetric sizing information | no display or display failure | Viewer
| DIR | The system shall provide the ability to support individual language packs. | inadequate language pack | inadequate user interface | Viewer
| DIR | The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language. | inadequate language pack | inadequate user interface | Viewer
| DIR | The system shall, when appropriate and available, use standard internationally recognized symbols. | unstandardized symbol | device markings issue | Viewer
| DIR | The system shall, when appropriate and available, use standard internationally recognized icons. | unstandardized icon | device markings issue | Viewer
| DIR | The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System. | fail to record operative data | data back-up problem | Viewer
| DIR | The System shall record all functions performed by the System pre, intra, and post operatively. | fail to record operative data | loss of data | Viewer
| DIR | The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel. | fail to record monitoring data | data back-up problem | Viewer
| DIR | The System shall record system functions and operational events generated by the System. | fail to record operational events | data back-up problem | Viewer
| CIR | The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning. | fail to create pre-op planning | application program issue | Viewer
| CIR | The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images. | incomplete 3D model | application program issue | Viewer
| CIR | All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated. | incorrect revision | device markings issue | Viewer
| CIR | 3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator. | fail to create 3D Model | poor quality image | Viewer
| CIR | 3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon. | fail to provide patient diagnosis info | use of device issue | Viewer
| CIR | Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components. | software corrupted | computer software issue | Viewer
| CIR | DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data. | image not in DICOM format | image display error | Viewer
| CIR | PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images. | fail to process medical imaging | application program issue | Viewer
| CIR | Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format. | image not in DICOM format | data issue | Viewer
| CIR | Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively | fail to refer pre-op plan intraoperatively | application program issue | Viewer
| CIR | 3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models | fail to upload 3D model | application program issue | Viewer
| CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | inadequate IFU | inadequate instructions for healthcare professional | Viewer
| CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | lack of video | nadequate instructions for non-healthcare professional | Viewer
| CIR | Marketing and Sales Representative: Internal users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies. | fail to monitor product use | Market Issue | Viewer
| CIR | View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to manipulate visualization parameters | application program issue | Viewer
| CIR | Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings. | unable to measure point to point | application program issue | Viewer
| CIR | 2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings | unable to develop surgical plan | application program issue | Viewer
| CIR | View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to review 3D model of pathology | application program issue | Viewer
| CIR | Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings | fail to manipulate visualization parameters | application program issue | Viewer
| CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | fail to provide simulation resection | application program issue | Viewer
| CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | incorrect volume calculation | parameter calculation error due to software problem | Viewer
| CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure image display | application program issue | Viewer
| CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure 3D model display | application program issue | Viewer
| CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display 3D model at OR monitor | incompatibility problem | Viewer
| CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display patient image at OR monitor | incompatibility problem | Viewer
| CIR | Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device. | clinical workflow interrupted | disinfection or sterilization issue at user location | Viewer
| CIR | Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice. | patient data breached | unauthorized access to computer system | Viewer
| CIR | Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently. | lack of local language | human-device interface issue | Viewer
| CIR | Indications for Use for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party. | Indications for Use misalignment | failure to align | Viewer
| CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain regulatory clerance | compliance issue | Viewer
| CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain registration | compliance issue | Viewer
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to capture data | data issue | Viewer
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to store data | data issue | Viewer
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to follow privacy regulation | compliance issue | Viewer
| CIR | Branding: As company marketing representative, I need the application to follow company defined branding, so that I can maximize consistency in the look and feel of company digital products that may be used by the same user base and reduce cognitive load on the users when using company digital products. | branding misaligned | compliance issue | Viewer
| CIR | Demo Capabilities: As company marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively. | unclear workflow | data issue | Viewer
| Clinical_HA | Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body. | incorrect procedure | bacterial contamination of device | Viewer
| Clinical_HA | Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| Clinical_HA | Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Viewer
| FDA | 3D projections may be incorrect when the pair of images used is a secondary capture generated from a series containing more than two images, where one of the images has been manually flipped. | image flipped | image reversal | Viewer
|  | Projection errors may lead to errors in calculation of 3D clinical parameters when adjusting the position, size and shape of the 3D objects used. | projection error | image display error
|  | Projection errors may lead to errors in calculation of 3D clinical parameters when adjusting the position, size and shape of the 3D objects used. | incorrect parameter | use of incorrect control settings
|  | Incorrect clinical parameters can be a contributing factor in incorrect clinical decisions, or sub-optimal surgical interventions. | incorrect clinical decision | poor quality image
| FDA | A customer experienced when using IMPAX CV Reporting software, specifically, when building a NIV report, the NIV Cardio report was showing incorrect findings for Aneurysm. | false output | computer software issue | Viewer
| FDA | A database handling error could occur during the image acquisition process affecting the completeness of acquired images with Centricity PACS-IW. | incomplete image | image display error | Viewer
|  | There is a potential that one or more image series (i.e. all images within an image set) may be missing from an exam without indication to the user. While this is rare, this can occur with imaging studies that consist of a very small number of images per series. E.g. CR Thorax exam with 1 image per series. | missing image series | image display error
| FDA | A defect component of the Compano reader power supply can cause the power supply to burn and some smoke to leave the housing. | power supply burned | Power Source Issue | Viewer
| FDA | A dentist found implant positions are not correctly exported from the implant planning software SICAT IMPLANT V2.0 - for the specific export format CMG.DXD. | incorrect position calculation | computer software issue | Viewer
| FDA | A migrated study that has annotations will display the annotations, but when a prior study is loaded that should have annotations, the second study will not display annotations. | fail to display annotations | no display or display failure | Viewer
| FDA | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | incorrect image measurement | incorrect measurement | Viewer
|  | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | fail to display calibration | image display error
|  | A product issues happens for RF projection images. If user measures on RF image, the .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | fail to display label | image display error
|  | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. | incorrect image measurement | incorrect measurement
|  | A product issues happens for RF projection images. The .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | fail to display calibration | image display error
| FDA | A report can be confirmed with the incorrect patient demographics resulting in it being saved under the incorrect patient record in the Electronic Medical Record (EMR). | incorrect patient record | patient data issue | Viewer
| FDA | A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | multiple accesses at once time | computer software issue | Viewer
|  | A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | missing read-only notification | device alarm system issue
| FDA | A software defect negatively affects the quality of images taken during diagnostic imaging (mammograms) which may result in misdiagnosis. | poor quality image | therapeutic or diagnostic output failure | Viewer
| FDA | A software defect that has the potential to result in image misdiagnosis and incorrect treatment of a patient. | software defect | therapeutic or diagnostic output failure | Viewer
| FDA | A software defect was discovered that causes images to be out of context with clinical information. | missing clinical information | image display error | Viewer
| FDA | A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. | washed-out grey image | poor quality image | Viewer
| FDA | A software functionality in the report sections Findings Information and Summary of Measured Findings is not performing properly. The system will use the original values instead of displaying and saving corrected/modified values. Outdated information could be sent to the referring physician when creating the formal report in the reporting system. | improper report information | application program issue | Viewer
| FDA | A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. | inconsistent diagnosis | application program issue | Viewer
| FDA | Adding an additional comparison study to the Clinical Side Bar in the Enterprise Imaging Diagnostics Desktop, can result in a change of the sort order of the studies in the Clinical Side Bar. The thumbnail images of the studies are no longer chronologically sorted with the most recent study at the top. | inappropriate sort order | application program issue | Viewer
| FDA | Affected software product versions were found during internal audit to have been marketed for sale without sufficient documentation pertaining to certain required quality procedures. | insufficient documentation | device markings issue | Viewer
| FDA | After a period of time running Fusion Workstation, the Hounsfield measurement tool will report incorrect values. | incorrect measurement value | incorrect measurement | Viewer
| FDA | An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. | data calculated incorrectly | incorrect software programming calculations | Viewer
|  | An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. | decreased sensitivity | decreased sensitivity
| FDA | An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | inappropriate labeling orientation | device markings issue | Viewer
|  | An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | image flipped | image reversal
| FDA | An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | incorrect grid selected | use of device issue | Viewer
|  | An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | incorrect biopsy | unexpected therapeutic results
| FDA | Any change made to current ultrasound machine software or hardware upgrades, adding or modifying custom fields, or adding a new Ultrasound Machine, must be tested prior to the use in production environment to ensure that ultrasound measurements are crossing over correctly to the AS system | lack of release test | false device output | Viewer
| FDA | Cardio study list does not show STAT studies without refreshing. | inappropriate list displayed | incorrect display | Viewer
| FDA | Carestream Health received a complaint related to CARESTREAM Image Suite 4 from a foreign hospital stating that the annotation on the overlay is displayed as "L (Left)", when it should be "R(Right)". | incorrect image annotation | image display error | Viewer
| FDA | Centricity Universal Viewer measurements saved into a DICOM Grayscale Presentation State are incorrect in subsequent views for exams containing series with different pixel sizes and may lead to a potential misdiagnosis. | inappropriate pixel size | image resolution poor | Viewer
|  | Centricity Universal Viewer measurements saved into a DICOM Grayscale Presentation State are incorrect in subsequent views for exams containing series with different pixel sizes and may lead to a potential misdiagnosis. | inappropriate pixel size | therapeutic or diagnostic output failure
| FDA | Customers may experience an issue with the software study preferences when changes are made to the study protocol, resulting in incorrect patient follow-up. | inappropriate software release | application program version or upgrade problem | Viewer
| FDA | Cut lines on the image may present horizontally rather than vertically. | incorrect image cut line | image orientation incorrect | Viewer
| FDA | Digital templates were created with the incorrect files. | incorrect file applied | device markings issue | Viewer
| FDA | ON 07/14/2016, MERGE WAS NOTIFIED THAT THE STATION CRASHED WHILE DICTATING A REPORT, SUBSEQUENTLY, THE REPORT WAS MISSING. THE CUSTOMER REPORTED THAT THE SYSTEM STATED THERE WAS AN EXISTING REPORT WHEN THE USER TRIED TO RE-DICTATE THE REPORT, HOWEVER NO EXISTING REPORT COULD BE FOUND. AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT THE MOST PROBABLE CAUSE OF MISSING REPORT WAS DUE TO THE REPORT FILE NOT UPLOADING TO THE SERVER. | missing report | device markings issue | Viewer
| FDA | ON (B)(6) 2016 MERGE HEALTHCARE'S CUSTOMER, GE HEALTHCARE, REPORTED THAT ONE OF THEIR INSTALL SITES (B)(4) CONTACTED GE HEATHCARE TO NOTIFY THEM THAT NO LICENSE IS AVAILABLE WHEN LAUNCHING (B)(4). AS A RESULT OF THIS MALFUNCTION, USERS AT (B)(4) WERE NOT ABLE TO READ MAMMOGRAPHY IMAGES UNTIL THE MORNING OF (B)(6) 2016. THIS MALFUNCTION CAN RESULT IN DELAY IN DIAGNOSIS OR TREATMENT. | unable to read image | application program issue | Viewer
|  | ACCORDINGLY THE RESPONSE ON 5/19/2016 "...SITES HAVE BEEN UPGRADED FROM (B)(4) VERSION 3.7.3.7 TO (B)(4) VERSION 3.7.3.9SP3. THIS HAS ENABLED THE CUSTOMERS TO BYPASS THE ERROR, BECAUSE THE LICENSE ARCHITECTURE USING (B)(4) VERSION 3.7.3.9SP3 HAS A NEW DESIGN THAT USES WORKSTATION-BASED LICENSES. THE OLD SYSTEM USED SERVER-BASED LICENSES. (B)(4)'S SERVER UPGRADE WAS STARTED FRIDAY EVENING APRIL 22, AND BY MONDAY MORNING APRIL 25, 2016, ALL OF THEIR MAMMO STATIONS WERE ABLE TO VIEW MAMMO IMAGES ONCE AGAIN." THE PRODUCT IS FUNCTIONING AS EXPECTED AS OF 4/25/2016 AM ACCORDING TO GE HEALTHCARE. (B)(4). | incompatible license | application program issue
| FDA | ON (B)(6) 2016 MERGE HEALTHCARE'S CUSTOMER, GE HEALTHCARE, REPORTED THAT ONE OF THEIR INSTALL SITES (B)(4) CALLED GE HEALTHCARE TO NOTIFY THAT THEY WERE GETTING LICENSE ERRORS WHEN OPENING STUDIES IN (B)(4). AS A RESULT OF THIS MALFUNCTION, USERS AT (B)(6) WERE NOT ABLE TO READ MAMMOGRAPHY IMAGES UNTIL THE EVENING OF (B)(6) 2016. THIS MALFUNCTION CAN RESULT IN DELAY IN DIAGNOSIS OR TREATMENT OF THE PATIENT. (B)(4). | license error | application program issue | Viewer
|  | Event Description: (B)(4) IS MEDICAL IMAGE REVIEW WORKSTATION SOFTWARE. THE PRODUCT CONSISTS OF FEATURES THAT ALLOW QUALIFIED MEDICAL PROFESSIONALS TO VIEW PATIENT MEDICAL IMAGES WITH THE DESIRED VIEWING PROTOCOL AND WORKFLOW IN ORDER TO OPTIMIZE THE EFFICIENT USE OF THEIR TIME. | unable to read image | application program issue
| FDA | ON (B)(6)2017, MERGE HEALTHCARE TECHNICAL SUPPORT WAS NOTIFIED BY A (B)(4) CUSTOMER THAT PROCEDURES WERE NOT ABLE TO BE PERFORMED CAUSING A DELAY IN TREATMENT FOR 2 PATIENTS. PER THE CUSTOMER SITE, EMERGENCY DEPARTMENT STUDIES ARE TYPICALLY READ WITHIN 30 MINUTES HOWEVER, THERE WERE TWO STUDIES THAT WERE NOT READ UNTIL THE NEXT MORNING. WITH MERGE (B)(4) NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. THERE WAS NO REPORTED HARM TO EITHER PATIENT A OR PATIENT B AS A RESULT OF THE DELAY. (B)(4). | unable to read image | therapeutic or diagnostic output failure | Viewer
| FDA | Event Description: (B)(4). MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. | fail to create 2d images | application program issue | Viewer
|  | Event Description: (B)(4). MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. | fail to create 3d images | application program issue
|  | ON (B)(6) 2017, A CUSTOMER REPORTED THAT WHEN TWO STUDIES WERE MERGED USING THE QC EDITOR, THE STUDY COMMENTS MADE ON THE SOURCE STUDY WERE NOT RETAINED AND MOVED TO THE TARGET STUDY. THE LOSS OF THE COMMENTS MAY LEAD TO A DELAY IN PATIENT CARE OR MISDIAGNOSIS. HOWEVER, THE IMAGES ARE AVAILABLE FOR DICTATION AND THERE HAS BEEN NO REPORT OF HARM TO A PATIENT DUE TO THIS ISSUE. (B)(4). | lost of annotation | image display error
| FDA | Event Description: (B)(6) 2017 CUSTOMER REPORTED THAT THEY WERE UNABLE TO OPEN ALL IMAGES OF A CERTAIN STUDY IN THE SYNGO.PLAZA VIEWER. THE AFFECTED STUDY CONTAINED 240 IMAGES, HOWEVER, ONLY 52 WERE READABLE. ACCORDING TO THE SYNGO.PLAZA'S LOG FILES, THE IMAGE FILES WERE WRITTEN SUCCESSFULLY TO THE SHORT TERM STORAGE (STS) DISKS, HOWEVER THE STS'S FILESYSTEM CONTAINED ONLY EMPTY FILES (NO INFORMATION INCLUDING DICOM HEADER IS PRESENT). THE SYNGO.PLAZA SYSTEM RELIES ON WINDOWS CONFIRMATION TO DETERMINE WHETHER A FILE WAS SUCCESSFULLY STORED. SINCE NO ERRORS HAVE BEEN LOGGED IN THE SYNGO.PLAZA DURING THE STORAGE PROCESS TO THE STS AND THERE ARE KNOWN INTERMITTENT HARD DRIVE ISSUES IDENTIFIED IN WINDOWS EVENT LOGS, THEREFORE IT IS ASSUMED THAT THE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE CORRUPT STS FILES. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THIS INCIDENT OCCURRED IN (B)(6). | file corrupted | application program issue | Viewer
| FDA | Event Description: (B)(6) WAS ENROLLED ON PROTOCOL (B)(6) OF RESPONSE ADAPTED THERAPY BASED ON POSITRON EMISSION TOMOGRAPHY (PET) FOR BULKY STATE I AND STAGE II CLASSICAL HODGKIN LYMPHOMA ON (B)(6) 2015. THE PT PRESENTED TO CLINIC FEELING WELL ON (B)(6) 2015 FOR CONSIDERATION TO BEGIN CYCLE 3. THE LAB RESULTS DEMONSTRATED GRADE 4 NEUTROPENIA. (B)(6), 2015 ANEUT 0.33; (B)(6) 2015 WBC 1.71 (L), (B)(6) 2015 RBC 4.44, (B)(6) 2015 HGB 12.1, (B)(6) 2015 HGB 12.1, (B)(6) 2015 HCT 36.7, ON (B)(6) 2015 CYCLE 3, NEUTROPENIA REDUCED TO GRADE 1. THE PT IS FEELING WELL AND WILL PROCEED WITH PROTOCOL TREATMENT. (B)(6) 2015 ANEUT 1.51, (B)(6) 2015 WBC 5.23, (B)(6) 2015 RBC 4.18, (B)(6) 2015 HGB 11.5 (L), (B)(6) 2015 HCT 34.8 (L).ATTRIBUTE TO GR, 4 NEUTROPHIL COUNT DECREASED. (B)(6). PHASE II TRIAL OF RESPONSE-ADAPTED THERAPY BASED ON POSITRON EMISSION TOMOGRAPHY (PET) FOR BULKY STAGE 1 AND STAGE II CLASSICAL HODGKIN LYMPHOMA (HL). | inappropriate therapy | therapeutic or diagnostic output failure | Viewer
| FDA | Event Description: A CUSTOMER REPORTED A PATIENT DATA ERROR WHILE USING THE TOMTEC-ARENA. THE SOFTWARE ALLOWED PATIENT DATA ACQUIRED FROM ONE DATA SET TO BE INCLUDED IN THE NEXT PATIENT STUDY. THERE WAS NO MISDIAGNOSIS OR INJURY ASSOCIATED WITH THIS EVENT. | incorrect patient information | patient data issue | Viewer
| FDA | Event Description: A CUSTOMER REPORTED AN EMERGENCY ROOM (ER) PATIENT WAS IMAGED USING A GENERIC NAME PRIOR TO SCHEDULING THE STUDY WITH THE PT'S ACTUAL NAME RESULTING IN TWO PATIENTS IN HMI. THE TWO STUDIES WERE MERGED IN HMI USING THE WA MANAGER TOOL. IN THIS INSTANCE, MERGING OF THE SCHEDULED STUDY AND PERFORMED STUDY WAS INCOMPLETE WITH THE SCHEDULED STUDY NOT BEING REMOVED FROM THE SCHEDULED STUDIES LIST. SUBSEQUENTLY, THE ER PT WAS INADVERTENTLY SUBJECTED TO A SECOND CHEST X-RAY IMAGING PROCEDURE. THE DATE OF THIS ADVERSE EVENT AND PT INFO WERE NOT PROVIDED BY THE REPORTING FACILITY. | fail to merge multiple images | application program issue | Viewer | SECOND CHEST X-RAY IMAGING PROCEDURE
| FDA | Event Description: A CUSTOMER REPORTED MEASUREMENT DATA CAN APPEAR FROM A PREVIOUS STUDY ONTO A NEW STUDY IF AN ERROR MESSAGE APPEARS DURING LAUNCH WHILE USING THE TOMTEC-ARENA TTA 2. THIS EVENT OCCURRED DURING A RESEARCH STUDY. THERE WAS NO INJURY OR CLINICAL USE ASSOCIATED WITH THIS EVENT. | inappropriate measurement data displayed | incorrect display | Viewer
| FDA | Event Description: A CUSTOMER REPORTED THAT WHEN WORKING WITH MULTIPLE STUDIES AND USING THUMBNAIL PANELS, THE THUMBNAIL TAB CHANGED UNEXPECTEDLY AND THE RADIOLOGIST ALMOST DICTATED A REPORT ON THE BASIS OF LOOKING AT THE WRONG STUDY. THE PROBLEM WAS DISCOVERED AND NO MISDIAGNOSIS OCCURRED. | tab changed unexpectedly | application program issue | Viewer
| FDA | Event Description: A CUSTOMER USED A SURGIGUIDE TO CREATE THE OSTEOTOMY FOR TWO KEYSTONE DENTAL IMPLANTS ON TOOTH POSITION #7 AND #10. THE CENTRAL AXIS OF THE OSTEOTOMY FOR BOTH IMPLANTS DID DEVIATE 12 DEGREES TO THE BUCCAL SIDE FROM THE ORIGINAL PLANNING. THE INCORRECT IMPLANT PLACEMENT LEADS TO THE FACT THAT THE APICAL PART OF BOTH IMPLANTS ARE PERFORATING THE BUCCAL PLATE AND NEED GRAFTING OR THE IMPLANTS NEED TO BE REMOVED AND REPLACED DURING A SECOND CORRECTIVE SURGERY. | incorrect implant placement | positioning issue | Viewer | second corrective surgery
| FDA | Event Description: A DATA LOSS IN THE SYNGO IMAGING APPLICATION OCCURRED FOR 62 OUT OF 601 IMAGES IN ONE SERIES ON (B)(6) 2017. NO DATA MIX-UP OCCURRED; NO PATIENT RESCAN WAS REQUIRED. DUE TO CHANGES ON THE POWER SUPPLY OF THE SYNGO IMAGING ENVIRONMENT PERFORMED BY THE CUSTOMER (B)(6) 2017, AN IMPROPER SHUTDOWN OF THE STORAGE TERM RAID AND THE SYNGO IMAGING APPLICATION OCCURRED. THIS RESULTED IN THE AFFECTED IMAGES BEING CREATED WITH 0 KB FILE SIZE IN THE SYNGO IMAGING'S SHORT TERM STORAGE PRIOR TO THEM BEING ARCHIVED IN THE LONG TERM ARCHIVE. THE AFFECTED SERIES WERE AVAILABLE IN THE SYNGO IMAGING VIEWER, HOWEVER, ONLY 539 IMAGES CAN BE DISPLAYED AND 62 IMAGES WERE CORRUPT. THESE CORRUPT IMAGES WERE FOUND DUE TO MAINTENANCE TASKS ON THE SYSTEM (B)(6) 2017. THE REPORTED INCIDENT OCCURRED IN (B)(6). | image with 0 kb size | data issue | Viewer
| FDA | Event Description: A DATA LOSS OF APPROXIMATELY 200 CLINICAL IMAGE STUDIES WHILE ARCHIVING WAS REPORTED FROM A SITE. SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE AND FOUND THAT USER ACTIONS CAUSED THE DATA LOSS AND THAT THE DATA LOSS WAS UNRECOVERABLE. THE HISTORICAL IMAGES ARE NO LONGER AVAILABLE FOR COMPARISON WITH NEWER IMAGES HOWEVER THE CLINICAL REPORTS GENERATED FROM THE READ OF THE STUDIES PRIOR TO ARCHIVAL ARE STILL PRESENT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | image loss | loss of data | Viewer
| FDA | Event Description: A HAZARD MAY ARISE WHEN NON-DICOM COMPLIANT JPEG-2000-COMPRESSED IMAGES ARE RECEIVED BY THE (B)(6) COMPONENT WITH DICOM BIT DEPTH (BITS STORED TAG) THAT DOES NOT MATCH THE BIT DEPTH IN THE PIXEL DATA CODESTREAM. THESE IMAGES ARE DECOMPRESSED WITH A LOSS OF PRECISION IN THE IMAGE DATA WHICH CAN LEAD TO FAILING IMAGE REGISTRATION. | non-dicom format image | data issue | Viewer
|  | Event Description: A HAZARD MAY ARISE WHEN NON-DICOM COMPLIANT JPEG-2000-COMPRESSED IMAGES ARE RECEIVED BY THE (B)(6) COMPONENT WITH DICOM BIT DEPTH (BITS STORED TAG) THAT DOES NOT MATCH THE BIT DEPTH IN THE PIXEL DATA CODESTREAM. THESE IMAGES ARE DECOMPRESSED WITH A LOSS OF PRECISION IN THE IMAGE DATA WHICH CAN LEAD TO FAILING IMAGE REGISTRATION. | loss of precision | data issue
| FDA | Event Description: A LITHIUM ION BATTERY USED FOR OUR PORTABLE COMPUTERS SPARKED, OVERHEATED AND MELTED. AN ENTIRE FLOOR OF OUR HOSPITAL HAD TO BE EVACUATED IN ORDER TO PROTECT OUR PATIENTS FROM SMOKE INHALATION. | battery melted | battery issue | Viewer
| FDA | Event Description: A MERGE CARDIO CUSTOMER REPORTED THAT THEIR STUDY LIST DOES NOT SHOW NEW STUDIES UNLESS MANUALLY REFRESHED. THE CLIENTS ARE GETTING NOTIFICATIONS REGARDING REPORT AND STUDY STATUS UPDATES BUT IF THEY HAVE A FILTER SET FOR ALL QC'D REPORTS FOR TODAY AND DON'T MANUALLY REFRESH THE STUDY LIST BETWEEN REVIEWING STUDIES, THEY DON'T SEE STAT STUDIES UNLESS THE STUDY LIST IS MANUALLY REFRESHED. (B)(4). | inappropriate list displayed | incorrect display | Viewer
| FDA | Event Description: A MERGE ORTHOCASE CUSTOMER REPORTED THAT THE DISPLAYED VALUE OF THE MEASUREMENT IS CHANGING AFTER SAVING A PLAN WITH A MEASUREMENT RESULT. | meassurement changed | application program issue | Viewer
|  | Event Description: WHEN USING THE PERPENDICULAR LINE TOOL, THERE IS AN OPTION TO "SNAP" THE LINES TOGETHER, WHICH PREVENTS THE PERPENDICULAR LINES FROM EXTENDING PAST THE STRAIGHT/BASE LINE. THE SYSTEM DEFAULT IS TO SNAP THESE LINES. TO REMOVE THE SNAPPING FUNCTION FROM THE CURRENT ANNOTATION, SIMPLY RIGHT CLICK NEAR THE LINE AND DESELECT "SNAP TO." WHILE THE SYSTEM DEFAULT IS SET TO SNAP THE LINES TOGETHER, THE ACTUAL "SNAP" FUNCTION MAY NOT OCCUR UNTIL THE PLAN IS SAVED, AT WHICH TIME THE LINE END POINT WITHIN 20 PIXELS OF THE BASE LINE WILL ADJUST ITS LOCATION TO SNAP PERFECTLY TO THE BASE LINE. THIS MAY ADJUST THE LENGTH AND CORRESPONDING MEASUREMENT OF THE PERPENDICULAR LINE WHEN THE PLAN IS REOPENED. (B)(4). | snap function unexpected | application program issue
| FDA | Event Description: A NEUROSURGEON INITIALLY MISINTERPRETED ANCHOR IMAGES DISPLAYED IN THE VIEWPORT OF THE CHANGE HEALTHCARE ENTERPRISE VIEWER (CHEV) AS A PRIOR STUDY. IT WAS REPORTED THIS CAUSED THE NEUROSURGEON TO ALMOST USE THE WRONG IMAGE FOR SURGERY. NO ADVERSE CONSEQUENCES WERE REPORTED. | image misinterpreted | use of device issue | Viewer
| FDA | Event Description: A POTENTIAL LOSS OF CLINICALLY-RELEVANT DATA WHICH MAY BE UNRECOVERABLE WAS REPORTED. THE INVESTIGATION INTO THE ROOT CAUSE AND RECOVERABILITY OF THE DATA IS ONGOING. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES ARE STILL AVAILABLE IN THE SITES HIS. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. | -fail to identify root cause -data breached -fail to record operative data | -compliance issue -application security issue -data issue | Viewer
| FDA | Event Description: A RADIOLOGIST WAS REVIEWING A TUMOR CASE IN RADSUITE AND NOTICED THAT THE PIXEL VALUE TOOL WAS DISPLAYING DIFFERENT VALUES THEN THE PET MODALITY'S SOFTWARE. TO AID IN THE INTERPRETATION OF PET IMAGES, THE PIXEL VALUE TOOL WITHIN THE RADSUITE SOFTWARE CAN BE USED TO VIEW THE ABSORPTION OF AN INJECTED ISOTOPE IN BODY TISSUE BASED ON STANDARDIZED UPTAKE VALUES (SUV). THE PATIENT WEIGHT VALUE IS USED FOR CALCULATING THE SUV. IF THE PATIENT WEIGHT IS MISSING, THE VALUE IS NULL. IF A STUDY WITH A PATIENT WEIGHT VALUE OF NULL IS VIEWED AFTER A STUDY FOR A DIFFERENT PATIENT WITH A PATIENT WEIGHT VALUE GREATER THAN 0, THE WEIGHT FROM THE PREVIOUS STUDY WILL BE USED TO CALCULATE THE SUVS FOR THE STUDY WITH THE NULL WEIGHT VALUE. THIS MAY RESULT IN INCORRECT SUVS. IF OTHER CLINICAL FACTORS ARE NOT AVAILABLE OR TAKEN INTO DUE CONSIDERATION, INCORRECT SUVS COULD POTENTIALLY LEAD TO OVERSTATEMENT OR UNDERSTATEMENT OF THE EFFECTIVENESS OF A CANCER TREATMENT, OR TO A MISDIAGNOSIS OF A CANCER. | -incorrect display value -incorrect message display -incorrect display format | -electrical issue -particulates -Contamination During Use | Viewer
| FDA | Event Description: A RADSUITE CUSTOMER REPORTED THAT WHEN THE RADIOLOGIST IS SCROLLING THROUGH THE IMAGES IN THE VIEWER THERE HAVE BEEN DELAYS OF BETWEEN 5 AND 10 MINUTES. USER ERROR IN SETTING UP A CUSTOM MACRO AND LACK OF VERIFICATION OF THAT MACRO RESULTED IN THIS ISSUE. THIS MACRO CONTAINS A PERSISTABLE THAT NEEDS TO QUERY THE DATABASE EVERYTIME IT IS CALLED ON. WHEN THE USER IS SCROLLING THRU IMAGES THE MACRO IS MAKING A DATABASE CALL AND NEEDS TO HAVE THE INFORMATION FROM THE DATABASE BEFORE IT CAN GO ANY FURTHER THUS CAUSING THE VIEWPORT TO "FREEZE" AS IT IS WAITING ON THE INFORMATION FROM THE DATABASE. MOST OF THE TIME IT IS NOT A PROBLEM BUT THERE ARE TIMES WHERE THERE SEEMS TO BE CONTENTION IN THE DATABASE OR A NETWORK OR STORAGE ISSUE THAT IS CAUSING A DELAY. (B)(4). | -delivery delayed -setup delayed -delayed data | -human-device interface issue -failure of device to self-test -filling problem | Viewer
| FDA | Event Description: A SERVICE CALL WAS PLACED FOR THE SMART CR UNIT (S/N (B)(4)) AT ABOUT 2AM CST ON (B)(6) 2016. CUSTOMER ALLEGES SMOKE COMING FROM THE UNIT. HOSPITAL STAFF UNPLUGGED THE UNIT AND AWAITED SERVICE. SERVICE ENGINEER ARRIVED ON-SITE LATER THAT DAY AND DETERMINED THE POWER SUPPLY TO BE THE ISSUE. A NEW POWER SUPPLY WAS ORDERED (B)(4) AND REPLACED THE FOLLOWING DAY, (B)(6) 2016. THE UNIT WAS TURNED OVER TO THE HOSPITAL IN GOOD WORKING ORDER. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | -part damaged -exposure to gas -device burned | -Detachment of Device Component -fluid leak -detachment of device component | Viewer
| FDA | Event Description: A SIMPLE REQUEST FOR SERVICE SUPPORT TO RECOVER DATA AFTER A HARDWARE FAILURE IN THE SITE'S LONG TERM ARCHIVE SYSTEM WAS ESCALATED TO SIEMENS. HOWEVER. IT WAS IDENTIFIED THAT AN UNKNOWN AMOUNT OF IMAGE STUDIES WERE MISSING IN THE SITES LONG TERM ARCHIVE. THE HISTORICAL STUDY IMAGES ARE NO LONGER . HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. | -hardware failure -fail to record intraoperative data -lack of data validation | -device maintenance issue -compliance issue -no device output | Viewer
| FDA | Event Description: A SINGLE IMAGE IN A CT SERIES WAS FLIPPED AND SAVED. WHEN THE RADIOLOGIST MADE HIS FINAL INTERPRETATION, THERE WAS NOT A SIGNIFICANT VISUAL INDICATION THAT THE IMAGE HAD BEEN MANIPULATED. | -loss of indication -unclear image -loss of visual | -Device Handling Issue -no display or display failure -failure of device to self-test | Viewer
| FDA | Event Description: A SITE USING MERGE PACS REPORTED THAT A MIGRATED STUDY THAT HAS ANNOTATIONS WILL DISPLAY THAT STUDY'S ANNOTATIONS, BUT WHEN A PRIOR IS LOADED THAT SHOULD HAVE ANNOTATIONS THEN THE SECOND STUDY WILL NOT DISPLAY ANNOTATIONS. SITE EXPRESSED CONCERNS THAT SEEING NO ANNOTATIONS COULD POSSIBLY INDICATE TO THE PROVIDER THAT NOTHING SIGNIFICANT WAS NOTED WHEN THE IMAGE WAS READ INITIALLY, AND THEREFORE COULD POTENTIALLY IMPACT THE TREATMENT PLAN OR COULD POTENTIALLY RESULT IN A FAILURE TO TREAT ACCORDINGLY. | -inappropriate annotation tool -missing content description -system migrated | -device maintenance issue -component missing -no display or display failure | Viewer
| FDA | Event Description: A SOFTWARE CODING ERROR RESULTED IN AN INCORRECT REGISTRY SETTING WHERE THE UNITS THAT ARE DISPLAYED FOR ORGAN DOSE ARE MISMATCHED WITH THE NUMERICAL VALUE DISPLAYED. THE PROBLEM WAS NOT DISCOVERED FOR A LONG TIME, LIKELY BECAUSE CUSTOMERS HAD NOT CHANGED THE DEFAULT REGISTRY SETTING FOR RADIATION DOSAGE (WHICH IS CORRECT) TO DISPLAY ORGAN DOSE UNTIL NOW. IF A CLINICIAN DID NOT NOTICE THAT THE DISPLAYED INDIVIDUAL IMAGE DOSAGE WAS WELL OUTSIDE OF THE NORMAL RANGE, THEY COULD THINK THAT THE PATIENT HAD RECEIVED LESS RADIATION THAN THEY ACTUALLY HAD AND SO ANY ADDITIONAL IMAGING STUDIES WOULD EXPOSE THE PATIENT TO MORE CUMULATIVE RADIATION THAN EXPECTED. THE INCORRECT DOSAGE VALUE IS NOT PRESERVED OR SAVED UNLESS A SCREEN CAPTURE IS GENERATED OR THE "COPY IMAGE" OR "SAVE IMAGE" FEATURE IS USED. FOR ORGANIZATIONS THAT CALCULATE A CUMULATIVE RADIATION DOSAGE VALUE FOR PATIENTS, THE CORRECT DOSAGE DATA FROM ALL TYPES OF SCANS (NOT JUST MAMMOGRAPHY) SHOULD BE IN A PACS OR OTHER CENTRAL LOCATION OR SYSTEM, NOT IN THE MAMMO APPLICATION. THE ACTUAL DOSE VALUES SAVED IN THE PACS IN THE DICOM IMAGE HEADERS ARE CORRECT IN CASE THEY ARE USED TO CALCULATE CUMULATIVE DOSE. | -incorrect display value -incorrect display format -incorrect message display | -device alarm system issue -material disintegration -medical gas supply problem | Viewer
| FDA | Event Description: A SPECIFIC WORKFLOW IN ONE CUSTOMER'S ENVIRONMENT IS LIKELY CONTRIBUTING TO THE FILLING UP OF THE RADSUITE CACHE. THE SITE TYPICALLY CONFIGURES LESS STORAGE FOR THE LOCAL EA (DSS PARTITIONS) AND MORE STORAGE FOR THE RADSUITE CACHE (UV_DSS PARTITIONS). WITH THEIR CONFIGURATION, IT IS LIKELY THAT MANY OF THEIR STUDIES WILL NOT QUALIFY FOR WATERMARKING TO REMOVE THEM FROM RADSUITE'S STORAGE AND FREE UP SPACE THERE. WHEN THE RADSUITE CACHE IS FULL, USERS WILL OBSERVE STORAGE FAILURE MESSAGES FROM THE MODALITY DEVICES. THOSE STUDIES COULD BE VIEWED ON THE MODALITIES OR STORED ON ANOTHER RADSUITE NODE FOR VIEWING THERE. IF NOT, THEN THERE IS A POTENTIAL THAT SOME STUDIES MIGHT NOT BE VIEWED, ALTHOUGH WE HAVE NOT RECEIVED ANY REPORTS OF THAT ACTUALLY OCCURRING. THE PROBLEM WAS FOUND AT ONE CUSTOMER SITE BY MERGE'S SYSTEM MONITORING PROCESS. | -fail to configure infusion parameters -fail to setup configuration -insufficient storage space | -device maintenance issue -data back-up problem -medication error | Viewer
| FDA | Event Description: A SRS TREATMENT FOR 2 METASTASES IN THE LEFT PARIETAL AND LEFT OCCIPITAL LOBE IN THE BRAIN, PERFORMED ON (B)(6) 2020, HAD BEEN PLANNED WITH THE AID OF BRAINLAB IMAGE FUSION 4.0 (AND BRAINLAB MULTIPLE BRAIN METS SRS 2.0). DURING PLANNING/TREATMENT THE USER: PERFORMED A RIGID FUSION OF A MR AND CT IMAGE SET (TO COMBINE THE INFORMATION CONTAINED WITHIN THE INDIVIDUAL IMAGE SETS). PROCEEDED WITH CALCULATION OF AN ELASTIC FUSION FOR THIS MR/CT FUSION PAIR (A NEW MR IMAGE SET IS GENERATED USING A DEFORMABLE REGISTRATION ALGORITHM, BY BRAINLAB CRANIAL DISTORTION CORRECTION). REVIEWED AND ACCEPTED THE RESULTING ELASTICALLY DEFORMED / DISTORTION CORRECTED MR (INCL. DISTORTION CORRECTED OBJECTS). CREATED A DOSE PLAN BASED ON THE DISTORTION CORRECTED MR AND CT IMAGE SET IRRADIATED THE METASTASES ACCORDING TO THE DOSE PLAN. SUSPECTED ON (B)(6) 2020, FOR ONE METASTASIS, A DEVIATION OF ACTUAL METASTASIS LOCATION/SIZE VS. IRRADIATED TARGET VOLUME, BY COMPARING AFOREMENTIONED IMAGE SETS AND A POST-TREATMENT MR VIA RIGID FUSION. ACCORDING TO THE HOSPITAL: THE TREATMENT WAS A SRS TREATMENT (1 FRACTION), 2 BRAIN METASTASES, 20-24 GY. ONE METASTASIS (LEFT PARIETAL) WAS NOT IRRADIATED AS INTENDED, THE SECOND METASTASIS (LEFT OCCIPITAL) TREATED WITH THIS SAME PLAN SEEMS TO BE RESOLVED. FOR THE AFFECTED METASTASIS, A LARGER VOLUME WAS TREATED THAN WHAT CORRESPONDED TO THE REAL METASTASIS SIZE (THE METASTASIS RECEIVED 50%-71% OF THE INTENDED DOSE DISTRIBUTION, AND APPROX. 0.35 CC OF HEALTHY BRAIN TISSUE RECEIVED UNINTENDED RADIATION DOSE). THE INCORRECT DOSAGE EXCEEDED CLINICALLY ACCEPTABLE LIMITS, AND THERE WAS A NEGATIVE CLINICAL EFFECT TO THE PATIENT (THE AFFECTED METASTASIS CONTINUED TO GROW). IF FURTHER REMEDIAL ACTIONS ARE NEEDED IS CURRENTLY UNDECIDED. 8 MORE METASTASES WERE TREATED ON (B)(6) 2020 AND THE PHYSICIAN WILL MONITOR THE PATIENT AND THE AFFECTED METASTASIS). | -fail to toggle image visibility -inappropriate therapy -loss of therapy | -application program issue -device difficult to setup or prepare -inadequate user interface | Viewer
| FDA | Event Description: A USER REPORTED THAT APPROXIMATELY 1000 STUDIES WERE LOST FROM THE PACS STORAGE AND A REFERRING PHYSICIAN WAS UNABLE TO VIEW IMAGES. RADIOLOGIST REPORTS FOR THESE DELETED STUDIES REMAINED AVAILABLE. PRELIMINARY ANALYSIS DETERMINED THE PROBLEM IS DUE TO A CONFIGURATION KEY THAT WAS ACTIVATED IN COMBINATION WITH A SOFTWARE BUG. WHEN THE FEATURE FORCEGLOBALSTUDYDELETION IS ENABLED, ANY SERIES DELETIONS THAT OCCUR ON A SERVER ARE AUTOMATICALLY ELEVATED TO A SYSTEM-LEVEL STUDY DELETION. THIS SYSTEM-LEVEL DELETION BYPASSES EXISTING MITIGATION POLICIES (I.E., ARCHIVE POLICIES) ON OTHER SERVERS. | -fail to interpret images -insufficient storage space -fail to view | -Device Handling Issue -component missing -loose or intermittent connection | Viewer
| FDA | Event Description: ACCORDING TO THE LITERATURE STUDY. A RETROSPECTIVE ANALYSIS OF A SINGLE CENTER EXPERIENCE IN ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY MICROWAVE ABLATION IN HYBRID OPERATING ROOM. A TOTAL OF 30 LUNG NODULES FROM 25 PATIENTS WERE TREATED. PNEUMOTHORAX OCCURRED IN 2 CASES (6.67%) AND BOTH REQUIRED CHEST DRAIN INSERTION. | -fail to interpret images -fail to visualize spatial relationship -fail to calculate tidal volume | -use of device issue -improper or incorrect procedure or method -inadequate user interface | Viewer
| FDA | Event Description: AFTER A PERIOD OF TIME RUNNING THE FUSION WORKSTATION (FW) 3.0.2 P3 AND LATER, THE HOUNSFIELD MEASUREMENT TOOL WILL REPORT INCORRECT VALUES. IF THE EXAM IS CLOSED, AND USER EXITS FW AND LOGS BACK IN TO REOPEN THE EXAM, THE MEASUREMENT WILL THEN BE REPORTED CORRECTLY. THERE APPEARS TO BE NO CONSISTENCY TO THE METHOD TO INVOKE THE ERROR. THE FACILITY REPORTED THE MEASUREMENTS TO BE CLEARLY AND OBVIOUSLY INCORRECT UPON VIEWING THE RESULTS, AND IT HAPPENS ABOUT 1 OUT OF EVERY 500 STUDIES. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS OR MISDIAGNOSES OF A PATIENT, HOWEVER THIS WAS DEEMED A POTENTIAL SAFETY ISSUE. | -incorrect measurement -incorrect value -incorrect display value | -use of device issue -inadequate user interface -failure to read input signal | Viewer
| FDA | Event Description: AFTER A PERIOD OF TIME RUNNING THE FUSION WORKSTATION (FW) 3.0.2 P3 AND LATER, THE HOUNSFIELD MEASUREMENT TOOL WILL REPORT INCORRECT VALUES. IF THE EXAM IS CLOSED, AND USER EXITS FW AND LOGS BACK IN TO REOPEN THE EXAM, THE MEASUREMENT WILL THEN BE REPORTED CORRECTLY. THERE APPEARS TO BE NO CONSISTENCY TO THE METHOD TO INVOKE THE ERROR. THE FACILITY REPORTED THE MEASUREMENTS TO BE CLEARLY INCORRECT UPON VIEWING THE RESULTS, AND IT HAPPENS ABOUT 1 OUT OF EVERY 500 STUDIES. CUSTOMER NEEDS TO BE USING FW WITH 3RD PARTY SOFTWARES (EPIC AND POWERSCRIBE) TO OBSERVE THIS ISSUE. THE USER NEEDS TO LINK THE STUDIES IN EPIC. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS OR MISDIAGNOSES OF A PATIENT, HOWEVER THIS WAS DEEMED A POTENTIAL SAFETY ISSUE. | -incorrect measurement -incorrect value -incorrect display value | -use of device issue -inadequate user interface -failure to read input signal | Viewer
| FDA | Event Description: AFTER PERFORMING A SUCCESSFUL 5DSA DV 3D DURING A CEREBRAL ANGIOGRAM, THE IMAGES WERE SENT TO THE SYNGO WORKSTATION OF IPP 8 FOR INTERPRETATION, SOFTWARE VERSION (B)(4). THE 3D WAS UNABLE TO SUCCESSFULLY RECONSTRUCT. ERROR CODE "THE SETTINGS OF THE USED ACQUISITION PROGRAM ARE NOT SUITABLE FOR THE REQUESTED 3D RECONSTRUCTION." A DEFAULT 3D RECONSTRUCTION IS PERFORMED. SEVERAL ATTEMPTS WERE MADE TO SALVAGE THE DATA FOR USE BUT WERE UNSUCCESSFUL. A SECOND RECONSTRUCTION WAS PERFORMED UNDER THE 3D 5DSA HEAD AND THE IMAGES WERE SUCCESSFULLY RECONSTRUCTED. | -fail to interpret images -loss of image -distorted images | -improper or incorrect procedure or method -Device Handling Issue -incorrect or inadequate result | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 12 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: A NEW FINDING WAS ADDED TO COMMENT ON THE LEFT GREATER SAPHENOUS VEIN, BUT THE SENTENCE PRODUCED READS "RIGHT GREATER SAPHENOUS VEIN (DIST CALF): VENOUS INCOMPETENCE NOTED." POTENTIAL IMPACT: WHEN THE USER GOES TO REPORT ON THE LEFT GREATER SAPHENOUS VEIN (DIST CALF), THE SENTENCE WILL REFER TO THE RIGHT SAPHENOUS VEIN. THIS COULD POTENTIALLY LEAD TO MISDIAGNOSIS AND INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE # Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | -incorrect device orientation -device malfunction -occlusion notified | -Device displays error message -break -fitting problem | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 12 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: A NEW FINDING WAS ADDED TO COMMENT ON THE LEFT GREATER SAPHENOUS VEIN, BUT THE SENTENCE PRODUCED READS "RIGHT GREATER SAPHENOUS VEIN (DIST CALF): VENOUS INCOMPETENCE NOTED." POTENTIAL IMPACT: WHEN THE USER GOES TO REPORT ON THE LEFT GREATER SAPHENOUS VEIN (DIST CALF), THE SENTENCE WILL REFER TO THE RIGHT SAPHENOUS VEIN. THIS COULD POTENTIALLY LEAD TO MISDIAGNOSIS AND INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE # Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | -incorrect device orientation -device malfunction -occlusion notified | -Device displays error message -break -fitting problem | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 2 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, ONLY TWO REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: LABEL WAS CHANGED FROM TRUNCAL V REGURG P 1/2-TIME TO PULMONARY INSUFFICIENCY PRESSURE HALF TIME (PI PHT). POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED THOSE REPORTS PRIOR TO THE CHANGE COULD DISPLAY INCORRECT CLINICAL DATA IF VIEWED IN RM OR IF USED AS A TEMPLATE FOR A REPEAT STUDY ON THAT PATIENT - IN THIS EXAMPLE THE TWO DISEASE STATES ARE VERY DIFFERENT, LEADING TO POTENTIAL INCORRECT DIAGNOSIS. DESCRIPTION OF MEASUREMENT DOES NOT CORRELATE WHEN USING HOVER TOOL WHICH CAN CAUSE CONFUSION. TRUNCUS ARTERIOSUS IS A RARE HEART DEFECT THAT'S PRESENT AT BIRTH (CONGENITAL). IN TRUNCUS ARTERIOSUS, IT MEANS THAT ONE LARGE BLOOD VESSEL LEADS OUT OF THE HEART. NORMALLY, THERE ARE TWO SEPARATE VESSELS COMING OUT OF THE HEART. AFTER A NEWBORN IS TREATED FOR TRUNCUS ARTERIOSUS, HE/SHE CAN DEVELOP TRUNCAL VALVE STENOSIS AND OR TRUNCAL VALVE REGURGITATION. PULMONARY VALVE INSUFFICIENCY (OR INCOMPETENCE, OR REGURGITATION) IS A CONDITION WHERE THE PULMONARY VALVE IS NOT STRONG ENOUGH TO PREVENT BACKFLOW TO THE RIGHT VENTRICLE. IF IT IS SECONDARY TO PULMONARY HYPERTENSION IT IS REFERRED TO AS A GRAHAM STEELL MURMUR. THE THREE PRIMARY PATHOLOGICAL MECHANISMS CAUSING PULMONARY VALVE INSUFFICIENCY ARE DILATATION OF THE PULMONIC VALVE RING, ACQUIRED ALTERATION OF PULMONIC VALVE LEAFLET MORPHOLOGY, OR CONGENITAL ABSENCE OR MALFORMATION OF THE VALVE. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE # Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | -incorrect device orientation -device malfunction -occlusion notified | -Device displays error message -break -fitting problem | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 2 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, ONLY TWO REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: LABEL WAS CHANGED FROM TRUNCAL V REGURG P 1/2-TIME TO PULMONARY INSUFFICIENCY PRESSURE HALF TIME (PI PHT). POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED THOSE REPORTS PRIOR TO THE CHANGE COULD DISPLAY INCORRECT CLINICAL DATA IF VIEWED IN RM OR IF USED AS A TEMPLATE FOR A REPEAT STUDY ON THAT PATIENT - IN THIS EXAMPLE THE TWO DISEASE STATES ARE VERY DIFFERENT, LEADING TO POTENTIAL INCORRECT DIAGNOSIS. DESCRIPTION OF MEASUREMENT DOES NOT CORRELATE WHEN USING HOVER TOOL WHICH CAN CAUSE CONFUSION. TRUNCUS ARTERIOSUS IS A RARE HEART DEFECT THAT'S PRESENT AT BIRTH (CONGENITAL). IN TRUNCUS ARTERIOSUS, IT MEANS THAT ONE LARGE BLOOD VESSEL LEADS OUT OF THE HEART. NORMALLY, THERE ARE TWO SEPARATE VESSELS COMING OUT OF THE HEART. AFTER A (B)(6) IS TREATED FOR TRUNCUS ARTERIOSUS, HE/SHE CAN DEVELOP TRUNCAL VALVE STENOSIS AND OR TRUNCAL VALVE REGURGITATION. PULMONARY VALVE INSUFFICIENCY (OR INCOMPETENCE, OR REGURGITATION) IS A CONDITION WHERE THE PULMONARY VALVE IS NOT STRONG ENOUGH TO PREVENT BACKFLOW TO THE RIGHT VENTRICLE. IF IT IS SECONDARY TO PULMONARY HYPERTENSION IT IS REFERRED TO AS A GRAHAM STEELL MURMUR. THE THREE PRIMARY PATHOLOGICAL MECHANISMS CAUSING PULMONARY VALVE INSUFFICIENCY ARE DILATATION OF THE PULMONIC VALVE RING, ACQUIRED ALTERATION OF PULMONIC VALVE LEAFLET MORPHOLOGY, OR CONGENITAL ABSENCE OR MALFORMATION OF THE VALVE. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE # Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | -incorrect device orientation -device malfunction -occlusion notified | -Device displays error message -break -fitting problem | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT # 1225058-2016-00001, 1225058-2016-00002, AND 1225058-2016-00003 TO THE FDA ON (B)(6) 2016 FOR A SITE IN THE US. THIS REPORT IS FOR A 4TH OCCURRENCE OF THE SAME ISSUE AT A DIFFERENT SITE AND INCLUDES ONE (1) REPORT AFFECTED ON (B)(6) 2016, WHEN THE ERROR OCCURRED. THIS WAS AN INTERNAL DISCOVERY BY AGFA DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN AS REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA (B)(4) (1225058-12-22-2016-001-C). REPORT ID DATE ERROR OCCURRED. (B)(4), (B)(6) 2016. BACKGROUND ON NOVEMBER 22, 2016, A CUSTOMER REPORTED TO AGFA THAT WHEN USING IMPAX CV REPORTING FOR NON-INVASIVE VASCULAR ULTRASOUND (NIV), WHEN BUILDING A NIV REPORT, IN THE CONCLUSION SECTION OF THE UPPER ARTERIAL DUPLEX, IF YOU SELECT ANY CHECK BOXES UNDER THE LEFT ANEURYSM IT PUTS IN THE REPORT THAT THERE WAS NO ANEURYSM. ON NOVEMBER 22, 2016, AGFA TECHNICAL SUPPORT CONTACTED THE CUSTOMER AND UPDATED HIM THAT ONLY THREE (3) REPORTS WERE AFFECTED AND THOSE THREE (3) REPORTS WERE ALL LISTED AS TEST REPORTS. THE CUSTOMER WAS PROVIDED THE LIST OF THE AFFECTED REPORTS. AGFA TECHNICAL SUPPORT RE-CONFIGURED THE REPORTS TO PRODUCE CORRECT STATEMENTS IN BOTH THE TEST AND PRODUCTION ENVIRONMENT. A PROBLEM RECORD WAS CREATED ON NOVEMBER 23, 2016 TO ADDRESS THE ISSUE VIA (B)(4). ON (B)(6) 2016 AGFA HAS IMPLEMENTED A CORRECTION VIA (B)(4) (1225058-12-22-2016-001-C). AGFA SENT URGENT SAFETY NOTICES TO AFFECTED CONSIGNEES FOR THIS ISSUE DESCRIBED IN (B)(4) (1225058-12-22-2016-001-C) ON (B)(4) 2016. THE LETTER INFORMED THE CONSIGNEES THAT AGFA SERVICES WILL CONTACT THE AFFECTED CONSIGNEES TO: EXECUTE A QUERY TO FIND ALL AFFECTED REPORTS AT THE SITE AND TO PROVIDE THE RESULTS. THE QUERY WILL NOT REQUIRE DOWNTIME AND WILL NOT IMPACT SYSTEM PERFORMANCE. CORRECT THE NIV REPORTING CLINICAL CONTENT. THE CORRECTION ACTIVITY WILL NOT REQUIRE DOWNTIME. THE LETTER ALSO INFORMED CUSTOMERS TO CONTACT AGFA IF THEY HAD ANY QUESTIONS RELATED TO THE COMMUNICATION LETTER. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. ANY ADDITIONAL REPORTING FOR THIS CORRECTION ACTIVITY WILL BE REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA (B)(4) (1225058-12-22-2016-001-C). THERE HAVE BEEN NO REPORTS OF PHYSICAL HARM OR HEALTH DAMAGE RESULTING FROM THIS ISSUE. | -occlusion notified -component open -incorrect device orientation | -disconnection -break -device damaged prior to use | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS OCCURRENCE ARE 1737 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER/S (MRN) IDENTIFIED. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 14TH OCCURRENCE ARE 97 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 936 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. | -incorrect device orientation -device malfunction -occlusion notified | -component missing -break -Device displays error message | Viewer
| FDA | Event Description: ALL IMAGES ON THE SYNGO.PLAZA SYSTEM WERE NOT ACCESSIBLE IN THE ONLINE STORAGE DISKS. THERE WERE MULTIPLE POWER BLACKOUTS AND NO UNINTERRUPTABLE POWER SUPPLY (UPS) IS INSTALLED, WHICH COULD WITHSTAND VOLTAGE VARIATIONS. THE DAMAGE ON THE STS ARRAYS WAS UNRECOVERABLE AND THE SYSTEM WAS REINSTALLED FROM SCRATCH TO MAKE STORING OF NEW IMAGES POSSIBLE AGAIN. DURING ATTEMPTS TO RESTORE DATA FROM SYNGO.PLAZA'S ARCHIVE ALSO A MAJOR DAMAGE ON ITS DISKS WAS IDENTIFIED BECAUSE OF THE SAME POWER OUTAGES. ALL DATA ON THE ARCHIVE (5375 STUDIES) CANNOT BE ACCESSED ON ITS FILESYSTEM AS WELL. THE RESTORE AND CHECKS ON THE ARCHIVE ARE ONGOING, IF THE FILESYSTEM CAN BE MADE ACCESSIBLE AGAIN AND THUS USE IT FOR THE RESTORE. THE CUSTOMER WAS INFORMED ABOUT AFFECTED PATIENTS/STUDIES. BASED ON QUANTUM, OF CURRENTLY NOT ACCESSIBLE DATA, THIS IS PRELIMINARY TREATED AS A DATA LOSS. | -fail to open images -unable to access -unable to scroll image | -no display or display failure -audible noise -Power Source Issue | Viewer
| FDA | Event Description: ALLEGEDLY, WHILE PREPARING AN OR ROOM FOR THE FIRST CASE OF THE DAY, A STAFF MEMBER HEARD A LOUD POP COMING FROM THE MONITOR. SHE WENT TO NOTIFY HER SUPERVISOR OF THE POP; HOWEVER, WHEN SHE RETURNED TO THE OR ROOM, FLAMES WERE COMING FROM THE BACK OF THE MONITOR. STAFF WAS ABLE TO EXTINGUISH THE FIRE USING TWO CO2 EXTINGUISHERS. | -loud fan noise -device alarmed -noise emitted | -improper or incorrect procedure or method -fitting problem -fluid leak | Viewer
| FDA | Event Description: ANNULUS MEASUREMENT WAS PERFORMED BY A RADIOLOGIST WITH ASSISTANCE OF ADVANTAGE WORKSTATION TAVI (TRANSCATHETER AORTIC VALVE REPLACEMENT) SOFTWARE ON A PATIENT CT SCAN OBTAINED (B)(6) 2018. A TAVI VALVE IMPLANTATION PROCEDURE WAS PERFORMED ON (B)(6) 2018 WHERE A MEDTRONIC EVOLUTE PRO 29MM WAS PLACED, BASED ON RADIOLOGIST MEASUREMENT OF 465.4 MM2. SEVERE PERIPROSTHETIC VASCULAR INSUFFICIENCY WAS IDENTIFIED FOLLOWING VALVE PLACEMENT. A TAVI VALVE IN VALVE PROCEDURE WAS THEN PERFORMED IN (B)() 2018 FOLLOWING THE SAME RADIOLOGIST MEASUREMENT OF ANNULUS USING A MORE RECENT CT IMAGE AND DIFFERENT METHOD THAT RESULTED IN A MEASUREMENT OF 547 MM2, WHERE AN EDWARDS SAPIEN3 29MM VALVE WAS PLACED WITH NORMAL POST IMPLANT OUTCOME. | -incorrect measurement -loss of measurement -difficult to calibrate | -compliance issue -fluid leak -fail-safe design failure | Viewer
| FDA | Event Description: ANNULUS MEASUREMENT WAS PERFORMED BY A RADIOLOGIST WITH ASSISTANCE OF ADVANTAGE WORKSTATION TAVI (TRANSCATHETER AORTIC VALVE REPLACEMENT) SOFTWARE ON A PATIENT CT SCAN OBTAINED (B)(6) 2018. A TAVI VALVE IMPLANTATION PROCEDURE WAS PERFORMED IN (B)(6) 2018 WHERE A MEDTRONIC EVOLUTE PRO 26MM WAS PLACED, BASED ON RADIOLOGIST MEASUREMENT OF 377 MM2. WITHIN 5 MINUTES OF VALVE PLACEMENT, THE VALVE MIGRATED TO ASCENDING AORTA. DURING THIS TAVI PROCEDURE, FOLLOWING THE SAME RADIOLOGIST RE-MEASUREMENT OF ANNULUS USING A DIFFERENT METHOD THAT RESULTED IN A MEASUREMENT OF 390 MM2, AN EDWARDS SAPIEN3 29MM VALVE WAS PLACED WITH NORMAL POST IMPLANT OUTCOME. | -incorrect measurement -loss of measurement -difficult to calibrate | -compliance issue -fluid leak -fail-safe design failure | Viewer
| FDA | Event Description: ANNULUS MEASUREMENT WAS PERFORMED BY A RADIOLOGIST WITH ASSISTANCE OF ADVANTAGE WORKSTATION TAVI (TRANSCATHETER AORTIC VALVE REPLACEMENT) SOFTWARE ON A PATIENT CT SCAN OBTAINED (B)(6) 2018. A TAVI VALVE IMPLANTATION PROCEDURE WAS PERFORMED IN (B)(6) 2018 WHERE A MEDTRONIC EVOLUTE PRO 26MM WAS PLACED, BASED ON RADIOLOGIST MEASUREMENT OF 377 MM2. WITHIN 5 MINUTES OF VALVE PLACEMENT, THE VALVE MIGRATED TO ASCENDING AORTA. DURING THIS TAVI PROCEDURE, FOLLOWING THE SAME RADIOLOGIST RE-MEASUREMENT OF ANNULUS USING A DIFFERENT METHOD THAT RESULTED IN A MEASUREMENT OF 390 MM2, AN EDWARDS SAPIEN3 29MM VALVE WAS PLACED WITH NORMAL POST IMPLANT OUTCOME. | -incorrect measurement -loss of measurement -difficult to calibrate | -compliance issue -fluid leak -fail-safe design failure | Viewer
| FDA | Event Description: AS PART OF A CHANGE OF EHR AT THIS INSTITUTION, A NEW PACS SYSTEM WAS INTRODUCED REPLACING ONE THAT WAS FUNCTIONING WELL. NONE OF THE HOSPITAL IT EXPERTS, NOR EXPERTS FROM THE VENDOR, CHECKED THE COMPUTERS TO DETERMINE IF THE IMAGING SYSTEM'S LAUNCHER COULD BE LOADED. AS IT TURNS OUT, INNUMERABLE COMPUTERS DO NOT HAVE THE STORAGE TO LOAD THE THIN CLIENT LAUNCHER AND THUS, NO RADIOGRAPH IMAGES COULD BE SEEN ON THESE SYSTEMS LINKED TO THE HOSPITAL'S EHR. THIS HAS ADVERSE IMPACT ESPECIALLY AT SITES THAT HAVE ONE OR TWO COMPUTERS CAUSING DELAYS IN CARE AND DIAGNOSIS. THIS IS ESPECIALLY DANGEROUS IN PTS WITH CARDIOPULMONARY DISEASE AND HEART FAILURE OR PNEUMONIA. THERE HAS BEEN NO REMEDY PUT FORTH AND NO ONE HAS ASSUMED LEADERSHIP TO SOLVE THE DYSFUNCTION. | -fail to replace manually -upgrade failed -lack of implementation process | -spark -protective measure issue -fluid leak | Viewer
| FDA | Event Description: BEFORE GUIDED IMPLANT SURGERY WAS CONDUCTED, THE DOCTOR IDENTIFIED THAT THE IMPLANT LENGTHS BETWEEN THE REFERENCE CHART PROVIDED WITH THE GUIDE AND HIS TREATMENT PLANNING FILE DID NOT MATCH. THE REFERENCE CHART STATED THAT THE IMPLANT LENGTH WAS 6 MM BUT THE PLANNING FILE STATED IT WAS 8 MM. HE ALSO NOTICED THAT THERE WAS NO 6 MM OPTION AVAILABLE FOR HIS SELECTED IMPLANT TYPE. AS THE DOCTOR WAS UNCERTAIN WHICH IMPLANT LENGTH WAS CONSIDERED DURING GUIDE MANUFACTURE, HE CHOSE TO FREE-HAND THE SURGERY. THERE HAVE BEEN NO ADVERSE EVENTS REPORTED FOR THE PATIENT. ALTHOUGH THE ISSUE WAS FIRST REPORTED AS A SURGICAL GUIDE COMPLIANT, THE INVESTIGATION IDENTIFIED THAT THE CAUSE WAS RELATED TO A DISCREPANCY BETWEEN INVIVODENTAL/TXSTUDIO SOFTWARE VERSIONS WHICH WAS USED TO CREATE THE TREATMENT PLAN. THEREFORE, THE SUSPECT MEDICAL DEVICE SECTION REFLECTS THE SOFTWARE PRODUCT INFORMATION. | -incorrect device orientation -incorrect procedure -part out of specification | -loose or intermittent connection -disconnection -communication or transmission issue | Viewer
| FDA | Event Description: CADATREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING A STUDY WITH INCORRECT PATIENT LATERALITY. THE PATIENT'S LATERALITY WAS ENTERED AT THE SCANNER INCORRECTLY. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER: COMPLAINT-(B)(4). | -mri image distorted -fail to enable mri setting -fail to place annotation on image | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2017, A RADIOLOGY TECHNOLOGIST CONTACTED MERGE TECHNICAL SUPPORT FOR ASSISTANCE WITH STUDY PROCESSING. DURING TROUBLESHOOTING IT WAS DISCOVERED THAT THE PATIENT ORIENTATION HAD BEEN INCORRECTLY ENTERED AT THE MRI SCANNER. CADSTREAM IS UNABLE TO EDIT IMAGE DICOM DATA. PER THE CUSTOMER'S REQUEST, ADJUSTMENTS WERE MADE TO REACH A 'READY TO READ' STATUS WITH THE INCORRECT DICOM REFLECTED ON ALL IMAGES. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED ((B)(4)) THAT FOLLOWING A ROUTINE SYSTEM POWER DOWN, THE TIME WITHIN THE SYSTEM IS INCORRECT AND THE DRIVES ARE NOT AVAILABLE. | -mri image distorted -magnet mode discrepancy -fail to follow delivery schedule | -protective measure issue -device maintenance issue -detachment of device component | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED ((B)(6)) THAT IMAGES ARE PROCESSING FINE WITH THE EXCEPTION OF THE RIGHT AND LEFT DESIGNATION ON THE SUBTRACTION IMAGES. CUSTOMER RECENTLY UPGRADED TO CADSTREAM 6.0.1. WHEN CADSTREAM BREAKS UP THE BREAST SEQUENCES INTO RIGHT AND LEFT, THE SEQUENCES WERE LABELED AS R POST IMAGES, OR R SUB 2, ETC. USING VERSION 5.0.7 OF CADSTREAM. NOW ONLY THE POST IMAGES ARE LABELED. | -mri image distorted -image not in dicom format -fail to enable mri setting | -no display or display failure -protective measure issue -break | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. A CUSTOMER REPORTED (CASE (B)(4)) THEY ARE HAVING ISSUES SENDING IMAGES TO CADSTREAM. THIS COULD CAUSE A DELAY IN PATIENT CARE. | -mri image distorted -magnetic interference -mri scan restriction confused | -fluid leak -protective measure issue -Device Handling Issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). CADSTREAM ALSO INCLUDES USER-CONFIGURABLE FEATURES FOR REPORTING ON FINDINGS IN BREAST OR GENERAL MRI STUDIES. ADDITIONALLY, CADSTREAM ASSISTS USERS IN PLANNING MRI GUIDED INTERVENTIONAL PROCEDURES. WHEN INTERPRETED BY A SKILLED PHYSICIAN, THIS DEVICE PROVIDES INFORMATION THAT MAY BE USED FOR SCREENING, DIAGNOSIS, AND INTERVENTIONAL PLANNING. PATIENT MANAGEMENT DECISIONS SHOULD NOT BE MADE BASED SOLELY ON THE RESULTS OF CADSTREAM. CADSTREAM MAY ALSO BE USED AS AN IMAGE VIEWER OF MULTI-MODALITY, DIGITAL IMAGES, INCLUDING ULTRASOUND AND MAMMOGRAPHY. CADSTREAM IS NOT INTENDED FOR PRIMARY INTERPRETATION OF DIGITAL MAMMOGRAPHY IMAGES. ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT IMAGES WERE SLOW TO LOAD DURING A BIOPSY PROCEDURE. THE CUSTOMER REPORTED COMPLETING THE BIOPSY BY ESTIMATING THE TARGET LOCATION OF THE BIOPSY BASED ON LANDMARKS AND NOT VISUAL ENHANCEMENT ON IMAGES. ESTIMATING TARGETS HAS THE POTENTIAL TO LEAD TO AN INCORRECT FOLLOW UP WHICH MAY LEAD TO A DELAY IN DIAGNOSIS AND TREATMENT. | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST ASSISTANCE WITH RE-PROCESSING IMAGES THAT WERE BELIEVED TO BE PROCESSED INCORRECTLY. SUPPORT REVIEWED THE INCOMING IMAGES AND SYSTEM REMOTELY. SUPPORT FOUND THAT THE IMAGES APPEARED TO BE SCANNED WITH INCORRECT PATIENT ORIENTATION. THE CUSTOMER CONFIRMED THE ORIENTATION OF THE IMAGES. CADSTREAM DOES NOT ALLOW FOR THE EDITING OR MODIFICATION OF THE ORIGINAL IMAGE. THE CUSTOMER HAS THE ABILITY TO MAKE ANNOTATIONS REGARDING THE ORIENTATION, HOWEVER, CADSTREAM WILL CONTINUE TO REFLECT THE ORIENTATION AS ORIGINALLY SCANNED ON THE PROCESSED IMAGES. SUPPORT INSTRUCTED THE CUSTOMER REGARDING THESE IMAGES AND ORIENTATION.. THE INCORRECT ORIENTATION IS NOT A MALFUNCTION OF THE CADSTREAM DEVICE NOR IS IT A FAILURE OF THE DEVICE TO MEET SPECIFICATION. THE ISSUE IS USE ERROR CAUSED ORIGINATING AT THE MRI SCANNER DUE TO IMPROPER DATA ENTRY. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST ASSISTANCE WITH RE-PROCESSING IMAGES THAT WERE BELIEVED TO BE PROCESSED INCORRECTLY. SUPPORT REVIEWED THE INCOMING IMAGES AND SYSTEM REMOTELY. SUPPORT FOUND THAT THE STUDIES WERE PROCESSED AS EXPECTED; HOWEVER, IT WAS NOTED THAT THE IMAGES APPEARED TO BE SCANNED WITH INCORRECT PATIENT ORIENTATION. THE CUSTOMER CONFIRMED THE IMAGES WERE SCANNED SUPINE INSTEAD OF PRONE SO ALL OF THE IMAGES WERE UPSIDE DOWN. CADSTREAM DOES NOT ALLOW FOR THE EDITING OR MODIFICATION OF THE ORIGINAL IMAGE. THE CUSTOMER HAS THE ABILITY TO MAKE ANNOTATIONS REGARDING THE ORIENTATION, HOWEVER, CADSTREAM WILL CONTINUE TO REFLECT THE ORIENTATION AS ORIGINALLY SCANNED ON THE PROCESSED IMAGES. SUPPORT INSTRUCTED THE CUSTOMER REGARDING THESE IMAGES AND ORIENTATION. THE INCORRECT ORIENTATION IS NOT A MALFUNCTION OF THE CADSTREAM DEVICE NOR IS IT A FAILURE OF THE DEVICE TO MEET SPECIFICATION. THE ISSUE IS USE ERROR CAUSED ORIGINATING AT THE MRI SCANNER DUE TO IMPROPER DATA ENTRY. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING A STUDY WITH INCORRECT PATIENT ORIENTATION. THE PATIENT'S ORIENTATION WAS ENTERED INCORRECTLY AT THE SCANNER. CADSTREAM IS PROHIBITED, PER DESIGN, FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY OF IMAGE DATA AS RECEIVED FROM OTHER APPLICATION ENTITIES. PATIENT IMAGES AND STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND THOSE RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT. THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM DUE TO THE INCORRECT PATIENT ORIENTATION AS DISPLAYED IN CADSTREAM. HOWEVER, THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING A STUDY SCANNED WITH INCORRECT PATIENT LATERALITY. THE PATIENT'S LATERALITY WAS ENTERED AT THE SCANNER INCORRECTLY. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER: (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH PROCESSING A STUDY SCANNED WITH INCORRECT PATIENT LATERALITY. THE PATIENT'S LATERALITY WAS ENTERED AT THE SCANNER INCORRECTLY. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER: (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2020, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING A STUDY WITH INCORRECT PATIENT ORIENTATION. THE PATIENT'S ORIENTATION WAS ENTERED AT THE SCANNER INCORRECTLY. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENT'S PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2019, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE PROCESSING A STUDY WITH INCORRECT PATIENT LATERALITY, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2019, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT ORIENTATION, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. CADSTREAM SUPPORT ADVISED THE CUSTOMER THE PATIENT ORIENTATION COULD NOT BE CORRECTED IN THE CADSTREAM SYSTEM. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2019, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE PROCESSING A STUDY WITH INCORRECT PATIENT ORIENTATION, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. THE CUSTOMER WAS ADVISED THE PATIENT ORIENTATION COULD NOT BE CORRECTED IN CADSTREAM. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2018, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE PROCESSING A STUDY WITH INCORRECT PATIENT LATERALITY, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. THE CUSTOMER WAS ADVISED THE PATIENT LATERALITY COULD NOT BE CORRECTED IN CADSTREAM. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2018, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE PROCESSING A STUDY WITH INCORRECT PATIENT LATERALITY, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. THE CUSTOMER WAS ADVISED THE PATIENT LATERALITY COULD NOT BE CORRECTED IN CADSTREAM. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2018, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT ORIENTATION, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. SUPPORT SUGGESTED ADJUSTMENTS COULD BE MADE TO THE STUDY TO IMPROVE ACCURACY, HOWEVER, THE ORIENTATION WOULD REMAIN INCORRECT. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON MAY 1, 2018, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT ORIENTATION, WHICH WAS ENTERED INCORRECTLY AT THE MR SCANNER. SUPPORT SUGGESTED ADJUSTMENTS COULD BE MADE TO THE STUDY TO IMPROVE ACCURACY, HOWEVER, THE ORIENTATION WOULD REMAIN INCORRECT. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2018, MERGE TECHNICAL SUPPORT WAS CONTACTED BY A USER REPORTING THAT IMAGES IN CADSTREAM DISPLAYED THE INCORRECT PATIENT ORIENTATION. CADSTREAM IS PROHIBITED FROM EDITING OR MODIFYING IMAGE DICOM DATA, INCLUDING PATIENT ORIENTATION AND LATERALITY, RECEIVED FROM OTHER APPLICATION ENTITIES. THE CUSTOMER'S ISSUE IS SUSPECTED OF ORIGINATING AT THE MR SCANNER DURING THE PATIENT'S SCAN ACQUISITION. WHILE THERE IS NO INDICATION THAT THE ISSUE REPORTED BY THE CUSTOMER RESULTED IN ANY HARM TO A PATIENT, STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT; THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. HOWEVER, THERE IS NO INDICATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS (SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS). ON (B)(6) 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH STUDY PROCESSING. WHILE TROUBLESHOOTING, MERGE SUPPORT DISCOVERED THE STUDY REFLECTED AN INCORRECT PATIENT ORIENTATION. THE PATIENT'S ORIENTATION IS ENTERED AT THE SCANNER (MODALITY). CADSTREAM DOES NOT ALLOW FOR EDITING OF THIS DICOM INFORMATION. THE STUDY WAS PROCESSED "AS IS" WHICH LEAD TO ALL IMAGES REFLECTING INCORRECT DICOM ON ALL IMAGES PROCESSED BY CADSTREAM. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER 130128. | -mri image distorted -fail to enable mri setting -fail to monitor power event | -protective measure issue -use of device issue -device markings issue | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. ON (B)(6) 2014 A CUSTOMER CALLED TO REPORT THAT A STUDY WOULD NOT PROCESS. THE ERROR WAS THAT CADSTREAM COULDN'T CREATE REFORMATS OF THE RIGHT/LEFT DUE TO SERIES BEING OBLIQUED. THE TECH USED A SKIN MARKER THAT WAS TOO LARGE, GENERALLY USED FOR EXTREMITIES, AND THE CADSTREAM SOFTWARE IDENTIFIED IT AS THE CHEST WALL. MERGE CUSTOMER SUPPORT TRIED TO TURN OFF EVERY PROCESSING WITHIN THE PROCESS TAB AND CREATED A MATCH REFORMAT WHICH ALLOWED THE CHEST WALL TO BE MOVED BACK, BUT WHEN THE CADSTREAM REPROCESSED THE STUDY THE 'FADED' AREA WAS STILL IN ITS' ORIGINAL LOCATION. MERGE CUSTOMER SERVICE WAS UNABLE TO PROCESS THE STUDY. THE PATIENT MAY NEED TO BE RESCHEDULED FOR A SECOND SCAN. (B)(4) Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -mri image distorted -image not in dicom format -magnetic interference | -leak -break -disconnection | Viewer
| FDA | Event Description: CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. ON (B)(6) 2016 A CUSTOMER REPORTED AN IMAGE EXPORTED FROM CADSTREAM WAS SHOWING UP UNDER A DIFFERENT PATIENT IN PACS. WHILE THE PACS SYSTEM FLAGGED THE INCOMING IMAGE AND ALERTED THE CUSTOMER, IT IS UNCLEAR AT THIS TIME IF THE INCORRECT IMAGE ASSIGNMENT IS ATTRIBUTED TO THE CADSTREAM SYSTEM OR THE PACS SYSTEM. A PATIENT'S IMAGE SHOWING UP UNDER A DIFFERENT PATIENT HAS THE POTENTIAL OF LEADING TO A DELAY IN CARE AND/OR AN INCORRECT TREATMENT OR DIAGNOSIS. HOWEVER, THERE IS NO INDICATION OF ANY HARM TO THE PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | -mri image distorted -image not in dicom format -fail to enable mri setting | -application program issue -detachment of device component -fluid leak | Viewer
| FDA | Event Description: CALCIUM SCORING IN MERGE UNITY PACS IS A NON-INVASIVE WAY OF OBTAINING INFORMATION ABOUT THE PRESENCE, LOCATION, AND EXTENT OF CALCIFIED PLAQUE ON CORONARY ARTERIES. DURING DEVELOPMENT OF THE NEXT RELEASE OF UNITY (VERSION 11.0), THE DEVELOPMENT TEAM FOUND THAT THE CURRENT ALGORITHM OVERESTIMATES THE CALCIUM SCORE FOR SITES THAT OBTAIN CALCIUM SCORING SCANS AT SLICE THICKNESS OF <3MM AND A SLICE INTERVAL OF 0. IT WAS PRESUMED THAT THE SCORE FOR SITES WITH A SLICE THICKNESS OF >3MM AND A THICKNESS OF 0 WOULD BE UNDERESTIMATED, HOWEVER PERFORMING CORONARY CALCIUM SCORING CTS USING A SLICE THICKNESS OF >3MM IS UNLIKELY. THE ERROR IN THE CALCULATION MINIMALLY IMPACTS EVIDENCE AND MILD EVIDENCE OF CORONARY ARTERY DISEASE, AND WAS DEEMED UNLIKELY TO CAUSE ADVERSE HEALTH CONSEQUENCES. THERE WERE NO REPORTED PATIENTS WHO WERE AFFECTED BY THE ERROR. UNITY 11.0.2 WAS RELEASED ON 11/25/2015 TO ACCOUNT FOR SLICES NOT EQUAL TO 3MM AND TO DISPLAY A WARNING WHENEVER SLICE THICKNESS IS NOT 3MM. | -unable to obtain readings -inaccurate information -missing information | -device markings issue -protective measure issue -data back-up problem | Viewer
| FDA | Event Description: CUSTOMER CALLED ABOUT A MERGE CARDIO REPORT FOR ONE PATIENT THAT IS ATTACHED TO THE STUDY OF ANOTHER. ONE PATIENT HAD US STUDY DONE ON (B)(6) 2012. IT HAS A CONFIRMED REPORT, HOWEVER, THE REPORT BELONGS TO ANOTHER PATIENT WHO HAD US STUDY ON (B)(6) 2012. WHEN TRYING TO VIEW THE REPORT, THE USER GETS THE SAME GLOBAL EXCEPTION HANDLER ERROR AS BEFORE: EXCEPTION INFORMATION: EXCEPTION TYPE: HTTPEXCEPTION. EXCEPTION MESSAGE: CANNOT HAVE MULTIPLE ITEMS SELECTED IN A DROPDOWN LIST. THE PDF GENERATION FUNCTIONALITY WAS NOT DESIGNED TO BE EXCLUSIVE PER PATIENT. IF TWO OR MORE USERS WERE REPORTING ON DIFFERENT PATIENTS AND REQUESTED THE GENERATION OF A PDF REPORT FILE AT THE SAME TIME, THE SYSTEM WOULD INCORRECTLY ATTACH THE FIRST PDF REPORT FILE GENERATED TO THE OTHER PATIENT(S). Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. | -patient incident -lack of intended patient population -fail to follow process | -break -disconnection -detachment of device component | Viewer
| FDA | Event Description: CUSTOMER REPORT ALLEGES SMOKE COMING FROM SMARTCR COMPUTED RADIOGRAPHY (CR) READER UNIT AROUND 23:00 AFTER FIRST NOTICING A BURNING SMELL AROUND 21:00. UNIT WAS REMOVED FROM NICU AND TAKEN TO PLANT OPS. THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | -burning odor -device burned -component burned | -Leak -leak -Detachment of Device Component | Viewer
| FDA | Event Description: CUSTOMER REPORTED AN ISSUE OF HEMO DATA CREATING INCORRECT PATIENT STUDY WHEN IMPORTED TO CARDIO. IN THE VERY RARE COINCIDENCE WHERE TWO PATIENTS ARE ADDED TWO HEMO DATA SETS ARE IMPORTED INTO CARDIO DATABASE AT ALMOST EXACTLY THE SAME TIME, CARDIO WILL POTENTIALLY ASSOCIATE THE HEMO DATA WITH THE INCORRECT PATIENT. THE HEMO DATA WOULD NOT SHOW WITH THE EXPECTED PATIENT BUT WITH ANOTHER PATIENT. THIS ISSUE CAN POTENTIALLY RESULT IN ONE OF TWO RISKS: DELAY IN PATIENT CARE SINCE THE PHYSICIAN WILL NEED TO SEARCH FOR THE HEMO DATA THAT WAS EXPECTED OR ASK THAT DATA GET REIMPORTED BEFORE THE REPORT CAN BE FINALIZED. IN THE VERY RARE OCCURRENCE THAT THE HEMO DATA IS SAVED UNDER A PATIENT THAT HAS A HEMO STUDY SCHEDULED OR COMPLETED WITHIN THE SAME TIME FRAME. THE PHYSICIAN CAN LOOK AT THE HEMO DATA OF THE INCORRECT PATIENT AND ISSUE THE WRONG TREATMENT OR DIAGNOSIS. (B)(4). | -incorrect patient information -patient experienced adverse event -patient data breached | -application security issue -device alarm system issue -Protective Measure Issue | Viewer
| FDA | Event Description: CUSTOMER REPORTED FRAME SKIPPING ISSUES WHEN PLAYING CINE SERIES. THE RADIOLOGIST RECENTLY MISSED A BRAIN BLEED ON A PATIENT WHICH CAUSED A NEGATIVE OUTCOME, BUT DID NOT RESULT IN DEATH. THE BRAIN BLEED APPARENTLY SHOWED ON ONLY ONE IMAGE AND IT WAS THE IMAGE THAT WAS SKIPPED. INFORMATION REPORTED INDICATES A REPORTABLE MALFUNCTION DATA DISPLAY ISSUE (NON-IMAGE QUALITY), WHICH MAY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY REQUIRING MEDICAL INTERVENTION; THEREFORE, A REPORTABLE MALFUNCTION WILL BE FILED. | -tissue bleed -part missing -skip failed | -failure of device to self-test -no flow -Device Alarm System Issue | Viewer
| FDA | Event Description: CUSTOMER REPORTED SMOKE COMING FROM UNIT. UNIT WAS POWERED DOWN AND TAKEN OUT OF SERVICE. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | -fail to read input signal -device burned -part burned | -fluid leak -air leak -leak | Viewer
| FDA | Event Description: CUSTOMER REPORTED SMOKE RISING OUT OF SLOT AREA OF THE SMARTCR READER. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | -incorrect readings -invisible indication -unable to read | -Device displays error message -air leak -Inaccurate Flowrate | Viewer
| FDA | Event Description: CUSTOMER REPORTED THAT A SMARTCR COMPUTED RADIOGRAPHY (CR) READER UNIT IN THE EMERGENCY DEPARTMENT WAS EMITTING SMOKE FROM THE UNDERSIDE. THE UNIT WAS REMOVED FROM ITS LOCATION AND TAKEN OUT OF SERVICE UNTIL A FUJI SERVICE REPRESENTATIVE COULD ARRIVE. THE UNIT WAS POWERED ON AT THE TIME OF THE EVENT, ALTHOUGH THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. | -patient experienced adverse event -exposure to gas -part ejected | -Detachment of Device Component -fluid leak -fire | Viewer
| FDA | Event Description: CUSTOMER REPORTED THAT STUDIES ARE BEING RECEIVED BUT NOT MOVING FROM CACHE. CUSTOMER REPORTED POTENTIAL DELAY IN PATIENT CARE, AS THE DATE OF THE STUDY WAS (B)(6) 2021 AND THE STUDY WAS NOT RECEIVED UNTIL (B)(6) 2021. | -positioning delayed -inadequate clinical history data -insufficient clinical evidence | -Detachment of Device Component -Connection Issue -leak | Viewer
| FDA | Event Description: CUSTOMER REPORTED THAT SWITCHING STUDIES CAUSES THE MEASUREMENT LIST TO KEEP MEASUREMENTS FROM THE PREVIOUS STUDY, AND MIXES THEM WITH THOSE FROM THE NEW ONE. MEASUREMENT LIST CONTAINS MEASUREMENTS FROM ANOTHER STUDY. | -incorrect measurement -inaccurate measurement item -inadequate measurement unit | -disconnection -imprecision -use of device issue | Viewer
| FDA | Event Description: CUSTOMER REPORTED THAT SYNAPSE WAS UNABLE TO DISPLAY IMAGE FILES AND THE "IMAGE NOT LOADED" MESSAGE WAS DISPLAYED. SUBSEQUENT INSPECTION BY SERVICE PERSONNEL IDENTIFIED THE IMAGES WERE NO LONGER IN THE DATABASE. | -fail to display message -incorrect message display -inadequate message display | -Crack -component missing -No Display or Display Failure | Viewer
| FDA | Event Description: CUSTOMER REPORTED THAT WHEN MEASURING A LESION ON AN UNMAGNIFIED MAMMOGRAPHY IMAGE USING MERGE PACS, AND THEN PERFORMING THE SAME MEASUREMENT ON THE IMAGE MAGNIFIED BY THE HOLOGIC IMAGER, THE MEASUREMENT WAS NOT THE SAME. IF DOING MEASUREMENTS ON MAGNIFIED AND UNMAGNIFIED IMAGES IT IS RELATIVELY OBVIOUS THAT THE MEASUREMENTS ARE NOT CONSISTENT. HOWEVER, IF PERFORMING MEASUREMENT ONLY ON A MAGNIFIED IMAGE IT MAY NOT BE OBVIOUS THAT THE MEASUREMENTS ARE NOT ACCURATE. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | -fail to interpret images -incorrect measurement -fail to measure | -application program issue -no display or display failure -leak | Viewer
| FDA | Event Description: CUSTOMER REPORTED WHEN MEASURING WITH THE MERGE CARDIO TOOLBAR, IT DISPLAYS CM ON THE WORKSTATION, BUT MM ON THE REPORT. EX: THE RVOT VALUE SENT FROM THE CART IS 1.5CM. THE UNITS OF MEASURE IN THE (B)(4) TABLE ARE CORRECT (CM) HOWEVER THE UNITS OF MEASURE IN THE 17 MEASUREMENTS. XSLT FILE ARE INCORRECT (MM). THIS CAUSES THE NUMBER IN THE REPORT TO DISPLAY AS 0.15CM. WHEN TAKING MEASUREMENTS FROM IMAGES ON THE CARDIO WORKSTATION OR FROM THE US CART, NUMBERS ARE NOT CROSSING TO THE REPORT IN THE CORRECT UNIT OF MEASURE. SPECIFICALLY, TWO MEASUREMENTS: RIGHT VENTRICULAR OUTFLOW TRACT (RVOT) DIAMETER AND MAIN PULMONARY ARTERY ED DIAMETER. SINCE THE TWO VALUES ARE SMALLER THAN THEY ARE WHICH CAN MISLEAD THE CARDIOLOGIST THAT THERE ARE ABNORMALITY. THE CARDIOLOGIST COULD POTENTIALLY PRESCRIBE UNNECESSARY MEDICATIONS OR TREATMENT. (B)(4). | -inappropriate measurement unit -incorrect measurement -incorrect message display | -no display or display failure -human-device interface issue -deployment issue | Viewer
| FDA | Event Description: CUSTOMER REPORTED WHEN MEASURING WITH THE MERGE CARDIO TOOLBAR, IT DISPLAYS CM ON THE WORKSTATION, BUT MM ON THE REPORT. EX: THE RVOT VALUE SENT FROM THE CART IS 1.5CM. THE UNITS OF MEASURE IN THE DDOCRREPORTDATATOCPMAP TABLE ARE CORRECT (CM) HOWEVER THE UNITS OF MEASURE IN THE 17MEASUREMENTS.XSLT FILE ARE INCORRECT (MM). THIS CAUSES THE NUMBER IN THE REPORT TO DISPLAY AS 0.15CM. WHEN TAKING MEASUREMENTS FROM IMAGES ON THE CARDIO WORKSTATION OR FROM THE US CART, NUMBERS ARE NOT CROSSING TO THE REPORT IN THE CORRECT UNIT OF MEASURE. SPECIFICALLY, TWO MEASUREMENTS: RIGHT VENTRICULAR OUTFLOW TRACT (RVOT) DIAMETER AND MAIN PULMONARY ARTERY ED DIAMETER. SINCE THE TWO VALUES ARE SMALLER THAN THEY ARE WHICH CAN MISLEAD THE CARDIOLOGIST THAT THERE ARE ABNORMALITY. THE CARDIOLOGIST COULD POTENTIALLY PRESCRIBE UNNECESSARY MEDICATIONS OR TREATMENT. (B)(4). | -inappropriate measurement unit -incorrect measurement -incorrect message display | -no display or display failure -human-device interface issue -deployment issue | Viewer
| FDA | Event Description: DISTANCE MEASUREMENTS MAY NOT BE CALCULATED ACCURATELY BY THE UNIVERSAL VIEWER WHEN APPLIED TO IMAGES ACQUIRED WITH A MAGNIFICATION FACTOR ON SIEMEN'S XA MODALITY (ANGIOGRAPHY X-RAY IMAGES) . DURING A PRODUCT DEMONSTRATION OF THE UNIVERSAL VIEWER WEB 6.0 , A RADIOLOGIST MEASURED VALUES ONSCREEN EQUALING TO 1.36CM AND 1.31CM, WHEN THE ACTUAL MEASUREMENTS SHOULD HAVE BEEN AROUND 1CM. IF NOT NOTICED BY THE CAREGIVER, THIS COULD LEAD TO AN OVERESTIMATE OF THE ANATOMY OF INTEREST. WHEN A MODALITY SENDS IMAGES TO PACS THAT ONLY CONTAIN VALUES IN THE IMAGER PIXEL SPACING (0018, 1164) , AND MAGNIFICATION FACTOR (0018, 1114) DICOM TAGS, BUT DO NOT INCLUDE A VALUE IN THE PIXEL SPACING (0020,0030) DICOM TAG THE MEASUREMENTS MAY BE INCONSISTENT WITH THE MAGNIFICATION FACTOR RESULTING A DISCREPANCY IN THE MEASUREMENTS DISPLAYED ON THE VIEWER. GE RECOMMENDS THAT USERS INCLUDE AN APPROPRIATELY PLACED OBJECT OF KNOWN SIZE IN THE IMAGE TO BE USED FOR CALIBRATION THAT WILL ALLOW USERS TO OBTAIN ACCURATE DISTANCE MEASUREMENTS. UPON CALIBRATION THE MEASUREMENTS ARE CORRECT. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: INVESTIGATION REVEALED THERE ARE MULTIPLE PRIMARY ROOT CAUSES THAT LED TO THE REPORTED ISSUE: 1. DIFFERENCE IN DICOM IMPLEMENTATION ACROSS VENDORS AND PRODUCTS 2. LOGIC ERROR IN THE USAGE OF DICOM TAGS IN THE UNIVERSAL VIEWER PRODUCT FOR DIGITAL PROJECTION MODALITIES. 3. REQUIREMENTS ARE LACKING DETAILS. 4. CROSS-FUNCTIONAL COMMUNICATION. 5. INSUFFICIENT DICOM CONFORMANCE STATEMENT. | -incorrect distance measurement -incorrect measurement -inaccurate measurement | -Image Resolution Poor -sticking -Communication or Transmission Issue | Viewer
| FDA | Event Description: DOING A MR FUSION PROSTATE BIOPSY WITH ANESTHESIA AND DURING THE PLANNING STAGE OF THE STUDY, THE MACHINE RESET/TURNED OFF. THE ENTIRE MAPPING PROCESS/PROCEDURE HAD TO BE RESTARTED. | -device reset incorrectly -fail to reset sensor -device failsafe failed | -incorrect or inadequate result -human-device interface issue -crack | Viewer
| FDA | Event Description: DUE TO AN IDENTIFIED HARD DRIVE FAILURE ON THE SYNGO.PLAZA SERVER, ALL STUDIES AVAILABLE IN SHORT TERM STORAGE (STS) WERE LOST. AS THERE IS NO ARCHIVE SOLUTION IN THIS SITE, IS NOT POSSIBLE TO RESTORE THE DATA IN THE STS. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED STUDIES AND THE POSSIBILITY TO RECOVER THE DELETED DATA; HOWEVER, THE AFFECTED DATA IS CONSIDERED AS CLINICALLY RELEVANT. THE REPORTED EVENT OCCURRED IN (B)(6). | -insufficient storage space -hardware failure -fail to record data | -use of device issue -device operates differently than expected -device inoperable | Viewer
| FDA | Event Description: DUE TO HARDWARE ISSUES WITH THE SYNGO.PLAZA'S HARD DRIVES, A YET UNIDENTIFIED AMOUNT OF MEDICAL IMAGES IS NOT ACCESSIBLE ANYMORE. THE SYSTEM WAS CONFIGURED AS A WORKSTATION AND NOT AS A PACS. NO LONG TERM STORAGE ARCHIVE IS CONFIGURED, THUS THERE IS NO POSSIBILITY TO RESTORE DATA FROM THERE. ONCE THE RESTORE ATTEMPTS OF THE AFFECTED HARD DRIVES ARE COMPLETED, AN UPDATED REPORT WILL FOLLOW CONTAINING THE DETAILED INFORMATION ABOUT THE DATA LOSS. | -loss of image -unable to access -fail to access | -unstable -use of device issue -device operates differently than expected | Viewer
| FDA | Event Description: DURING A BILATERAL TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 10 STEM FOR BOTH PATIENT'S LEFT AND RIGHT HIPS. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED SIZE 10 AND SIZE 12 STEMS. IT IS UNKNOWN WHICH SIDE RECEIVED WHICH SIZED STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -incorrect sample size -incorrect device orientation -inappropriate interface design | -fitting problem -defective component -mechanical issue | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THAT THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A STEM 3 TO 8 SIZES SMALLER THAN THE STEM SIZE THAT WAS IMPLANTED. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 11 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -size out of specification -inappropriate size -inappropriate software interface | -component missing -leak -inadequate storage | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 10 STEM. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 16 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -inappropriate display size -size out of specification -incorrect sample size | -fitting problem -improper or incorrect procedure or method -failure to deploy | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 12 STEM. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 14 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -inappropriate display size -incorrect sample size -size out of specification | -failure to deploy -component missing -fitting problem | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 14 STEM FOR THE PATIENT'S LEFT HIP AND A SIZE 11 STEM FOR THE PATIENT'S RIGHT HIP. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 16 STEM IN THE PATIENT'S LEFT HIP AND A SIZE 16 STEM IN THE PATIENT'S RIGHT HIP. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -incorrect patient information -incorrect sample size -incorrect device orientation | -fitting problem -Contamination During Use -Device Handling Issue | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 7 STEM. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 10 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -inappropriate display size -incorrect sample size -size out of specification | -fitting problem -improper or incorrect procedure or method -failure to deploy | Viewer
| FDA | Event Description: DURING A TOTAL HIP ARTHROPLASTY, IT WAS DISCOVERED THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A STEM 3 TO 8 SIZES SMALLER THAN THE STEM SIZE THAT WAS IMPLANTED. THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED A SIZE 13 STEM. THERE WAS NO PATIENT INJURY OR DELAY IN PROCEDURE. | -inappropriate size -size out of specification -incorrect device orientation | -component missing -inadequate storage -kinked | Viewer
| FDA | Event Description: DURING AN OPEN SURGERY, THE FACILITY USED THE SUBJECT OTV-S7PRO FOR OBSERVATION OF THE OPERATIONAL FIELD AND RECORDING THE IMAGES. UPON CONNECTING A CAMERA HEAD, OTV-S7PRO-HD-E12 TO OTV-S7PRO, A MALFUNCTION OF THE WHOLE IMAGE TURNING GREEN OCCURRED PERIODICALLY. THE FACILITY ATTEMPTED TO TURN THE POWER SWITCH ON AND OFF BUT THE DEVICE DID NOT COME TO NORMAL. THE PROCEDURE WAS COMPLETED BY REPLACING THE SUBJECT OTV-S7PRO TO CV-190. NO HEALTH DAMAGE TO THE PATIENT WAS REPORTED OTHER THAN REPLACING THE DEVICE. IN RESPONSE TO THE FAILURE, THEY HAD BEEN SUBSTITUTED FOR THE OLYMPUS-OWNED OTV-S7PRO AND OTV-S7PROH-HD-E12 LATER. HOWEVER, THE SIMILAR PHENOMENON OCCURRED AGAIN. | -loss of image -image orientation incorrect -inadequate image manipulation | -improper or incorrect procedure or method -accessory incompatible -fitting problem | Viewer
| FDA | Event Description: DURING ARTHROSCOPY AND ORIF ORTHOPEDIC PROCEDURE, THE COMPUTER ON THE TOWER (WHERE YOU CAN ENTER THE PATIENT'S INFORMATION AND WHERE IT HAS THE CONTROL RECORDINGS) TURNED OFF SPONTANEOUSLY. IT HAD TO BE RESET AND TURNED BACK ON. WHEN IT WAS RESET, THE PATIENT'S DATA WAS MISSING FROM THE SYSTEM. NO INJURY TO THE PATIENT. | -loss of control -loss of indication -loss of conduction | -fluid leak -data back-up problem -leak | Viewer
| FDA | Event Description: DURING INTERNAL REVIEW IT WAS FOUND THAT IN THE ADVANCED REPORTING (AR) SYSTEM OF SYNAPSE CARDIOVASCULAR SOFTWARE, WHEN UPDATING A SAVED VOLUME MEASUREMENT ON AN ECHO REPORT, THE VALUES FOR AREA AND VOLUME APPEARED TO BE SWAPPED IN THE RESULTS PANEL. THIS ISSUE AFFECTS AR VERSIONS 6.0.4 THROUGH 6.2.0. THE IMPACT IS IN ECHO REPORTS AND NON-INVASIVE VASCULAR REPORTS AND AFFECTS 10 OF 17 AR SITES. THERE WAS NO ADVERSE EVENT OR INJURY REPORTED; THIS INCIDENT IS BEING REPORTED IN AN ABUNDANCE OF CAUTION. NO OTHER INFORMATION WAS PROVIDED. | -incorrect volume calculation -inappropriate software release -fail to update configuration | -leak -incorrect measurement -no display or display failure | Viewer
| FDA | Event Description: DURING MIGRATION FROM THE SIEMENS SYNGO IMAGING TO SECTRA PACS, IT WAS FOUND THAT 13 STUDIES HAD NOT MIGRATED TO SECTRA PACS. THE DATA LOSS WAS OF CLINICAL RELEVANCE; HOWEVER, NO PATIENT RESCANS WERE REQUIRED. THE LOST DATA IS OLD AND THE CLINICAL RELEVANCE OF LOST DATA IS ASSESSED BY SIEMENS EXPERTS AS VERY LOW. READINGS AND REPORTS FOR ALL STUDIES HAVE BEEN SUCCESSFULLY GENERATED PRIOR TO THE INCIDENT. ONLY ONE IMAGE DATED 2017 IS CONSIDERED CLINICALLY RELEVANT; IT WAS CORRUPTED DUE TO COMPRESSION AFTER RECEIVING. ALL OTHER AFFECTED STUDIES (9 IMAGES FROM 2008 AND ONE IMAGE FROM 2013) ARE BETWEEN 5-10 YEARS OLD AND THEIR CLINICAL RELEVANCE WAS ASSESSED AS LOW. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN GREAT BRITAIN. | -device migrated -missing maintenance -patient incident | -data back-up problem -data issue -device remains activated | Viewer
| FDA | Event Description: DURING MIGRATION FROM THE SYNGO IMAGING PACS TO THE SYNGO.PLAZA PACS IT WAS IDENTIFIED THAT 297,808 IMAGES BELONGING TO 1,434 STUDIES COULD NOT BE TRANSFERRED TO THE NEW PACS. THEREFORE, THESE IMAGES ARE NOT AVAILABLE IN THE NEW ARCHIVE FOR COMPARISON WITH PRIOR STUDIES. THE AFFECTED DATA IS DATED WITHIN THE TIME FRAME 2003 TO 2017. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE INCIDENT OCCURRED IN (B)(6). | -unable to scroll image -fail to transfer data -difficult to transfer through | -leak -unstable -failure to read input signal | Viewer
| FDA | Event Description: DURING TWENTY-THREE TOTAL HIP PROCEDURES, THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A STEM 3 TO 8 SIZES SMALLER THAN THE STEM PLANNED TO BE IMPLANTED. IN EACH PROCEDURE, THE SURGEON BROACHED TO THE APPROPRIATE SIZE AND IMPLANTED THE PROPER STEM. THERE WERE NO PATIENT INJURIES OR DELAYS IN PROCEDURE. | -inappropriate display size -size out of specification -incorrect device orientation | -use of device issue -improper or incorrect procedure or method -compliance issue | Viewer
| FDA | Event Description: ECHO MACHINE PROBE LOCKED UP TWICE. MACHINE REBOOTED. A XSPROBE FROM ANOTHER MACHINE OBTAINED AND TESTING COMPLETED. FDA SAFETY REPORT ID # (B)(4). | -fail to lock -fail to activate alarm -device alarmed | -charging issue -device contamination with biological material -leak | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). ON (B)(6) 2016, A CUSTOMER EMAILED SUPPORT ALLEGING THAT THE MEASUREMENTS ON A DIGITAL RADIOGRAPHY FLUOROSCOPY (RF) IMAGE WERE INACCURATE. MERGE HEALTHCARE INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT MEASUREMENTS ARE DISPLAYED IN CENTIMETERS, UNTIL CALIBRATION (CAL) IS PERFORMED. ONCE CALIBRATION IS PERFORMED ON THE IMAGE, THE IMAGE WILL DISPLAY IN CAL, NOT CENTIMETERS. THE RF IMAGE TYPE THAT THE CUSTOMER WAS USING IS MEASURING IN CENTIMETERS AFTER THE IMAGE IS CALIBRATED WHEN IT SHOULD BE MEASURED IN CAL. THE CUSTOMER CONFIRMED THAT NO PATIENTS WERE HARMED DUE TO THIS ISSUE. HOWEVER THERE IS A POTENTIAL FOR A MIS-TREATMENT OR MIS-DIAGNOSIS DUE TO UNEXPECTED MEASUREMENT UNITS ON AN IMAGE. (B)(4). | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(4)) THAT THEY WERE UNABLE TO SEND IMAGES TO THE SERVER. THE INVESTIGATION SHOWED THIS WAS AN ISSUE WITH THE PERMISSIONS THE USER SET WHEN THEY BUILT THE DATABASE. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(4)) THEY ARE UNABLE TO SEND IMAGES TO THE SERVER. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(4)) THEY WERE UNABLE TO SEND IMAGES TO THE SERVER. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(4)) THEY WERE UNABLE TO RETRIEVE STUDIES FROM REMOTE DEVICES TO SEND TO THEIR PRIMARY WORKSTATION. THE CUSTOMER REPORTED THAT THEY WERE ABLE TO RESOLVE THE ISSUE. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED (CASE (B)(4) THEY WERE UNABLE TO SEND STUDIES TO THE PACS SERVER. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(4)) THAT AFTER SETTING UP V 4.0.3 OF THE EFILM SERVER, THEY WERE UNABLE TO RECEIVE IMAGES UNLESS EFILM SERVICE IS RESTARTED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. THE INVESTIGATION DISCOVERED THAT IMAGES WERE BEING STORED ON A NETWORK SHARE. MERGE INFORMED THE CUSTOMER THAT THE SYSTEM DOES NOT SUPPORT NETWORK STORAGE. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED ((B)(6)) THEY WERE UNABLE TO RETRIEVE IMAGES TO THE EFILM SYSTEM. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). A CUSTOMER REPORTED (CASE (B)(4)) IMAGES WERE FREEZING ON THE EFILM SYSTEM. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW. | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: EFILM WORKSTATION IS A SOFTWARE APPLICATION THAT IS USED FOR VIEWING MEDICAL IMAGES. EFILM WORKSTATION RECEIVES DIGITAL IMAGES AND DATA FROM VARIOUS SOURCES (INCLUDING BUT NOT LIMITED TO CT, MR, US, RF UNITS, COMPUTED AND DIRECT RADIOGRAPHIC DEVICES, SECONDARY CAPTURE DEVICES, SCANNERS, IMAGING GATEWAYS OR IMAGING SOURCES). ON (B)(6) 2017, A CUSTOMER WAS CONCERNED ABOUT A SOFTWARE DEFECT IDENTIFIED IN EFILM V. 3.4 WHICH THE CUSTOMER WAS CURRENTLY RUNNING. THE DEFECT WAS PART OF RECALL 2183926-09/20/2016-074-C (FDA CLASSIFICATION #: Z-0707-2017) REPORTED TO THE FDA AND HAS THE POTENTIAL TO LEAD TO INACCURATE MEASUREMENTS. THE ISSUES DOES NOT IMPACT ALL IMAGES, ONLY SPECIFIC PROJECTION IMAGES. CR, DX, MG AND XA IMAGES ARE THE MOST COMMON TYPES OF IMAGES AND THERE WERE NO ISSUES RELATED TO THESE. USE OF THE RECALLED VERSION OF THIS PRODUCT MAY RESULT IN INCORRECT MEASUREMENTS HOWEVER, A USER WILL LIKELY BE ABLE TO NOTICE THAT THE MEASUREMENTS ARE NOT AS EXPECTED. THE CUSTOMER WAS INFORMED THE DEFECT WAS CORRECTED IN EFILM VERSION 4.1.1 AND ABOVE. THE CUSTOMER IS UNABLE TO UPGRADE TO CURRENT VERSION OF EFILM 4.2 BECAUSE THE OPERATING SYSTEM THE CUSTOMER IS RUNNING IS NOT COMPATIBLE WITH THE CURRENT VERSION OF EFILM. WHEN THE CUSTOMER UPGRADES TO THE MINIMUM SYSTEM REQUIREMENTS, THE CUSTOMER WILL BE ABLE TO UPDATE THEIR SYSTEM TO THE FIXED VERSION OF SOFTWARE. THE CUSTOMER DID NOT INDICATE ANY INACCURATE MEASUREMENTS WERE OBTAINED WITH THIS PRODUCT. INACCURATE MEASUREMENTS MAY RESULT IN A DELAY IN PATIENT CARE AND/OR CHANGE IN RECOMMENDATION FOR TREATMENT, HOWEVER, THERE IS NO INDICATION THAT THE ISSUE, REPORTED BY THE CUSTOMER, RESULTED IN A DEATH OR SERIOUS INJURY. (B)(4). | -fail to process medical imaging -inadequate image types -inadequate image manipulation | -data issue -problem with software installation -application security issue | Viewer
| FDA | Event Description: ENDOSCOPIC CAMERA FAILED TO TURN ON DURING CASE; UNABLE TO STORE DATA. | -unable to record event -fail to provide video encoded data -fail to record data | -incorrect or inadequate result -data issue -insufficient flow or underinfusion | Viewer
| FDA | Event Description: E-PRESCRIPTIONS WERE SENT USING ARIA EHR SOFTWARE PRODUCED BY (B)(4). THEY WERE NEVER RECEIVED BY (B)(6) PHARMACY ON (B)(6). EHR VENDOR INFORMED, BUT INSISTS THAT PRESCRIPTIONS WERE SENT. PHARMACY INSISTS THEY NEVER RECEIVED THEM. | -inadequate labeling -inadequate identification -missing logo | -component missing -Biocompatibility Issue -Failure to Cycle | Viewer
| FDA | Event Description: ER CALLED AN AMI. PT WAS BROUGHT TO LAB AND LOADED ONTO TABLE. PT WAS CLEANED AND DRAPED PER PROTOCOL. ARTERIAL SHEATH WAS PLACED. WHEN TECH TRIED TO GUIDE C-ARM INTO PLACE TO OBTAIN IMAGES OF ARTERIES THE C-ARM DID NOT WORK. TECH RAN TO BACK AND REBOOTED MACHINE IN ATTEMPT TO CORRECT ERROR. MACHINE CAME BACK UP WITH SAME ERROR MESSAGE. PHYSICIAN DECIDED IT WAS BEST INTEREST OF PT TO CARE FLIGHT TO A HIGHER LEVEL OF CARE FACILITY TO FINISH HEART CATH. BIOMED AND APPROPRIATE PARTIES WERE NOTIFIED. | -inappropriate indication -unclear setup procedures -difficult to move | -Occlusion within Device -Detachment of Device Component -Mechanical Issue | Viewer
| FDA | Event Description: FOLLOWING AN UPGRADE TO CHRS 14.0 A CUSTOMER SITE, THE CUSTOMER REPORTED ISSUES ON (B)(6) 2020 WHERE SOME NEWLY IMPORTED ULTRASOUND IMAGES WEREN'T IN ORDER AND SOME OF THEM WERE CUT OFF AND BLACK. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. | -loss of image -unable to scroll image -incorrect display format | -break -Disconnection -application program issue | Viewer
| FDA | Event Description: FOLLOWING THE UPDATE FROM V3.3 TO V3.5 OF THE DYNACAD SOFTWARE IT WAS NOTED THAT THE KTRANS MAP WAS NOT RENDERED CORRECTLY ON THE DYNACAD CLIENT. INVESTIGATION BY THE INVIVO FIELD TEAM DISCOVERED THAT THE KTRANS MAP WAS DISPLAYED CORRECTLY WHEN THE DYNACAD CLIENT WAS INSTALLED ON THE SERVER MACHINE, BUT NOT WHEN THE DYNACAD CLIENT WAS INSTALLED ON A REMOTE MACHINE. | -fail to set value correctly -incorrect data chart -incorrect placement | -device operates differently than expected -leak -improper device output | Viewer
| FDA | Event Description: FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) RECENTLY COMPLETED A RETROSPECTIVE REVIEW OF COMPLAINT FILES FROM 2013 THROUGH 2015 AND HAS DETERMINED THAT THIS COMPLAINT MAY MEET THE THRESHOLD FOR REPORTING BECAUSE IMAGES WERE REPORTED AS LOST. FMSU IS THEREFORE REPORTING IT AT THIS TIME. THE SITE (B)(6) REPORTS THAT THEY HAVE LOST TWO IMAGES DURING OPERATION OF A CLUSTERED (B)(4) CONSOLE RUNNING V6.2. THESE IMAGES WERE TAKEN BETWEEN 1:00 AND 1:10 PM ON (B)(6) 2013. THESE IMAGES HAD BECOME INACCESSIBLE AND THE (B)(4) WAS THEN REBOOTED. REBOOTING DELETES THE IMAGES FROM THE DATABASE. ONCE THE REBOOT PROCESS WAS COMPLETE THE IMAGES WERE NO LONGER PRESENT ON THE (B)(4) CONSOLE. THE FACILITY DID NOT PROVIDE ANY INFORMATION INDICATING THIS AS AN ADVERSE EVENT, I.E., THAT THE PATIENT WAS RE-IMAGED. FUJIFILM THEREFORE IS CONSIDERING THIS AS A POSSIBLE MALFUNCTION. | -fail to document feedback -lack of problem investigation -fail to assess feedback | -Inflation Issue -Pressure Issue -contamination during use | Viewer
| FDA | Event Description: FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) RECENTLY COMPLETED A RETROSPECTIVE REVIEW OF COMPLAINT FILES FROM 2013 THROUGH 2015 AND HAS DETERMINED THAT THIS COMPLAINT MEETS THE THRESHOLD FOR REPORTING. FMSU IS THEREFORE REPORTING IT AT THIS TIME. AN EXAM AT (B)(6) HOSPITAL IN (B)(6), WAS ENTERED MANUALLY ON THE MODALITY AND THE ACCESSION NUMBER WAS ENTERED INCORRECTLY BY THE TECHNOLOGIST. AS A RESULT, THE EXAM WAS SENT FROM THE MODALITY TO THE PATIENT EXAM FOLDER IN SYNAPSE MATCHING THE ACCESSION NUMBER THAT WAS ENTERED. THIS PATIENT FOLDER HOWEVER WAS NOT FOR THE CORRECT PATIENT. SYNAPSE HAS FUNCTIONALITY DESIGNED TO MITIGATE THIS TYPE OF EVENT. WHEN STUDIES ARE RECEIVED FROM A MODALITY, SYNAPSE PULLS THE PATIENT INFORMATION FROM THE DICOM HEADER. SYNAPSE THEN TRIES TO MATCH THE PATIENT INFORMATION, E.G., THE PATIENT ID, PATIENT NAME, ETC. WITH THE ACCESSION NUMBER. WHEN SYNAPSE FINDS A POSSIBLE MISMATCH IT CREATES AN ANOMALY ALERT. ANYTIME THE EXAM IS OPENED AN ANOMALY ALERT WILL OPEN ON THE USER'S MONITOR SCREEN INDICATING THAT AN ANOMALY EXISTS FOR THIS EXAM. THE ANOMALY ALERT IDENTIFIED IN THE EVENT ABOVE WOULD INDICATE "MISMATCH PATIENT NAME." A SCREENSHOT OF AN ANOMALY ALERT IS PROVIDED FOR YOUR REFERENCE (THIS IS NOT FROM THE PARTICULAR EVENT). THE USER HAS THE ABILITY TO SELECT A CHECKBOX NOT TO SHOW ANOMALIES WHILE IN THEIR CURRENT SESSION WHICH WAS DONE IN THIS INSTANCE. AFTER BEING NOTIFIED OF THE ANOMALY AND TURNING OFF THE WARNING THE RADIOLOGIST STILL WENT AHEAD AND DICTATED THE RESULTS OF PATIENT A'S EXAM TO THE PATIENT B'S EXAM FOLDER; AS A RESULT THERE WAS AN ALLEGED MISDIAGNOSIS OF PATIENT B. FMSU IS NOT AWARE OF WHETHER TREATMENT WAS PROVIDED ON THE BASIS OF THE MISDIAGNOSIS OR OF THE OUTCOME. EXAMS WITH ANOMALIES ARE ALSO COPIED TO AN ANOMALY FOLDER. THE PACS ADMINISTRATOR SHOULD BE REVIEWING THE ANOMALY FOLDER PERIODICALLY TO RESOLVE ANOMALIES. OR, AS IN THIS SITUATION, THE RADIOLOGIST SHOULD HAVE ALERTED THE PACS ADMIN THAT AN ANOMALY NEEDED TO BE RESOLVED. A COPY OF THE PAGE FROM THE OPERATION MANUAL SHOWING THE ADMIN FOLDER AND ANOMALY SUB-FOLDER HAS BEEN ATTACHED. THE SPECIFIC ACTION OF MANUALLY ENTERING AN INCORRECT ACCESSION NUMBER IS CLEARLY USER ERROR AND SYNAPSE CANNOT PREVENT SUCH AN ERROR FROM OCCURING. SYNAPSE OPERATED EXACTLY AS DESIGNED. | -fail to document feedback -fail to assess feedback -inadequate document | -Inflation Issue -Break -Pressure Issue | Viewer
| FDA | Event Description: FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) RECENTLY COMPLETED A RETROSPECTIVE REVIEW OF COMPLAINT FILES FROM 2013 THROUGH 2015 AND HAS DETERMINED THAT THIS COMPLAINT MEETS THE THRESHOLD FOR REPORTING. FMSU IS THEREFORE REPORTING IT AT THIS TIME. THE DATE OF OCCURRENCE IS UNKNOWN, HOWEVER ONE CAN PRESUME IT WAS CLOSE TO (B)(6) 2014, THE DATE FUJIFILM WAS NOTIFIED BY THE USER.. ORDER ENTERED IN THE (B)(4) WAS SENT TO THE (B)(4) AND THE (B)(4) CONSOLE WORKSTATION WORKLIST. TECHNOLOGIST IMAGED PATIENT AND SENT EXAM IMAGES TO (B)(4). TECHNOLOGIST WENT BACK INTO THE (B)(4) CONSOLE WORKSTATION AND CHANGED THE ACCESSION NUMBER OF THE ACQUIRED EXAM AND TRIED TO RESEND THE IMAGES TO (B)(4). THE IMAGES WITH THE NEW ACCESSION NUMBER WERE NOT ACCEPTED BY THE (B)(4) BECAUSE THE STUDY ID WAS ALREADY ATTACHED TO THE ACCESSION NUMBER ORIGINALLY ASSIGNED TO THE PATIENT (AS ENTERED IN THE (B)(4)). THEREFORE THE IMAGES BECAME STUCK IN THE QUEUE OF THE (B)(4) CONSOLE BECAUSE THEY WERE REJECTED BY THE (B)(4). TECHNOLOGIST MAY HAVE REINITIALIZED EITHER THE (B)(4) CONSOLE OR THE OUTPUT QUEUE. TECHNOLOGIST COULD NOT FIND THE IMAGES AND CONCLUDED THE IMAGES HAD BEEN ERASED FROM THE LOCAL DATABASE. THE PATIENT WAS REIMAGED AND THE TECHNOLOGIST REPORTED THE ORIGINAL IMAGES AS LOST. HOWEVER, THE IMAGES WERE ALREADY IN THE (B)(4) SO THE TECHNOLOGIST DID NOT NEED TO REIMAGE THE PATIENT. | -fail to document feedback -fail to assess feedback -inadequate document | -Inflation Issue -Break -Pressure Issue | Viewer
| FDA | Event Description: FUSION PACS IS INTENDED FOR THE TRANSMISSION, ARCHIVING AND REVIEW OF RADIOLOGICAL IMAGES. ON (B)(6) 2017, A CUSTOMER CALLED MERGE TECHNICAL SUPPORT ALLEGING THAT A HARD DRIVE ON THEIR FUSION PACS SERVER FAILED AND THE FUSION PACS SOFTWARE NEEDED TO BE REINSTALLED. THE CUSTOMER'S IMAGES AND FUSION PACS DATABASE WERE STORED ON A DIFFERENT HARD DRIVE THEREFORE, THERE WAS NO LOSS OF PATIENT DATA. HOWEVER, THE CUSTOMER WAS UNABLE TO USE THEIR FUSION PACS FOR SEVEN DAYS WHILE MERGE SUPPORT WORKED TO INSTALL THE SOFTWARE. DUE TO THE LOSS OF ACCESS, THERE IS A POTENTIAL THAT A DELAY COULD LEAD TO HARM HOWEVER, THERE IS NO INDICATION THAT ANY PATIENTS WERE HARMED DUE TO THIS ISSUE. (B)(4).Event Description: FURTHER INVESTIGATION WAS PERFORMED AND IT WAS FOUND THAT CUSTOMER'S FUSION PACS WAS DOWN FOR APPROXIMATELY SEVEN DAYS. DURING THAT DOWNTIME, THE FACILITY HAD TO CLOSE DOWN FOR TWO DAYS AND SIX PROCEDURES HAD TO BE RESCHEDULED. RADIOLOGISTS DID READ SOME EXAMS FROM THE MODALITY AS A WORKAROUND DURING THE DOWNTIME. NO PATIENTS WERE HARMED DUE TO THE DOWNTIME. | -ram broken or damaged -fail to display data -driver heavy | -use of device issue -missing test results -infusion or flow issue | Viewer
| FDA | Event Description: GE PACS UNIVERSAL VIEWER SOFTWARE IS SUBSTANTIALLY WORSE THAN THE PREVIOUS CLIENT CALLED CENTRICITY. THIS PT HAD A BREAST SONOGRAM AND MAMMOGRAM. THE NEW VIEWER CANNOT SHOW MAMMOGRAMS AND SONOGRAMS AT THE SAME TIME. THIS CREATES INTERPRETIVE DIFFICULTIES THAT DIDN'T EXIST WITH THE OLDER CLIENT CALLED CENTRICITY. | -unable to scroll image -inadequate message display -inappropriate user | -improper or incorrect procedure or method -failure to deploy -failure to sterilize | Viewer
| FDA | Event Description: HAD MAMMOGRAM AT (B)(6) HOSPITAL AND BECAME VERY ILL WITHIN 3 WEEKS. I HAD SILICONE BREAST IMPLANTS AND HAD NO IDEA THEY BOTH GOT RUPTURED. I STARTED IN (B)(6) GOING TO DOCTOR. UNABLE TO WALK. SEVERE PAIN IN FEET. MANY TEST TO FOLLOW. WAS SENT TO PODIATRIST AND RHEUMATOLOGIST. WAS IMMEDIATELY PUT ON MELOXICAM FOR INFLAMMATION. TESTED POSITIVE ANA FOR POSSIBLE LUPUS. HAD MANY BLOOD TEST AND I HAVE ALL OF THE RESULTS. AS OF(B)(6) I NO LONGER WAS ABLE TO WORK. DIAGNOSED WITH MID FOOT OSTEOARTHRITIS AND SURGERY WAS SUGGESTED. I WAS ALSO DIAGNOSED WITH LOWER BACK OSTEOARTHRITIS AND IN (B)(6) WITH FIBRO. I APPLIED FOR SOCIAL SECURITY DISABILITY AND AS APPROVED. I DID NOT CORRELATE THE BREAST IMPLANTS TILL APPROX (B)(6). I REQUESTED AN MRI AND IT WAS SHOWN BOTH IMPLANTS WERE RUPTURED. THEN UPON INVESTIGATING, I FOUND THAT ALL OF MY SYMPTOMS WERE RELATED TO SILICONE POISONING DUE TO RUPTURED IMPLANTS. I GOT SO SICK SO FAST AND DID NOT KNOW WHY. SINCE THE REMOVAL OF THE IMPLANTS, I HAVE HAD 2 ADD'L RECONSTRUCTION SURGERIES INCLUDING A LATISSIMUS FLAP. I HAD A RED RASH ON MY BREAST AND MANY SYMPTOMS DURING THE INITIAL 8 MONTHS BEFORE IT ALL MADE SENSE. I HAVE PICTURES, NOTES AND MANY BILLS. THE HOSPITAL HAS DENIED THE MAMMOGRAM CAUSED THE RUPTURE BUT BOTH AT THE SAME TIME IS SO OBVIOUS. THE IMPLANTS WERE SILICONE MADE BY MENTOR 600CC. | -part ruptured -silicone cracked -silicone stuck | -Break -moisture or humidity problem -Crack | Viewer
| FDA | Event Description: HOUNSFIELD MEASUREMENTS ARE USED TO MEASURE TISSUE DENSITY AND MAY BE USED AS AN INPUT TO CLINICAL DECISION MAKING. AFTER A PERIOD OF TIME RUNNING FUSION WORKSTATION 3.0.2 P3 AND LATER, THE HOUNSFIELD MEASUREMENT TOOL (AND REPORTEDLY OTHER MEASUREMENT TOOLS) WILL REPORT INCORRECT VALUES. IF THE EXAM IS CLOSED AND THE USER EXITS FUSION WORKSTATION AND THEN LOGS BACK IN AND REOPENS THE EXAM, THE MEASUREMENT WILL THEN BE REPORTED CORRECTLY. THE CUSTOMER REPORTED THAT MEASURING IN MATRIX IS NOT ALWAYS ACCURATE ON TWO DIFFERENT WORKSTATIONS. CLIENT DOES NOT CONSIDER IT A MAJOR ISSUE BUT DOES RESULT IN THE DOCTORS HAVING TO REBOOT THESE WORKSTATIONS. ONCE THIS IS DONE THEN THEY FUNCTION PROPERLY. | -inadequate measurement unit -incorrect measurement -measurement low | -incorrect measurement -chemical issue -incorrect or inadequate result | Viewer
| FDA | Event Description: ICONNECT ACCESS DISPLAYED INCORRECT PRIORS REPORT(S) IN THE VIEWPORT AREA ONLY WHEN MORE THAN ONE PRIOR STUDY (2 OR MORE) WAS VIEWED. NO INJURIES WERE REPORTED. ON THE STUDY HISTORY TAB WHERE THE PRIOR STUDIES ARE LISTED, THE CORRECT PRIORS REPORT WAS ALWAYS DISPLAYED ON CLICKING THE STUDY NAME. THE ISSUE HAPPENED ONLY IN THE VIEWER. THERE IS ONLY ONE COMPLAINT REPORTED SO FAR, OVER THE DEPLOYMENT PERIOD FROM (B)(6) 2012 (ICONNECT ACCESS 3.0 RELEASE) UNTIL THE PRESENT. THE LONE COMPLAINT WAS NOT REPORTED FROM AN ACTUAL CLINICAL USER. IT WAS ENCOUNTERED BY A PACS ADMINISTRATOR AT A MERGE CUSTOMER SITE WHILE EVALUATING THE SYSTEM IN PREPARATION FOR AN UPCOMING GO-LIVE. | -incorrect display value -incorrect message display -incorrect display format | -failure to read input signal -incorrect display -no display or display failure | Viewer
| FDA | Event Description: ICONNECT ACCESS IS A ZERO-FOOTPRINT, ZERO-DOWNLOAD DICOM AND XDS VIEWER THAT PROVIDES DIAGNOSTIC QUALITY IMAGE REVIEW FOR REFERRING PHYSICIANS. BY PROVIDING EASY ACCESS TO IMAGES AND CONTENT, ICONNECT ACCESS BENEFITS PROVIDERS BY PROVIDING QUALITY SERVICE LEVELS TO REFERRING PHYSICIANS. CURRENTLY THE PATIENT COMPARISON STRATEGY USED IN THE IMAGE SHARING SOLUTION IS PID (PATIENT ID)+IPID (ISSUER OF PATIENT ID). THE RESULT OF THIS IS THAT WHEN TWO SEPARATE PATIENTS WHO HAVE THE SAME PID ARE MERGED IN THE HOLDING PEN, SPECIFICALLY THE DEMOGRAPHICS ON THE FIRST EXAM UPLOADED ARE UPDATED TO THE DEMOGRAPHICS ON THE SECOND EXAM UPLOADED. SCENARIO: A USER WITH ADMINISTRATIVE SYSTEM PRIVILEGES, (B)(6) UPLOADS EXAM 1 FOR PATIENT (B)(6) (WITH ADDITIONAL DEMOGRAPHICS), (B)(6). A USER WITH THE PROPER CREDENTIALS CAN SEARCH FOR (B)(6) AND FIND EXAM 1 WITH THE DEMOGRAPHICS THAT WERE ASSOCIATED WITH THE EXAM WHEN (B)(6) UPLOADED THE EXAM. (B)(6) NOW UPLOADS EXAM 2 FOR (B)(6) (WITH ADDITIONAL DEMOGRAPHICS), (B)(6). DESIRED BEHAVIOR: (B)(6) CAN SEARCH FOR (B)(6) AND FIND EXAM 2 AND SEARCH FOR (B)(6) AND FIND EXAM 1. IT IS UNDERSTOOD AND ACCEPTED IN THE INDUSTRY THAT VARIATIONS IN THE SEARCH TERMS WILL IMPACT THE ACCURACY AND COMPLETENESS OF THE SEARCH, E.G. SEARCHING FOR (B)(6) WOULD RETURN EXAM 1 WHILE SEARCHING FOR (B)(6) WOULD NOT FIND EXAM 1. ACTUAL BEHAVIOR: IF (B)(6) SEARCHES FOR (B)(6), THE SYSTEM WILL RETURN BOTH EXAM 1 AND EXAM 2. AND IF (B)(6) SEARCHES FOR (B)(6), THE SYSTEM WILL RETURN NO RESULTS. NOTE THAT THIS ISSUE ONLY ARISES WHEN TWO (OR MORE) EXAMS HAVE BEEN IMPORTED THAT HAVE THE SAME IPID / PID COMBINATIONS, AND THE SECOND EXAM IS UPLOADED BEFORE THE FIRST EXAM HAS BEEN RECONCILED AGAINST A PATIENT IN THE ARCHIVE OR AGAINST AN ORDER. BOTH CUSTOMERS, (B)(6), WHO ARE AWARE OF THIS BEHAVIOR HAVE STATED THAT IT IS NOT ACCEPTABLE AND THAT THEY REQUIRE A CHANGE TO BE MADE TO THE SYSTEM TO OBTAIN THE DESIRED BEHAVIOR. THIS ISSUE WAS DISCOVERED FIRST BY (B)(6) WHICH COMMUNICATED THIS ISSUE TO (B)(6). WHILE DISCUSSIONS AROUND THIS SOLUTION ARE ONGOING WITH (B)(6), (B)(6) HAS INDICATED THAT THEY VIEW THIS AS POTENTIALLY IMPACTING PATIENT CARE THAT NEEDS TO BE CORRECTED AS SOON AS POSSIBLE. | -inadequate image types -lack of review function -fail to open images | -human-device interface issue -poor quality image -inadequate user interface | Viewer
| FDA | Event Description: ICONNECT ACCESS IS A ZERO-FOOTPRINT, ZERO-DOWNLOAD DICOM AND XDS VIEWER THAT PROVIDES DIAGNOSTIC QUALITY IMAGE REVIEW FOR REFERRING PHYSICIANS. BY PROVIDING EASY ACCESS TO IMAGES AND CONTENT, ICONNECT ACCESS ENHANCES STICKINESS FOR PROVIDERS BY PROVIDING QUALITY SERVICE LEVELS TO REFERRING PHYSICIANS. CUSTOMER USES ICONNECT ACCESS TO VIEW IMAGES FROM THEIR EMR VIA AN INTEGRATION LINK. THE IMAGES ARE THEN DISPLAYED IN ICONNECT ACCESS. ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHEN LAUNCHING IMAGES FOR PATIENT A, PATIENT B'S IMAGES WERE DISPLAYED. THE RADIOLOGIST NOTICED THAT THE STUDY DESCRIPTION FOR PATIENT A'S EXAM DID NOT MATCH THE IMAGES THAT WERE DISPLAYED. DUE TO ICONNECT ACCESS NOT PRESENTING PATIENT DATA AS EXPECTED THERE IS A POTENTIAL FOR EITHER A DELAY IN TREATMENT OR AN INCORRECT TREATMENT OR DIAGNOSIS. THE ISSUE WAS DETECTABLE BY USER. THERE WAS NO INDICATION OF PATIENT HARM OR MISDIAGNOSIS DUE TO THIS ISSUE. (B)(4). | -inadequate image types -lack of review function -fail to open images | -human-device interface issue -poor quality image -inadequate user interface | Viewer
| FDA | Event Description: ICONNECT ACCESS IS A ZERO-FOOTPRINT, ZERO-DOWNLOAD DICOM AND XDS VIEWER THAT PROVIDES DIAGNOSTIC QUALITY IMAGE REVIEW FOR REFERRING PHYSICIANS. BY PROVIDING EASY ACCESS TO IMAGES AND CONTENT, ICONNECT ACCESS ENHANCES STICKINESS FOR PROVIDERS BY PROVIDING QUALITY SERVICE LEVELS TO REFERRING PHYSICIANS. CURRENTLY THE PATIENT COMPARISON STRATEGY USED IN THE IMAGE SHARING SOLUTION IS PID+IPID. THE RESULT OF THIS IS THAT WHEN TWO SEPARATE PATIENTS WHO HAVE THE SAME PID ARE MERGED IN THE HOLDING PEN, SPECIFICALLY THE DEMOGRAPHICS ON THE FIRST EXAM UPLOADED ARE UPDATED TO THE DEMOGRAPHICS ON THE SECOND EXAM UPLOADED. SCENARIO A USER WITH ADMINISTRATIVE PRIVILEGES, BILL SMITH UPLOADS EXAM 1 FOR PATIENT TOM JONES (WITH ADDITIONAL DEMOGRAPHICS), IPID A, PID 1. A USER WITH THE PROPER CREDENTIALS CAN SEARCH FOR TOM JONES AND FIND EXAM 1 WITH THE DEMOGRAPHICS THAT WERE ASSOCIATED WITH THE EXAM WHEN BILL SMITH UPLOADED THE EXAM. BILL SMITH NOW UPLOADS EXAM 2 FOR SARA WILLIAMS (WITH ADDITIONAL DEMOGRAPHICS), IPID A, PID 1. DESIRED BEHAVIOR: BILL SMITH CAN SEARCH FOR SARA WILLIAMS AND FIND EXAM 2 AND SEARCH FOR TOM JONES AND FIND EXAM 1. IT IS UNDERSTOOD AND ACCEPTED IN THE INDUSTRY THAT VARIATIONS IN THE SEARCH TERMS WILL IMPACT THE ACCURACY AND COMPLETENESS OF THE SEARCH, E.G. SEARCHING FOR TOM WOULD RETURN EXAM 1 WHILE SEARCHING FOR THOMAS WOULD NOT FIND EXAM 1. ACTUAL BEHAVIOR: IF BILL SMITH SEARCHES FOR SARA WILLIAMS, THE SYSTEM WILL RETURN BOTH EXAM 1 AND EXAM 2. AND IF BILL SMITH SEARCHES FOR TOM JONES, THE SYSTEM WILL RETURN NO RESULTS. NOTE THAT THIS ISSUE ONLY ARISES WHEN TWO (OR MORE) EXAMS HAVE BEEN IMPORTED THAT HAVE THE SAME IPID / PID COMBINATIONS, AND THE SECOND EXAM IS UPLOADED BEFORE THE FIRST EXAM HAS BEEN RECONCILED AGAINST A PATIENT IN THE ARCHIVE OR AGAINST AN ORDER. BOTH CUSTOMERS, UNIVERSITY OF MIAMI AND UC SAN DIEGO, WHO ARE AWARE OF THIS BEHAVIOR HAVE STATED THAT IT IS NOT ACCEPTABLE AND THAT THEY REQUIRE A CHANGE TO BE MADE TO THE SYSTEM TO OBTAIN THE DESIRED BEHAVIOR. GIVEN THEIR RESPONSE, IT CAN BE EXPECTED THAT OTHER CUSTOMERS WOULD HAVE A SIMILAR RESPONSE IF THEY KNEW OF THE ISSUE. THIS ISSUE WAS DISCOVERED FIRST BY UNIVERSITY OF MIAMI WHICH COMMUNICATED THIS ISSUE TO UC SAN DIEGO. WHILE DISCUSSIONS AROUND THIS SOLUTION ARE ONGOING WITH UM, UCSD HAS INDICATED THAT THEY VIEW THIS AS IMPACTING PATIENT CARE THAT NEEDS TO BE CORRECTED AS SOON AS POSSIBLE. | -inadequate image types -lack of review function -fail to open images | -human-device interface issue -poor quality image -inadequate user interface | Viewer
| FDA | Event Description: ICONNECT ENTERPRISE ARCHIVE A VENDOR NEUTRAL ARCHIVE FOR STORAGE AND COMMUNICATIONS OF MEDICAL IMAGES AND DATA. CUSTOMER REPORTED THAT THEIR PREFETCHER IS NOT WORKING AS EXPECTED. DURING A READ OF A STUDY ((B)(6) 2015: CT CHEST SCAN WITHOUT CONTRAST), THE RADIOLOGIST DID NOT HAVE A RELEVANT PRIOR ((B)(6) 2013: CT CHEST SCAN WITH CONTRAST) THAT HE WAS EXPECTING. PACS ADMINISTRATOR WAS ABLE TO MANUALLY RETRIEVE THE PRIOR NOT AVAILABLE LOCALLY. THE SITE REVIEWED MERGE PACS AND FOUND A THIRD RELEVANT PRIOR THAT WAS NOT RETRIEVED ((B)(6) 2012: CT SCAN). BOTH STUDIES WERE NOT RETRIEVED FALL UNDER THE "LAST 3 IDENTICAL MODALITY" RULE. (B)(4). | -fail to store medical imaging -fail to enter patient data -loss of communication | -over-sensing -therapeutic or diagnostic output failure -Shelf Life Exceeded | Viewer
| FDA | Event Description: IMAGES FROM PHILIPS PET MODALITY DISPLAY INCORRECT SUV VALUE (STANDARDIZED UPTAKE VALUE), WHEN THE IMAGES ARE LOADED TO THE EMBEDDED TRUED APPLICATION FROM THE SYNGO IMAGING PATIENT LIST. IN THE INITIAL CASE, THE CORRECT VALUE IS '3', WHICH IS VALIDATED ON THE PHILIPS PET MODALITY WORKSTATION, HOWEVER, IN TRUED THE VALUE SHOWS AS '15'. THE INVESTIGATION IS ON-GOING. THE REPORTED INCIDENT OCCURRED IN (B)(6). | -incorrect display value -image orientation incorrect -incorrect display format | -incorrect display -use of device issue -device alarm system issue | Viewer
| FDA | Event Description: IMAGES WERE SENT FROM DAR - 8000I TO PACS AFTER FLUOROSCOPIC EXAMINATION, BUT ONLY ONE WAS STORED IN PACS. EVEN WHEN CHECKING THE DB OF THE DAR - 8000I DEVICE, THERE WAS ONLY ONE IMAGE. THEREFORE, RE-EXAMINATION WAS CARRIED OUT. | -inadequate image manipulation -inadequate device information -inadequate labeling | -fluid leak -compliance issue -fitting problem | Viewer
| FDA | Event Description: IN RESPONSE TO THE FAILURE WITH THE FACILITY'S OTV-S7PRO AND OTV-S7PROH-HD-E12 ON (B)(6), THEY HAD BEEN SUBSTITUTED FOR THE OLYMPUS-OWNED EQUIPMENT. DURING AN OPEN SURGERY, THE FACILITY USED THE SUBSTITUTED OTV-S7PRO, WHICH WAS THE SUBJECT DEVICE, AND THE SUBSTITUTED OTV-S7PROH-HD-E12 FOR OBSERVATION OF THE OPERATIONAL FIELD AND RECORDING THE IMAGES. HOWEVER, THE SIMILAR PHENOMENON OCCURRED AGAIN AS THE IMAGE MALFUNCTION TURNING GREEN PERIODICALLY WHICH WAS CAUSED BY THE FACILITY'S EQUIPMENT. THE PROCEDURE WAS COMPLETED BY REPLACING THE SUBJECT OTV-S7PRO TO CV-190. NO HEALTH DAMAGE TO THE PATIENT WAS REPORTED OTHER THAN REPLACING THE DEVICE. | -incorrect electrode placement -fail to reuse -lack of implementation process | -fitting problem -disconnection -incorrect error code | Viewer
| FDA | Event Description: IN THE PACS SUITE FOR MAMMOGRAPHY IMAGES, IN THE DIAGNOSTIC VIEW (LM VIEW)THE PACS WILL READ AND LABEL THE LEFT 3D TOMO INDICATOR BAR AS REVERSED IN ORIENTATION ON PACS ON BREAST TOMOS (I.E. LATERAL BREAST REGISTERS AS "M" FOR MEDIAL, AND VICE VERSA) FOR THE HOLOGIC AFFIRM PRONE BIOPSY SYSTEM AND THE GE ESSENTIAL WITH SENOCLAIRE. THIS HAS BEEN RECOGNIZED BY THE RADIOLOGISTS AND THEY ARE AWARE OF ITS OCCURRENCE. PHILIPS WAS FIRST CONTACTED ABOUT THIS OBSERVATION IN (B)(6) 2020. | -fail to provide indicator -fail to list 3d models -fail to label symbol | -no display or display failure -application program issue -data back-up problem | Viewer
| FDA | Event Description: IN THIS CASE A DENTIST UTILIZED A WAFER TO SUPPORT THE CORRECTION OF THE PATIENT'S OCCLUSAL SITUATION SURGICALLY. HOWEVER HE DID NOT ORDER THE CORRECT TYPE OF WAFER FOR THE PLANNED SURGERY. THE SURGERY RESULTED IN AN OPEN BITE AND THE PATIENT IS REFUSING ANOTHER SURGICAL INTERVENTION. | -fail to save surgical plan -inappropriate direct visualization surgical technique -part misoriented | -use of device issue -improper or incorrect procedure or method -medication error | Viewer
| FDA | Event Description: IN THIS EVENT IT WAS REPORTED THAT A SURGIGUIDE WAS USED FOR IMPLANT PLACEMENT IN REGION #20. THE IMPLANT WAS PLACED TOO DEEP AND TOO CLOSE TO THE ROOT OF THE NEIGHBOURING TOOTH. THE SURGERY WAS COMPLETED, NO ADDITIONAL TREATMENT WAS NECESSARY. | -implant placed at improper location -incorrect device orientation -inappropriate circuit configuration | -fluid leak -loose or intermittent connection -compliance issue | Viewer
| FDA | Event Description: IN THIS EVENT IT WAS REPORTED THAT A SURGIGUIDE WAS USED FOR IMPLANT PLACEMENT IN REGION #7. THE IMPLANT WAS PLACED TOO MUCH TO THE DISTAL SIDE AND WAS HITTING THE ROOTS OF PERMANENT TEETH. THE SURGERY WAS COMPLETED, NO ADDITIONAL TREATMENT WAS NECESSARY. | -implant placed at improper location -device damaged -part splitted | -fluid leak -leak -electrical issue | Viewer
| FDA | Event Description: IN THIS EVENT IT WAS REPORTED THAT A SURGIGUIDE WOULD NOT FIT A PATIENT PROPERLY. THE DENTIST STATED IT WAS OFF APPROXIMATELY 3MM, SHIFTED TO THE BUCCAL. THE PROBLEM WAS NOTICED DURING SURGERY. THE PATIENT WAS NOT SEDATED. BONE GRAFT MATERIAL AND A MEMBRANE WERE PLACED AT THE SURGICAL SITE. THE IMPLANTS WERE NOT PLACED, BECAUSE THE OSTEOTOMY WAS TOO FAR BUCCAL, THERE WOULD HAVE BEEN NO PRIMARY STABILITY. AN ADDITIONAL SURGERY IS NOW REQUIRED. | -inappropriate fit -part mismatched -increased opacity | -fluid leak -device operates differently than expected -fitting problem | Viewer
| FDA | Event Description: IN THIS EVENT, IT WAS REPORTED THAT A DENTIST HAD ORDERED A MUCOSA-SUPPORTED SURGIGUIDE. HOWEVER, THE GUIDE HE RECEIVED WAS ALSO SUPPORTING ON THE TWO REMAINING DISTAL TEETH, CAUSING A GAP BETWEEN THE GUIDE AND THE MUCOSA SURFACE. THE GUIDE WAS UNSTABLE AS THERE WAS A SPACE OF 1-2MM BETWEEN THE FLANGE AND MUCOSA BOTH BETWEEN THE BUCCAL AND THE PALATAL SIDES. ALL THIS WAS NOTICED WHEN THE PATIENT WAS IN THE CHAIR. THE DOCTOR TRIED TO REMOVE THE PORTION OF THE GUIDE THAT SUPPORTED ON THE TEETH BUT THIS COMPROMISED THE FIT OF THE GUIDE. HE STILL TRIED TO USE IT BY DRILLING FOR IMPLANT ONE (AT TOOTH POSITION 16) BUT HE ENTERED THE SINUS, WHICH WAS NOT PLANNED. | -teeth shift -inadequate guide -inappropriate method | -Detachment of Device Component -Decrease in Pressure -split | Viewer
| FDA | Event Description: IN THIS EVENT, IT WAS REPORTED THAT A SURGIGUIDE WAS ROCKING AND NOT STABLE. CHANGING AND TRIMMING THE GUIDE DID NOT FIX THE ISSUE ENTIRELY. THE DOCTOR THEN MADE A FLAP TO POSITION THE GUIDE ON THE BONE, BUT THE GUIDE DIDN'T FIT ON THE BONE. NO DRILLING WAS DONE, HOWEVER A SECOND SURGERY IS NEEDED. | -device handling inappropriately -unstable interface -device malfunction | -Inflation Issue -Device Operates Differently than Expected -failure to cycle | Viewer
| FDA | Event Description: INDICATIONS FOR USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. MERGE CARDIO IS SOFTWARE COMPRISED OF MODULES THAT PERFORM UNDER STANDARD "OFF THE SHELF" PERSONAL COMPUTERS AND SERVERS RUNNING THE (B)(4). MERGE CARDIO IS IMAGE DATA STORAGE AND DISPLAY SOFTWARE THAT ACCEPTS DICOM (DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE) IMAGE DATA FILES FROM MULTIPLE MODALITIES. IT ACCEPTS TEXT DATA USING OTHER STANDARDS-BASED FORMATS INCLUDING BUT NOT LIMITED TO HL7 AND XML. MERGE CARDIO IS AN INTERNET/INTRANET NETWORK SYSTEM THAT IS DESIGNED FOR SMALL AND LARGE, MULTI-USER ENVIRONMENTS. THE MERGE CARDIO NETWORK STRUCTURE (INCLUDING SERVER AND WORKSTATIONS) PROVIDES FOR THE SYSTEM'S DATABASE MANAGEMENT, STORAGE, PRINTING, AND ALL DICOM/HL-7 INTERFACE SERVICES. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THEY WERE MISSING MEASUREMENTS AND SOME MEASUREMENTS WERE INCORRECT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE (B)(4), COMPLAINT-(B)(4). | -fail to access clinical history -inadequate clinical history data -lack of documentation indication | -data issue -use of device issue -data back-up problem | Viewer
| FDA | Event Description: INDICATIONS FOR USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. MERGE CARDIO IS SOFTWARE COMPRISED OF MODULES THAT PERFORM UNDER STANDARD "OFF THE SHELF" PERSONAL COMPUTERS AND SERVERS RUNNING THE MICROSOFT WINDOWS 2000/2003/XP OPERATING SYSTEMS. MERGE CARDIO IS IMAGE DATA STORAGE AND DISPLAY SOFTWARE THAT ACCEPTS DICOM (DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE) IMAGE DATA FILES FROM MULTIPLE MODALITIES. IT ACCEPTS TEXT DATA USING OTHER STANDARDS-BASED FORMATS INCLUDING BUT NOT LIMITED TO HL7 AND XML. MERGE CARDIO IS AN INTERNET/INTRANET NETWORK SYSTEM THAT IS DESIGNED FOR SMALL AND LARGE, MULTI-USER ENVIRONMENTS. THE MERGE CARDIO NETWORK STRUCTURE (INCLUDING SERVER AND WORKSTATIONS) PROVIDES FOR THE SYSTEM'S DATABASE MANAGEMENT, STORAGE, PRINTING, AND ALL DICOM/HL-7 INTERFACE SERVICES. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MULTIPLE DIAGNOSTIC MEASUREMENTS WERE POPULATING THE CLINICAL REPORT INCORRECTLY. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE (B)(4), COMPLAINT-(B)(4). | -fail to access clinical history -inadequate clinical history data -lack of documentation indication | -data issue -use of device issue -data back-up problem | Viewer
| FDA | Event Description: INDICATIONS FOR USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. MERGE CARDIO IS SOFTWARE COMPRISED OF MODULES THAT PERFORM UNDER STANDARD "OFF THE SHELF" PERSONAL COMPUTERS AND SERVERS RUNNING THE MICROSOFT WINDOWS 2000/2003/XP OPERATING SYSTEMS. MERGE CARDIO IS IMAGE DATA STORAGE AND DISPLAY SOFTWARE THAT ACCEPTS DICOM (DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE) IMAGE DATA FILES FROM MULTIPLE MODALITIES. IT ACCEPTS TEXT DATA USING OTHER STANDARDS-BASED FORMATS INCLUDING BUT NOT LIMITED TO HL7 AND XML. MERGE CARDIO IS AN INTERNET/INTRANET NETWORK SYSTEM THAT IS DESIGNED FOR SMALL AND LARGE, MULTI-USER ENVIRONMENTS. THE MERGE CARDIO NETWORK STRUCTURE (INCLUDING SERVER AND WORKSTATIONS) PROVIDES FOR THE SYSTEM'S DATABASE MANAGEMENT, STORAGE, PRINTING, AND ALL DICOM/HL-7 INTERFACE SERVICES. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE MEASUREMENT FOR STRAIN IS NOT CORRECT (GLS, 2D). DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE-(B)(4), COMPLAINT-(B)(4). | -fail to access clinical history -inadequate clinical history data -lack of documentation indication | -data issue -use of device issue -data back-up problem | Viewer
| FDA | Event Description: INDICATIONS FOR USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. MERGE CARDIO IS SOFTWARE COMPRISED OF MODULES THAT PERFORM UNDER STANDARD "OFF THE SHELF" PERSONAL COMPUTERS AND SERVERS RUNNING THE MICROSOFT WINDOWS 2000/2003/XP OPERATING SYSTEMS. MERGE CARDIO IS IMAGE DATA STORAGE AND DISPLAY SOFTWARE THAT ACCEPTS DICOM (DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE) IMAGE DATA FILES FROM MULTIPLE MODALITIES. IT ACCEPTS TEXT DATA USING OTHER STANDARDS-BASED FORMATS INCLUDING BUT NOT LIMITED TO HL7 AND XML. MERGE CARDIO IS AN INTERNET/INTRANET NETWORK SYSTEM THAT IS DESIGNED FOR SMALL AND LARGE, MULTI-USER ENVIRONMENTS. THE MERGE CARDIO NETWORK STRUCTURE (INCLUDING SERVER AND WORKSTATIONS) PROVIDES FOR THE SYSTEM'S DATABASE MANAGEMENT, STORAGE, PRINTING, AND ALL DICOM/HL-7 INTERFACE SERVICES. ON 12/30/2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT AFTER UPGRADE, MULTIPLE DIAGNOSTIC MEASUREMENTS WERE POPULATING THE CLINICAL REPORT INCORRECTLY. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE CASE (B)(4)/ COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -lack of documentation indication | -data issue -use of device issue -data back-up problem | Viewer
| FDA | Event Description: INDICATIONS FOR USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. MERGE CARDIO IS SOFTWARE COMPRISED OF MODULES THAT PERFORM UNDER STANDARD "OFF THE SHELF" PERSONAL COMPUTERS AND SERVERS RUNNING THE MICROSOFT WINDOWS 2000/2003/XP OPERATING SYSTEMS. MERGE CARDIO IS IMAGE DATA STORAGE AND DISPLAY SOFTWARE THAT ACCEPTS DICOM (DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE) IMAGE DATA FILES FROM MULTIPLE MODALITIES. IT ACCEPTS TEXT DATA USING OTHER STANDARDS-BASED FORMATS INCLUDING BUT NOT LIMITED TO HL7 AND XML. MERGE CARDIO IS AN INTERNET/INTRANET NETWORK SYSTEM THAT IS DESIGNED FOR SMALL AND LARGE, MULTI-USER ENVIRONMENTS. THE MERGE CARDIO NETWORK STRUCTURE (INCLUDING SERVER AND WORKSTATIONS) PROVIDES FOR THE SYSTEM'S DATABASE MANAGEMENT, STORAGE, PRINTING, AND ALL DICOM/HL-7 INTERFACE SERVICES. ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS THAT WERE NEVER TAKEN ON THE MODALITY WERE POPULATING THE CARDIO CLINICAL REPORT AND DIAGNOSTIC MEASUREMENT MITRAL VALVE PEAK E WAS POPULATING WITH THE INCORRECT VALUE. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to access clinical history -inadequate clinical history data -lack of documentation indication | -data issue -use of device issue -data back-up problem | Viewer
| FDA | Event Description: INITIALLY PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT WHEN RECONSTRUCTING DATA BY THE INTELLISPACE PORTAL AUTOSPECT_PRO APPLICATION SOME OF THE RESULTS MAY BE DIFFERENT THAN EXPECTED. THE FEEDBACK INDICATE THAT THE AFFECTED DATA IS CREATED ON LEGACY AXIS/IRIX GAMMA CAMERAS MANUFACTURED BY MARCONI MEDICAL SYSTEMS USING 102 DEGREE ACQUISITION PROTOCOL. FURTHER INFORMATION FROM CUSTOMER RECEIVED ON (B)(6) 2015, LEADING OUR ORGANIZATION TO REPORT IT. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. | -incorrect display format -incorrect device orientation -incorrect patient information | -unauthorized access to computer system -filling problem -Connection Issue | Viewer
| FDA | Event Description: INTELERAD HAS BEEN MADE AWARE OF A PROBLEM WHICH COULD BE A POSSIBLE PATIENT SAFETY: THE ISSUE IS TRIGGERED BY THE PRESENCE OF THE LESS THAN CHARACTER (<) IN THE REPORT. WHEN VIEWING A REPORT, THE LESS THAN CHARACTER IS NOT PRESENT WHILE IT WAS THERE IN THE ORIGINAL REPORT RECEIVED BY INTELEPACS. IN OTHER CASES, WHEN THE CHARACTER IS PRESENT IN A COMBINATION WITH OTHER CHARACTERS, PARTS OF THE REPORT TEXT FOLLOWING THE LESS THAN CHARACTER ARE NOT DISPLAYED DEPENDING ON THE INTELERAD APPLICATION USED BY THE END-USER TO REVIEW THE REPORTS. NO PATIENT WAS REPORTED IMPACTED BY THE PROBLEM. THIS REPORT IS SUBMITTED PREVENTIVELY. TICKET CASE ASSOCIATED: (B)(4). INVESTIGATION DETAILS IS DOCUMENTED IN (B)(4): IN INTELEVIEWER, WHEN VIEWING A REPORT CONTAINING THE LESS THAN (<) CHARACTER WHICH IS NOT PROPERLY FORMATTED FOR HTML, THE SYMBOL WILL NOT BE DISPLAYED. THIS CAN CHANGE THE MEANING OF THE REPORT. ROOT CAUSE: WHEN PLAIN TEXT REPORTS ARE RECEIVED FROM AN EXTERNAL SOURCE, ESCAPING OF HTML CHARACTERS SHOULD BE PERFORMED. THIS BEHAVIOUR IS CONTROLLED BY CONFIGURATION THAT WAS NOT SET CORRECTLY FOR THIS INTEGRATION. CATEGORIES OF REPORTS: INTELEPACS CAN RECEIVE REPORTS FROM MULTIPLE SOURCES. AS SUCH, THE REPORTS CAN BE STORED IN THREE DIFFERENT CATEGORIES: HTML REPORTS GENERATED BY INTELERAD - FOR THIS CATEGORY, THE LESS THAN CHARACTER IS VISIBLE IN THE REPORTS AS EXPECTED BY THE USER. HTML REPORTS FROM EXTERNAL SOURCES - THESE REPORTS DEPEND ON THE HTML FORMATTING PROVIDED BY THE EXTERNAL SOURCE. IF YOU ARE AFFECTED BY THIS CATEGORY, YOU SHOULD PROCEED WITH A REVIEW OF CONFIGURATIONS OF THE EXTERNAL SOURCE. PLAIN TEXT REPORTS FROM EXTERNAL SOURCES THAT HAVE BEEN MODIFIED TO BE COMPATIBLE WITH HTML - THE REPORTS ASSOCIATED WITH THIS PROBLEM ARE FROM THIS CATEGORY. RISK ASSESSMENT: DESCRIPTION OF HAZARDOUS SITUATION IN USE OF THE DEVICE: ú WHEN VIEWING A REPORT, THE LESS THAN, I.E. "<" CHARACTER, IS NOT PRESENT WHILE IT WAS THERE IN THE ORIGINAL REPORT. IN OTHER CASES, PARTS OF REPORT TEXT FOLLOWING THE LESS THAN CHARACTER ARE NOT DISPLAYED. ú IN BOTH SCENARIOS, THERE IS A RISK OF MISDIAGNOSIS. IS THIS PROBLEM CURRENTLY DOCUMENTED AS A KNOWN RISK? ú YES. IS THERE CURRENTLY A MITIGATION IN PLACE TO REDUCE THIS RISK? YES/NO ú YES. IF YES, HAS THE MITIGATION WORKED TO PREVENT RISK TO PATIENT? YES/NO ú NO. THE MITIGATION APPLIED TO A DIFFERENT CATEGORY OF REPORT TRANSFORMATION. CAN THIS RISK BE FURTHER REDUCED? DESIGN FACTORS THAT MIGHT MITIGATE RISK: ú YES. ENSURE APPLICATIONS THAT RETRIEVE REPORTS FOR RENDERING AUTOMATICALLY DETECT AND CORRECT UNESCAPED HTML. PROBABILITY OF RISK? BEFORE/AFTER MITIGATION ú BEFORE MITIGATION: REMOTE PROBABILITY BECAUSE IT REQUIRES MISSING CONFIGURATION FOR ESCAPING HTML, THE MIS-RENDERING HAVING CLINICAL SIGNIFICANCE AND THE ABSENCE OF THE AFFECTED TEXT NOT BEING EVIDENT TO EXHIBIT THE BEHAVIOR. SEVERITY OF RISK? BEFORE/AFTER MITIGATION ú BEFORE MITIGATION: POTENTIAL MISDIAGNOSIS LEADING TO POSSIBLE INJURY REQUIRING TREATMENT. DETECTABILITY (USER AWARENESS. IF DEVICE FAILURE OCCURS, IS IT EASILY RECOGNIZED BY USER) ú DEPENDING ON THE POSITIONING OF THE CORRUPTION WITHIN THE TEXT, IT MAY NOT BE READILY DETECTABLE BY A TRAINED USER. RISK VS. BENEFIT COMPARISON (TO HELP DETERMINE THE RECALL STRATEGY AND TYPE OF THE RECALL) ú BENEFIT OF AN INCREASE IN AVERAGE TIME SAVINGS, DIAGNOSTIC ACCURACY, DATA INTEGRITY AND IMAGE ACCESSIBILITY OUTWEIGHS THE RISK OF LOW PROBABILITY EVENTS LEADING TO INCORRECT DIAGNOSIS. HOW OFTEN DOES THE PROBLEM OCCURS? HOW OFTEN CAN THE PROBLEM CAUSE DIRECT OR INDIRECT HARM TO THE PATIENT? ú <1% OF ALL THE REPORTS, INCLUDING MIGRATED PRIOR REPORTS. THERE HAVE BEEN NO REPORTED INCIDENTS OF HARM TO PATIENT IN MORE THAN 15 YEARS. THIS REPRESENTS OVER 500 MILLION REPORTS. DESCRIBE THE FACTORS THAT MAY CONTRIBUTE TO PRODUCT RISK TO THE PATIENT ú MISCONFIGURATION OF INTERFACE. ú USE OF INCORRECTLY FORMATTED REPORT TEXT. FACTORS THAT MAY CONTRIBUTE TO PRODUCT RISK (I.E. DEVICE DESIGN, MANUFACTURING PROBLEMS, OR USE ERROR) ú AS ABOVE. (B)(6). RATIONALE AND RECOMMENDED ACTIONS: ú ADJUSTING THE CONFIGURATION FOR INBOUND REPORTS WILL ENSURE THAT ALL NEW REPORTS ARE NOT AFFECTED BY THIS ISSUE. ROLLING OUT AN INTELEPACS SOFTWARE UPDATE WILL ENSURE THAT ALL PRIOR REPORTS AFFECTED BY THIS ISSUE WILL DISPLAY CORRECTLY. PROPOSED CORRECTIVE ACTIONS: INTELERAD WILL: INFORM ALL CLIENTS OF THE ISSUE. VERIFY CONFIGURATION OF REPORT INBOUND CHANNELS AT ALL CLIENT SITES AND APPLY CORRECTION IF NEEDED. THE CONFIGURATION SHOULD CATCH AND ESCAPE THE LESS THAN CHARACTER ("<") IN TEXT REPORTS. PROVIDE ALL CLIENTS A FIXED VERSION OF THE DEVICE WITH AN INTELEPACS SOFTWARE UPDATE. VERSIONS 4-1-1 AND 4-2-1 ARE NO LONGER SUPPORTED AND WILL NOT BE UPDATED. CORRECTIVE ACTION CALENDAR: FOR THE CONFIGURATION REVIEW AND CHANGE, INTELERAD REQUIRES A FORMAL APPROVAL FROM THE CLIENTS. THE TIME TO COMPLETE THIS CORRECTIVE ACTION THEREFORE TOTALLY DEPENDS ON IF AND WHEN THE CLIENT RESPONDS TO THE CHANGE REQUEST. FOR THE SOFTWARE UPDATE, INTELERAD REQUIRES A FORMAL APPROVAL FROM THE CLIENTS AND AGREE ON A SCHEDULE FOR THE UPGRADE. THE TIME TO COMPLETE THE CORRECTIVE ACTION THEREFORE TOTALLY DEPENDS ON IF AND WHEN THE CLIENT RESPONDS TO THE SOFTWARE UPGRADE REQUEST. AFFECTED VERSIONS: THE INTELEPACS SYSTEM IS POTENTIALLY AFFECTED IF THE VERSION "R" IDENTIFIER HAS A NUMBER BELOW THE FOLLOWING: PACS 4-1-1 R28, PACS 4-3-1 R49, PACS 4-3-2 R11, PACS 4-3-4 R39, PACS 4-6-1 R107, PACS 4-8-1 R65, PACS 4-9-1 R33, PACS 4-10-1 R61, PACS 4-11-1 R58, PACS 4-12-1 R45, PACS 4-14-1 R24, PACS 4-15-1 R7. | -inappropriate safety factor -fail to activate alarm -system security compromised | -device alarm system issue -failure to read input signal -failure of device to self-test | Viewer
| FDA | Event Description: INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM IN ITS (B)(6) SOFTWARE. THE ISSUE IS TRIGGERED BY THE PRESENCE OF THE PIPE CHARACTER ( ¦ ) IN THE REPORT. IF THE TEXT OF AN INTELEPACS-GENERATED REPORT CONTAINS AN UNQUOTED PIPE CHARACTER (I.E. ¦) THEN QUERYING FOR IT VIA HL7 WILL RESULT IN A REPORT THAT APPEARS TO BE TRUNCATED AT THAT POINT. ONLY CLIENTS WITH VERSIONS 4-11-1 AND LATER USING MULTI METHOD REPORTING OR INTELEDICTATION AS WELL AS THE HL7GATEWAYSERVER DYNAMIC REPORT RENDERER ARE AFFECTED. THIS ISSUE COULD IMPACT THE PATIENT SAFETY BECAUSE IT CAN RESULT IN MISSING REPORT TEXT. WHILE IN SOME CASES IT MAY BE OBVIOUS, IN OTHER CASES IT COULD RESULT IN LOST ADDENDA, WHICH WOULD NOT BE VISIBLE TO THE HEALTH PROFESSIONAL WHO COULD MAKE A DIAGNOSIS WITH INCOMPLETE OR MISSING DATA. ROOT CAUSE ANALYSIS: ROOT CAUSE ANALYSIS: THE DYNAMIC REPORT RENDERER OF HL7GATEWAYSERVER APPLICATION FAILED TO ESCAPE A SPECIAL CHARACTER IN A REPORT, CAUSING THE HL7 OBSERVATION FIELD TO BE UNEXPECTEDLY TRUNCATED WHEN RETRIEVING REPORTS THROUGH HL7. THIS ONLY AFFECTS REPORTS STORED IN THE REPORT ACCESS API (I.E., MMR OR INTELEDICTATION), WHERE THE CLIENT IS USING A DYNAMIC REPORT RENDERER. AFFECTED SOFTWARE VERSIONS SYSTEM IS POTENTIALLY AFFECTED IF THE VERSION "R" IDENTIFIER HAS A NUMBER BELOW THE FOLLOWING: PACS 4-11-1 R69, PACS 4-12-1 R81, PACS 4-14-1 R65, PACS 4-15-1 R47, PACS 4-16-1 R46, PACS 4-17-1 R21. RISK ASSESSMENT: DESCRIPTION OF HAZARDOUS SITUATION IN USE OF THE DEVICE: O WHEN VIEWING A REPORT, PARTS OF THE REPORT TEXT FOLLOWING THE UNESCAPED PIPE CHARACTER ( ¦ ) ARE NO LONGER DISPLAYED. O THIS PROBLEM IS CURRENTLY DOCUMENTED AS A KNOWN RISK. THE MITIGATION IN PLACE TO REDUCE THIS RISK DID NOT WORK TO PREVENT RISK TO PATIENT BECAUSE IT APPLIED TO A DIFFERENT CLASS OF REPORT TRANSFORMATION. O THIS RISK CAN BE FURTHER REDUCED BY ENSURING THAT ANY CHARACTER ESCAPING FOR SERIALIZATION IS PERFORMED CLOSER TO THE PERSISTENCE LAYER IN A GENERIC WAY THAT APPLIES INDEPENDENTLY OF CONFIGURATION. O THE PROBABILITY OF RISK BEFORE MITIGATION IS REMOTE BECAUSE THE MIS-RENDERING MUST HAVE CLINICAL SIGNIFICANCE AND THE ABSENCE OF THE AFFECTED TEXT MUST NOT BE EVIDENT TO EXHIBIT THE BEHAVIOR. O THE SEVERITY OF RISK BEFORE MITIGATION IS POTENTIAL MISDIAGNOSIS LEADING TO POSSIBLE INJURY REQUIRING TREATMENT O THE DETECTABILITY OF THIS PROBLEM BEING EASILY RECOGNIZED BY USER WILL DEPEND ON THE POSITIONING OF THE CORRUPTION WITHIN THE TEXT, IT MAY NOT BE READILY DETECTABLE BY A TRAINED USER. RISK VS. BENEFIT COMPARISON (TO HELP DETERMINE THE RECALL STRATEGY AND TYPE OF THE RECALL) O BENEFIT OF AN INCREASE IN AVERAGE TIME SAVINGS, DIAGNOSTIC ACCURACY, DATA INTEGRITY AND IMAGE ACCESSIBILITY OUTWEIGHS THE RISK OF LOW PROBABILITY EVENTS LEADING TO INCORRECT DIAGNOSIS. MAIN FACTORS FOR THIS PROBLEM TO OCCUR IS THE USAGE OF SPECIAL CHARACTERS IN REPORT TEXT. FUNCTION: REPORT DISPLAY FAILURE MODE/DEFECT: REPORT CHARACTERS MISSING LOCAL EFFECT: MISLEADING DISPLAY OF REPORT TEXT SYSTEM EFFECT: NONE ADVERSE HEALTH CONSEQUENCES: POTENTIAL MISDIAGNOSIS SEVERITY: DELAY IN TREATMENT (1 OUT OF 5) PROBABILITY: REMOTE (2 OUT OF 5) RISK RATING: LOW CORRECTIVE ACTION : INTELERAD WILL INFORM ALL CLIENTS WHOSE SYSTEM HAS THE AFFECTED SOFTWARE VERSIONS (4-11-1 AND UP). INTELERAD WILL INSTALL A SOFTWARE UPDATE ON ALL CLIENTS' SYSTEMS WITH THE AFFECTED SOFTWARE VERSIONS, IN THE FOLLOWING SEQUENCE: I. SITES WHOSE SYSTEM MEETS THE DESCRIBED CONDITIONS ABOVE (I.E. INTELEDICTATION AND MULTI-METHOD REPORTING) - ONLY ON SITE IN UK HAS THE CONFIGURATION CAUSING THIS PROBLEM. II. SITES WHOSE SYSTEM CURRENTLY DO NOT HAVE THE CONDITIONS DESCRIBED ABOVE. INTELERAD WILL PROVIDE A LIST OF POTENTIALLY AFFECTED REPORTS TO THE CLIENTS WHOSE SYSTEMS MEET THE DESCRIBED CONDITIONS. INTELERAD WILL CREATE TICKETS IN ITS SERVICE PORTAL TO TRACK AND TO SCHEDULE THE SOFTWARE UPGRADE REQUIRED FOR EACH CLIENT SYSTEM. CORRECTIVE ACTION CALENDAR: CORRECTION OF THE SOFTWARE: ESTIMATED END OF (B)(6) 2019. SOFTWARE UPDATE OF CLIENTS SYSTEMS WHOSE SYSTEMS HAVE THE CONFIGURATION CAUSING THE PROBLEM: (B)(6) 2019. SOFTWARE UPDATE OF CLIENTS SYSTEMS WHOSE SYSTEMS DO NOT HAVE THE CONFIGURATION CAUSING THE PROBLEM: END OF (B)(6) 2019. | -software installation problem -incorrect software settings -incorrect device orientation | -device alarm system issue -use of device issue -output below specifications | Viewer
| FDA | Event Description: INTENDED USE: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2020 A CUSTOMER CALLED MERGE CARDIO SUPPORT STATING THAT A PATIENT'S REPORT WAS IN A CONFIRMED STATUS WHEN IT WAS ACTUALLY ONLY IN A QC STATUS OR A STATUS WHERE A CLINICIAN HAS NOT READ THE IMAGES AND REPORTED ON THEIR FINDINGS. QC REPORTS WITH THE STATUS OF CONFIRMED, IF SENT TO THE PATIENT'S PHYSICIAN, COULD LEAD TO DELAY OF TREATMENT AS THE PATIENT'S IMAGES HAD NOT ACTUALLY BEEN REVIEWED AND THE FINDINGS DOCUMENTED IN THE REPORT. THE LACK OF THE REVIEW AND REPORTING MAY NOT BE READILY APPARENT TO A PHYSICIAN. REFERENCE COMPLAINT-(B)(4). Manufacturer Narrative: CUSTOMER CALLED CARDIO TECHNICAL SUPPORT AND STATED THAT A REPORT IN MERGE CARDIO LOOKED AS IF IT WAS CONFIRMED BUT IT WAS ACTUALLY JUST QC'D. TECHNICAL SUPPORT WAS ABLE TO RESOLVE THIS ISSUE BY FLIPPING THE CONFIRMATION STATE FLAG, IN THE DATABASE, BACK TO QC IN ORDER FOR THE REPORT TO BE READ BY A PHYSICIAN. A REVIEW OF THE CUSTOMER'S DEVICE DETERMINED THAT THERE WERE NO DETAILED ACTIVITY LOGS, THE ONLY ACTIVITY THAT COULD BE SEEN FROM THE LOGS WAS THAT THE REPORT WAS OPENED AND THEN CLOSED. WITH LIMITED INFORMATION AVAILABLE AT THE CUSTOMER'S SITE, TECHNICAL SUPPORT WAS UNABLE TO TROUBLESHOOT THIS ISSUE FOR A PRIMARY CAUSE. THE CUSTOMER DID CONFIRM THAT NO PATIENT HARM RESULTED FROM THIS ISSUE. THE ISSUE WAS RESOLVED WITH A FLAG STATUS CHANGE IN THE DATABASE WHICH ALLOWED A CLINICIAN TO COMPLETE READING AND REPORTING OF THE PATIENT'S IMAGES. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -use of device issue -incorrect display -use of incorrect control settings | Viewer
| FDA | Event Description: INTERNALLY FOUND ISSUE. DISTANCE MEASUREMENTS MAY NOT BE CALCULATED ACCURATELY BY ZERO FOOTPRINT VERSIONS 5.0 SP2 TO SP7 WHEN APPLIED TO IMAGES ACQUIRED WITH A MAGNIFICATION FACTOR. | -incorrect distance measurement -incorrect measurement -fail to calculate area | -Image Resolution Poor -Device Inoperable -calibration issue | Viewer
| FDA | Event Description: INVIVO OFFERS THE ELBS BREAST COILS WITH A GRID THAT HAS TWO ADDITIONAL GRID ROWS WHICH ARE OFF-SET FROM THE REST OF THE GRID. SURELOC DID NOT INCLUDE A CONFIGURATION FOR THIS GRID AND THE GENERIC GRID OPTION IN SURELOC ASSUMES THE GRID PLACEMENT IS WITH-OUT ANY OFF SETS. IF THE USER IS NOT EXPERIENCED ENOUGH TO REALIZE THAT THE TOP TWO GRID ROWS ARE OFFSET, THEN THE BIOPSY MAY END UP MISSING THE DESIRED AREA. INVIVO DID NOT MAKE MERGE HEALTHCARE AWARE OF THIS NEW GRID, AND THIS IS THE REASON FOR IT NOT BEING A RECOGNIZED GRID FOR SURELOC. THE USERS ARE TAUGHT TO DO A CONFIRMATION SCAN PRIOR TO PERFORMING BIOPSY. | -fail to setup configuration -incorrect configuration -fail to add line | -use of device issue -medication error -improper or incorrect procedure or method | Viewer
| FDA | Event Description: IT HAS BEEN REPORTED TO PHILIPS THAT DURING A CEREBRAL ARTERY ANEURYSM TREATMENT, THE CUSTOMER NOTICED A DIFFERENCE BETWEEN THE MEASUREMENT OF THE ANEURYSM SIZE ON THE 3DRA INTERVENTIONAL TOOL AND THE LIVE X-RAY IMAGE. NO HARM TO THE PATIENT OR USER WAS REPORTED. PHILIPS HAS STARTED AN INVESTIGATION OF THIS COMPLAINT. | -incorrect measurement -inadeqaute measurement -incorrect sample size | -fluid leak -disconnection -leak | Viewer
| FDA | Event Description: IT HAS BEEN REPORTED TO PHILIPS THAT DURING AN EMERGENCY PROCEDURE FOR A STROKE PATIENT, THE CUSTOMER WAS NOT ABLE TO USE THE XPERCT FUNCTIONALITY OF THE SYSTEM. THE PROCEDURE WAS COMPLETED USING REGULAR CT. NO HARM TO THE PATIENT OR USER HAS BEEN REPORTED TO PHILIPS. PHILIPS HAS STARTED AN INVESTIGATION FOR THIS COMPLAINT. | -unable to use -loss of function -difficult to use | -fitting problem -improper or incorrect procedure or method -disconnection | Viewer
| FDA | Event Description: IT WAS IDENTIFIED AT CUSTOMER'S SITE THAT 80 STUDIES PRESUMED TO BE MISSING FROM THE SYNGO DYNAMICS (SD) PRODUCT DATING FROM JUNE 2011 THROUGH APRIL 2014. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH WOULD RESULT IN THE NEED FOR A PATIENT RESCAN. THE REPORTED EVENT OCCURRED IN (B)(6). | -haptic missing -accessory missing -missing splint | -leak -no device output -disconnection | Viewer
| FDA | Event Description: IT WAS REPORTED BY THE CUSTOMER SITE THAT ON (B)(6) 2015, A PATIENT WAS IMAGED FOR A NUCLEAR MEDICINE (NM) STUDY AT 18:10. THE PATIENT EXPIRED ON (B)(6) 2015 AT 21:50. ON (B)(6) 2015, THE CUSTOMER SITE CLINICAL IMAGING ADMINISTRATOR CONTACTED MCKESSON REQUESTING ASSISTANCE, IN DETERMINING,THROUGH THE REVIEW OF THE PACS LOGS, THE DATE AND TIME THE NM STUDY WAS REVIEWED BY THE PERFORMING TECHNOLOGIST AND MOVED TO THE RADIOLOGIST'S WORKLIST FOR DIAGNOSIS. | -patient expired -patient experienced adverse event -patient hospitalized | -Detachment of Device Component -Loss of Power -break | Viewer
| FDA | Event Description: IT WAS REPORTED IN MULTIPLE CASES THE ORTHOSIZE SOFTWARE TEMPLATED A SMALLER SIZE STEM THAN APPROPRIATE FOR THE PATIENT. THE TEMPLATED SIZE WAS IMPLANTED AND SUBSIDENCE ISSUES RESULTED. | -inappropriate size -inadequate size -incorrect device orientation | -Incomplete or Missing Packaging -Contamination During Use -component missing | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A CLINICIAN PERFORMED A PROCEDURE TO PLACE DENTAL IMPLANTS IN REGIONS #18, #19 AND #31 USING A SURGIGUIDE. THE SURGIGUIDE FIT PERFECTLY AND IT WAS STATED THAT THE CLINICIAN FOLLOWED THE DRILLING PROTOCOL. IMPLANTS 18 AND 19 WENT APPROXIMATELY 7.5MM INTO THE 9MM OSTEOTOMY AND COULD NOT BE SEATED COMPLETELY. IMPLANT #31 WENT 6MM INTO THE 9MM OSTEOTOMY. THE CLINICIAN REMOVED ALL IMPLANTS AND VERIFIED WITH AN X-RAY THAT OSTEOTOMY WAS DRILLED TO DEPTH. AFTER A SECOND TRY THE IMPLANTS WERE REMOVED AND ALL SITES WERE RE-GRAFTED. A SECOND SURGERY IS NEEDED TO CONTINUE WITH THE IMPLANT PLACEMENT. DURING INVESTIGATION IT WAS NOTICED THAT THE FINAL DRILL STEP FOR TWO OF THE IMPLANTS WAS INCORRECT ON THE DRILLING PROTOCOL. THIS LEAD TO UNDER PREPARATION OF THE OSTEOTOMY AND INABILITY TO INSERT IMPLANTS TO THE CORRECT DEPTH. | -incorrect procedure -incorrect device orientation -inappropriate technique | -Mechanical Issue -fluid leak -fitting problem | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A CUSTOMER USED A SURGIGUIDE COMPATIBLE WITH THE ASTRA TECH IMPLANT SYSTEM TX TO PLACE TWO IMPLANTS IN THE MANDIBLE. THE PATIENT IS PARTIALLY EDENTULOUS AND THE GUIDE WAS PLACED ON THE REMAINING TEETH DURING SURGERY TO PERFORM THE OSTEOTOMIES AND PLACE THE IMPLANTS. THE OSTEOTOMY FOR THE FIRST IMPLANT WAS NOT CREATED AT THE DESIRED POSITION IN THE PATIENT MOUTH AS PLANNED. THE ROOT OF THE NEIGHBORING TOOTH WAS DAMAGED DURING SURGERY. | -incorrect device orientation -device malfunction -device damaged | -detachment of device component -fluid leak -disconnection | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A DENTAL IMPLANT HAD TO BE REMOVED BECAUSE IT WAS 2MM ABOVE CRESTAL BONE AFTER USING A SURGIGUIDE TO PLACE THE IMPLANT. | -device damaged -device blocked -hole in part | -use of device issue -device inoperable -device operates differently than expected | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A DIFFERENT PATIENT PROFILE WILL LAUNCH WHEN SELECTING A PATIENT IN THE DEXIS IMAGING SUITE PATIENT SELECTION WINDOW. THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT'S NAME IN DEXIS ADMINISTRATION SCREEN. NO INJURIES WERE REPORTED. THERE WAS NO IMPACT TO PATIENT CARE. | -incorrect patient information -launch delayed -incorrect user manual | -High Readings -Hole in Material -material disintegration | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SERVER ROOM AT THE CUSTOMER'S FACILITY HAD A POWER OUTAGE IN OCTOBER 2016, WHICH RESULTED IN THE FAILURE OF THE RAID (HARD DISK UNIT) ATTACHED TO THE SIEMENS SYNGO IMAGING SYSTEM. A PREVIOUS POWER OUTAGE THAT OCCURRED EARLIER (SEPTEMBER 2016) HAD ALREADY DEACTIVATED THE RAID MIRRORING FUNCTION. THEREFORE, THE DATA WAS NOT PROTECTED BY ANY BACK-UP OPTION. AFTER THE ASSESSMENT WAS CONDUCTED, IT WAS DETERMINED THAT 15 POTENTIALLY CLINICALLY RELEVANT STUDIES DATED BETWEEN OCTOBER 13TH AND 15TH WERE LOST. THIS DATA WILL NOT BE AVAILABLE FOR DIAGNOSIS IF REQUIRED BY PHYSICIAN. THE SYNGO IMAGING SYSTEM DID NOT CONTRIBUTE TO THE REPORTED DATA LOSS. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. THE REPORTED EVENT OCCURRED IN (B)(6). | -power outage -power interrupted -illumination power out of specification | -optical issue -device markings issue -Material Puncture | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE BROKE DURING SURGERY. THE PATIENT WAS ANAESTHETIZED, HOWEVER THE CLINICIAN DELAYED DENTAL IMPLANT PLACEMENT. Manufacturer Narrative: MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO OBTAIN THE DEVICE FOR EVALUATION. | -device damaged -surgical graft failure -device broken or damaged | -device inoperable -material disintegration -Air Leak | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE BROKE WHILE THE DOCTOR WAS TRYING TO PLACE GUIDE, AFTER REFLECTING THE TISSUE. THE SURGERY WAS ABORTED AND A SECOND SURGERY IS NEEDED. | -tissue damaged -patient experienced adverse event -device damaged | -device operates differently than expected -fluid leak -insufficient information | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT FIT IN A PATIENT'S MOUTH. THE DENTAL IMPLANT PLACEMENT SURGERY HAS BEEN DELAYED. Manufacturer Narrative: THERE HAS BEEN A PREVIOUS REPORT RECEIVED WHERE THIS MALFUNCTION RESULTED IN A SERIOUS INJURY. THEREFORE, IT MUST BE PRESUMED THAT RECURRENCE OF THIS MALFUNCTION COULD POSSIBLY CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR REQUIRE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE SUCH. AS SUCH, THIS EVENT IS REPORTABLE PER 21CFR PART 803. WE COMPARED THE BUILD MODEL WITH THE PLASTER MODEL. THERE IS A DEVIATION AT THE TEETH 17 AND 28 WHICH PROBABLY CAUSES THE NON FIT IN THE PATIENT MOUTH. THE MODEL IN PLANNING WASN'T THE SAME AS SITUATION IN THE PATIENT'S MOUTH. | -inappropriate fit -incorrect placement -part mismatched | -fitting problem -Device Handling Issue -fluid leak | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT FIT IN A PATIENT'S MOUTH. THE DOCTOR STATED THAT HE OPENED A FLAP TO TRY TO OBTAIN A BETTER FIT AND DRILLED, BUT THEN PERFORATED THE LINGUAL PLATE. THE IMPLANT PLACEMENT SURGERY WAS ABORTED AND A SECOND CORRECTIVE SURGERY IS NEEDED. | -inappropriate fit -component open -part mismatched | -fitting problem -Device Handling Issue -Human-Device Interface Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE DID NOT PROPERLY FIT IN A PATIENT'S MOUTH. THE DENTAL IMPLANT PLACEMENT SURGERY HAS BEEN DELAYED. | -inappropriate fit -incorrect placement -fail to replicate | -fitting problem -fluid leak -improper or incorrect procedure or method | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE THAT WAS MEANT TO BE USED TO PLACE THREE IMPLANTS, COULD ONLY BE USED TO PLACE TWO IMPLANTS. THE SURGERY WAS ABORTED. THE DOCTOR COULDN'T FOLLOW THROUGH WITH THE DRILL SEQUENCE BECAUSE OF THE INTEROCCLUSAL SPACE. THE DOCTOR STATED THAT THE SLEEVES IN THE GUIDE WERE TOO HIGH AND VERY LONG AND THEY SIT OFF OF THE PATIENT'S RIDGE. | -incorrect device orientation -fail to enter data -inappropriate interface design | -crack -fluid leak -Detachment of Device Component | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED FOR PREPARATION OF THREE IMPLANTS. THE GUIDE SEATED CORRECTLY AND THE DRILLING SEQUENCE WAS UNEVENTFUL AND IDENTICAL FOR ALL THREE OSTEOTOMIES. THE IMPLANTS WERE PLACED WITHOUT THE GUIDE THROUGH A TISSUE PUNCH. POST OP X-RAY REVEALED THAT ONE OF THE IMPLANTS WAS 3 MM APICAL TO PLANNED LOCATION. THE DOCTOR WAS UNABLE TO RETRIEVE THE IMPLANT. | -incorrect device orientation -device malfunction -part mismatched | -fluid leak -decrease in pressure -Device Handling Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO CREATE THE OSTEOTOMY FOR THREE DENTAL IMPLANTS (TOOTH POSITION 5, 6 AND 7). AFTER THE FIRST DRILLING STEP THE SURGICAL GUIDE WAS REMOVED TO COMPLETE THE ADDITIONAL DRILLING STEPS FREE HANDED ACCORDING TO THE PILOT GUIDE DRILLING PROTOCOL. AFTER PLACING THE IMPLANTS FREE HANDED, THE CUSTOMER DISCOVERED THAT THE THREE IMPLANTS WERE POSITIONED TOO BUCCAL. A SECOND SURGERY IS NEEDED TO CORRECT THE POSITIONING OF THE IMPLANTS. | -incorrect surgical decision -surgical graft failure -inappropriate direct visualization surgical technique | -component missing -fluid leak -Device Handling Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO CREATE THE OSTEOTOMY FOR TWO DENTAL IMPLANTS (TOOTH POSITION 7 AND 10). AFTER THE FIRST DRILLING STEP THE SURGICAL GUIDE WAS REMOVED TO COMPLETE THE ADDITIONAL DRILLING STEPS FREE HANDED ACCORDING TO THE PILOT GUIDE DRILLING PROTOCOL. AFTER PLACING THE IMPLANTS FREE HANDED, THE CUSTOMER DISCOVERED THAT THE IMPLANT IN #7 WAS POSITIONED CORRECTLY BUT THE IMPLANT IN #10 WAS POSITIONED TOO DISTAL. A SECOND SURGERY IS NEEDED TO CORRECT THE POSITIONING OF IMPLANT IN #10. | -inappropriate direct visualization surgical technique -surgical graft failure -incorrect surgical decision | -Air Leak -component missing -no device output | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO PLACE A DENTAL IMPLANT. THE IMPLANT TRANSFER WAS TOO SHALLOW AND THE GUIDE DID NOT FIT. THE CLINICIAN REPORTED THAT ADDITIONAL STEPS DURING THE SURGERY WERE NEEDED TO PLACE THE IMPLANT DEEPER. A SECOND CORRECTIVE SURGERY WAS NOT NEEDED. | -incorrect device orientation -part mismatched -inappropriate direct visualization surgical technique | -fluid leak -Bleed Back -Detachment of Device Component | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO PLACE A DENTAL IMPLANT. THE SURGIGUIDE SEATED PERFECTLY AND THE DOCTOR FOLLOWED WITH THE DRILLING SEQUENCE. WHEN THE DOCTOR ATTEMPTED TO PLACE THE IMPLANT, IT SEATED ONLY TO THE CREST OF GINGIVA INSTEAD OF THE CREST OF THE BONE. THE DOCTOR STATED THAT HE HAD TO REMOVE THE IMPLANT AND PLACE BONE GRAFTING MATERIAL IN THE IMPLANT SITE AND THAT A CORRECTIVE SURGERY WILL BE NEEDED IN APPROXIMATELY FOUR MONTHS. ADDITIONAL INFORMATION IS PENDING. | -improper use -improper stent implanted -incorrect device orientation | -fluid leak -Bleed Back -failure of device to self-test | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WAS USED TO PLACE FOUR DENTAL IMPLANTS. TWO OF THE IMPLANTS WERE PLACED WITHOUT ISSUES. HOWEVER, THE IMPLANTS PLANNED IN THE AREAS OF #20 AND 21 WERE PLACED SO FAR BUCCAL THAT THEY LACKED SUFFICIENT PRIMARY STABILITY. A CORRECTIVE SURGERY IS NOW REQUIRED. | -incorrect device orientation -device malfunction -defective module | -Decrease in Pressure -device expiration issue -kinked | Viewer
| FDA | Event Description: IT WAS REPORTED THAT A SURGIGUIDE WOULD NOT FIT PROPERLY IN A PATIENT'S MOUTH. THE CLINICIAN REPORTED THAT THE FIT WAS BAD ON THE LEFT SIDE, BUT IT FIT PROPERLY ON THE RIGHT SIDE. THEREFORE, THE CLINICIAN SPLIT THE GUIDE INTO TWO DIFFERENT PIECES. HE USED ONLY THE RIGHT SIDE OF THE GUIDE AND RESCHEDULED THE SURGERY FOR THE LEFT SIDE. | -inappropriate fit -part mismatched -inappropriate personnel competent | -fitting problem -Loss of Power -detachment of device component | Viewer
| FDA | Event Description: IT WAS REPORTED THAT DUE TO A SOFTWARE ERROR IN THE SYNGO.PLAZA WITH SW VERSIONS VB10A (ALL HOTFIX LEVELS) AND VB10B INCOMPLETE SERIES WERE ARCHIVED IN DICOM LONG TERM STORAGE (LTA) ARCHIVES. A REVIEW OF THE SYNGO.PLAZA SYSTEMS WITH DICOM LTA ARCHIVES WAS TRIGGERED BY A CORRECTIVE ACTION "SYNGO.PLAZA VB10B NON-MEDICAL SOFTWARE APPLICATION CHECK TOOL V1.0 (DISTRIBUTED VIA UPDATE INSTRUCTIONS SY024/17/P AND SY028/17/P, REPORTED TO FDA UNDER C&R # 2240869-05/25/17-0015-C) INTENDED TO DISCOVER INCONSISTENCIES BETWEEN THE SYNGO.PLAZA DATABASE AND THE FILES ARCHIVED IN THE DICOM LTA. DURING THIS CHECK IT WAS DISCOVERED THAT 9249 STUDIES WERE MISSING SERIES IN THE ARCHIVE FOR THIS PARTICULAR CUSTOMER. DUE TO THE HIGH AMOUNT OF AFFECTED STUDIES IT WILL BE VERIFIED IF THIS IDENTIFIED SOFTWARE ERROR IS THE ONLY ROOT CAUSE. THE REPORTED ISSUE OCCURRED IN (B)(6). | -software defect -software corruption -user error | -compliance issue -device inoperable -material disintegration | Viewer
| FDA | Event Description: IT WAS REPORTED THAT DURING AN INITIAL TOTAL HIP ARTHROPLASTY PROCEDURE ON AN UNKNOWN DATE THE ORTHOSIZE V1.2.6 SOFTWARE DID NOT CORRECTLY TEMPLATE THE STEM. THE SURGEON BROACHED AND ATTEMPTED TO IMPLANT THE STEM, BUT IT DID NOT FIT. THE STEM WAS REMOVED AND THE SURGEON CONTINUOUSLY BROACHED UNTIL THE APPROPRIATE SIZE WAS REACHED.Event Description: IT WAS REPORTED THAT DURING AN INITIAL TOTAL HIP ARTHROPLASTY PROCEDURE ON (B)(6), 2015 THE ORTHOSIZE V1.2.6 SOFTWARE TEMPLATED A SIZE 8 STEM. THE SURGEON BROACHED TO SIZE 8 AND ATTEMPTED TO IMPLANT THE STEM; HOWEVER IT DID NOT FIT. THE STEM WAS REMOVED AND THE SURGEON CONTINUOUSLY BROACHED UNTIL THE APPROPRIATE SIZE WAS REACHED. A SIZE 12 STEM WAS IMPLANTED. | -fail to set value correctly -incorrect procedure -incorrect specification | -fail-safe design failure -use of device issue -leak | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT FEMORAL STEM TEMPLATES MADE BY THE SOFTWARE WERE SMALLER THAN EXPECTED AND THERE HAVE BEEN REPORTS OF FEMORAL SUBSIDENCE POST-OPERATIVELY. THE NUMBER OF EVENTS ARE UNKNOWN AND NO REVISION PROCEDURES HAVE BEEN REPORTED TO DATE. | -size out of specification -incorrect product information -inappropriate display size | -device markings issue -improper or incorrect procedure or method -use of device issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT IN THE COURSE OF SEVERAL YEARS THE CONCERNED CUSTOMER SITE HAD EXPERIENCED SEVERAL INCIDENTS RELATED TO THEIR NETWORK ATTACHED STORAGE (NAS). FIRST REPORTED OCCURRENCE WAS A HARDWARE CRASH AT THE NAS LEVEL WHILE DATA MIGRATION TO THE NAS WAS IN PROGRESS. SECOND REPORTED OCCURRENCE WAS AN INCORRECT MAINTENANCE OF A PROVISIONALLY ESTABLISHED SET OF SMALLER RAIDS (REDUNDANT ARRAY OF INDEPENDENT DISKS ) USED INSTEAD OF THE CRASHED NAS (LEADING TO WRONG FILE SYSTEM REFERENCES IN THE ARCHIVE DATABASE). THERE ARE NO INJURIES RELATED TO THIS EVENT. THIS EVENT OCCURRED IN (B)(6). | -inappropriate storage environment or condition -insufficient storage space -connection broken or damaged | -Detachment of Device Component -disconnection -detachment of device component | Viewer
| FDA | Event Description: IT WAS REPORTED THAT IN THE COURSE OF SEVERAL YEARS THE CONCERNED CUSTOMER SITE HAD EXPERIENCED SEVERAL INCIDENTS RELATED TO THEIR NETWORK ATTACHED STORAGE (NAS). FIRST REPORTED OCCURRENCE WAS A HARDWARE CRASH AT THE NAS LEVEL WHILE DATA MIGRATION TO THE NAS WAS IN PROGRESS. SECOND REPORTED OCCURRENCE WAS AN INCORRECT MAINTENANCE OF A PROVISIONALLY ESTABLISHED SET OF SMALLER RAIDS (REDUNDANT ARRAY OF INDEPENDENT DISKS ) USED INSTEAD OF THE CRASHED NAS (LEADING TO WRONG FILESYSTEM REFERENCES IN THE ARCHIVE DATABASE). THERE ARE NO INJURIES RELATED TO THIS EVENT. THIS EVENT OCCURRED IN (B)(6). | -inappropriate storage environment or condition -insufficient storage space -connection broken or damaged | -Detachment of Device Component -disconnection -detachment of device component | Viewer
| FDA | Event Description: IT WAS REPORTED THAT IN THE SYNGO IMAGING XS THERE IS A POTENTIAL FOR MALFUNCTION WHEN USING THE SYNGO IMAGING XS, VERSION VA70A SOFTWARE AND HIGHER. WITH VA70A APPLICATIONS ROI (REGION OF INTEREST) ARE SHOWING INCORRECT VALUES AND DO NOT MATCH TO SYNGO.VIA AND OSIRIX. THE ISSUE IS ISOLATED TO IMAGES OF PIXEL DEPTH GREATER THAN 12 BIT (99% OF ALL STORED IMAGES). ANALYSIS SHOWED THAT THIS BEHAVIOR IS CAUSED BY A SOFTWARE BUG. A MISCALCULATION DONE FOR THE GREY SCALE VALUES IN THE FUNCTIONS "REGION OF INTEREST", "PIXEL LENS", "EDGE ENHANCEMENT", AND "HISTOGRAMS" MAY OCCUR IN SOME SITUATIONS. THE MIN/MAX VALUE OF PIXEL DENSITY IN THE FUNCTIONS "REGION OF INTEREST", "PIXEL LENS", "EDGE ENHANCEMENT", AND "HISTOGRAMS" IS NOT CORRECT FOR IMAGES WITH A PIXEL DEPTH GREATER THAN 12 BITS. IT HAS TO BE NOTED THAT THE MENTIONED MEASUREMENTS SHOWN IN PRIOR IMAGES ARE ALSO NOT CORRECT. ALL VALUES THAT ARE CALCULATED FOR IMAGES WITH A PIXEL DEPTH LESS THAN 12 BITS ARE CORRECT. THIS INCIDENT OCCURRED IN (B)(6). | -device malfunction -malfunction -image display error | -fluid leak -component missing -Dislodged or Dislocated | Viewer
| FDA | Event Description: IT WAS REPORTED THAT PATIENT DEMOGRAPHIC DATA UPDATES IN MERGE PACS ARE NOT REFLECTING IN ICONNECT ACDESS 4.0. THE IMAGES ARE BEING PULLED WITH STALE DEMOGRAPHIC DATA. THE CORRECT STUDY LAUNCHES, THE STUDY'S TITLE BAR IS CORRECT, BUT THE DEMOGRAPHICS WITHIN THE VIEWPORT MAY CONFLICT. THE USER CAN GO BACK TO MERGE PACS TO CHECK THE INFORMATION. THE MANUFACTURER REVIEWED THE COMPLAINTS DATABASE TO IDENTIFY ANY ADDITIONAL CUSTOMER COMPLAINTS, AND NONE WERE IDENTIFIED. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | -fail to reset patient data -incorrect patient information -incorrect data format | -therapeutic or diagnostic output failure -device markings issue -material disintegration | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE IMAGE BECAME GRAINY AND HAD TO CONVERT TO OPEN SURGERY. | -distorted images -loss of image -incorrect device orientation | -Device Operates Differently than Expected -device inoperable -protective measure issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE ISSUE: ONE TRIAL ANALYSIS CHART DOES NOT CORRECTLY REFLECT THE CHANGE TO LEG LENGTH WHEN CHANGING BETWEEN PROXIMAL BODY HEIGHTS. | -incorrect data chart -incorrect graph -incorrect display format | -component missing -loose or intermittent connection -aborted charge | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE RADIOLOGY SERVER AT THE MEDICAL FACILITY WAS INFECTED BY THE WANNACRY VIRUS. THE SEVER WAS DOWN FOR THREE WEEKS WITH NO REQUEST FROM THE USER FOR RECOVERY ASSISTANCE. IT WAS DETERMINED THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED. HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. | -patient infected -clinic infection -cross infection | -fluid leak -Air Leak -spark | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE SYNGO IMAGING XS SERVICES WOULD NOT START. THE STATION HAD TO BE REBOOTED, HOWEVER, THE SAME ERROR MESSAGE "NETWORK SERVICE STOPPED" KEPT BEING DISPLAYED ON THE SCREEN. THE ONLY INFORMATION THAT COULD BE EXTRACTED FROM THE LOGS WAS "E:\ DRIVE HAS NO ENOUGH SPACE" AND "MAGIC VIEW WAS NOT ABLE TO ACCESS FILES ON IT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE INCIDENT WAS REPORTED IN (B)(6). | -error message -incorrect message display -system fault displayed | -Device displays error message -Device Inoperable -break | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE SYSTEM FAILED TO BOOT UP. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. | -fail to boot up -system failure -fail to setup configuration | -Device Operates Differently than Expected -Device Alarm System Issue -failure to calibrate | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE SYSTEM LOCKED UP. THERE WAS NO PT INJURY OR DEATH REPORTED. | -unable to lock screen -lock to incorrect pressure -fail to lock | -Device displays error message -reset problem -inappropriate or unexpected reset | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE SYSTEM WOULD NOT BOOT UP ALL THE WAY. THERE WAS NO PT INJURY OR DEATH REPORTED. | -fail to boot up -fail to enter inoperative state -unable to compress | -Device Alarm System Issue -decrease in pressure -reset problem | Viewer
| FDA | Event Description: IT WAS REPORTED THAT THE SYSTEM WOULD NOT BOOT UP. THIS CAUSED A TOTAL LOSS OF NAVIGATION FUNCTION. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY WITH THIS CASE. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE COMPUTER WAS EVALUATED AND REPLACED. THE SYSTEM SOFTWARE WAS RELOADED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -fail to boot up -fail to setup configuration -unable to compress | -device operates differently than expected -protective measure issue -Device Alarm System Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT TWENTY TO THIRTY PATIENTS UNDERWENT TOTAL HIP ARTHROPLASTY PROCEDURES ON UNKNOWN DATES. DURING THE PROCEDURES, THE ORTHOSIZE V1.2.6 SOFTWARE DID NOT CORRECTLY TEMPLATE THE STEM. THE SURGEON BROACHED AND ATTEMPTED TO IMPLANT THE STEM, BUT IT DID NOT FIT. THE STEM WAS REMOVED AND THE SURGEON CONTINUOUSLY BROACHED UNTIL THE APPROPRIATE SIZE WAS REACHED. | -incorrect procedure -incorrect patient information -incorrect test procedure | -fluid leak -leak -device operates differently than expected | Viewer
| FDA | Event Description: IT WAS REPORTED THAT WHILE USING A SURGIGUIDE TO PLACE FOUR IMPLANTS, AT IMPLANT 43 (#27) THERE WAS A HUGE FENESTRATION OF THE CORTICAL LINGUAL BONE. THE SURGERY COULD NOT BE COMPLETED. A NEW SURGERY WILL BE PERFORMED IN TWO MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT IS NOT YET AVAILABLE. | -fail to enter data -part fractured -part ejected | -fluid leak -insufficient information -Device Handling Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THAT WHILE USING A SURGIGUIDE TO PLACE SIX IMPLANTS, THE GUIDE SHIFTED DURING THE DRILLING IN THE VESTIBULAR AREA (CLOSE TO SITE 32 AND 42) RESULTING IN A FRACTURE IN THE BONE. THE SURGERY WAS NOT COMPLETED. ALSO IMPLANT 36 DOES NOT SEEM TO BE PLACED IN THE PLANNED AREA. | -component fractured -part fractured -incorrect surgical decision | -fluid leak -device unsafe to use in environment -disconnection | Viewer
| FDA | Event Description: IT WAS REPORTED THAT WHILE USING TOMTEC-ARENA ZERO IF TWO TABS ARE OPEN ON THE BROWSER AT THE SAME TIME, IF THE USER SWITCHES FROM ONE TAB TO THE OTHER, THEN THE PATIENT INFORMATION FROM THE FIRST TAB IS COPIED TO THE PATIENT STUDY ON THE SECOND TAB. TO DATE NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS ISSUE. | -incorrect patient information -incorrect user manual -inappropriate user | -leak -fluid leak -failure to read input signal | Viewer
| FDA | Event Description: IT WAS REPORTED THE SYS EXHIBITED AN UNCOMMANDED SHUT DOWN. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. | -device shut down -unintended program shut down -unexpected shutdown | -Device Operates Differently than Expected -device operates differently than expected -device inoperable | Viewer
| FDA | Event Description: IT WAS REPORTED THE SYSTEM DISPLAYED A FIELD DISTORTION MESSAGE AND LOCKED UP. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY ASSOCIATED W/THE COMPLAINT. | -unable to lock screen -error message -locking fragments or breaks | -reset problem -Device displays error message -Device Alarm System Issue | Viewer
| FDA | Event Description: IT WAS REPORTED THE SYSTEM FAILED TO BOOT UP. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. | -fail to boot up -fail to setup configuration -system failure | -reset problem -device inoperable -Device Operates Differently than Expected | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS BY A USER THAT A SIMPLE FAILURE TO LOAD AN IMAGE STUDY FROM THE LONG TERM ARCHIVE OCCURRED. HOWEVER, SIEMENS IDENTIFIED THAT AN UNKNOWN AMOUNT OF IMAGE STUDIES WAS MISSING IN THE SITE'S LONG TERM ARCHIVE. THE HISTORICAL STUDY IMAGES ARE NO LONGER AVAILABLE AND CONSIDERED TO BE UNRECOVERABLE AT THIS POINT OF TIME. HOWEVER, THE PREVIOUS CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE IN THE SYNGO DYNAMICS FOR CLINICAL COMPARISON. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH RESULT IN THE NEED FOR A PATIENT WOULD RESCAN. THE REPORTED INCIDENT OCCURRED IN (B)(6). | -fail to interpret images -unable to scroll image -fail to verify image | -application program issue -device alarm system issue -fitting problem | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X WORKPLACE SYSTEM. DURING A TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR), WHEN LOADING AN OUTSIDE CTA (COMPUTED TOMOGRAPHY ANGIOGRAPHY) WITH BOOKMARKS, THE 4D VIEWER CLOSED BEFORE OVERLAY FOR 2D3D FUSION COULD BE SELECTED. THE RESULT OF THE ERROR WAS A DELAY IN PROCEDURE. THERE IS NO REPORT OF IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. | -device malfunction -malfunction -inappropriate syringe barrel | -fluid leak -Human-Device Interface Issue -leak | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X WORKPLACE SYSTEM. DURING AN INTERVENTIONAL PROCEDURE, THE USER REPORTED THAT WHILE ATTEMPTING TO SETUP FOR 3D, THE SYSTEM WOULD NOT BOOT CORRECTLY. THE PATIENT WAS SAFELY REMOVED FROM THE SYSTEM AND TRANSFERRED TO AN ALTERNATE SYSTEM WHERE THE PROCEDURE WAS COMPLETED. WE ARE UNAWARE OF ANY IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. | -device malfunction -malfunction -patient experienced adverse event | -Human-Device Interface Issue -leak -disconnection | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X WORKPLACE SYSTEM. THE USER REPORTED THAT THE 3D IMAGE RECONSTRUCTION HAS DEVIATIONS. THERE IS NO REPORT OF IMPACT TO THE STATE OF HEALTH OF ANY PATIENT OR USER INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | -distorted images -image orientation incorrect -loss of image | -fluid leak -Device displays error message -reset problem | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X- WORKPLACE. DURING AN INTERVENTIONAL PROCEDURE, A 3D DYNACT SPIN FAILED TO RECONSTRUCT. THE PROCEDURE WAS CONTINUED AND COMPLETED ON AN ALTERNATE SYSTEM. WE ARE UNAWARE OF ANY IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | -device malfunction -fail to deliver -fail to generate indication | -leak -Human-Device Interface Issue -disconnection | Viewer
| FDA | Event Description: IT WAS REPORTED TO SIEMENS THAT A MALFUNCTION OCCURRED WHILE OPERATING THE SYNGO X-WORKPLACE SYSTEM. THE USER REPORTED POOR IMAGE QUALITY DURING A PROCEDURE. WE ARE UNAWARE OF ANY IMPACT TO THE STATE OF HEALTH OF THE PATIENT INVOLVED. SIEMENS HAS REQUESTED ADDITIONAL INFORMATION IN ORDER TO CONDUCT AN INVESTIGATION OF THE REPORTED EVENT. | -poor quality image -poor image -low image accuracy | -fluid leak -Device Handling Issue -Human-Device Interface Issue | Viewer
| FDA | Event Description: IVUS MACHINE NOT HOLDING A CHARGE OR WORKING WHILE PLUGGED IN. BIOMED WAS CALLED YESTERDAY TO HELP TROUBLE SHOOT. MACHINE HAS BEEN PLUGGED INTO 3 DIFFERENT OUTLETS ALL DAY; WRITER PUSHED IN ALL CONNECTIONS TO ELECTRICAL PACK. MACHINE POWERS ON, RED BATTERY LIGHT FLASHES, AND WHEN MACHINE TURNS ON, IT IMMEDIATELY COUNTS DOWN AND SHUTS OFF. | -fail to conduct -fail to lock -unable to use | -Charging Issue -temperature issue -environmental control or utility issue | Viewer
| FDA | Event Description: MALE PATIENT WITH HISTORY OF PROSTATE ADENOCARCINOMA, STAGE IV. INITIALLY WAS DIAGNOSED APPROXIMATELY 17 YEARS AGO, AND STATUS POST (S/P) PROSTATECTOMY WITH SALVAGE RADIATION THERAPY. THE PATIENT ALSO RECEIVED HORMONAL THERAPY FOR SALVAGE TREATMENT WITH RISING PSA LEVELS BONE SCAN APPROXIMATELY 17 MONTHS AGO, DEMONSTRATED INCREASED RADIOTRACER UPTAKE IN RIGHT MID TO DISTAL TIBIA, STRONGLY SUGGESTIVE OF BLASTIC METASTASIS. SYSTEMIC TREATMENT CONTINUED. THE PATIENT PRESENTED TO HIS MEDICAL ONCOLOGIST WITH INTERMITTENT RIGHT LOWER TIBIA PAIN (10/10 AT WORST). FOLLOW UP X-RAY OF HIS RIGHT LOWER LEG APPROXIMATELY ONE YEAR AGO WAS ABNORMAL. PMH/PSH: URETHRAL STRICTURE, PROSTATECTOMY, URETERAL STENT X5, NEPHROSTOMY, CATARACT REMOVAL THE PATIENT WAS IMAGED ON THE MOBILE MRI UNIT ON APPROXIMATELY 10 MONTHS AGO BY 2 MRI TECHS. ONE TECH ATTENDED TO THE PATIENT, THE SECOND TECH ATTENDED TO THE SIEMENS MRI SCREEN. PRIOR FILMS NOT AVAILABLE FOR REFERENCE. DURING MRI REGISTRATION, THE PATIENT ORIENTATION (IN THE SCANNER) WAS SELECTED INCORRECTLY AS PRONE RATHER THAN SUPINE. (PRONE SELECTION DIRECTLY UNDER SUPINE SELECTION). PER PROTOCOL, A VITAMIN E CAPSULE WAS PLACED (TAPED) BY THE STAFF/PATIENT ON THE CORRECT (RIGHT) LEG. THE FLEXIBLE COIL WAS WRAPPED AROUND RIGHT LEG. THE PATIENT WENT INTO THE MRI SCANNER SUPINE AND FEET FIRST. IMAGING SEQUENCE PERFORMED. IMAGE OF THE LOCALIZER NOT SAVED IN PACS (TECH WAS TOLD RECENTLY THAT IT WAS NO LONGER NECESSARY). THE MRI TECH ATTENDING TO THE SCANNER WAS FOCUSED ON FINDING THE TUMOR ON THE SCAN AND DID NOT NOTICE THE DISCREPANCY IN THE ANATOMY. THE QUALITY OF THE IMAGE WAS NOT AS CLEAR AS IT COULD HAVE BEEN, BUT THE TECH HAD RECALLED SEEING SIMILAR QUALITY SCANS IN THE PAST. THE ORIGINAL INTERPRETING RADIOLOGIST DID NOT NOTICE THE DISCREPANCY. MRI LOW EXT W/WO CONT RT. COMPARISON: NONE AVAILABLE. CORRELATION: RIGHT TIBIA AND FIBULA FILMS OF APPROXIMATELY 7 WEEKS EARLIER. TECHNIQUE: AXIAL T1, FAT-SUPPRESSED T2, PRE-AND POSTCONTRAST FAT-SUPPRESSED T1, CORONAL AND SAGITTAL T1, IR AND POSTCONTRAST FAT-SUPPRESSED T1. TEN CC IV MULTIHANCE. IMPRESSION: NO EVIDENCE OF OSSEOUS METASTASIS INVOLVING THE RIGHT TIBIA OR FIBULA. NO EVIDENCE OF ACUTE FRACTURE OR SUBLUXATION. APPROXIMATELY 7 WEEKS LATER: THE EVENT WAS NOT IDENTIFIED, UNTIL THE CHART WAS RANDOMLY SELECTED FOR IMAGING PEER REVIEW. SUBSEQUENTLY, AN ADDENDUM TO THE ORIGINAL INTERPRETATION/FINDINGS WILL BE MADE BY ORIGINAL INTERPRETING RADIOLOGIST: THE MRI IMAGES HAVE BEEN RE-REVIEWED. THE ORDER WAS FOR AN MRI OF THE RIGHT LOWER EXTREMITY. IT WAS NOTED THAT THE MRI WAS CONDUCTED ON THE LEFT LOWER EXTREMITY. RECOMMEND PATIENT TO HAVE MRI OF RIGHT LOWER EXTREMITY. THE FINDINGS WERE DISCUSSED WITH ORDERING PHYSICIAN. APPROXIMATELY 8 WEEKS AFTER EARLIER MRI - PATIENT RETURNS FOR MRI OF RIGHT LEG. ABNORMAL FINDING IMPRESSION: METASTASIS (1ST PARAGRAPH) AND CONSEQUENT REACTIVE CHANGE (2ND PARAGRAPH). OSTEONECROSIS (3RD PARAGRAPH). CONSULT FOR PALLIATIVE RADIATION THERAPY. DIAGNOSIS: PROSTATE ADENOCARCINOMA METASTATIC TO RIGHT TIBIA, STAGE IV.Event Description: MALE PATIENT WITH HISTORY OF PROSTATE ADENOCARCINOMA, STAGE IV. INITIALLY WAS DIAGNOSED APPROXIMATELY 17 YEARS AGO, AND STATUS POST (S/P) PROSTATECTOMY WITH SALVAGE RADIATION THERAPY. THE PATIENT ALSO RECEIVED HORMONAL THERAPY FOR SALVAGE TREATMENT WITH RISING PSA LEVELS BONE SCAN APPROXIMATELY 17 MONTHS AGO, DEMONSTRATED INCREASED RADIOTRACER UPTAKE IN RIGHT MID TO DISTAL TIBIA, STRONGLY SUGGESTIVE OF BLASTIC METASTASIS. SYSTEMIC TREATMENT CONTINUED. THE PATIENT PRESENTED TO HIS MEDICAL ONCOLOGIST WITH INTERMITTENT RIGHT LOWER TIBIA PAIN (10/10 AT WORST). FOLLOW UP X-RAY OF HIS RIGHT LOWER LEG APPROXIMATELY ONE YEAR AGO WAS ABNORMAL. PMH/PSH: URETHRAL STRICTURE, PROSTATECTOMY, URETERAL STENT X5, NEPHROSTOMY, CATARACT REMOVAL THE PATIENT WAS IMAGED ON THE MOBILE MRI UNIT ON APPROXIMATELY 10 MONTHS AGO BY 2 MRI TECHS. ONE TECH ATTENDED TO THE PATIENT, THE SECOND TECH ATTENDED TO THE SIEMENS MRI SCREEN. PRIOR FILMS NOT AVAILABLE FOR REFERENCE. DURING MRI REGISTRATION, THE PATIENT ORIENTATION (IN THE SCANNER) WAS SELECTED INCORRECTLY AS PRONE RATHER THAN SUPINE. (PRONE SELECTION DIRECTLY UNDER SUPINE SELECTION). PER PROTOCOL, A VITAMIN E CAPSULE WAS PLACED (TAPED) BY THE STAFF/PATIENT ON THE CORRECT (RIGHT) LEG. THE FLEXIBLE COIL WAS WRAPPED AROUND RIGHT LEG. THE PATIENT WENT INTO THE MRI SCANNER SUPINE AND FEET FIRST . IMAGING SEQUENCE PERFORMED. IMAGE OF THE LOCALIZER NOT SAVED IN PACS (TECH WAS TOLD RECENTLY THAT IT WAS NO LONGER NECESSARY). THE MRI TECH ATTENDING TO THE SCANNER WAS FOCUSED ON FINDING THE TUMOR ON THE SCAN AND DID NOT NOTICE THE DISCREPANCY IN THE ANATOMY. THE QUALITY OF THE IMAGE WAS NOT AS CLEAR AS IT COULD HAVE BEEN, BUT THE TECH HAD RECALLED SEEING SIMILAR QUALITY SCANS IN THE PAST. THE ORIGINAL INTERPRETING RADIOLOGIST DID NOT NOTICE THE DISCREPANCY. MRI LOW EXT W/WO CONT RT COMPARISON: NONE AVAILABLE. CORRELATION: RIGHT TIBIA AND FIBULA FILMS OF APPROXIMATELY 7 WEEKS EARLIER. TECHNIQUE: AXIAL T1, FAT-SUPPRESSED T2, PRE-AND POSTCONTRAST FAT-SUPPRESSED T1, CORONAL AND SAGITTAL T1, IR AND POSTCONTRAST FAT-SUPPRESSED T1. 10 CC IV MULTIHANCE. IMPRESSION: NO EVIDENCE OF OSSEOUS METASTASIS INVOLVING THE RIGHT TIBIA OR FIBULA. NO EVIDENCE OF ACUTE FRACTURE OR SUBLUXATION. APPROXIMATELY 7 WEEKS LATER: THE EVENT WAS NOT IDENTIFIED, UNTIL THE CHART WAS RANDOMLY SELECTED FOR IMAGING PEER REVIEW SUBSEQUENTLY, AN ADDENDUM TO THE ORIGINAL INTERPRETATION/FINDINGS WILL BE MADE BY ORIGINAL INTERPRETING RADIOLOGIST: THE MRI IMAGES HAVE BEEN RE-REVIEWED. THE ORDER WAS FOR AN MRI OF THE RIGHT LOWER EXTREMITY. IT WAS NOTED THAT THE MRI WAS CONDUCTED ON THE LEFT LOWER EXTREMITY. RECOMMEND PATIENT TO HAVE MRI OF RIGHT LOWER EXTREMITY. THE FINDINGS WERE DISCUSSED WITH ORDERING PHYSICIAN. APPROXIMATELY 8 WEEKS AFTER EARLIER MRI - PATIENT RETURNS FOR MRI OF RIGHT LEG. ABNORMAL FINDING. IMPRESSION: METASTASIS (1ST PARAGRAPH) AND CONSEQUENT REACTIVE CHANGE (2ND PARAGRAPH). OSTEONECROSIS (3RD PARAGRAPH). CONSULT FOR PALLIATIVE RADIATION THERAPY. DIAGNOSIS: PROSTATE ADENOCARCINOMA METASTATIC TO RIGHT TIBIA, STAGE IV. | -tumor removed inappropriately -inappropriate mri scan -delayed monitoring | -compliance issue -device contamination with biological material -Device Operates Differently than Expected | Viewer
| FDA | Event Description: MALFORMED HL7 PATIENT IDS WERE SENT TO THIS MEDICAL DEVICE RESULTING IN DIFFERENT PATIENT'S STUDIES BEING IDENTIFIED AS THE SAME PATIENT. NONE OF THESE STUDIES WERE READ OR USED FOR DIAGNOSIS. THE ISSUE WAS DISCOVERED DURING A ROUTINE SERVICE CHECK, CUSTOMER WAS CONTACTED, AND THE ISSUE WAS CONTAINED. Manufacturer Narrative: THE HL7 STANDARD AND IHE TESTING FOR THE STANDARD COMBINE TO STATE THAT THE PID FIELDS WHICH CONTAIN PATIENT IDENTIFIERS ARE EXPECTED TO CONTAIN UNIQUE IDENTIFYING INFORMATION IN THE FIRST 15 DIGITS OF THE FIELD, A CARAT SEPARATOR, THEN 4 DIGITS FOR AN ASSIGNING AUTHORITY (TYPICALLY A FACILITY ID), FOR A TOTAL OF 20 DIGITS PER PID (THERE ARE THREE: PID 2, PID 3, PID 4). IN KEEPING WITH THE HL7 STANDARD AND IHE GUIDELINES/TESTING PARAMETERS, HEALTHMYNE IDENTIFIES PATIENTS BASED ON THE FIRST 15 DIGITS OF THE PID FIELDS. IF A FACILITY USES ALL 20 DIGITS AS THE UNIQUE IDENTIFYING INFORMATION AND THE DIGITS THAT UNIQUELY IDENTIFY THE PATIENT ARE WITHIN THE LAST FIVE DIGITS, MULTIPLE PATIENTS COULD BE MERGED TOGETHER AS ONE PATIENT IN HEALTHMYNE. THIS IS ONLY A PROBLEM IF HEALTHMYNE IS RECEIVING MALFORMED HL7 PIDS PER THE IHE GUIDELINES/TESTING PARAMETERS. HEALTHMYNE PARSES THE FIELD WITHOUT CHECKING IT FOR CONFORMANCE. FUTURE RELEASES WILL BE ENHANCED TO CHECK FOR MALFORMED/NON-CONFORMING PATIENT ID FIELDS TO FLAG AND CONTAIN THEM TO PREVENT THEIR USE UNTIL THE MALFORMATION OR NON-CONFORMANCE IS ADDRESSED BY THE FACILITY. ALL CUSTOMERS HAVE BEEN CONTACTED AND CHECKED. THIS OCCURRED AT ONLY ONE CUSTOMER WITH A SUBSET OF STUDIES, & NONE OF THE STUDIES WERE ACCESSED FOR INTERPRETATION. | -patient experienced adverse event -incorrect device orientation -device damaged | -improper or incorrect procedure or method -device markings issue -decrease in pressure | Viewer
| FDA | Event Description: MCKESSON IDENTIFIED THAT 8 STUDIES WERE MISSING ON THE SHORT TERM STORAGE LOCATION PRIOR TO PERFORMING A DATA MIGRATION. ALL 8 STUDIES WERE PREVIOUSLY REPORTED . 7 OF THESE REPORTED STUDIES WERE DETERMINED TO BE LOST AND ONE STUDY WAS IDENTIFIED AS BEING INDEXED TO AN INCORRECT PATIENT. NO PATIENT HARM WAS REPORTED AS A RESULT OF THIS ISSUE. | -insufficient storage location -missing maintenance -lack of storage document | -component missing -device markings issue -improper or incorrect procedure or method | Viewer
| FDA | Event Description: MEDTRONIC O-ARM WITH STEALTH SOFTWARE NOT CONNECTING TO PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS) PROPERLY AFTER UPGRADE, DELAYING PATIENT CARE. MANUFACTURER RESPONSE FOR O-ARM, MEDTRONIC STEALTH (PER SITE REPORTER). WORKED TO CONNECT SYSTEM. | -software corrupted -loss of communication -incorrect software settings | -use of device issue -device operates differently than expected -device disinfection or sterilization issue | Viewer
| FDA | Event Description: MERGE (B)(4) APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON 23 JUNE 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING STATING THAT "FOLLOWING DATABASE INTEGRITY CHECK, EXAMS STILL MISSING." WITH EXAMS NOT BEING AVAILABLE AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM FOR A PATIENT. THERE HAVE BEEN NO REPORTS OF DELAY IN TREATMENT OR DIAGNOSIS TO A PATIENT, HARM OR POTENTIAL HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | -fail to distribute medical imaging -fail to process medical imaging -fail to provide diagnostic test | -data issue -use of device issue -incorrect or inadequate result | Viewer
| FDA | Event Description: MERGE CADSTREAM IS INTENDED TO BE USED FOR VIEWING, ANALYZING, REPORTING, COMMUNICATING AND STORING DIGITAL IMAGES DERIVED FROM AN MR IMAGE MODALITY. CADSTREAM FUNCTIONALITY INCLUDES RECORDING, ANNOTATING, CALCULATIONS, PATIENT MEDICATION/INTERACTION INFORMATION, MONITORING AND TOOLS TO MODIFY/ENHANCE IMAGE AND DEMOGRAPHIC INFORMATION. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT IMAGES WERE TAKEN OF A PATIENT WITH THE INCORRECT LINEALITY AT THE MODALITY AND WERE APPEARING THAT WAY IN MERGE CADSTREAM. DUE TO A USER ERROR, IMAGES ACQUIRED DURING MRI SCANNING WERE LABELED WITH THE INCORRECT ORIENTATION ON THE MR SCANNER. PER THE CUSTOMER'S REQUEST, MANUAL PROCESSING OF STUDY BY MERGE SUPPORT RESULTED IN ALL IMAGES REFLECTING THE INCORRECT LATERALITY AND ANTERIOR/POSTERIOR POSITIONING EMBEDDED IN EVERY IMAGE. STUDY RESULTS HAVE THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND HAVE THE POTENTIAL TO IMPACT PATIENT TREATMENT WHICH COULD LEAD TO DEATH OR SERIOUS INJURY. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -fail to interpret images -fail to store data -unable to view display | -use of device issue -device markings issue -fluid leak | Viewer
| FDA | Event Description: MERGE CADSTREAM IS INTENDED TO BE USED FOR VIEWING, ANALYZING, REPORTING, COMMUNICATING AND STORING DIGITAL IMAGES DERIVED FROM AN MR IMAGE MODALITY. CADSTREAM FUNCTIONALITY INCLUDES RECORDING, ANNOTATING, CALCULATIONS, PATIENT MEDICATION/INTERACTION INFORMATION, MONITORING AND TOOLS TO MODIFY/ENHANCE IMAGE AND DEMOGRAPHIC INFORMATION. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE SUPPORT INDICATING THAT THERE WAS A USER THAT LABELED IMAGES ACQUIRED DURING MRI WITH THE INCORRECT ORIENTATION VIA THE MR SCANNER. I.E. PATIENT WAS PUT IN THE SCANNER FEET FIRST WHEN IT WAS EXPECTED THAT THE HEAD WOULD HAVE BEEN PUT IN FIRST. DUE TO THIS ERROR OF PATIENT POSITIONING AT THE MODALITY, THE CUSTOMER REQUESTED THAT MERGE SUPPORT MANUALLY PROCESS THE STUDY AND CHANGE THE LATERALITY TO ATTEMPT TO RECTIFY THE USER ERROR. THIS THEN CAUSES THE IMAGES TO REFLECT A DIFFERENT LATERALITY AND ANTERIOR/POSTERIOR POSITIONING EMBEDDED IN THE IMAGES VERSUS WHAT WAS ORIGINALLY TAKEN BY THE MODALITY. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -fail to interpret images -fail to store data -unable to view display | -use of device issue -device markings issue -fluid leak | Viewer
| FDA | Event Description: MERGE CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. MERGE CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. MERGE CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS. MERGE CADSTREAM ALSO INCLUDES USER-CONFIGURABLE FEATURES FOR REPORTING ON FINDINGS IN BREAST OR GENERAL MRI STUDIES. ADDITIONALLY, MERGE CADSTREAM ASSISTS USERS IN PLANNING MRI GUIDED INTERVENTIONAL PROCEDURES. WHEN INTERPRETED BY A SKILLED PHYSICIAN, THIS DEVICE PROVIDES INFORMATION THAT MAY BE USED FOR SCREENING, DIAGNOSIS, AND INTERVENTIONAL PLANNING. PATIENT MANAGEMENT DECISIONS SHOULD NOT BE MADE BASED SOLELY ON THE RESULTS OF MERGE CADSTREAM. MERGE CADSTREAM MAY ALSO BE USED AS AN IMAGE VIEWER OF MULTI-MODALITY, DIGITAL IMAGES, INCLUDING ULTRASOUND AND MAMMOGRAPHY. MERGE CADSTREAM IS NOT INTENDED FOR PRIMARY INTERPRETATION OF DIGITAL MAMMOGRAPHY IMAGES. ON DECEMBER 1, 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH PROCESSING A STUDY WHICH ENCOUNTERED A PROCESSING ERROR. THROUGH TROUBLESHOOTING, SUPPORT DISCOVERED THE PATIENT ORIENTATION HAD BEEN ENTERED INCORRECTLY AT THE SCANNER THE CUSTOMER REQUESTED ADJUSTMENTS TO THE PROCESSING INSTRUCTIONS TO REACH A READY TO READ STATUS, HOWEVER, THE PATIENT ORIENTATION COULD NOT BE CORRECTED. CADSTREAM IS UNABLE TO EDIT THE ORIGINAL IMAGE DICOM DATA, AS SUCH, THE PROCESSED STUDY WILL REFLECT THE INCORRECT PATIENT POSITIONING ON ALL POST PROCESSING IMAGES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD, AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to receive prescribed mri scan | -use of device issue -electrical shorting -protective measure issue | Viewer
| FDA | Event Description: MERGE CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. MERGE CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. MERGE CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS. MERGE CADSTREAM ALSO INCLUDES USER-CONFIGURABLE FEATURES FOR REPORTING ON FINDINGS IN BREAST OR GENERAL MRI STUDIES. ADDITIONALLY, MERGE CADSTREAM ASSISTS USERS IN PLANNING MRI GUIDED INTERVENTIONAL PROCEDURES. WHEN INTERPRETED BY A SKILLED PHYSICIAN, THIS DEVICE PROVIDES INFORMATION THAT MAY BE USED FOR SCREENING, DIAGNOSIS, AND INTERVENTIONAL PLANNING. PATIENT MANAGEMENT DECISIONS SHOULD NOT BE MADE BASED SOLELY ON THE RESULTS OF MERGE CADSTREAM. MERGE CADSTREAM MAY ALSO BE USED AS AN IMAGE VIEWER OF MULTI-MODALITY, DIGITAL IMAGES, INCLUDING ULTRASOUND AND MAMMOGRAPHY. MERGE CADSTREAM IS NOT INTENDED FOR PRIMAR INTERPRETATION OF DIGITAL MAMMOGRAPHY IMAGES. ON (B)(4) 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT THE FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT POSITIONING WHICH WAS ENTERED INCORRECTLY AT THE MODALITY. CADSTREAM IS UNABLE TO EDIT THE ORIGINAL IMAGE DICOM DATA, AS SUCH, THE PROCESSED STUDY WILL REFLECT THE INCORRECT PATIENT POSITING ON ALL POST PROCESSING IMAGES. STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD, AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THEREFORE THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). | -mri image distorted -fail to enable mri setting -fail to receive prescribed mri scan | -use of device issue -electrical shorting -protective measure issue | Viewer
| FDA | Event Description: MERGE CADSTREAM IS INTENDED TO BE USED IN THE VISUALIZATION, ANALYSIS, AND REPORTING OF MAGNETIC RESONANCE IMAGING (MRI) STUDIES. MERGE CADSTREAM SUPPORTS EVALUATION OF DYNAMIC MR DATA ACQUIRED DURING CONTRAST ADMINISTRATION. MERGE CADSTREAM PERFORMS OTHER USER SELECTED PROCESSING FUNCTIONS SUCH AS IMAGE REGISTRATION, SUBTRACTIONS, MEASUREMENTS, 3D RENDERINGS, AND REFORMATS. ON (B)(6) 2017, MERGE SUPPORT WAS CONTACTED BY A USER AT A FACILITY FOR ASSISTANCE WITH CORRECTING THE PATIENT POSITIONING WHICH WAS ENTERED INCORRECTLY AT THE MODALITY. CADSTREAM IS UNABLE TO EDIT THE ORIGINAL IMAGE DICOM DATA, AS SUCH, THE PROCESSED STUDY WILL REFLECT THE INCORRECT PATIENT POSITIONG ON ALL POST PROCESSING IMAGES. DUE TO STUDY RESULTS HAVING THE POTENTIAL TO BECOME PART OF THE PATIENTS PERMANENT RECORD AND RECORDS HAVE THE POTENTIAL TO IMPACT A PATIENT'S TREATMENT, THERE IS A POSSIBILITY FOR A MISDIAGNOSIS OR MISTREATMENT THAT COULD LEAD TO HARM. THERE IS NO INDICATION THAT THIS ISSUE AS REPORTED BY THE CUSTOMER HAS RESULTED IN ANY HARM TO A PATIENT. (B)(4). | -mri image distorted -fail to enable mri setting -fail to receive prescribed mri scan | -use of device issue -electrical shorting -protective measure issue | Viewer
| FDA | Event Description: MERGE CARDIO 10.1 REPORTS DO NOT DISPLAY PATIENT IDENTIFIERS IN THE HEADER OF EACH PAGE, THEY ARE ONLY DISPLAYED ON THE FIRST PAGE. IF REPORTS ARE PRINTED TO HARD COPY, PAGES COULD POTENTIALLY BECOME MIXED UP BETWEEN DIFFERENT PATIENTS. PRINTING REPORTS TO HARD COPY SO THEY CAN BE FAXED IS NOT A COMMON METHOD OF DELIVERY BUT IS USED BY THE CUSTOMER THAT REPORTED THIS ISSUE. USER ERROR IS THE MAIN CONTRIBUTING FACTOR. THE FREQUENT USE OF MASS FAXES AND PRINTOUTS IS MAKING IT MORE LIKELY THAT PAGES GET MIXED UP. THIS WOULD NOT BE AN ISSUE IF ALL REPORTS ARE ELECTRONICALLY DISTRIBUTED AS PDFS. | -incorrect display value -inadequate message display -fail to display message | -incorrect display -device alarm system issue -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT WHEN USING THE MEASUREMENT TOOL TO MEASURE RIGHT VENTRICLE TAPE, IN M-MODE THAT THE DECIMAL PLACE IS SHIFTED CAUSING THE MEASUREMENT TO LOOK INCORRECT IN THE REPORT AND WORKSTATION. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REQUEST RE-MAPPING OF 2 MEASUREMENTS: MITRAL REGURGITATION MEASUREMENT IS POPULATING THE MITRAL VALVE GRADIENT FIELD, AND WHEN MEASURING BOTH MITRAL REGURGITATION AND MITRAL VALVE GRADIENT, THE INCORRECT MR MEASUREMENT IS OVERWRITTEN WITH THE CORRECT MV GRADIENT MEASUREMENT. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT PREND VMAX WAS MAPPED TO PR PEAK VELOCITY AND SHOULD BE MAPPED TO PR PEAK IN END DIASTOLE AND THAT PREND PG WAS MAPPED TO PR PEAK GRADIENT AND SHOULD BE MAPPED TO PR GRADIENT IN END DIASTOLE. WORK BETWEEN MERGE HEALTHCARE AND THE CUSTOMER IS ONGOING AND A SUPPLEMENTAL REPORT WILL BE FILED WHEN MROE INFORMATION BECOMES AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THE LV EF (LEFT VENTRICULAR EJECTION FRACTION), LV ID ES PLAX (LEFT VENTRICULAR INNER DIMENSION AT END SYSTOLE, PARASTERNAL LONG-AXIS VIEW), AND RV PEAK PRESSURE (RIGHT VENTRICULAR PEAK PRESSURE) MEASUREMENTS WERE DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. THESE MAPPING STRINGS WERE UPDATED AND THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON 1/11/2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON 1/15/2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THE LEFT ATRIAL INTERNAL DIAMETER MEASUREMENT WAS DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. THIS MAPPING STRING WAS UPDATED AND THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, AN INTERNAL INVESTIGATION REVEALED THAT DUE TO TWO MEASUREMENTS BEING MAPPED TO THE SAME CODE, ONE MEASUREMENT COULD OVERWRITE THE OTHER IN THE CLINICAL CARDIO REPORT. THE "MR FLOW RATE" MEASUREMENT AND THE "MAX VELOCITY (MR)" MEASUREMENT WERE BOTH MAPPED TO CPCODE (10607). INTERNAL RESOLUTION IS IN PROCESS. A SUPPLEMENTAL REPORT WILL BE FILED WHEN MORE INFORMATION IS AVAILABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO CLIENT REPORTS OF THIS ISSUE. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE E/E MEASUREMENT (THE RATIO BETWEEN EARLY MITRAL INFLOW VELOCITY AND MITRAL ANNULAR EARLY DIASTOLIC VELOCITY) WAS DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. THIS MAPPING STRING FOR E/E WAS UPDATED AND THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS REQUIRED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE MITRAL VALVE PEAK E VELOCITY WAS DISPLAYING INCORRECTLY ON THE CARDIO CLINICAL REPORT. UPON FURTHER INVESTIGATION, MEASUREMENT MAPPING STRINGS WERE ANALYZED BY MERGE HEALTHCARE SUPPORT AND IT WAS DETERMINED THAT THE E/E' AVERAGE WAS BEING MAPPED TO THE MITRAL VALVE PEAK E VELOCITY FIELD. THIS MAPPING STRING WAS REMOVED THE ISSUE WAS CONFIRMED TO BE RESOLVED. NO FURTHER ACTION IS REQUIRED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT STATING THAT A RECENT REPORT WAS UNAVAILABLE IN THEIR EMR AND THAT THIS HAD RESULTED IN A DELAY OF CARE FOR A PATIENT. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT FAILED TO EXPORT TO THE EMR BECAUSE HOSPITAL STAFF ENTERED "/WIFE" IN THE PHONE NUMBER FIELD OF THE REPORT AND THE HL7 OUTBOUND MESSAGE WILL FAIL IF IT ENCOUNTERS AN UNEXPECTED FORMAT, SUCH AS LETTERS AND/OR SPECIAL CHARACTERS IN A FIELD DESIGNED FOR ONLY NUMBERS. MERGE HEALTHCARE SUPPORT WAS ABLE TO REMOVE THESE CHARACTERS FROM THE FIELD IN QUESTION AND CONFIRMED THAT THE EXPORT OF THE REPORT TO THE EMR WAS SUCCESSFUL. ADDITIONALLY, MERGE HEALTHCARE SUPPORT WORKED WITH THE CLIENT TO CHANGE THE HL7 FIELDS SO THAT PHONE NUMBERS ARE NO LONGER INCLUDED, SO AS TO PREVENT THIS ISSUE FROM CAUSING FURTHER POSSIBLE DELAYS TO PATIENT CARE. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON 10/6//2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, DURING INTERNAL TESTING OF AN UPCOMING VERSION, DUPLICATE MEASUREMENT MAPPING STRINGS WERE DISCOVERED. UPON FURTHER INVESTIGATION, MERGE HEALTHCARE INTERNAL STAFF DISCOVERED THAT FOR GE VIVID 7 MODALITIES WITH XLS FILE DROPS, 2 MEASUREMENTS (AAO AP DIAM, S" AND INFERIOR VENA CAVA DIAM) HAVE DUPLICATE MAPPING STRINGS AND THE 2 STRINGS HAVE DIFFERENT UNITS OF MEASURE. ONE STRING HAS MM AND THE OTHER STRING HAS CM. THE CONFLICT OF UNITS CAUSES THE MEASUREMENTS TO IMPORT INCORRECTLY. THE NUMERICAL VALUE OF THE MEASUREMENT ROUNDS UP AND THE DECIMAL POINT MOVES. FOR EXAMPLE 3.7CM ROUNDS UP TO 4 AND THEN DISPLAYS ON THE REPORT AS 0.4CM. TO RESOLVE THIS ISSUE, MERGE HEALTHCARE INTERNAL STAFF DELETED ALL ROWS IN THE MAPPING TABLE, POPULATED NEW MAPPINGS, AND VERIFIED THAT THE CORRECT MAPPINGS WERE RETAINED FOR EACH OF THE FORMERLY DUPLICATED REPORT TEXTS. NO FURTHER ACTION IS REQUIRED. DUE TO INCORRECT MEASUREMENTS DISPLAYING ON A DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR MISDIAGNOSIS AND INCORRECT TREATMENT THAT MAY LEAD TO HARM. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO CUSTOMER REPORTS OF THIS ISSUE. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE MITRAL REGURG PISA RADIUS WAS MAPPED INCORRECTLY. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, THE MEASUREMENT WAS REMOVED AND REMAPPED. IT WAS CONFIRMED TO THEM BE MAPPED CORRECTLY IN THE CUSTOMER'S SYSTEM. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT ONE MEASUREMENT WAS INCORRECTLY MAPPED. MERGE HEALTHCARE IS CURRENTLY WORKING WITH THE CUSTOMER TO RESOLVE THIS AND AWAITING THE CUSTOMER'S RETURN FROM VACATION. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT STRESS DATA WAS LABELED INCORRECTLY. MERGE HEALTHCARE SUPPORT HAS BEEN IN CONTACT WITH THE CUSTOMER MULTIPLE TIMES SINCE THIS REPORT AND IS CURRENTLY WORKING ON GATHERING FURTHER INFORMATION. AT THIS TIME, THERE ARE NO REPORTS OF PATIENT HARM OR INJURY DUE TO THIS ISSUE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED AS SOON AS MORE INFORMATION IS AVAILABLE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THERE WERE ISSUES WITH MEASUREMENT MAPPINGS IN THE CARDIO CLINICAL REPORT. THEY STATED THAT THE IVC (INFERIOR VENA CAVA) MEASUREMENT WAS POPULATING THE IVS (INTERVENTRICULAR SEPTUM) FIELD, THAT THE AORTIC REGURGITATION DECELERATION MEASUREMENT HAD TOO MANY DIGITS, AND THAT THE DESCENDING AORTA ID APPEARED TO BE MISMAPPED. AT THIS TIME, INVESTIGATION BETWEEN THE CUSTOMER AND MERGE HEALTHCARE IS ONGOING. DUE TO INCORRECT VALUES DISPLAYING IN THE CARDIO CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM (B)(4) STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT FROM ONE PHILIPS MODALITY, THE LEFT VENTRICULAR OUTFLOW TRACK (LVOT) MEASUREMENT WAS POPULATING IN THE CLINICAL CARDIO REPORT INCORRECTLY AND THAT THE CORRECT LVOT MEASUREMENT WAS POPULATING IN THE TR (TRICUSPID REGURGITATION) VENA CONTRACTA FIELD. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM (B)(4) STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE STATING THAT IVC (INFERIOR VENA CAVA) MEASUREMENTS WERE MISSING IN CLINICAL REPORTING AND THAT THE IVC MEASUREMENT WAS POPULATING THE IVS (INTRAVENTRICULAR SEPTUM) FIELD ON THE CLINICAL CARDIO REPORT. DUE TO INCORRECT VALUES DISPLAYING ON THE CLINICAL CARDIO REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. IN (B)(6) 2020 MERGE HEALTHCARE INITIATED AN INVESTIGATION FOR AN INTERNALLY IDENTIFIED DISCREPANCY IN PEDIATRIC ECHOGRAPHY MICHIGAN Z-SCORE REFERENCE RANGES. THESE REFERENCE RANGES ARE CALCULATED BY A SUPPLIERS'S DATABASE THAT IS INSTALLED WITH THE MERGE CARDIO PRODUCT. BOTH MERGE HEALTHCARE AND THE PEADIATRIC ECHO KNOWLEDGE BASE SUPPLIER ASCEND'S INVESTIGATION CONCLUDED THAT THERE WAS AN ERROR AND THAT THE MICHIGAN REFERENCE RANGES WERE NOT CALCULATING CORRECTLY. ASCEND'S 3.X VERSION OF THEIR PEDIATRIC ECHO KNOWLEDGE BASE CALCULATES MICHIGAN REFERENCE RANGES INCORRECTLY. THIS DEFECT RESULTS IN WIDER THAN EXPECTED UPPER AND LOWER LIMIT REFERENCE RANGES. THEREFORE, WHEN A PATIENT'S Z-SCORE, FOR A PARTICULAR MEASUREMENT, IS EVALUATED AGAINST THE RANGE, A FALSE NEGATIVE MAY BE INDICATED BY THE WIDER RANGE OF ACCEPTABLE REFERENCE RANGE VALUES. AS A RESULT OF THIS REFERENCE RANGE MISCALCULATION, THERE IS A POTENTIAL FOR A MISDIAGNOSIS WHICH MAY LEAD TO PATIENT HARM. HOWEVER, THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. DUE TO THE MICHIGAN Z-SCORE REFERENCE RANGE INVESTIGATION AND SUPPLIER ASCEND CONFIRMING A DEFECT IN THEIR PEDIATRIC ECHO DATABASE, A HISTORICAL ANALYSIS OF SERVICE REQUESTS AND COMPLAINTS SUBMITTED TO MERGE HEALTHCARE WAS PERFORMED. A RECORD WITH SIMILAR INFORMATION TO THE REFERENCE RANGE INVESTIGATION WAS IDENTIFIED AND REASSESSED. ON (B)(6) 2017, A REQUEST FOR INFORMATION CAME IN FROM A MERGE CARDIO CUSTOMER. THE CUSTOMER WAS REQUESTING ADDITIONAL INFORMATION FOR THE REFERENCE RANGES RELATED TO Z-SCORES. DUE TO THE CUSTOMER REQUESTING INFORMATION ABOUT HOW THE REFERENCE RANGES WERE CALCULATED, THIS RECORD WAS CLASSIFIED, AT THE TIME, AS NOT A COMPLAINT. HOWEVER, WITH THE CONFIRMATION FROM SUPPLIER ASCEND THAT THERE IS A DEFECT IN THEIR 3.X DATABASE'S CALCULATION OF MICHIGAN REFERENCE RANGES, THIS HISTORICAL RECORD HAS BEEN RE-EVALUATED AND DETERMINED TO MEET THE DEFINITION OF A COMPLAINT AND THEREFORE, MEETS MALFUNCTION REPORTING REQUIREMENTS. THE CUSTOMER DID NOT ALLEGE ANY INJURY OR IMPACT TO PATIENTS. MERGE HEALTHCARE TECHNICAL SUPPORT WORKED WITH THE CUSTOMER AND CONFIGURED THEIR DATABASE TO PRODUCE THE EXPECTED MICHIGAN REFERENCE RANGES. THE CUSTOMER'S REQUEST WAS RESOLVED. THIS HISTORICAL RECORD IS BEING REPORTED BECAUSE OF AN INVESTIGATION DETERMINED THAT CONFIRMED A DEFECT RESIDES IN SUPPLIER ASCEND'S DATABASE. MERGE HEALTHCARE'S SUPPLIER ASCEND CONFIRMED THE DEFECT ON (B)(6) 2020. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM (B)(4) STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, MERGE HEALTHCARE WAS PERFORMING INTERNAL TESTING OF A NEW 5.X PEDIATRIC ECHO KNOWLEDGE BASE FILE FOR MICHIGAN Z-SCORES. IT WAS FOUND THAT WHILE TESTING THIS VERSION 5.X FILE AND COMPARING WITH THE PREVIOUSLY RELEASED VERSION 3.X FILE FOR MICHIGAN Z-SCORE CALCULATIONS, THAT THERE WAS A DIFFERENCE BETWEEN THE TWO. BOTH THE INTERNAL MERGE HEALTHCARE INVESTIGATION AND THE PAEDIATRIC ECHO KNOWLEDGE BASE SUPPLIER ASCEND'S INVESTIGATION CONCLUDED THAT THERE WAS AN ERROR AND THAT THE MICHIGAN REFERENCE RANGES AND Z-SCORES ARE NOT CALCULATING CORRECTLY. FOR CUSTOMERS USING A 3.X VERSION OF THE ASCEND PEDIATRIC ECHO KNOWLEDGE BASE TO CALCULATE MICHIGAN Z-SCORES, THERE IS A POTENTIAL FOR A MISDIAGNOSIS WHICH MAY LEAD TO PATIENT HARM. HOWEVER, THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT-(B)(6). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, DURING INTERNAL TESTING OF A NEW RELEASE BY MERGE HEALTHCARE, IT WAS FOUND THAT THE LENGTH MEASUREMENTS IMPORT TO THE DEPTH LOCATION IN THE VEIN MEASUREMENT TABLE FOR THE VEIN ANATOMY OF THE LEG. THIS OCCURS IN THE 5.0 VERSION OF VASCULAR REPORTING WHICH IS CURRENTLY IN CARDIO VERSIONS 11.0-12.0.1. DEPTH MEASUREMENTS CURRENTLY DO NOT MAP SO THEY WERE NOT OVERWRITTEN. THE ISSUE IS THAT LENGTH MEASUREMENTS POPULATE THE DEPTH FIELDS. THIS WAS FOUND TO IMPACT ALL OF THE VEINS FOR LOWER EXTREMITIES. THERE HAVE BEEN ZERO REPORTS OF THIS ISSUE BY A CUSTOMER. DUE TO LENGTH MEASUREMENTS POPULATING THE DEPTH FIELD, THERE IS A POTENTIAL FOR MISDIAGNOSIS. HOWEVER, THERE ARE ZERO REPORTS OF IMPACT TO A PATIENT AND NO CUSTOMERS HAVE REPORTED THIS ISSUE. REFERENCE COMPLAINT-(B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE, AND OTHER VENDORS SYSTEMS, INCLUDING IMAGES, HEMODYNAMIC STUDIES, AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES), AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE, AND STATED THAT THREE MEASUREMENTS WERE MAPPED INCORRECTLY (MEAN VELOCITY RATIO, AORTIC VALVE AREA MEAN VELOCITY, AND AORTIC VALVE AREA BSA MEAN VELOCITY). MERGE HEALTHCARE IS CURRENTLY WORKING WITH THE CUSTOMER TO RESOLVE THIS ISSUE WITH THE CUSTOMER . THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY, OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to enter patient data | -use of device issue -data issue -incorrect or inadequate result | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE. HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY. SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE. STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT A RECENT REPORT WAS NOT AVAILABLE FOR VIEWING. MERGE HEALTHCARE SUPPORT FOUND THAT THE REPORT WAS NEVER CONFIRMED BUT WAS IN CONFIRMED STATUS. THE LOGS HAD BEEN TRUNCATED SO THEY WERE UNAVAILABLE FOR TROUBLESHOOTING PURPOSES. SUPPORT MANUALLY RESET THE STATUS OF THE REPORT AND IT WAS THEN ABLE TO BE READ AND CONFIRMED BY THE PHYSICIAN. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to access clinical history -inadequate clinical history data -fail to record diagnosis data | -fluid leak -use of device issue -electrical issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM (B)(6) STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, AT THIS TIME THE CUSTOMER HAS NOT ALLEGED ANY HARM OR IMPACT TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT FETAL MEASUREMENTS TAKEN BY THEIR 3RD PARTY MODALITIES WERE POPULATING THE MERGE CARDIO CLINICAL REPORTING FETAL KNOWLEDGEBASE WITH INCORRECT Z-SCORE VALUES. WITH MERGE CARDIO INCORRECTLY POPULATING Z-SCORE VALUES THAT ARE USED FOR DIAGNOSIS, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. AT THIS TIME NO INFORMATION HAS BEEN PROVIDED THAT INDICATES THAT A DEATH OR SERIOUS INJURY HAS OCCURRED AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING WITH THE INCORRECT UNIT OF MEASURE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MERGE CARDIO'S PDF REPORTS WERE POSTING UNDER THE WRONG PATIENT IN THEIR HOSPITAL INFORMATION SYSTEM. CLINICAL REPORTS POSTING UNDER THE WRONG PATIENT HAS POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT WHICH MAY LEAD TO PATIENT HARM. THE CUSTOMER WAS CONTACTED AND STATED THAT NO PATIENT HARM WAS CAUSED DUE TO THIS ISSUE AND THAT IT IS READILY APPARENT TO THE USER. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THEY COULD NOT FIND A STUDY. THIS STUDY WAS COMPLETED UNDER A DIFFERENT PATIENT/ORDER IN ERROR AND THEN THE PHYSICIAN REPORTED ON THE IMAGES WITH THE INCORRECT PATIENT INFORMATION AND PATIENT ORDER. THE CLINICAL REPORT WAS CONFIRMED ON THE WRONG PATIENT WITH THE WRONG PATIENT INFORMATION, THIS HAS THE POTENTIAL FOR INCORRECT DIAGNOSIS OR TREATMENT OF THE PATIENT WHICH MAY LEAD TO PATIENT HARM. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER ALLEGED TO MERGE HEALTHCARE THAT A SOFTWARE MALFUNCTION OCCURRED WHICH ALLOWED A CLINICAL REPORT TO BE CONFIRMED WITH AN INCORRECT PATIENT IDENTIFIER. THIS MALFUNCTION, IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MERGE CARDIO'S PDF REPORTS WERE POSTING UNDER THE WRONG PATIENT IN THEIR HOSPITAL INFORMATION SYSTEM (HIS). CLINICAL REPORTS POSTING UNDER THE WRONG PATIENT HAS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT WHICH MAY LEAD TO HARM. THE CUSTOMER REPORTED THAT NO PATIENT HARM WAS CAUSED AS A RESULT OF THIS ISSUE AND THAT THE ISSUE IS READILY APPARENT TO THE USER. (B)(4) | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT HEMODYNAMIC DATA WAS INCORRECTLY IMPORTED TO THE WRONG PATIENT'S STUDY. INCORRECT DIAGNOSTIC DATA BEING IMPORTED TO A STUDY HAS THE POTENTIAL TO CAUSE INCORRECT TREATMENT OF THE PATIENT, WHICH COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE PULMONARY ARTERY Z-SCORES WERE NOT CALCULATING AS EXPECTED IN THEIR ECHO.PEDS CLINICAL REPORT. WITH AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THE CUSTOMER INDICATED THAT THERE WERE NO PATIENTS THAT WERE TREATED INCORRECTLY BASED ON THE OUTCOME OF THE PERCEIVED INCORRECT VALUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE ALLEGING THAT THE LEFT CORONARY ARTERY Z-SCORES WERE NOT CALCULATING CORRECTLY IN THEIR ECHO.PEDS CLINICAL REPORT. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THE CUSTOMER WAS CONTACTED AND NO PATIENTS WERE TREATED INCORRECTLY BASED ON THE OUTCOME OF THE INCORRECT VALUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, MERGE HEALTHCARE RECEIVED AN ALLEGATION FROM A CUSTOMER INDICATING THAT THE MEASUREMENT TOOL ON THE CARDIO WORKSTATION PROVIDED INACCURATE RESULTS FOR MEAN GRADIENT. MERGE INVESTIGATED THIS ALLEGATION AND DETERMINED THAT THE CUSTOMER WAS QUESTIONING THE NO VALUE MEAN GRADIENT CALCULATION. ALL OTHER MEAN GRADIENT CALCULATIONS WERE CONSISTENT WITH THE CUSTOMER'S EXPECTATIONS. POTENTIALLY INACCURATE CALCULATIONS CAN LEAD TO CONFUSION BY USER(S) AND MAY LEAD TO INCORRECT TREATMENT AND/OR A DELAY IN CARE THAT MAY CAUSE HARM TO A PATIENT. ACCORDING TO THE CUSTOMER, THE PATIENT IN WHICH THIS ISSUE WAS FOUND, WAS NOT INCORRECTLY DIAGNOSED AS THEY RECOGNIZED THE DISCREPANCY AND MODIFIED THE VALUE PRIOR TO COMPLETING THE REPORT. THE PHYSICIAN STATED THAT THE MEASUREMENT TOOL WAS ONLY USED RARELY TO MANUALLY CALCULATE MEAN GRADIENTS. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THE INCORRECT VALUE WAS READILY APPARENT TO PHYSICIANS AND THERE IS NO INFORMATION TO SUGGEST THAT A PATIENT WAS HARMED AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT CLINICAL REPORTS THAT HAD NOT BEEN REVIEWED BY A PHYSICIAN WERE DISPLAYING WITH A "CONFIRMED" STATUS. AN INCORRECT CONFIRMATION STATUS ON A CLINICAL REPORT THAT HAS NOT BEEN REVIEWED BY A PHYSICIAN COULD CAUSE A DELAY IN PATIENT CARE, WHICH COULD POTENTIALLY HARM A PATIENT. HOWEVER, THE CUSTOMER STATED THAT NO PATIENTS WERE ADVERSELY IMPACTED DUE TO THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT THE HEIGHT MEASUREMENT WAS IMPORTING INCORRECTLY TO THE CLINICAL REPORT CAUSING MISCALCULATION OF THE BSA. DUE TO THE INCORRECT BSA CALCULATION IN THE CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT THAT COULD RESULT IN HARM. THERE IS NO INDICATION OF ANY PATIENT HARM THAT HAS OCCURRED AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT ALL PROXIMAL AORTIC ARCH MEASUREMENTS ARE BEING REPORTED AS DISTAL AORTIC ARCH MEASUREMENTS AND DISTAL AORTIC ARCH MEASUREMENTS ARE BEING REPORTED AS PROXIMAL AORTIC ARCH IN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE IS NO INDICATION THAT ANY PATIENTS HAVE BEEN HARMED AS A RESULT OF THIS MEASUREMENT MAPPING ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT THE HEADERS ON ECHO AND STRESS ECHO REPORTS CONTAINING UNIQUE IDENTIFYING INFORMATION (PATIENT NAME, PATIENT ID, AND STUDY DATE/TIME) WAS ONLY VIEWABLE ON THE FIRST PAGE OF THE REPORT AND WAS NOT INCLUDED ON ANY SUBSEQUENT PAGES. WHEN MERGE CARDIO DOES NOT PROVIDE THE CORRECT UNIQUE IDENTIFIER IN THE HEADER OF A PATIENT'S REPORT, THERE IS A POTENTIAL FOR PATIENT DATA TO BE COMPROMISED WHICH COULD POTENTIALLY LEAD TO AN INCORRECT TREATMENT OR DIAGNOSIS THAT MAY RESULT IN HARM TO THE PATIENT. INFORMATION OBTAINED REVEALED THAT THE USER WAS AWARE OF THE PROBLEM, NO PATIENT DATA WAS COMPROMISED, AND NO PATIENT HARM OCCURRED. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT DATA WAS NOT IMPORTING INTO THEIR STRESS TABLE WITHIN CLINICAL REPORTING, FOR A STUDY PERFORMED ON (B)(6) 2016. (B)(6) 2016 WHICH REVEALED THAT THE INCORRECT TARGET HEART RATE IS BEING USED FOR STRESS ECHO CALCULATIONS. IF THE WRONG TARGET HEART RATE IS USED FOR DIAGNOSIS, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING WITH THE INCORRECT UNIT OF MEASURE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE IS NO INDICATION THAT ANY ACTUAL PATIENT HARM. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: THE ISSUE WAS BEING MANUALLY CORRECTED BY USERS AT THE FACILITY. FURTHER INVESTIGATION OF THE ISSUE DETERMINED THAT THE MEASUREMENT WAS INCORRECTLY MAPPED AT THE TIME OF IMPLEMENTATION. MERGE HEALTHCARE SUPPORT CORRECTED THE MAPPING, AND CONFIRMED WITH THE FACILITY THAT THE CORRECT VALUE IS NOW IMPORTED. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE CLINICAL REPORTING TOOL WITH THE INCORRECT UNIT OF MEASURE. DUE TO THE INCORRECT UNIT OF MEASURE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. THERE IS NO INDICATION OF ANY ACTUAL HARM AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHILE READING IMAGES OF STRESS STUDIES THE SPEED INDICATOR DOES NOT MATCH THE ACTUAL SPEED IN FRAMES PER SECOND (FPS) AND ARE INCONSISTENT. DUE TO THE INCONSISTENT FRAMES PER SECOND READING, THERE IS A POTENTIAL FOR A MISDIAGNOSIS OR MIS-TREATMENT FOR A PATIENT. THERE IS NO INDICATION THAT THERE HAS BEEN ANY PATIENT HARM DUE TO THIS INCONSISTENCY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A FETAL PATIENT REPORT WAS AUTOMATICALLY PULLING PRIOR MEASUREMENT DATA FOR A PRIOR FETUS; HOWEVER, THE ATTENDING MEDICAL STAFF IMMEDIATELY NOTICED THE PROBLEM BEFORE THE PATIENTS' DATA WAS COMPROMISED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THAT THE WRONG FETAL MEASUREMENTS GET REFERENCED, WHICH CAN LEAD TO A POTENTIAL FOR OVER OR UNDER DIAGNOSIS AND/OR TREATMENT OF A PATIENT. (B)(4). Manufacturer Narrative: DUE TO THE POTENTIAL FOR MULTIPLE CUSTOMERS TO HAVE THE SAME REPORTED PROBLEM, A RECALL IS BEING INITIATED. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A FETAL PATIENT REPORT WAS AUTOMATICALLY PULLING PRIOR MEASUREMENT DATA FOR A PRIOR FETUS; HOWEVER, THE ATTENDING MEDICAL STAFF IMMEDIATELY NOTICED THE PROBLEM BEFORE THE PATIENTS' DATA WAS COMPROMISED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THAT THE WRONG FETAL MEASUREMENTS GET REFERENCED, CAUSING OVER OR UNDER DIAGNOSIS AND/OR TREATMENT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT MEASUREMENTS WERE NOT BEING POPULATED INTO THE APPLICATION AND THE REPORT. WHILE NOT SPECIFIC IN THIS CASE, THIS MAY IMPACT THE PHYSICIAN'S ABILITY TO COMPLETE THE REPORT AND/OR FINALIZE DIAGNOSIS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. CUSTOMER REPORTED THAT MEASUREMENTS ARE NOT POPULATING IN THE REPORT. THERE WERE 4 IMAGES THAT HAVE MEASUREMENTS ON THEM, HOWEVER THE MEASUREMENTS ARE NOT POPULATING FROM THE IMAGES FROM THE REPORT. TECHNICAL SUPPORT DETERMINED THAT THE SEQUIA MODALITIES SHOULD NOT SEND IN SR FILES. THIS RESULTED IN A LOT OF UNMAPPED MEASUREMENTS THAT WERE PROPRIETORY MEASUREMENTS SENT FROM THE MODALITY. TECHNICAL SUPPORT ASKED FOR A TEST PATIENT IN ORDER TO MAP SOME THE MEASUREMENTS TO MEET THE MODALITY. THE SITE LATER REPORTED THAT THE MODALITY IS NO LONGER BEING USED. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT AN UPDATE STATEMENT WAS RUN ON THE REPORTS TABLE WHICH IN TURN UPDATED THE STUDY TAG FOR ALL ROWS TO ONE STUDY TAG. CUSTOMER SUPPORT UPDATED THE DATABASE TO/FROM BACKUP AND CONFIRMED FUNCTIONALITY. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT THE DEVICE SOFTWARE IS INCORRECTLY CALCULATING THE MEAN GRADIENT ACROSS THE AORTIC VALVE. WHEN USING THE GE VIVID E9 ECHO CART, THE MEASURED VALUE WAS 38.16. WHEN USING THE DEVICE SOFTWARE, THE MEASURED VALUE WAS 24.93. THIS HAS BEEN VERIFIED WITH OTHER STUDIES. TECHNICAL SUPPORT WORKED WITH THE SITE TO REPRODUCE THE ISSUE. IT WAS FOUND THAT THEY WERE TAKING THE MEASUREMENT ON THE WORKSTATION AND IT WAS CALCULATING CORRECTLY. THE SITE WAS GOING TO FIND OTHER EXAMPLES TO TEST ON BUT NEVER CALLED BACK IN. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT REPORTING WOULD CRASH AND RESULT IN INFORMATION FROM REPORTS AND QC TO GO MISSING. THIS IN TURN IMPACTED WORKFLOW AND CAUSED REPEAT DATA ENTRY FOR PHYSICIANS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT THE IE33 CARTS ARE SENDING MV VELOCITY MEASUREMENTS WITH A DISCRIMINATOR FOR LATERAL VS MEDIAL, BUT THE DISCRIMINATOR IS DROPPED IN TRANSLATION RESULTING IN DUPLICATE STRINGS FOR DIFFERENT MEASUREMENTS. THIS IN TURN CAUSES THE FOLLOWING: ONLY HALF THE MEASUREMENTS ARE MAPPING; THE VELOCITY MEASUREMENTS ARE NOT MAPPING. WHEN ONLY HALF THE MEASUREMENTS ARE MAPPING, THIS WAS TRACED TO A STREAMING OF THE WORKFLOW, WHICH IS RESOLVED BY SENDING THE STUDY AT THE END. WHEN THE VELOCITY MEASUREMENTS ARE NOT MAPPING, THIS WAS TRACED TO THE WAY THAT THE STRINGS ARE TRANSLATED BY THE APPLICATION, WHICH RESULTS IN DUPLICATE STRINGS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT INCORRECT DATA WAS RECEIVED FROM THE INITIAL HEMO DATA EXPORT. THE REPORT WAS CONFIRMED PRIOR TO THE SECOND HEMO EXPORT. IT WAS FURTHER REPORTED THAT THE HEMO DATA WAS LOCKED IN. A RECOMMENDATION TO CREATE A NEW STUDY WAS MADE. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT CURRENT MEASUREMENTS WERE NOT CROSSING FROM THE IE33'S AND CX50 ACQUISITION DEVICES. BOTH ARE PHILIPS ECHO MACHINES. IT WAS FURTHER REPORTED THAT THERE WERE A FEW ISSUES WITH MEASUREMENTS THAT WERE CROSSING OVER THAT WERE EITHER INCORRECT OR DO NOT MATCH. NEW MAPPINGS WERE ADDED TO ADDRESS THE ISSUE. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT THE CUSTOMER WAS HAVING A REPORTING ISSUE WHERE THE KNOWLEDGE-BASE POPULATES MODERATE PULMONARY EDEMA INFORMATION FOR SEVERE PULMONARY EDEMA. IT WAS FURTHER REPORTED THAT THE MACRO WAS FIXED AND A NEW KNOWLEDGE BASE (VERSION 2.16.06) WAS LOADED TO THE PRODUCTION VAS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT AGE, SEX, DOB, AND THE REFERRING MD ARE NOT POPULATING THE REPORT FROM THE ORDER THAT IS BEING SENT TO VERICIS. IT WAS DETERMINED BY CUSTOMER SUPPORT THAT IF TECHNICIANS ARE NOT SENDING IN ADT MESSAGES FROM THE HIS SYSTEM, THEN INFORMATION SUCH AS DOB, SEX, AND AGE WILL NOT BE INCLUDED IN THE REPORT. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT THE REPORT IS MISSING A PIECE OF INFORMATION THAT IT SAYS SHOULD BE THERE. SPECIFICALLY, THE REPORT SAYS TO LOOK AT 3RD LESION INTERVENTION, BUT IN THAT INFORMATION WAS NOT AVAILABLE TO THE USER. TECHNICAL SUPPORT DISCUSSED THIS WITH THE CUSTOMER AND EXPLAINED WHAT TO DO TO GET THE 3RD INTERVENTION SECTION TO DISPLAY. STEPS (INCLUDING A SCREEN SHOT) WERE PROVIDED ON HOW TO DO THIS (E.G., PROCEDURE NARRATIVE>INTERVENTIONS AND SELECT THE INTERVENTION NUMBER). THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT THE ECHOIMS DOES NOT SAVE INFORMATION ON A (B)(4) WORKSTATION. THE INFORMATION FROM QUANTIFY AND SUMMARY REPORTING LOCATIONS DO NOT HOLD FROM WORKSTATIONS BUT WILL SAVE FROM THE HOSPITAL'S PC'S WHEN USED IN CONJUNCTION WITH VERICIS WEB. IT WAS FURTHER REPORTED THAT WHEN THE POP-UP BLOCKER WAS TURNED OFF, THE INFORMATION WAS GETTING SAVED IN THE ECHOIMS. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED WITH USERS RIGHTS ACTIVATING GROUP NOT PROPAGATING PROPERLY. THIS COULD PREVENT A USER FROM LOGGING INTO THE SYSTEM TO COMPLETE NECESSARY FUNCTIONS WITHIN THE SYSTEM. IT WAS DETERMINED THAT THE CUSTOMER ACCIDENTALLY ADDED TECHNICIANS TO THE REPORT FELLOW GROUP. ONCE REMOVED, IIS AND REPORTING NEEDED TO BE RESTARTED. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT MEASUREMENTS WERE NOT POPULATING THE REPORT. (B)(4) REPORT MONITOR AND (B)(4) EVENT WERE BOTH RESTARTED. THE SR FILE WAS RECEIVED AND WAS REPROCESSED IN (B)(4). THE FILE WAS THEN RESENT FROM B)(4). MEASUREMENTS NOW POPULATED THE REPORT. THIS WAS CONFIRMED WITH THE CUSTOMER. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT STUDIES DID NOT CONTAIN HEMO DATA. IT WAS FURTHER REPORTED THAT MULTIPLE STUDIES SHOWED THIS PROBLEM. IF THERE IS AN AUTOMATIC ASSOCIATION OF IMAGES AND HEMO DATA, THEN IT SHOWS ON THE WEBVIEWER AND THE WORKSTATION THAT THERE IS HEMO DATA FOR THE STUDY. IF THERE IS A MANUAL ASSOCIATION OF THE TWO, IT SAYS THERE IS NO HEMO DATA FOR THE STUDY. EVEN WHEN THE REPORT IS STARTED, AND EVEN IF IT SAYS THERE IS HEMO DATA, THERE IS NO HEMO DATA IN THE REPORT. IT WAS DETERMINED BY CUSTOMER SUPPORT THAT THERE WEREN'T ANY HEMO TRANSFERS IN PLACE. THE ISSUE WAS RESOLVED AND DATA ARE POPULATING THE REPORT. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT HEMO DATA WAS NOT CROSSING OVER TO SR. CUSTOMER SUPPORT RESOLVED THE ISSUE BY DISASSOCIATING THE HEMO DATA FROM IMAGES FOLLOWED BY REASSOCIATING THEM AGAIN. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. IT WAS REPORTED THAT STRUCTURED REPORT IS SHOWING AS QC STATUS IN CARDIO AS CONFIRMED. THIS ISSUE COULD POTENTIALLY RESULT IN A REPORT NOT BEING REVIEWED BY A PHYSICIAN PRIOR TO SENDING TO EMR OR FAXED. CUSTOMER SUPPORT IDENTIFIED CONFLICTING SQL CONFIRMATION STATES AND CORRECTED THE ISSUE. THERE WERE NO REPORTS OF PATIENT INJURY. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE RECEIVED COMPLAINT OF REPORTING TIMING OUT AND SENDING USER BACK TO STUDY LIST WHICH CAUSES ALL REPORTING DATA BEING ENTERED TO BE LOST. THIS OCCURS IF THE USER HAS THE REPORTING OPEN FOR MORE THAN AN HOUR, THE SYSTEM WILL CLOSE THE REPORTING AND WILL NOT SAVE THE INFORMATION ENTERED. THE SYSTEM IS DESIGNED TO TIMEOUT AND CLOSE THE REPORTING AFTER AN HOUR OF INACTIVITY. AFTER ONE HOUR, THE SERVICE WILL CLEAN UP ALL RESOURCES OF THE INACTIVE SESSION. THIS RESULTS IN THE REPORT NOT SAVING ANY OF THE INFORMATION ENTERED. THIS WORKFLOW IS MOST COMMON WITH PEDIATRIC CASES AND EP STUDIES. IN PEDIATRIC CASES, THE PHYSICIANS REQUIRE MORE THAN AN HOUR FOR REPORTING DUE TO THE LEVEL OF DETAIL AND RESEARCH REQUIRED FOR SUCH REPORTS. EP STUDIES TYPICALLY LAST MORE THAN 3 HOURS AND REPORTING IS OPEN FOR THE DURATION OF THE STUDIES. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. WHILE TESTING ON A CUSTOMER'S TEST ENVIRONMENT WITH CARDIO 10.1, MERGE CUSTOM ENGINEERING DEVELOPER FOUND IF THERE ARE TWO PATIENTS WITH THE SAME LAST NAME, SAME FIRST NAME AND MRN BUT OTHER DIFFERENT ATTRIBUTES, SUCH AS IPID, THE PHYSICIAN SEES STUDIES FROM BOTH PATIENTS IN STUDY COMPARISON VIEW. IN THE STUDY LIST IF TWO PATIENTS HAVE THE SAME FIRST NAME, LAST NAME AND MRN BUT DIFFERENT IPIDS THE SYSTEM TREATS ALL STUDIES AS BEING FROM THE SAME PATIENT. NOT KNOWING THAT THE STUDIES AVAILABLE ARE NOT ALL FOR THE SAME PATIENT WHEN COMPARING THEM IN STUDY COMPARISON MODE, THE PHYSICIAN CAN INTERPRET THE WRONG STUDIES WHICH MAY LEAD TO MISDIAGNOSIS. NOTE THAT THIS ONLY OCCURS WHEN TWO PATIENTS HAVE THE SAME FIRST NAME, LAST NAME AND MRN WHICH IS VERY RARE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE RECEIVED COMPLAINT (B)(4) THAT THE REPORTING FUNCTIONALITY TIMES OUT AND SENDS THE USER BACK TO THE STUDY LIST WHICH CAUSES ALL REPORTING DATA THAT WAS ENTERED TO BE LOST. THIS OCCURS IF THE USER HAS THE REPORTING FUNCTION OPEN FOR MORE THAN AN HOUR, THE SYSTEM WILL TIME OUT AND CLOSE THE REPORTING FUNCTION AND NOT SAVE THE INFORMATION ENTERED. THIS SCENARIO CAN POTENTIALLY LEAD TO DELAY IN CARE. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON 5/5/15, A CUSTOMER REPORTED ((B)(4)) THAT TWO PHYSICIANS FROM ONE CENTER WERE WORKING ON ONE CASE AT THE SAME TIME, AND THEY WERE NOT ALERTED THAT THE CASE WAS OPENED BY THE OTHER PHYSICIAN. THIS PROBLEM WAS IDENTIFIED AS A POTENTIAL SAFETY ISSUE IN THE EVENT CHANGES TO THE SYSTEM AFFECT THE DIAGNOSIS OR TREATMENT OF A PATIENT. NO KNOWN PATIENTS WERE HARMED DUE TO THIS COMPLAINT. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. A CUSTOMER REPORTED ((B)(4)) THAT THEY WENT INTO A STUDY TO VIEW THE IMAGES AND ACCIDENTALLY CLICKED ON THE MULTI STUDY REVIEW BUTTON INSTEAD OF THE REPORTING ICON. WHEN THE CUSTOMER CLICKED BACK ON IMAGE REVIEW, THE APPLICATION HAD STOPPED CACHING THE STUDY WHICH MADE IT LOOK LIKE THE STUDY ONLY HAD A FEW IMAGES. THERE IS THE POTENTIAL SAFETY RISK THAT THE USER BEGINS TO REVIEW A STUDY THAT IS NOT DISPLAYING ALL OF THE AVAILABLE IMAGES. | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. A MERGE CARDIO PRODUCT PROBLEM WAS IDENTIFIED INTERNALLY. THE PROBLEM COULD LEAD TO INCORRECT ACCESSION NUMBERS ON REPORTS. IF IMAGES ARE SENT WITHOUT AN ORDER, THE IMAGES WILL BE MATCHED WITH THE ORDER BASED ON THE CURRENT PATIENT/MODALITY MATCHING. THE REPORT THEN REFLECTS THE INCORRECT ACCESSION NUMBER FOR THE ORDER AND DOES NOT FILE PROPERLY IN THE EMR. AN INCORRECTLY FILED REPORT IN THE EMR COULD LEAD TO DELAY IN CARE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMETS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM DATA). A CUSTOMER DISCOVERED THAT A HEMO-ONLY STUDY CAN HAVE A CONFIRMED REPORT, AND THEN BE ASSOCIATED TO AN IMAGE-STUDY, WHICH RESULTS IN THE HEMO AND IMAGES TO BE MERGED WITH THE ASSOCIATED REPORT. A USER AT A SITE USED THIS WORKFLOW, BUT ACCIDENTALLY CHOSE THE WRONG IMAGE-STUDY WHEN MERGING THE HEMO STUDY. USERS SHOULD NOT BE ABLE TO ADD IMAGES TO A CONFIRMED REPORT VIA MERGE. IT WAS FOUND THAT 9.X USERS CAN MERGE DEVICE IMPORT DATA WITH AN IMAGE-STUDY THAT ALREADY HAS A CONFIRMED REPORT. THIS RESULTS IN ADDING DEVICE IMPORT DATA TO AN IMAGE STUDY WITH A CONFIRMED REPORT THEREFORE IT MAY INTRODUCE INFORMATION THAT WAS NOT AVAILABLE DURING THE INTERPRETATION. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS (MR VENA CONTRACTA AND TR VENA CONTRACTA) THAT WERE NEVER TAKEN ON THE MODALITY WERE POPULATING THE CARDIO CLINICAL REPORT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019, A CUSTOMER ALLEGED THAT CLINICAL REPORTING WAS LOGGING OFF WITHOUT SAVING PREVIOUSLY ENTERED REPORT DATA. THIS COULD LEAD TO A DELAY IN PATIENT CARE WHILE REDOING THE CLINICAL REPORT. DELAY IN TREATMENT HAS A POTENTIAL TO LEAD TO HARM. HOWEVER, THERE IS NO ALLEGATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019 A CUSTOMER CALLED MERGE CARDIO SUPPORT STATING THERE WAS A REPORT WAS IN CONFIRMED STATUS WHEN IT WAS ACTUALLY ONLY QC STATUS. QC REPORTS WITH THE STATUS OF CONFIRMED COULD LEAD TO DELAY OF TREATMENT AS THE REPORT HAS NOT ACTUALLY BEEN READ BY A PHYSICIAN. THE INCORRECT STATUS OF THE REPORT MAY NOT BE READILY APPARENT TO A USER. THERE HAS BEEN NO ALLEGATION OF AN INJURY OR IMPACT TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT-(B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019, A CUSTOMER REPORTED THAT THE MAIN PULMONARY ARTERY (MPA) END SYSTOLE (ES) MEASUREMENT IS POPULATING THE MPA END DIASTOLE (ED) MEASUREMENT FIELD IN CARDIO CLINICAL REPORTING. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS INCORRECT AND NEEDED TO BE DISABLED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2019, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS INCORRECT AND NEEDED TO BE DISABLED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(4) 2018, IT WAS FOUND INTERNALLY WHILE PATCHING THE CUSTOMERS 3.5 REPORTING KNOWLEDGE BASE TO 5.0 REPORTING KNOWLEDGE BASE THAT SOME OF THE MICHIGAN Z-SCORES WERE CALCULATING INCORRECTLY. THIS IS DUE TO THE FORMULA USING 'B' VS 'BSA' FOR BODY SURFACE AREA. THIS VARIABLE HAS BEEN CAUSING Z-SCORES TO CALCULATE INCORRECTLY. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER AT THIS TIME, THERE HAVE BEEN NO ALLEGATIONS OF PATIENT HARM AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018 A CUSTOMER REPORTED THAT THE IMAGES OF A STUDY DID NOT MATCH THE NAME OF THE REPORT FOR THAT STUDY. MERGE HEALTHCARE SUPPORT IS WORKING WITH THE CUSTOMER FOR MORE EXAMPLES OF THIS HAPPENING AS THEY ARE UNABLE TO INVESTIGATE THIS ISSUE AS THERE WAS NOT ENOUGH DATA LEFT FOR THEM TO KNOW WHAT HAPPENED DURING THIS INCIDENT. REFERENCE CASE (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, THE CUSTOMER ALLEGED THE REPORT WAS NOT GOING TO THEIR EMR. UPON INVESTIGATION, IT WAS FOUND THAT THE REASON THE REPORT NEVER EXPORTED TO THE CUSTOMER'S EMR, EPIC. THIS WAS DUE TO IT BEING INCORRECTLY IN CONFIRMED STATUS INSTEAD OF BEING IN QC STATUS AND ACTUALLY CONFIRMED. MERGE CARDIO SUPPORT CORRECTED THE ISSUE BY CHANGING THE STATUS IN THE DATABASE QC REPORTS WITH THE STATUS OF CONFIRMED COULD LEAD TO DELAY OF TREATMENT AS IT HAS NOT ACTUALLY BEEN READ BY A PHYSICIAN AND AS A RESULT MAY NOT BE READILY APPARENT TO A CLINICIAN. HOWEVER, THERE HAVE BEEN NO ALLEGATIONS OF HARM AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT WRONG RA VOLUME WAS BROUGHT FORWARD. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, THERE HAVE BEEN NO REPORTS OF HARM OR INJURY AS A RESULT OF THIS ISSUE. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. THE MERGE CARDIO AND TOMTEC INTEGRATION ALLOWS USERS TO VIEW MEDICAL IMAGES IN THE TOMTEC CLIENT (IMAGECOM) ON A MERGE CARDIO WORKSTATION WITHOUT HAVING TO MANUALLY ENTER PATIENT OR STUDY INFORMATION INTO THE TOMTEC APPLICATION. IMAGECOM IS A CLINICAL APPLICATION PACKAGE FOR HIGH PERFORMANCE PC PLATFORMS BASED ON MICROSOFT WINDOWS OPERATING SYSTEM STANDARDS. IMAGECOM IS SOFTWARE FOR THE RETRIEVING, REVIEWING, MEASURING, AND REPORTING OF DICOM DATA, FOR THE US. IN THIS INTEGRATION, THE MODALITY SENDS ALL STUDIES CONTAINING A VALID ACCESSION NUMBER TO THE ENTERPRISE ARCHIVE (EA). IMAGECOM IS INVOKED FROM THE MERGE CARDIO WORKSTATION. THE MODALITY SENDS ALL STUDIES TO THE ENTERPRISE ARCHIVE. SECONDARY CAPTURE (SC) OR STRUCTURED REPORT (SR) IMAGES ARE SENT BACK AN INTEGRATION APPLICATION, I.E. TOMTEC, FROM TO THE EA. SC IMAGES CAN BE VIEWED IN MERGE CARDIO. ON (B)(6) 2018, IT WAS FOUND THAT SECONDARY CAPTURES IN TOMTEC WERE SAVING UNDER THE WRONG PATIENT IN MERGE CARDIO. IMAGES SAVING UNDER THE WRONG PATIENT TO LEAD TO PATIENT HARM BY INCORRECT OR MISTREATMENT OF THE PATIENT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED CORRECTED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE MODIFIED IN MERGE CARDIO. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON 04/26/2018, A MERGE CARDIO CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT LEFT VENTRICULAR OUTFLOW TRACT (LVOT) MEASUREMENTS WERE INCORRECT. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION. DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY WERE NOT APPEARING IN MERGE CARDIO WITH THE SAME MEASUREMENT VALUE. THE CUSTOMER HAS NOT INDICATED THAT ANY PATIENT WAS IMPACTED PR HARMED AS A RESULT OF THIS ISSUE. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018 , A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT THE HEIGHT MEASUREMENT WAS IMPORTING INCORRECTLY TO THE MERGE CARDIO CLINICAL REPORT CAUSING MISCALCULATION OF OTHER MEASUREMENTS. MERGE HEALTHCARE TECHNICAL SUPPORT'S INITIAL INVESTIGATION FINDINGS APPEAR TO BE RELATED TO SOME NEW ECHOCARDIOGRAM IMAGING CARTS AT THE CUSTOMER'S FACILITY. IT IS POSSIBLE THAT THE NEW ECHO CARTS MAY BE SENDING DATA TO MERGE CARDIO PACS IN AN UNEXPECTED FORMAT THEREFORE RESULTING IN INCORRECT HEIGHT MEASUREMENTS. DUE TO THE MISCALCULATION OF MEASUREMENTS IN THE CLINICAL REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF THE PATIENT THAT COULD RESULT IN HARM. HOWEVER, THE CUSTOMER HAS NOT ALLEGED ANY IMPACT OR INJURY TO A PATIENT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT MEASUREMENT CALCULATIONS ARE NOT APPEARING PROPERLY IN THE DIAGNOSTIC REPORT. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED WITHIN THE MERGE CARDIO APPLICATION. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. REFERENCE COMPLAINT NUMBER -(B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT ON (B)(6) 2017 THE MERGE CARDIO OUTBOUND HL7 INTERFACE HAD STOPPED SENDING CLINICAL REPORTS ELECTRONICALLY TO THEIR HOSPITAL INTERFACE SYSTEM FOR APPROXIMATELY 18 HOURS. DURING THIS DOWNTIME, A PATIENT WAS DISCHARGED FROM THE FACILITY WITHOUT A FULL REVIEW OF THE PATIENT'S CLINICAL REPORT. THE CUSTOMER STATED THAT IT WAS NECESSARY FOR THE PATIENT TO RETURN TO THE FACILITY FOR FURTHER TESTING. THE MERGE CARDIO SOFTWARE PROVIDES USERS WITH THE ABILITY TO REVIEW AND PRINT CLINICAL REPORTS WHEN THEY ARE UNABLE TO BE DELIVERED TO THE CUSTOMER'S HOSPITAL INTERFACE SYSTEM. ADDITIONALLY, THE SYSTEM IS DESIGNED TO NOTIFY THE USER IF THE DELIVERY OF CLINICAL REPORTS FAILS. THE NOTIFICATION ALLOWS THE USER TO BE MADE AWARE OF REPORT COMMUNICATION ISSUES IN ORDER FOR OTHER METHODS OF COMMUNICATION TO BE USED (PRINT, SCAN, FAX, ETC.). WITH A PATIENT BEING DISCHARGED BEFORE STAFF COMPLETED A FULL REVIEW OF THAT PATIENT'S CLINICAL REPORT, THERE IS A POTENTIAL FOR A DELAY OR INCORRECT TREATMENT THAT COULD RESULT IN HARM TO THE PATIENT. HOWEVER, THE CUSTOMER REPORTED THAT NO HARM HAD OCCURRED. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON (B)(6) 2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT AN ISSUE OCCURRED THAT COULD LEAD TO INCORRECT ACCESSION NUMBERS ON CLINICAL REPORTS. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT WITH THEIR CLINICAL WORKFLOW, IF IMAGES ARE SENT WITHOUT AN ORDER, THE IMAGES WILL BE MATCHED WITH THE PATIENT'S ORDER BASED ON THE CURRENT PATIENT/MODALITY MATCHING. THE CLINICAL REPORT THEN REFLECTS THE INCORRECT ACCESSION NUMBER FOR THE ORDER AND DOES NOT FILE PROPERLY IN THE CUSTOMER'S EMR. WHILE THE ACCESSION NUMBER MAY NOT BE THE EXPECTED NUMBER THE PATIENT'S STUDY INFORMATION IS CORRECTLY PLACED WITHIN THE PATIENT'S RECORD ON THE EMR. AN INCORRECTLY FILED REPORT IN THE EMR COULD LEAD TO A DELAY IN CARE. HOWEVER, THE CUSTOMER STATED THAT THERE WAS THERE WAS NO HARM TO THE PATIENT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. ON 12/19/2016, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT DIAGNOSTIC MEASUREMENTS TAKEN BY THEIR MODALITY NEEDED TO BE CORRECTED. DUE TO AN INCORRECT VALUES DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, THE CUSTOMER HAS MADE NO INDICATION OF PATIENT SAFETY ISSUES WHILE WORKING WITH CARDIO SUPPORT. (B)(4). | -fail to store medical imaging -fail to access clinical history -inadequate clinical history data | -incompatibility problem -incorrect display -use of device issue | Viewer
| FDA | Event Description: MERGE CARDIO IS AN INTEGRATED CARDIOVASCULAR INFORMATION SYSTEM CLASSIFIED AS A PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). ON (B)(6) 2016 A CUSTOMER CALLED INTO MERGE HEALTHCARE AND ALLEGED THAT THERE WAS INCORRECT PROCEDURAL VERBIAGE IN A MERGE CARDIO REPORT. INACCURATE CATHERIZATION PROCEDURE DETAILS IN A REPORT COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS OR TREATMENT. MERGE HEALTHCARE IS INVESTIGATING THIS CUSTOMER'S ALLEGATION. THERE IS NO INDICATION FROM THE CUSTOMER THAT THERE WAS ANY CONFIRMED MIS-DIAGNOSIS OR MIS-TREATMENT OF A PATIENT DUE TO INACCURATE VERBIAGE IN A PATIENT'S CARDIOLOGY REPORT. REFERENCE (B)(4). | -loss of communication -insufficient information -inadequate information | -break -communication or transmission issue -reset problem | Viewer
| FDA | Event Description: MERGE CARDIO IS AN INTEGRATED CARDIOVASCULAR INFORMATION SYSTEM CLASSIFIED AS A PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). ON (B)(6) 2016 CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A PATIENT REPORT WAS CONFIRMED (SIGNED) WITH THE INCORRECT PATIENT DATA (SEE ORIGINAL REPORT 2183926-2016-00441). BASED ON THIS ALLEGATION FROM THAT CUSTOMER, MERGE HEALTHCARE INVESTIGATED THE ISSUE AND CONFIRMED THAT THERE WAS A MALFUNCTION OF THE MEDICAL DEVICE. THIS MALFUNCTION IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. MERGE HEALTHCARE WAS ABLE TO QUERY POTENTIALLY AFFECTED CUSTOMER'S DATABASES ON (B)(4) 2016 TO IDENTIFY IF THERE WERE ANY SUSPECT REPORTS STORED THERE. THE QUERY DETECTED A SUSPECT REPORT FOR THIS CUSTOMER AND THEY HAVE BEEN MADE AWARE OF THE ISSUE. (B)(4). | -unable to record event -inadequate information -inadequate clinical history data | -communication or transmission issue -break -reset problem | Viewer
| FDA | Event Description: MERGE CARDIO IS AN INTEGRATED CARDIOVASCULAR INFORMATION SYSTEM CLASSIFIED AS A PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). ON (B)(6) 2016 THIS CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A PATIENT REPORT WAS CONFIRMED (SIGNED) WITH THE INCORRECT PATIENT DATA. BASED ON THIS ALLEGATION FROM THAT CUSTOMER, MERGE HEALTHCARE INVESTIGATED THE ISSUE AND CONFIRMED THAT THERE WAS A MALFUNCTION OF THE MEDICAL DEVICE. THIS MALFUNCTION IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. MERGE HEALTHCARE WAS ABLE TO QUERY POTENTIALLY AFFECTED CUSTOMER'S DATABASES ON (B)(6) 2016 TO IDENTIFY IF THERE WERE ANY SUSPECT REPORTS STORED THERE. THE QUERY DETECTED A SUSPECT REPORT FOR THIS CUSTOMER AND THEY HAVE BEEN MADE AWARE OF THE ISSUE. (B)(4). | -unable to record event -fail to record operational events -confused interface | -break -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE CARDIO IS AN INTEGRATED CARDIOVASCULAR INFORMATION SYSTEM CLASSIFIED AS A PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A PATIENT REPORT WAS CONFIRMED (SIGNED) WITH THE INCORRECT PATIENT DATA (SEE ORIGINAL REPORT 2183926-2016-00441). BASED ON THIS ALLEGATION FROM THAT CUSTOMER, MERGE HEALTHCARE INVESTIGATED THE ISSUE AND CONFIRMED THAT THERE WAS A MALFUNCTION OF THE MEDICAL DEVICE. THIS MALFUNCTION IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. MERGE HEALTHCARE WAS ABLE TO QUERY POTENTIALLY AFFECTED CUSTOMER'S DATABASES ON FEBRUARY 6, 2016 TO IDENTIFY IF THERE WERE ANY SUSPECT REPORTS STORED THERE. THE QUERY DETECTED A SUSPECT REPORT FOR THIS CUSTOMER AND THEY HAVE BEEN MADE AWARE OF THE ISSUE. (B)(4). | -unable to record event -fail to record operational events -inadequate information | -communication or transmission issue -break -reset problem | Viewer
| FDA | Event Description: MERGE CARDIO IS AN INTEGRATED CARDIOVASCULAR INFORMATION SYSTEM CLASSIFIED AS A PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). ON (B)(6) 2016, A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A PATIENT REPORT WAS CONFIRMED (SIGNED) WITH THE INCORRECT PATIENT DATA (SEE ORIGINAL REPORT 2183926-2016-00441). BASED ON THIS ALLEGATION FROM THAT CUSTOMER, MERGE HEALTHCARE INVESTIGATED THE ISSUE AND CONFIRMED THAT THERE WAS A MALFUNCTION OF THE MEDICAL DEVICE. THIS MALFUNCTION IF NOT DETECTED, COULD POTENTIALLY LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT OF A PATIENT BASED ON THE INACCURATE INFORMATION IN THE REPORT. MERGE HEALTHCARE WAS ABLE TO QUERY POTENTIALLY AFFECTED CUSTOMER'S DATABASES ON (B)(6) 2016 TO IDENTIFY IF THERE WERE ANY SUSPECT REPORTS STORED THERE. THE QUERY DETECTED A SUSPECT REPORT FOR THIS CUSTOMER AND THEY HAVE BEEN MADE AWARE OF THE ISSUE. (B)(4). | -unable to record event -fail to record operational events -inadequate information | -communication or transmission issue -break -reset problem | Viewer
| FDA | Event Description: MERGE CARDIO IS USED TO ACQUIRE, STORE, PRINT, TRANSFER AND ARCHIVE CLINICAL INFORMATION FROM HEMODYNAMIC AND CARDIOLOGY IMAGING MODALITIES. IMAGES AND STUDIES CAN BE REVIEWED, ANNOTATED, DIGITAL CALCULATIONS AND QUANTITATIVE MEASUREMENTS CAN BE PERFORMED IN ORDER TO GENERATE AND STORE REPORTS. ON (B)(6) 2018, A CUSTOMER CONTACTED MERGE HEALTHCARE AND STATED THAT A DIAGNOSTIC MEASUREMENT TAKEN BY THEIR MODALITY WAS IMPORTING INTO THE INCORRECT NODE WITHIN THE CLINICAL REPORTING TOOL. DUE TO AN INCORRECT VALUE DISPLAYING IN THE DIAGNOSTIC REPORT, THERE IS A POTENTIAL FOR INCORRECT TREATMENT OF A PATIENT THAT COULD RESULT IN HARM. HOWEVER, THERE WAS NO IMPACT OR CONSEQUENCE TO THE PATIENT. (B)(4). | -fail to access clinical history -inadequate clinical history data -insufficient clinical evidence | -fluid leak -data issue -data back-up problem | Viewer
| FDA | Event Description: MERGE ICONNECT ENTERPRISE ARCHIVE IS INTENDED FOR USE AS A VENDOR NEUTRAL ARCHIVE FOR STORAGE AND COMMUNICATIONS OF MEDICAL IMAGES AND DATA. ICONNECT ENTERPRISE ARCHIVE PROVIDES WORKFLOW INTEGRATION CAPABILITIES FOR HEALTHCARE ENTERPRISES ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE ALLEGING THAT TWO (2) STUDIES WERE MISSING AND COULD NOT BE PULLED FROM THE ARCHIVE. MERGE HEALTHCARE TECHNICAL SUPPORT IS INVESTIGATING THE CUSTOMER'S ALLEGATION. AN ANALYSIS OF THE POTENTIAL SOURCES AND STORAGE OF THE STUDIES AND ANY OTHER POTENTIALLY RELATED STUDIES IS ONGOING. DUE TO STUDIES NOT BEING ABLE TO BE FOUND AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN PATIENT TREATMENT OR DIAGNOSIS OR PATIENT HARM. HOWEVER, THERE WERE NO REPORTS OF DELAYS IN PATIENT TREATMENT OR DIAGNOSIS. AT THIS TIME THERE ARE NO REPORTS OF PATIENT HARM. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ISSUE TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. DOES NOT APPLY TO SOFTWARE. | -fail to store medical imaging -loss of image -inadequate image manipulation | -use of device issue -therapeutic or diagnostic output failure -pacing inadequately | Viewer
| FDA | Event Description: MERGE ORTHOCASE DISPLAYS MEDICAL IMAGES FOR DIAGNOSTIC REVIEW. A PROBLEM WAS OBSERVED IN-HOUSE, WHERE ORIENTATION LABELS ARE NOT READ FROM DICOM FILE AND A DEFAULT ORIENTATION IS ASSUMED IN PLUG IN MODE. WHEN ORTHOCASE IS LAUNCHED IN PLUG IN MODE, ORIENTATION LABELS DISPLAYED IN THE VIEWER MAY BE INCORRECT (I.E. SHOW A DEFAULT VALUE). SAVED DATA (E.G. SECONDARY CAPTURE IMAGES) WILL HAVE THE INCORRECT LABELS BURNED IN. THIS ISSUE DOES NOT OCCUR IN STAND-ALONE MODE. THERE IS NO WORK AROUND. THE ORIENTATION LABELS CANNOT BE DISABLED THROUGH CONFIGURATION. THE ISSUE WAS ENCOUNTERED DURING INTEGRATION TESTING WITH ANOTHER PRODUCT. IN MOST CASES, SURGEONS AND PHYSICIANS WOULD USE OF LEAD MARKERS ON THE IMAGES. THE MARKERS WOULD IDENTIFY ORIENTATION OF THE IMAGE AND ANATOMY. | -inadequate message display -incorrect message display -incorrect display format | -therapeutic or diagnostic output failure -failure to read input signal -improper flow or infusion | Viewer
| FDA | Event Description: MERGE ORTHOCASE IS USED FOR ORTHOPEDIC APPLICATIONS AND CONSISTS OF FEATURES THAT ALLOW THE QUALIFIED MEDICAL PROFESSIONAL TO MAKE MEASUREMENTS THAT ARE COMMONLY REQUIRED WHEN DOING ORTHOPEDIC SURGICAL PLANNING. ON (B)(6) 2016 A CUSTOMER REPORTED INCORRECT MEASUREMENTS FROM THE TEMPLATE IN MERGE ORTHOCASE. THE CUSTOMER USED THE TEMPLATE TO MEASURE AN AREA FOR SURGERY AND WHAT SIZE IMPLANTS WILL BE USED. DURING THE SURGERY THEY FOUND THAT AREA THEY MEASURED ON THE IMAGE WAS OFF 4-6 MM. THIS EVENT WAS HIGHLY DETECTABLE BY THE SURGEON AND HE WAS ABLE TO RE-MEASURE AND USE CORRECT SIZE IMPLANT. THIS ISSUE HAS A POTENTIAL FOR USE OF AN INCORRECTLY SIZED IMPLANT DURING ORTHOPEDIC SURGERY. THE CUSTOMER CONFIRMED THE SURGERY WAS SUCCESSFUL AND THERE WAS NO HARM TO THE PATIENT. (B)(4). Manufacturer Narrative: AFTER FURTHER INVESTIGATION, IT WAS FOUND THE CUSTOMER WAS USING DIFFERENT SIZE BALLS WHICH CAUSED THE TEMPLATE TO NOT PRODUCE CORRECT MEASUREMENTS. NO ADDITIONAL CORRECTIVE ACTIONS ARE NEEDED. | -inadequate measurement unit -incorrect measurement -inappropriate measurement unit | -no display or display failure -insufficient flow or underinfusion -instruction for use issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. (B)(6) HAS REPORTED THAT MAMMO IMAGES FROM (B)(4) ARE OF POOR QUALITY. IN (B)(6) IT WAS A DETERMINED THAT (B)(4) NEEDED TO SUPPLY AN UPDATE TO THE FIRMWARE AND THE CASE WAS CLOSED. (B)(4) HAS TRIED FIX THE ISSUE AND HAS EXHAUSTED ALL OPTIONS AND THAT THE RADIOLOGIST IS REFUSING TO READ MAMMO STUDIES STATING THAT THE IMAGE QUALITY IS POOR. THIS IS ISOLATED TO ONE (B)(4) MODALITY AT THE CUSTOMERS FACILITY. MUCH RESEARCH HAS BEEN DONE BY DEVELOPMENT, SUPPORT AND SOLUTIONS AND AS OF 11/1/15 THERE IS NO ONE ROOT CAUSE. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. A CUSTOMER REPORTED (CASE (B)(4)) THAT UNVIEWED IMAGES WERE NOT DISPLAYING AFTER THE LAST HANGING PROTOCOL STEP. WHEN THE USER PREFERENCE "SHOW UNVIEWED IMAGES AFTER THE LAST HANGING PROTOCOL STEP" IS SELECTED, THE CUSTOMER EXPECTS THAT ANY IMAGES NOT MATCHED TO THE HANGING PROTOCOL WILL DISPLAY IN AN AUTO-GENERATED FINAL VIEW. HOWEVER, IF THE LAST DEFINED VIEW OF THE PROTOCOL WAS SKIPPED (BECAUSE NO IMAGES MATCHED), THE UNVIEWED IMAGES ARE NOT DISPLAYED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE OF THE RISK ASSOCIATED WITH NOT VIEWING ALL OF THE IMAGES IN THE STUDY. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. A CUSTOMER REPORTED (CASE (B)(4)) THE CUT LINES ARE NOT SHOWING CORRECTLY ON A STUDY IN PACS. THE CUT LINES WOULD GO FROM BEING VERTICAL AND THEN SHIFT ALMOST HORIZONTAL. THE DISPLAY IS VERY OBVIOUS TO THE USER AND SHOWS CONTRADICTORY VIEWS WHICH WOULD CAUSE THE USER TO QUESTION WHAT IS HAPPENING. THE IMAGES CAME FROM SIEMENS (B)(4) SCANNER. AN INVESTIGATION FOUND THAT SIEMENS CHANGED THE RESOLUTION OF THE VALUES OF SOME DICOM ATTRIBUTES CAUSING AN EDGE CASE WHICH MERGE PACS DID NOT HANDLE APPROPRIATELY. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE FOR THE PATIENT IN THE EVENT THERE WAS A MISDIAGNOSIS FROM VIEWING THE INCORRECT MEASUREMENTS OF A PATIENT'S IMAGES. | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO A 6.5.9 SYSTEM VIA TELMED. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO A 6.5.9 SYSTEM VIA TELMED. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. IN THE MEANTIME, A WORKAROUND WAS PROVIDED TO THE SITE TO ONLY SEND LOSSLESS AND NOT ALLOW LOSSY, SO THAT THERE IS A VALID FAILURE WHEN LOSSLESS ISN'T PRESENT, TO RETRIEVE FROM ARCHIVE AND RESEND. THE FAILURES WOULD NO LONGER BE SEEN WITH THIS CONFIGURATION. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO THE PATIENT. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO A 6.5.9 SYSTEM VIA TELMED. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. IN THE MEANTIME, A WORKAROUND WAS PROVIDED TO THE SITE TO ONLY SEND LOSSLESS AND NOT ALLOW LOSSY, SO THAT THERE IS A VALID FAILURE WHEN LOSSLESS ISN'T PRESENT, TO RETRIEVE FROM ARCHIVE AND RESEND. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. ON (B)(6) 2013 A CUSTOMER REPORTED THAT A STUDY STATUS HAS BEEN CHANGED INCORRECTLY. THE LOGS REPORTED THAT THE STUDY WAS CHANGED BY A STAFF MEMBER; HOWEVER THE STAFF MEMBER DID NOT HAVE SYSTEM PERMISSIONS TO DO SO. MERGE HEALTHCARE IDENTIFIED THE ROOT CAUSE TO BE A DESIGN FLAW WHERE ANY USER WITH THE CHANGE WORKFLOW STATUS BUTTON CAN PDE ANY STUDY, EVEN IF THE USER DOES NOT HAVE PERMISSIONS TO USE THE PDE TOOL. THERE WAS NO REPORT OF PATIENT INJURY. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. ON (B)(6) 2015 A CUSTOMER REPORTED THAT SOME PATIENT PRIORS INCLUDED INCORRECT INFORMATION AFTER A DATA MIGRATION. THE CUSTOMER WAS RECEIVING A PRINT OUT FROM THE RIS SYSTEM AND FOUND THAT MRN NUMBERS WERE MIGRATED OVER WITH THE OLD NUMBERS. THIS RESULTED IN EXAMS IN INCORRECT FOLDERS. MERGE PACS QUALIFIES STUDIES BY MRN WHEN PROCESSING A08 (PATIENT UPDATE) AND A34 (PATIENT MERGE) REQUESTS FROM A HIS/RIS. BY DEFAULT, IF AN A08 OR A34 REQUEST IS RECEIVED THAT HAS AN MRN THAT MATCHES A PATIENT BUT HAS THE WRONG PATIENT NAME, THE REQUEST WILL BE REJECTED BY THE MERGE PACS. MERGE PACS CAN BE CONFIGURED, HOWEVER, TO NOT CHECK TO SEE IF THE PATIENT NAME STORED IN ITS DATABASE MATCHES THAT SENT BY THE HIS/RIS AND WILL INSTEAD RETURN ALL STUDIES THAT HAVE AN MRN THAT MATCHES THAT SENT BY THE HIS/RIS (A PROCESS KNOWN AS "GLOBAL UPDATE"). MERGE WAS ABLE TO PROVIDE THE NECESSARY DATA SO THAT THE CUSTOMER COULD DO A COMPARISON AND IDENTIFY DUPLICATES TO MAKE SURE THE EXAMS WERE IN THE CORRECT FOLDER. THERE WAS NO REPORT OF PATIENT INJURY. REFERENCE: (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS (PICTURE ARCHIVING COMMUNICATION SYSTEM) IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES, INCLUDING DIGITAL MAMMOGRAPHY. THE CUSTOMER REPORTED THAT PRIORS ARE NOT GETTING SENT TO THE MERGE PACS SYSTEM VIA (B)(4) TRANSMITTER. THEY USE AN AFTER HOURS READING SERVICE AND ARE CALLING THIS A PATIENT CARE ISSUE. THERE IS NO INDICATION THAT THE STUDIES FAILED. RESEARCH NARROWED THE ISSUE TO SENDING ONLY LOSSY IMAGES WHICH HAVE BEEN QC'D AND ONLY SENDING TO AN OLDER VERSION OF PACS. DEVELOPMENT HAS DETERMINED THERE IS A DEFECT IN PACS AND WILL BE FIXING IT IN 6.6.3. IN THE MEANTIME, A WORKAROUND WAS PROVIDED TO THE SITE TO ONLY SEND LOSSLESS AND NOT ALLOW LOSSY, SO THAT THERE IS A VALID FAILURE WHEN LOSSLESS ISN'T PRESENT, TO RETRIEVE FROM ARCHIVE AND RESEND. THE FAILURES WOULD NO LONGER BE SEEN WITH THIS CONFIGURATION. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO A PATIENT. (B)(4). | -unable to read -unable to record event -inadequate labeling | -device contamination with biological material -communication or transmission issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER CALLED INTO MERGE SUPPORT AND ALLEGED THAT THEIR VIEWER WAS NOT RESPONDING. ON (B)(6) 2016 THE SAME CLIENT CALLED MERGE SUPPORT ALLEGING THAT THEIR VIEWER WAS CRASHING WHEN A WORKLIST WAS OPENED. BASED ON THE CUSTOMER'S ALLEGATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT DUE TO THE VIEWER NOT FUNCTIONING AS EXPECTED. THE CUSTOMER DID NOT INDICATE ANY PATIENT HARM AS A RESULT OF THE ISSUES WITH THE VIEWER. (B)(4). | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER CONTACTED SUPPORT ALLEGING THAT THERE WERE MISSING IMAGES WITHIN STUDIES FOR A NUMBER OF PATIENTS. MERGE HEALTHCARE BEGAN INVESTIGATING THE CUSTOMER'S ALLEGATION. IT WAS DETERMINED THAT ALL REPORTS THAT WERE REPORTED TO HAVE MISSING IMAGES HAD BEEN PREVIOUSLY READ AND REPORTED ON. ADDITIONALLY IT WAS DETERMINE THAT THE CUSTOMER HAD A PREVIOUS PRODUCTION DATABASE ISSUE AND HAD ELECTED TO MOVE DATA TO A CONTINGENCY SERVER. THE CLIENT HAD NOT BEEN PERFORMING RECOMMENDED DATABASE AND SERVER MAINTENANCE INCLUDING BACK UPS OF THE SYSTEM. INCONSISTENT AND IRREGULAR BACK-UPS OF CLINICAL DATA INCREASES THE LIKELIHOOD OF DATA LOSS. IT WAS DETERMINED THAT THE MISSING IMAGES WERE DUE TO THE CONTINGENCY SERVER BEING SET TO DO PERIODIC PURGES AS IT IS USED AS A TEMPORARY SOLUTION AND NOT INTENDED TO SUSTAIN LONG TERM USE. ONCE IT WAS DETERMINED THAT CONTINGENCY SERVER PURGING STUDIES, PURGE WAS DISABLED. DUE TO THE LOSS OF IMPAGES THAT CAN BE USED WHEN COMPARING PRIOR STUDIES TO RECENT STUDIES, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER REPORTED THAT THEY WERE NOT AWARE OF ANY INSTANCE WHERE A PATIENT WAS HARMED OR TREATMENT DELAYED. (B)(4). | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) CLIENT CALLED IN ISSUE "PRIOR SHOWS MAMO WITH A DIFFERENT PERSON". EXAMPLE, A DOCTOR COULD REACH PATIENT B AND BE ERRONEOUSLY COMPARING IT TO PATIENT A'S PRIOR. SUPPORT BEGAN INVESTIGATION AND ALSO WORKED WITH DEVELOPMENT. DISCUSSED WITH SUPPORT AND THEY EXPLAINED THIS WAS A WORKFLOW ISSUE: PACS ACCEPTS VARIOUS ADT MESSAGES LISTED BELOW. ANY ADT MESSAGE THAT THEY GENERATE WHETHER IT BE AN A01, A02, A03, ETC GETS TRANSLATED AT THE INTERFACE INTO AN A08 (CUSTOM TO THIS SITE/NOT USED AT ANY OTHER FACILITY TO MY KNOWLEDGE). SO WHAT HAS HAPPENED IN THIS CASE: A NEW PATIENT COMES INTO THE FACILITY, FIRST GETS REGISTERED AND THE INTERFACE STARTS WITH AN A04 (REGISTRATION) BUT CONVERTED INTO AN A08. WE RECEIVE THE A08 AND THE MRN + IPID MATCHES ONLY 1 PATIENT , WE UPDATE THE NAME AND ANY OTHER DEMOGRAPHIC INFORMATION IN THE A08 MESSAGE BECAUSE THAT IS THE FUNCTION OF AN A08 (PATIENT UPDATE). NOTE WE HAVE A FAIL SAFE IN THE SYSTEM TO ENSURE IF THERE ARE MULTIPLE PATIENTS WITH THE SAME MRN AND IPID THAT THE NAME IS CHECKED AS WELL. BUT IN THE CASES THEY HAVE SEEN, THE PATIENT THEY ARE REGISTERING IS YET TO EXIST IN PACS. THE SITE HAS ASKED WHY WE DO NOT CHECK ANY OTHER QUALIFIERS SUCH AS DOB. THAT WOULD DEFEAT THE PURPOSE OF AN A08 MESSAGE. FOR EXAMPLE A PATIENT COMES IN AND THE PERSON REGISTERING THE PATIENT ACCIDENTALLY PUTS IN AN INCORRECT DOB, THEY WOULD WANT TO BE ABLE TO CHANGE THAT. A01: ADMIT/VISIT NOTIFICATION, UPDATE PATIENT ADDRESS, PREGNANCY STATUS, CONFIDENTIALITY CODE, PATIENT CLASS, CURRENT LOCATION, REF PHYSICIAN, ADMISSION ID, ADMISSION DATE AND ADMISSION TIME. A02: TRANSFER A PATIENT, UPDATE CURRENT LOCATION FROM INCOMING MESSAGE. A03: DISCHARGE/END VISIT, UPDATE PATIENT STATUS AS DISCHARGED, DISCHARGE DATE, DISCHARGE TIME, PATIENT CLASS, CURRENT LOCATION. A04: REGISTER A PATIENT, UPDATE PATIENT ADDRESS, PREGNANCY STATUS, CONFIDENTIALITY CODE, PATIENT CLASS, CURRENT LOCATION, REF PHYSICIAN, ADMISSION ID, ADMISSION DATE AND ADMISSION TIME. A08: UPDATE PATIENT INFORMATION, UPDATE PATIENT DEMOGRAPHICS (PATIENT NAME, SEX, DATE OF BIRTH, SSN, ADDRESS, EMAIL, WORK PHONE, CELL PHONE, HOME PHONE). ALSO SUBMIT PDE COMMAND FOR STUDY UPDATE. A11: CANCEL ADMISSION, CHECK ALL PENDING ORDERS NOT ASSOCIATED WITH STUDIES AND DELETE ALL ORDER TABLES FOR THE PATIENT ID OF INCOMING MESSAGE. A12: CANCEL TRANSFER, UPDATE PATIENT LOCATION AND PATIENT CLASS FROM INCOMING MESSAGE. A13: CANCEL DISCHARGE, UPDATE PATIENT STATUS FROM "DISCHARGED" TO "ADMITTED", UPDATE CURRENT LOCATION AND PATIENT CLASS. A34: MERGE PATIENT - PATIENT MRN MERGE ALL ORDERS AND STUDIES (SUBMIT PDE REQUEST) FOR THE GIVEN PATIENT ID FROM INCOMING MESSAGE. A40: MERGE PATIENT - PATIENT IDENTIFIER LIST MERGE ALL ORDERS AND STUDIES (SUBMIT PDE REQUEST) FOR THE PATIENT ID LIST FROM INCOMING MESSAGE. A47: CHANGE PATIENT IDENTIFIER LIST, UPDATE ALL ORDERS AND STATUSES WITH NEW PATIENT ID LIST INCLUDING PATIENT TABLE. SUBMIT NEW PDE REQUEST TO UPDATE ALL STUDIES WITH NEW PATIENT ID LIST FROM INCOMING MESSAGE. THE REASON THE SITE HAD DONE THIS CUSTOMIZATION IS BECAUSE THEY WERE USING OTHER MESSAGES SUCH AS A01S AND INFORMATION IN THE ORDERS WERE BEING CHANGED. SO INSTEAD OF WORKING ON FOLLOWING OUR STANDARD AND MAKING ANY CORRECTIONS NEEDED ON THE INTERFACE END, THEY STARTED TO CONVERT THEIR MESSAGES TO A08S. SUPPORT EXPLAINED THAT THIS IS OUTLINED IN THE HL7 GUIDE (SEE ABOVE). AS PART OF INVESTIGATION ALSO SPOKE TO CLIENT AND THEY EXPLAINED THAT THE ISSUE WAS CAUGHT BY A RAD. THE RISK IS THAT YOU COULD READ THE WRONG PATIENT'S PRIOR. | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017 A CLIENT REPORTED THAT WHEN A RADIOLOGIST BEGINS VIEWING STUDY A IN THE PACS PRIMARY VIEWER AND THEN BEGINS VIEWING STUDY B IN THE SECONDARY VIEWER, IF STUDY A IS DICTATED AND MARKED AS READ, STUDY B IS SET TO A READ STATUS AS WELL. THE CUSTOMER IS EXPECTING THAT ONLY THE STUDY A IS MARKED AS READ, NOT BOTH OF THE STUDIES THAT THE CUSTOMER IS VIEWING (STUDY A AND STUDY B). WHEN A STUDY THAT HAS NOT BEEN DICTATED AND READ IS AUTOMATICALLY CHANGED TO A READ STATUS BY THE SYSTEM, THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS. THIS IS DUE TO THE CUSTOMER POTENTIALLY MISSING THE STUDY AS IT WILL NO LONGER APPEAR ON THEIR PENDING REVIEW WORKLIST. AT THIS TIME, THERE IS NO INDICATION OF DIRECT IMPACT OR HARM TO A PATIENT. (B)(4). | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THEY WERE HAVING ISSUES WITH THE UNVERIFIED-NEW IMAGE STATUS FOR PATIENT REPORTS. MERGE HEALTHCARE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORT STATUS WAS NOT CHANGING TO "UNVERIFIED NEW IMAGES" TO ALERT THE USER THAT ADDITIONAL IMAGES HAD BEEN RECEIVED AND ADDED/APPENDED TO THE PATIENT'S REPORT. THIS ISSUE WAS DETECTABLE TO THE USER AS A STUDY MARKED AS FINAL DID NOT CHANGE STATUS WHEN NEW IMAGES WERE RECEIVED. CURRENTLY, THERE IS NO KNOWN IMPACT TO SPECIFIC OR OVERALL PATIENTS; HOWEVER, THERE IS POTENTIAL FOR A SITUATION TO OCCUR WHERE NEW IMAGES ARE NOT READ RESULTING IN DELAYED OR INCORRECT TREATMENT. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017 A CUSTOMER REPORTED TO MERGE HEALTHCARE THAT A STAT EXAM HAD NOT BEEN READ IN A TIMELY MANNER. THE CUSTOMER REQUESTED THAT MERGE HEALTHCARE TECHNICAL SUPPORT PERFORM AN ANALYSIS OF THE ACTIVITIES FOR THE PATIENT'S STUDY. AN ANALYSIS OF THE FILE'S LOGS REVEALED MULTIPLE USERS ACCESSING THE FILE WITHOUT VALIDATING THE STUDY. ADDITIONAL ANALYSIS REVEALED THAT AFTER THE MULTIPLE USERS HAD ACCESSED THE NON-VALIDATED STUDY, A USER DID MANUALLY VALIDATE THE RECORD DURING A SCHEDULED DOWNTIME WHEN THE CUSTOMER'S IT PERSONNEL WERE PERFORMING MAINTENANCE ON THE SYSTEM. PERFORMING A STUDY VALIDATION DURING SYSTEM/SERVER DOWNTIME PREVENTS THE VALIDATION COMMAND FROM GOING THROUGH AND THE FILE IS PUT INTO A "FAILED" STATUS. THE LOGS INDICATED THAT THE NEXT ACTION WAS A USER MARKING THE STUDY AS "FINAL" WITHOUT VALIDATION BEING PERFORMED. FINALLY ON (B)(6) 2017, THE LOGS SHOWED A USER MANUALLY RE-OPENING THE STUDY, READING AND VALIDATING THE REPORT. MERGE HEATLHCARE PERSONNEL REVIEWED THE INVESTIGATION INFORMATION WITH THE CUSTOMER. IT IS THE RESPONSIBILITY OF THE SITE TO REVIEW AND MONITOR STATUS QUEUES FOR ANY FILES THAT HAVE FAILED TO PROCESS. THIS IS AN ACTIVITY THAT IS COMMON TO PACS DEVICES. WHILE THE SITE IS FAIRLY NEW TO MERGE PACS (OCTOBER 2016), IN PERSON TRAINING SESSIONS AND MERGE PACS DEVICE LABELING PROVIDE INFORMATION INSTRUCTING USERS TO MANAGE FILE QUEUES. THE CUSTOMER HAS CONFIRMED THAT NO SERIOUS INJURY OR DEATH RESULTED DUE TO THE DELAY IN VALIDATING THE STAT EXAM HOWEVER, THERE IS A POTENTIAL FOR INJURY SHOULD THIS ISSUE REOCCUR. MERGE HEALTHCARE WORKED WITH THE CUSTOMER TO ENHANCE THE SITE'S ABILITY TO MONITOR REPORT QUEUES. (B)(4). | -fail to display data -inadequate image types -unable to render 3d views | -device alarm system issue -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED THAT WHEN A USER EXITS OUT OF STUDY IN FLUENCY, MERGE PACS CHANGES THE STATUS OF THE STUDY TO "READ." WHEN A STUDY UNEXPECTEDLY GOES TO A READ STATUS WHEN EXITING OUT OF THAT STUDY, THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS DUE TO A CUSTOMER MISSING READING THE STUDY. HOWEVER, THIS ISSUE WAS DETECTED BY THE CUSTOMER AND THERE WAS NO REPORTED IMPACT TO PATIENT CARE. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER REPORTED THAT THEY WERE EXPERIENCING AN ISSUE WITH MERGING PATIENT DEMOGRAPHICS. IT WAS FOUND THAT THE GLOBAL PATIENT UPDATE WAS TURNED ON DURING IMPLEMENTATION AT THE CUSTOMER SITE AND WOULD UPDATE OR MERGE ALL PATIENTS ASSOCIATED WITH A SINGLE MRN/IPID. WHICH RESULTED IN PATIENT INFORMATION TO BE UPDATED TO THE INCORRECT PATIENT IN THEIR SYSTEM WHEN DUPLICATE MRN NUMBERS OCCUR. THIS ISSUE HAS A POTENTIAL TO LEAD TO PATIENT INFORMATION BEING COMPROMISED OR OVERWRITTEN BY A NEW STUDY SUBMISSION. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED AS A RESULT OF THIS ISSUE. ADDITIONALLY, IF GLOBAL PATIENT UPDATE WAS SET TO OFF, PATIENTS WITH THE SAME MRN NUMBERS WOULD NOT BE MERGED. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED STUDIES MARKED AS "READ" OR "FINAL" WERE NOT CHANGING BACK TO "UNREAD" OR "UNVERIFIED" WHEN A NEW DOCUMENT IS ADDED TO THE STUDY. THERE IS A POTENTIAL FOR DELAY IN DIAGNOSIS OR TREATMENT DUE TO THE CUSTOMER BEING UNAWARE OF DOCUMENT(S) ADDED TO A STUDY THAT HAS ALREADY BEEN READ. INFORMATION RECEIVED FROM THE CUSTOMER CONFIRMED THERE WAS NO IMPACT TO PATIENT CARE DUE TO THIS ISSUE. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER CONTACTED MERGE HEALTHCARE AND ALLEGED THAT AN INCIDENT HAD OCCURRED WHERE THE THIRD IMAGE ON A REPORT WAS NOT REVIEWED BY THE RADIOLOGIST. THE THIRD IMAGE SHOWED A FRACTURE THAT INITIALLY WENT UNDETECTED DUE TO NOT BEING REVIEWED. THIS RESULTED IN A DELAY IN DIAGNOSIS AND/OR TREATMENT FOR THE PATIENT. CURRENTLY THERE IS NO INDICATION OF AN ADVERSE EVENT TO THE PATIENT. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED A PATIENT WAS SENT HOME WITH A CERVICAL FRACTURE BECAUSE IMAGES WERE NOT READ. THE CUSTOMER FOUND THE IMAGES MISSED REVIEW AND CONTACTED THE PATIENT. THE PATIENT WAS ASKED TO GO TO THE CUSTOMER'S MAIN HOSPITAL FOR AN MRI. AFTER THE MRI WAS PERFORMED, THE CUSTOMER HAD SUCCESSFUL SURGERY. THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH CAN LEAD TO PATIENT HARM AS A RESULT OF CLINICAL STAFF NOT REVIEWING ALL IMAGES. IT WAS DETERMINED THAT THE PATIENT RECEIVED AN MRI SCAN AND SURGERY PRIOR TO ANY ADDITIONAL HARM BEYOND THE ORIGINAL CERVICAL FRACTURE. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON 04/15/2016 ORTHO (B)(4) REPORTED THAT EXISTING INFORMATION WAS GETTING OVERWRITTEN SINCE EMR INTERFACE WAS ACTIVATED. ORTHO (B)(4) WAS TO CONDUCT A CONVERSION OF THEIR MRN NUMBERS BEFORE EMR WENT LIVE SO THAT MRN NUMBERS WOULD MATCH ATHENA'S MRN NUMBERS. THIS CONVERSION WAS NOT COMPLETED BEFORE EMR WAS ACTIVATED CAUSING EXISTING PATIENT RECORDS TO BE UPDATED INCORRECTLY DUE TO OVERWRITING. OVERWRITTEN INFORMATION COULD POTENTIALLY LEAD TO AN INCORRECT TREATMENT OR DIAGNOSIS. CURRENTLY THERE IS NO INDICATION OF ACTUAL PATIENT HARM. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2018, THE CLIENT REPORTED A STUDY WAS NOT AVAILABLE ON THE WORKLIST DUE TO THE STUDY STATUS CHANGING TO PRELIMINARY. THE CUSTOMER CONFIRMED THERE WAS A THIRTY MINUTE DELAY IN DIAGNOSIS FOR THE STUDY. INFORMATION RECEIVED FROM THE CUSTOMER CONFIRMED THERE WAS NO DIRECT IMPACT TO PATIENT CARE DUE TO THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2018, A CUSTOMER STATED MERGE PACS CHANGED THE STUDY STATUS TO READ, WHICH MADE IT UNAVAILABLE ON THE WORKLIST. WHEN A STUDY AUTOMATICALLY GOES TO READ, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS BECAUSE THE CUSTOMER COULD MISS READING THE STUDY IN A TIMELY MANNER. HOWEVER, FOR THIS CUSTOMER, THE EVENT WAS DETECTED BY THE PHYSICIAN AND THERE WASN'T ANY IMPACT TO PATIENT CARE. OTHER CLINICAL DATA WAS AVAILABLE TO SUPPORT THE PROPER TREATMENT OF THE PATIENT. REFERENCE (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED A STUDY STATUS CHANGED FROM "UNREAD" TO "READ" WITHIN A FEW SECONDS OF IMPORTING INTO THE PACS SYSTEM. THIS CAUSED THE STUDY TO FALL OFF THE DOCTOR'S WORK LIST AND NOT BE READ WITHIN CUSTOMER'S TIMELINE. THE CUSTOMER CONFIRMED THE STUDY WAS READ WITHIN APPROXIMATELY ONE HOUR AND 20 MINUTES OF ENTERING THE PACS SYSTEM.ÿ WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW STUDY IS ADDED AND THEN REMOVED FROM THE WORKLIST, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS FOUND BY THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, DURING INTERNAL TESTING, A MERGE EMPLOYEE REPORTED AN ISSUE WITH VOICE CLIPS SAVING TO THE WRONG PATIENT. VOICE CLIPS ARE AUDIO ANNOTATIONS AND ARE USED DIFFERENTLY DEPENDING ON USER'S WORK FLOW. VOICE CLIPS COULD CONTAIN ESSENTIAL INFORMATION RELATED TO TREATMENT/DIAGNOSIS. THERE IS A POTENTIAL THAT THE WRONG VOICE CLIP MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITH THE WRONG VOICE CLIP, THERE IS A POTENTIAL FOR AN INCORRECT TREATMENT AND/OR DIAGNOSIS TO THE PATIENT. HOWEVER, THE UNDERLYING IMAGE IS PRESENT FOR THE USER TO READ. THERE WAS NO DIRECT IMPACT TO PATIENT CARE REPORTED. REFERENCE COMPLAINT-(B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED AN ISSUE WITH VOICE CLIPS SAVING TO THE WRONG PATIENT. VOICE CLIPS ARE AUDIO ANNOTATIONS AND ARE USED DIFFERENTLY DEPENDING ON USER'S WORK FLOW. VOICE CLIPS COULD CONTAIN ESSENTIAL INFORMATION RELATED TO TREATMENT/DIAGNOSIS. THERE IS A POTENTIAL THAT THE WRONG VOICE CLIP MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITH THE WRONG VOICE CLIP, THERE IS A POTENTIAL FOR AN INCORRECT TREATMENT AND/OR DIAGNOSIS TO THE PATIENT. HOWEVER, THE UNDERLYING IMAGE IS PRESENT FOR THE USER TO READ. THERE WAS NO DIRECT IMPACT TO PATIENT CARE REPORTED. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017, A CUSTOMER IDENTIFIED AN ISSUE WITH IMAGE LOAD ERRORS WHILE TESTING MERGE PACS VERSION 7.2.1 IN THEIR TEST ENVIRONMENT. IT WAS DETERMINED THAT IMAGES THAT FAILED TO LOAD PROPERLY WERE CORRUPTED. THE CORRUPTED IMAGES WERE SKIPPED DURING MANUAL SCROLLING THROUGH A STUDY. THERE IS NO WARNING OR ERROR THAT INDICATES IMAGES HAVE BEEN SKIPPED. THERE IS A POTENTIAL THAT A SKIPPED IMAGE MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITHOUT NECESSARY INFORMATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS &/OR TREATMENT OR AN INACCURATE DIAGNOSIS &/OR TREATMENT OF A PATIENT. AT THIS TIME, THERE IS NO EVIDENCE OF PATIENT HARM OR DELAY DUE TO THIS ISSUE. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON 07/10/2017, MERGE HEALTHCARE SUPPORT IDENTIFIED AN ISSUE DURING TROUBLESHOOTING ACTIVITIES WHERE CORRUPT IMAGES ARE SKIPPED DURING MANUAL SCROLLING. THERE IS NO WARNING OR ERROR THAT INDICATES IMAGES HAVE BEEN SKIPPED. THERE IS A POTENTIAL THAT A SKIPPED IMAGE MAY CONTAIN DATA THAT IS REQUIRED TO DETERMINE DIAGNOSIS AND/OR TREATMENT DECISIONS. WITHOUT NECESSARY INFORMATION, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS &/OR TREATMENT OR AN INACCURATE DIAGNOSIS &/OR TREATMENT OF A PATIENT. AT THIS TIME, THERE IS NO EVIDENCE OF PATIENT HARM OR DELAY DUE TO THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER CONTACTED MERGE PACS TECHNICAL SUPPORT ON (B)(6) 2017 AND ALLEGED THAT STUDIES WERE NOT GOING BACK TO AN UNVERIFIED STATE IF MORE IMAGES WERE SENT INTO THE SYSTEM AND APPENDED TO THE PATIENT'S STUDY AFTER THE STUDY WAS EVALUATED. THIS ISSUE IS DUE TO AN OVERWRITE STATUS FIELD BEING SET TO NO. WHEN THE OVERWRITE STATUS IS SET TO YES, IT CHANGES THE STUDY BACK TO UNVERIFIED TO ALERT THE USER OF ADDITIONAL IMAGES. WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW IMAGE IS ADDED TO A STUDY, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS APPARENT TO THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017, A CUSTOMER REPORTED AN ISSUE WITH ANNOTATIONS NOT SAVING WHEN THE STUDY IS DICTATED AND CLOSED. THE PHYSICIAN BECAME AWARE OF THE ISSUE WHEN HE WAS COMPARING A PRIOR IMAGE WITHOUT ANNOTATIONS TO A CURRENT, UNREAD IMAGE. AT THIS TIME, THE MEASUREMENTS WERE STILL AVAILABLE IN THE REPORT AND DID NOT CAUSE AN IMPACT TO THE CARE OF THE PATIENT. ON 06/12/2017, THE CUSTOMER REPORTED THIS ISSUE WAS NOW IMPACTING PATIENT CARE BECAUSE THE ANNOTATIONS WERE NOT AVAILABLE TO MAKE ACCURATE COMPARISONS ON THE CURRENT STUDY. THE CUSTOMER COMMUNICATED ADDITIONAL DETAILS REGARDING THE ANNOTATION ISSUE, IN THAT AN OUTSIDE PHYSICIAN HAD CALLED THEM ASKING WHERE THE ANNOTATIONS WERE LOCATED. THE PHYSICIAN REQUESTED THE INFORMATION FROM THE CUSTOMER BECAUSE THEY NEEDED TO REVIEW A PATIENT'S STUDY PRIOR TO SURGERY. THIS ANNOTATION ISSUE IS ONLY IMPACTING A CERTAIN TYPE OF STUDY FOR THE USER. BOTH THE MEASUREMENTS IN THE REPORT AND UNDERLYING PRIOR IMAGES ARE AVAILABLE TO THE USER. HOWEVER, THE ISSUE COULD POTENTIALLY CAUSE A DELAY AND/OR INCORRECT TREATMENT DUE TO THE TIME TO REVIEW THE IMAGES A SECOND TIME. (B)(4). Manufacturer Narrative: AFTER FURTHER INVESTIGATION, MERGE HEALTHCARE DEVELOPMENT DETERMINED THIS WAS A DEFECT IN THE CODE AND RESOLVED THIS ISSUE IN V 7.2. TO PREVENT THIS ISSUE, THE USER UNDERSTANDS THEY NEED TO DRAG THE SERIES IN FROM THE THUMBNAIL INSTEAD OF SERIES TOOLBAR. MERGE HEALTHCARE SUPPORT ALSO CREATED A HANGING PROTOCOL TO ASSIST WITH WORKFLOW CHANGES. THE CUSTOMER HAS NOT REPORTED ANY ADDITIONAL ISSUES SINCE CHANGING WORKFLOW. NO FURTHER CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017, A CUSTOMER ALLEGED AN ISSUE WITH A BREAST OUTLINE NOT BEING CORRECT, WHICH RESULTED THE SITE TO TAKING ADDITIONAL VIEWS OF THE PATIENT'S BREAST. THE CUSTOMER ALSO ALLEGED THAT THE CAD MARKINGS FOR MEASUREMENTS AND CALCIFICATION ARE INCORRECT ON 3D IMAGES. THE ALLEGED ISSUE WITH THE BREAST OUTLINE IS READILY APPARENT TO THE USER ON MAMMOGRAPHY IMAGES. HOWEVER, THE ISSUE COULD POTENTIALLY CAUSE A DELAY IN DIAGNOSIS OR INCORRECT TREATMENT DUE TO EITHER ADDITIONAL IMAGES BEING TAKEN AND/OR INCORRECT MEASUREMENTS. IT SHOULD BE NOTED THAT THE CUSTOMER HAS ACCESS TO CORRECT CAD MARKINGS IN 2D IMAGES AND IN SECONDARY CAPTURE. AT THIS TIME THERE IS NO INDICATION OF DIRECT IMPACT TO PATIENT CARE OR HARM AS A RESULT OF THIS ISSUE. | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THE NUMBER OF IMAGES ON THEIR LOCAL MAMMO SERVERS DO NOT MATCH THE NUMBER OF IMAGES ON THEIR PACS SERVER. THERE IS A POTENTIAL RISK OF THE CUSTOMER REMOVING IMAGES FROM THE MAMMO SERVER UNDER THE ASSUMPTION THAT THE IMAGES HAVE BEEN ARCHIVED PROPERLY TO THE PACS SERVER. AT THIS TIME, ALL IMAGES ARE BEING TRANSFERRED TO THE SERVER AND THERE IS NO EVIDENCE OF LOSS OF DATA/IMAGES. HOWEVER, THE ISSUE COULD POTENTIALLY CAUSE A DELAY OR INCORRECT TREATMENT DUE TO A PHYSICIAN NOT BEING ABLE TO VIEW A PATIENT'S PRIOR STUDIES FOR COMPARISON WITH THE CURRENT STUDY. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THAT TWO DIFFERENT PATIENT'S STUDIES WERE MERGED TOGETHER INTO A SINGLE PATIENT'S RECORD BY THE SYSTEM AUTOMATICALLY. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT TWO DIFFERENT PATIENTS HAD BEEN MERGED TOGETHER DUE TO THE PATIENT'S HAVING MATCHING MRN AND IPID IDENTIFIERS. EACH PATIENT'S DEMOGRAPHIC INFORMATION IS STORED AND VISIBLE ON THE IMAGES IN THE REPORT AND WOULD BE DETECTABLE TO THE USER WHEN REVIEWING DEMOGRAPHIC INFORMATION. THERE WAS NO INDICATION FROM THE CUSTOMER REGARDING ANY DIRECT IMPACT TO PATIENT CARE AT THIS TIME. HOWEVER, THIS ISSUE COULD POTENTIALLY CAUSE A DELAY IN PATIENT CARE DUE TO THE INABILITY TO FIND A PATIENT'S INFORMATION OR A MISDIAGNOSIS IF THE PATIENT'S DEMOGRAPHIC INFORMATION IS NOT READ AND A TREATMENT IS APPLIED TO THE INCORRECT PATIENT. (B)(4). | -fail to display data -inadequate image types -incorrect display format | -device unsafe to use in environment -application program issue -no display or display failure | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER CONTACTED AFTER HOURS MERGE HEALTHCARE SUPPORT ON (B)(6) 2016 AND REPORTED THAT A PATIENT'S ORDER WAS MERGED WITH ANOTHER PATIENT'S ORDER. MERGE HEALTHCARE SUPPORT BEGAN INVESTIGATING THE ISSUE AND WORKED WITH THE CUSTOMER OVER MULTIPLE WEEKS. THE INVESTIGATION CONCLUDED ON 3/25/2016 AND DETERMINED THAT THE SITE HAD BEEN IMPORTING PATIENT INFORMATION FROM 2 DIFFERENT RIS SYSTEMS. THIS WAS DETERMINED TO BE A USER ERROR THAT CONTRIBUTED TO THE MERGING OF 2 DIFFERENT PATIENT'S INFORMATION. THE SITE HAS DISCONTINUED SENDING PATIENT INFORMATION FROM THE 2ND RIS SYSTEM. MERGE PACS WAS FUNCTIONING AS DESIGNED FOR PATIENT DATA INTERACTIONS WITH ONE RIS SYSTEM; NOT MULTIPLE RIS SYSTEMS. THIS MDR IS BEING SUBMITTED DUE TO THE POTENTIAL FOR A POSSIBLE MISDIAGNOSIS OR MISTREATMENT OF A PATIENT DUE TO THE MERGING OF DIFFERENT PATIENT'S DATA. THE CUSTOMER DID NOT INDICATE ANY HARM OR IMPACT TO A PATIENT(S) DUE TO THE MERGING OF PATIENT INFORMATION BEING IMPORTED FROM THE CUSTOMER'S MULTIPLE RIS SYSTEMS INTO MERGE PACS. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED ON (B)(6) 2016 THAT THEY WERE EXPERIENCING SLOWNESS ON IMPORTING OF STUDIES AND QUERYING. MERGE HEALTHCARE SUPPORT BEGAN INVESTIGATING THE ISSUE AND DETERMINED THAT THE SITE HAD STUDIES THAT WERE STUCK WHILE IN IMPORT. THE ISSUE WAS FOUND TO BE DUE TO THE ORGANIZATION OF THE DICOM TABLES AND AS A RESULT THESE WERE REORGANIZED. THE ISSUES WITH SLOWNESS AND IMPORTING STUDIES COULD POTENTIALLY LEAD TO A DELAY IN DIAGNOSIS OR TREATMENT. SUPPORT DETERMINED NO ISSUE FOUND WITH THE MERGE PACS SOFTWARE. DATABASE RE-ORGANIZATIONS ARE PART OF NORMAL SERVER UPKEEP. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THE COMMUNICATION NOTES WERE NOT SAVING. WITH MERGE PACS NOT SAVING COMMUNICATION NOTES AS EXPECTED THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT LEADS TO HARM. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER REPORTED THAT WHEN THEY MERGED STUDY A AND STUDY B THE COMMENTS FROM STUDY A WERE LOST. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. HOWEVER, THERE IS A POTENTIAL THAT A USER MAY ADD CLINICAL INFORMATION INTO THE COMMENTS SECTION AND THEREFORE, THERE IS A POTENTIAL THAT INFORMATION THAT MAY IMPACT A TREATMENT OR DIAGNOSIS DECISION MAY NOT BE PRESENT ON A MERGED STUDY. (B)(4). Manufacturer Narrative: THE INVESTIGATION FOUND THAT MERGE PACS WAS FUNCTIONING AS DESIGNED. MERGE PACS VERSION 7.1 RELEASE NOTES WILL ADDRESS THE DELETION OF COMMENTS IN STUDY A WHEN MERGING WITH STUDY B. | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER ALLEGED TO MERGE HEALTHCARE THAT STUDIES WERE INTERMITTENTLY GOING TO A "READ" STATUS AFTER IMPORTING. WITH NON-READ OR NON-REVIEWED STUDIES INDICATING A READ STATUS COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THE INVESTIGATION FOUND THAT DURING IMPLEMENTATION OF THE UPGRADED VERSION OF MERGE PACS (7.0.2) THE OVERWRITE FUTURE STATUS SETTING WAS SET TO "NO." THIS SETTING WAS INTERMITTENTLY CAUSING NEWLY IMPORTED IMAGES TO NOT CHANGE THE STUDY STATUS TO "UNREAD." MERGE HEALTHCARE SUPPORT RESOLVED THE ISSUE BY CHANGING THE OVERWRITE FUTURE STATUS SETTING TO "NO." | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THEY WERE UNABLE TO RETRIEVE SOME IMAGES TAKEN IN 2012 FROM THEIR ARCHIVE. DUE TO THE LOSS OF IMAGES THAT CAN BE USED WHEN COMPARING PRIOR STUDIES TO RECENT STUDIES, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT WHILE VIEWING IMAGES OF CLAVICLE THEY OBSERVED THAT THE DICOM TAGS WERE IDENTIFYING THE IMAGES AS BEING THE PATIENT'S LEFT SIDE REGARDLESS OF WHETHER IT WAS AN IMAGE OF THE PATIENT'S LEFT OR RIGHT CLAVICLE. AS THE CUSTOMER ALSO PLACES MARKERS TO IDENTIFY THE PATIENT'S SIDE THE ISSUE IS EASILY NOTICED. HOWEVER, IF THE CUSTOMER DID NOT PLACE THE MARKER THE DICOM TAGS INCORRECTLY TAGGING THE IMAGES HAS THE POTENTIAL TO LEAD TO INCORRECT TREATMENT OR MISDIAGNOSIS. THERE WAS NO INDICATION FROM THE CUSTOMER THAT ANY PATIENTS WERE HARMED DUE TO THIS COMPLAINT. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT WHILE REVIEWING IMAGES WITH THE USE OF THE "OXFORD PARTIAL KNEE TWIN-PEG" TEMPLATE, FOUND THAT THE TEMPLATE SHOWED TEXT THAT STATED THE INCORRECT VIEW OF THE KNEE IMAGE. THE TEXT SHOULD HAVE READ RIGHT RATHER THEN LEFT. THIS ISSUE HAS A POTENTIAL TO LEAD TO AN INACCURATE DIAGNOSIS OR TREATMENT OF THE WRONG KNEE. THE CUSTOMER REPORTED THAT THE PHYSICIAN NOTICED THE ISSUE BEFORE MISDIAGNOSING THE INCORRECT KNEE. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED DUE TO THIS COMPLAINT. ADDITIONALLY, THE PHYSICIAN WOULD HAVE REFERENCED THE LEAD LATERALITY MARKER WITHIN THE IMAGE TO NOTE WHICH KNEE WILL BE DIAGNOSED. REFERENCE COMPLAINT NUMBER (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. THE CUSTOMER CALLED INTO MERGE SUPPORT IN (B)(6) 2016 BECAUSE THEY FOUND THAT SOME OUTSIDE STUDIES THAT THEY WERE BRINGING INTO THEIR SYSTEM WERE CONFLICTING AND MERGING TO SOME CURRENT STUDIES ON THEIR SYSTEM. MERGE SUPPORT RECOGNIZED THE ISSUE AND INFORMED CLIENT THEY NEEDED TO CREATE A DICOM HEADER MODIFICATION RULE TO PREFIX THE MRNS THAT ARE COMING FROM THAT FACILITY. THIS WOULD HAVE FIXED THE ISSUE. CLIENT STATED THAT THEY WOULD WORK WITH MERGE IMPLEMENTATION STAFF TO RESOLVE THE ISSUE. THE CUSTOMER REPORTED THAT THERE HAD BEEN CASES WHERE RAD HAD READ IMAGES OF PATIENT A THINKING THEY WERE PATIENT B AND THE WRONG REPORT WAS ATTACHED TO THE WRONG PATIENT. THE SITE WAS ABLE TO CATCH THESE VERY QUICKLY AND CORRECT THE ISSUE. THERE IS NO INFORMATION SUGGESTING THAT A PATIENT HAS BEEN ADVERSELY AFFECTED BY THIS. TO MAKE SURE THE INCORRECT MERGING DOES NOT HAPPEN, SITE IS MANUALLY QCING EVERYTHING THAT COMES INTO THE SYSTEM. HOWEVER, THERE IS A RISK THAT ONE COULD BE MISSED OR THAT THE STUDY WAS READ AFTER HOURS AND WAS NOT QC'ED. AGAIN, THERE IS NO EVIDENCE OF A SPECIFIC PATIENT BEING ADVERSELY EFFECTED BY THIS. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT AN AUTOMATED DAILY REPORT INDICATED THAT THREE OF THEIR ARCHIVED TAPES WERE SHOWING AS UNAVAILABLE. SHOULD A TAPE WITH ARCHIVED PATIENT DATA BE REQUESTED FOR VIEWING, THERE IS A POTENTIAL THAT THE INFORMATION MAY NOT BE AVAILABLE DUE TO THE PACS SYSTEM NOT BEING ABLE TO RETRIEVE/READ THE REQUESTED INFORMATION. THIS COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT OR DIAGNOSIS DUE TO UNAVAILABLE DATA. CURRENTLY, THERE IS NO KNOWN OR ALLEGED IMPACT TO SPECIFIC OR OVERALL PATIENTS. MERGE HEALTHCARE'S INITIAL INVESTIGATION RESULTS INDICATE THAT THE CUSTOMER IS RUNNING ON SERVER SOFTWARE/FIRMWARE THAT IS ANTIQUATED AND IS NO LONGER SUPPORTED BY THE OEM. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. MERGE HEALTHCARE RECEIVED AN ALLEGATION FROM A CUSTOMER ON (B)(6) 2016, CLAIMING THAT LINE MEASUREMENTS WERE NOT MATCHING THE OLDER VERSION OF PACS VERSUS THE NEWER VERSION. WHILE THE DIFFERENCE ALLEGED BY THE CUSTOMER WAS MINIMAL THERE IS A POTENTIAL THAT COMPARISONS MADE BETWEEN OLD IMAGES (~2011) AND NEWER IMAGES COULD RESULT IN AN INACCURATE DIAGNOSIS OR MISTREATMENT OF A PATIENT. MERGE HEALTHCARE IS INVESTIGATING THE ALLEGATION. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT WHEN A NEW STUDY WAS VIEWED, THE PATIENT'S PRIOR STUDIES WERE NOT BEING RETRIEVED FROM THE DATABASE AS EXPECTED. THE CUSTOMER DID NOT ALLEGE ANY DEATH, SERIOUS INJURY OR HARM TO ANY SPECIFIC OR OVERALL PATIENTS. THERE IS HOWEVER, A POTENTIAL FOR A DELAY OR INCORRECT TREATMENT DUE TO A PHYSICIAN IS NOT BEING ABLE TO VIEW A PATIENT'S PRIOR STUDIES FOR COMPARISON WITH THE CURRENT STUDY. | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(4) 2016, MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER THAT WHEN ADDITIONAL IMAGES ARE SENT TO A STUDY THAT HAS ALREADY BEEN READ, THERE WAS NO NOTIFICATION THAT ADDITIONAL STUDIES HAD BEEN RECEIVED. THE STATUS REMAINS AS "READ" OR "FINAL STATUS" RATHER THAN INDICATING "UNVERIFIED NEW IMAGE" HAS BEEN RECEIVED VIA AN ALERT NOTIFICATION. IT IS UNKNOWN TO THE CUSTOMER WHEN THE PROBLEM FOR THE ALERT NOT BEING RECEIVED FIRST OCCURRED. CURRENTLY, THERE IS NO KNOWN IMPACT TO SPECIFIC OR OVERALL PATIENTS; HOWEVER, THERE IS POTENTIAL FOR A SITUATION TO OCCUR WHERE NEW IMAGES ARE NOT READ RESULTING IN DELAYED OR INCORRECT TREATMENT. THE CUSTOMER WILL BE MODIFYING A CONFIGURABLE SOFTWARE SETTING WITHIN THE MERGE PACS DEVICE TO EVALUATE IF THE ALERT IS RECEIVED FOR STUDIES THAT HAVE BEEN READ BUT ADDITIONAL IMAGES ARE SENT FOR THE STUDY. | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(4) 2017 IT WAS REPORTED, "RADIOLOGIST NOTICED THAT SUV VALUES ON A FUSED PET CT ARE DIFFERENT COMPARED TO WHEN THE SUV IS MEASURED ON SOURCE IMAGES ALONE AND NOT FUSED IMAGES." THERE WAS NO ALLEGATION OF PATIENT INJURY, ADVERSE CONSEQUENCES, OR SIGNIFICANT DELAY IN PATIENT CARE. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6)2017 INTERNAL TESTING FOUND THAT THERE WERE INCONSISTENCIES BETWEEN THE SUV MEASUREMENTS FOR 2D IMAGES AND SUV MEASUREMENTS FOR SPINNING MIP (MAXIMUM INTENSITY PROJECTION). THIS ISSUE WAS FOUND BY MERGE HEATLHCARE ON (B)(6)2017 AND WAS NOT REPORTED BY A CUSTOMER. THE ISSUE OF POTENTIALLY INACCURATE SUV VALUES MAY HAVE AN IMPACT ON PATIENT TREATMENTS FOR CANCER AND THEREFORE MIGHT RESULT IN HARM. HOWEVER, THERE IS NO INDICATION THAT ANY PATIENTS HAVE BEEN NEGATIVELY IMPACTED BY THIS ISSUE AT THIS TIME (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017 A CUSTOMER REPORTED THAT THEY WERE UNABLE TO CONNECT TO THE SERVER FROM THE VIEWER. IT WAS FOUND THAT THE CUSTOMER HAD A BAD NETWORK INTERFACE CARD ON THEIR SERVER WHICH NEEDED TO BE REPLACED BY (B)(4). IT WAS REPORTED THAT SYSTEM DOWNTIME OCCURRED FROM (B)(6) 2017. THIS MAY POTENTIALLY CAUSE A DELAY IN TREATMENT BUT IS UNLIKELY TO CAUSE HARM AS THERE IS OTHER CLINICAL INFORMATION THAT WILL BE UTILIZED TO SUPPORT THE TREATMENT OF THE PATIENT. THE CUSTOMER DID NOT INDICATE ANY DELAY IN PATIENT CARE DURING ALLEGED DOWNTIME. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 A CUSTOMER REPORTED THAT A STUDY AND ITS IMAGES WERE DELETED AND ARE NO LONGER ABLE TO BE VIEWED. THE CUSTOMER DID NOT ALLEGE AN INJURY, THE NEED FOR MEDICAL INTERVENTION, OR ADVERSE CONSEQUENCES AS A RESULT OF THE UNRETRIEVABLE IMAGES. HOWEVER, WITH A PATIENT'S IMAGES OR STUDY BEING UNAVAILABLE FOR VIEWING THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED THEY USED THE QC EDITOR TOOL TO SPLIT A STUDY AND FOUND THE CHANGES WERE NOT SAVED AFTER SAVING AND REOPENING THE STUDY. THE QC EDITOR TOOL DID NOT GIVE THE CUSTOMER A WARNING THAT CHANGES WERE NOT SAVED. THE QC EDITOR TOOL IS USED TO SPLIT, MERGE AND/OR DELETE SERIES OF IMAGES IN A STUDY. THIS MAY POTENTIALLY CAUSE A DELAY IN TREATMENT BUT IS UNLIKELY TO CAUSE HARM AS THERE IS OTHER CLINICAL INFORMATION THAT WILL BE UTILIZED TO SUPPORT THE TREATMENT OF THE PATIENT. HOWEVER THERE IS NO REPORTED SERIOUS INJURY, DEATH OR HARM TO A PATIENT AS A RESULT OF THIS ISSUE. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED THEY USED THE QC EDITOR TOOL TO SPLIT A STUDY AND FOUND THE CHANGES WERE NOT SAVED AFTER SAVING AND REOPENING THE STUDY. THE QC EDITOR TOOL DID NOT GIVE THE CUSTOMER A WARNING THAT CHANGES WERE NOT SAVED. THE QC EDITOR TOOL IS USED TO SPLIT, MERGE AND/OR DELETE SERIES OF IMAGES IN A STUDY. THERE IS A POTENTIAL FOR A DELAY IN TREATMENT THAT MAY LEAD TO HARM HOWEVER THE CUSTOMER HAS NOT INDICATED THAT ANY HARM TO A PATIENT OCCURRED AS A RESULT OF THIS ISSUE. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED THEY WERE UNABLE TO LOGIN INTO MERGE PACS FOR APPROXIMATELY 8 HOURS. RATHER THAN USING MERGE PACS, PHYSICIANS AND RADIOLOGISTS WERE ABLE TO VIEW IMAGES AND PATIENT STUDIES AT THE IMAGING MODALITY. WITH CLINICIANS NOT BEING ABLE TO LOG INTO MERGE PACS THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS. HOWEVER, THE CUSTOMER CONFIRMED THERE WAS NO DIRECT IMPACT TO PATIENT CARE. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016, A CUSTOMER REPORTED IN THEIR PREVIOUS MERGE PACS VERSION THAT IF A DOCUMENT OR IMAGE WAS ADDED TO A STUDY AFTER IT WAS MARKED "READ", IT WOULD CHANGE THE STATUS BACK TO "UNVERIFIED". AFTER THEIR UPGRADE, THEY NOTICED THE STUDY STATUS WAS NOT CHANGING WHEN ADDITIONAL DOCUMENTATION WAS ADDED. WITH A CUSTOMER BEING POTENTIALLY UNAWARE WHEN A NEW IMAGE IS ADDED TO A STUDY, THERE IS A POTENTIAL FOR DELAY IN THE DIAGNOSIS AND/OR TREATMENT OF A PATIENT THAT COULD LEAD TO HARM. HOWEVER, THE ISSUE WAS APPARENT TO THE USER AND THERE IS NO INDICATION OF HARM TO A PATIENT. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. IN THE AFTERNOON HOURS OF (B)(6) 2016 A CUSTOMER REPORTED THEY WERE UNABLE TO LOGIN INTO MERGE PACS. PHYSICIANS AND RADIOLOGISTS WERE ABLE TO VIEW PATIENT STUDIES DIRECTLY AT THE MODALITY RATHER THAN THROUGH MERGE PACS. HOWEVER, AFTER 4 HOURS OF VIEWING STUDIES AT THE MODALITY, THE CUSTOMER ELECTED TO SEND PATIENTS SCHEDULED FOR IMAGING STUDIES TO OTHER FACILITIES. MERGE HEALTHCARE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND APPROXIMATELY SEVEN HOURS AFTER THE CUSTOMER INITIALLY CONTACTED MERGE HEALTHCARE ABOUT THE INABILITY TO LOGIN TO MERGE PACS, THE SITE REGAINED ACCESS. THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS DUE TO THE SITE ELECTING TO SEND SCHEDULED PATIENTS TO OTHER FACILITIES RATHER THAN VIEWING IMAGES ON THE MODALITY WHILE USERS WERE UNABLE TO LOG INTO MERGE PACS. FOLLOW-UP COMMUNICATION WITH THE CUSTOMER CONFIRMED THERE WAS NO HARM TO PATIENTS AS A RESULT OF THIS ISSUE. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2016 THE CUSTOMER REPORTED THAT THE HL7 INTERFACE IS OVERWRITING PATIENT INFORMATION. IT WAS FOUND THAT THE GLOBAL PATIENT UPDATE WAS TURNED ON BY AN EMPLOYEE ON SITE. THIS SETTING AUTOMATICALLY DEFAULTS TO OFF UNLESS THE SETTINGS ARE MANUALLY SWITCHED TO ON. THIS RESULTED IN PATIENT INFORMATION BEING UPDATED TO THE INCORRECT PATIENT IN THEIR SYSTEM WHEN A DUPLICATE MRN NUMBER OCCURRED. THIS ISSUE HAS A POTENTIAL TO LEAD TO PATIENT INFORMATION BEING COMPROMISED OR OVERWRITTEN BY A NEW STUDY SUBMISSION. THERE WAS NO INDICATION FROM THE CUSTOMER THAT THERE WERE ANY PATIENTS WHO WERE HARMED DUE TO THIS COMPLAINT. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED THAT WHEN MOVING OR CORRECTING STUDIES USING THE QC EDITOR IN MERGE PACS, THE STUDY COMMENTS MADE ON THE SOURCE STUDY WERE NOT RETAINED AND MOVED TO THE TARGET STUDY. THIS MAY CAUSE A DELAY IN PATIENT CARE OR MISDIAGNOSIS IF IT'S NOT READILY APPARENT TO THE USER AND THEY OVERLOOK IMPORTANT INFORMATION WITHIN THE COMMENTS. THE PATIENT'S IMAGES ARE AVAILABLE FOR DICTATION AND COMMENTS ARE SUPPLEMENTAL INFORMATION. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREEDIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2017, A CUSTOMER REPORTED WHEN TWO STUDIES ARE MERGED USING THE QC EDITOR, THE STUDY COMMENTS MADE ON THE SOURCE STUDY WERE NOT RETAINED AND MOVED TO THE TARGET STUDY. THIS MAY CAUSE A DELAY IN PATIENT CARE OR MISDIAGNOSIS IF IT'S NOT READILY APPARENT TO THE USER AND THEY MISS IMPORTANT INFORMATION. THE IMAGES ARE AVAILABLE FOR DICTATION AND COMMENTS ARE SUPPLEMENTAL INFORMATION. (B)(4). | -fail to display data -inappropriate image format -inadequate image types | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A PICTURE ARCHIVING COMMUNICATION SYSTEM THAT IS INTENDED TO CREATE AND DISPLAY TWO-DIMENSIONAL AND THREE-DIMENSIONAL IMAGES OF ANATOMY FROM A SERIES OF DIGITALLY ACQUIRED IMAGES. PACS IS DESIGNED FOR SOFT COPY READING, COMMUNICATION, AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. A CUSTOMER REPORTED ON (B)(6) 2016 THAT THEY WERE EXPERIENCING ENTERPRISE WIDE SLOWNESS. THEY WERE HAVING ISSUES PULLING ORDERS FROM THE WORKLIST. MERGE PACS WAS SLOW TO SCAN, SEND AND RETRIEVE AND THE CUSTOMER STATED THEY WERE NOT ABLE TO BURN A CD DUE TO THESE ISSUES. MERGE SUPPORT BEGAN INVESTIGATING THE ISSUE AND FOUND A RELATED SUPPORT CASE (B)(6). THAT CASE HAD BEEN OPENED ON (B)(6) AND ANOTHER MEMBER OF THE SUPPORT TEAM WAS ALREADY WORKING WITH SITE ON SLOWNESS ISSUES. THIS CASE WAS CLOSED AS A DUPLICATE AND INFORMATION WAS TRANSFERRED TO ORIGINAL CASE. THIS REPORT IS BEING ISSUED AS A RESULT OF THE CUSTOMER REPORTING AGAIN ISSUES WITH SLOWNESS PRIOR TO THE DICOM TABLES WERE BEING REORGANIZED. THE SLOWNESS HAS A POTENTIAL OF LEADING TO A DELAY IN A TREATMENT OF DIAGNOSIS OF A PATIENT. (B)(4). | -fail to display data -inappropriate image format -unable to zoom image | -no display or display failure -device unsafe to use in environment -incorrect display | Viewer
| FDA | Event Description: MERGE PACS IS A SCALABLE SOLUTION THAT SUPPORTS A DISTRIBUTED WORKFORCE AND DELIVERS OPERATIONAL EFFICIENCIES THROUGH WEB-BASED ACCESS, COMMON WORKLISTS, AND COMPREHENSIVE WORKFLOWS. THE CUSTOMER REPORTED THAT A NEW PATIENT CAME IN FOR A PROCEDURES. THE TECHNOLOGIST CLICKED ON PRIOR STUDIES AND NOTICED THERE WAS A BIG LIST AND SINCE IT WAS A NEW PATIENT, THERE SHOULD NOT HAVE BEEN PRIORS AT THAT POINT. SPECIFICALLY, THEY NOTICED A TEST PATIENT'S STUDY WAS ASSOCIATED WITH AN ACTUAL PATIENT'S STUDY. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE DUE TO THE POSSIBILITY OF IN ADVERTENTLY COMBINING STUDIES FROM DIFFERENT PATIENTS. ON 12/4/15, THE OBSERVATION TABLE WAS UPDATED WITH THE CORRECT PATIENT. SISU WILL SUPPRESS THE URN FROM COMING THROUGH ON FUTURE HL7 MESSAGES TO PACS TO PREVENT THIS ISSUE FROM HAPPENING AGAIN IN THE FUTURE. | -fail to access across devices and locations -fail to access -fail to deliver | -compliance issue -failure of device to self-test -filling problem | Viewer
| FDA | Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). MERGE PACS CAN BE ACCESSED FROM WITHIN THE HOSPITAL OR ENTERPRISE, OR FROM REMOTE LOCATIONS VIA WEB-BASED ACCESS. IMAGES VIEWED ON MOBILE DEVICES MUST NOT BE USED FOR DIAGNOSTIC REVIEW. ON (B)(6) 2021, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT THERE WERE DISCREPANT ROI (REGION OF INTEREST) VALUES BETWEEN 2D AND MPR IMAGES. MERGE HEALTHCARE SUPPORT WORKED WITH THE CUSTOMER TO REPRODUCE THE ISSUE. THE ISSUE IS CONFINED TO THE 3D FLAIR SERIES OF IMAGES ONLY. THIS ISSUE HAS BEEN CORRECTED IN THE CURRENT RELEASE OF MERGE PACS AND THE CUSTOMER WAS INFORMED OF THIS. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to read images -unable to read -fail to provide diagnostic test | -use of device issue -contamination during use -disconnection | Viewer
| FDA | Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS), SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO REPORT THAT SOME IMAGES IN A MAMMOGRAPHY SERIES WERE NOT DISPLAYING AND WERE MISSED OR SKIPPED. MERGE HEALTHCARE SUPPORT IS CURRENTLY IN THE PROCESS OF ENGAGING THE CUSTOMER TO DETERMINE THE ISSUE AND NEXT STEPS. WHEN FURTHER INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to read images -unable to read -fail to provide diagnostic test | -use of device issue -contamination during use -disconnection | Viewer
| FDA | Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STANDALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS ICONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, MERGE HEALTHCARE RECEIVED A REPORT FROM A CUSTOMER THAT HOUNSFIELD UNIT VALUES REPORTED IN THE 3D VIEW WERE INCONSISTENT WITH THE HOUNSFIELD UNIT VALUES REPORTED IN THE 2D VIEW OF MERGE PACS. UPON INVESTIGATION, IT WAS DISCOVERED THAT PIXEL VALUE MEASUREMENTS (INCLUDING HOUNSFIELD UNIT MEASUREMENTS AND SUV MEASUREMENTS) FROM THE PROBE, ROI, AND REGION ANALYSIS AREA TOOLS ARE INCORRECT ON MULTI-PLANAR REFORMATS (MPRS) GENERATED BY MERGE PACS. THE ISSUE IS NOTICEABLE WHEN A MODALITY ACQUIRES A SERIES OF IMAGES AND SENDS THE IMAGES TO MERGE PACS. THE RADIOLOGIST THEN GENERATES ADDITIONAL PLANES SUCH AS SAGITTAL OR CORONAL USING MPR. MEASUREMENTS DONE ON THE MERGE PACS GENERATED MPR'S MAY HAVE INCORRECT MEASUREMENTS. INCORRECT PIXEL MEASUREMENTS MAY RESULT IN THE IMPLEMENTATION OF AN INCORRECT OR UNNECESSARY TREATMENT PLAN THAT COULD RESULT IN THE PATIENT RECEIVING UNNECESSARY MEDICATION, CHANGE IN DOSAGE OF EXISTING MEDICATION, OR UNNECESSARY MEDICAL PROCEDURES. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. (B)(4). | -fail to read images -unable to read -fail to provide diagnostic test | -use of device issue -contamination during use -disconnection | Viewer
| FDA | Event Description: MERGE PACS IS A STANDARDS-BASED MEDICAL IMAGING DIAGNOSTIC WORKSTATION THAT SERVES AS AN ADJUNCT TO ASSIST THE CLINICIAN TO VIEW, READ, AND REPORT THEIR FINDINGS. MERGE PACS PROCESSES AND DISPLAYS MEDICAL IMAGES FROM DICOM-COMPLIANT MODALITIES. THE DEVICE IS DESIGNED TO ENABLE EFFICIENT WORKFLOWS BY MAINTAINING CLINICIANS' WORKLISTS AND RETRIEVING AND MANAGING STUDIES FOR READING, REPORTING, COMMUNICATION, AND STORAGE. MERGE PACS SOFTWARE RUNS ON OFF-THE-SHELF COMPUTER HARDWARE AND CAN BE CONFIGURED TO OPERATE STAND ALONE OR TO INTEGRATE WITH VENDOR-NEUTRAL IMAGING ARCHIVES (VNAS) SUCH AS I CONNECT ENTERPRISE ARCHIVE (ICEA) FOR IMAGE STORAGE, AND WITH RADIOLOGICAL AND HOSPITAL INFORMATION SYSTEMS (RIS AND HIS) AND MEDICAL RECORD SYSTEMS (EMR, EHR, ETC.). ON (B)(6) 2020, THE CUSTOMER REPORTED THAT A STUDY CAME IN AND THE CT TECH MANUALLY CHANGED FROM UNVERIFIED TO UNREAD AND SET THE CUSTOM STATUS TO ER. IT UPDATED EVERYTHING AS EXPECTED AND SHOWED ON THE RADIOLOGST'S ER WORKLIST. A LITTLE LATER, THE CT TECH SENT IN SCANNED PAPERWORK, WHICH REVERTED THE STATUS BACK TO UNVERIFIED, THEREBY, REMOVING THE STUDY FROM THE RADIOLOGIST'S ER WORKLIST (WHICH IS EXPECTED BEHAVIOR). HOWEVER, THE STATUS OF THE STUDY DID NOT UPDATE TO UNVERIFIED ON THE CT TECH'S WORKLIST AS EXPECTED. THE CT TECH CALLED THE RADIOLOGIST INQUIRING ABOUT THE STUDY AND THE RADIOLOGIST DID NOT HAVE THAT STUDY ON THEIR WORKLIST. THE CT TECH THEN FOUND THAT THE STUDY NEEDED TO BE RE-VERIFIED, THEREBY PLACING ONTO THE RADIOLOGIST'S ER WORKLIST. THE CT TECH WAS NOT ALERTED THAT THE STUDY HAD BEEN PLACED BACK IN THE UNVERIFIED STATUS AND NEEDED TO BE RE-VERIFIED TO UNREAD. BASED ON THIS, THERE IS A POTENTIAL FOR DELAY IN TREATMENT THAT COULD CAUSE HARM TO THE PATIENT. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. | -fail to read images -unable to read -fail to provide diagnostic test | -use of device issue -contamination during use -disconnection | Viewer
| FDA | Event Description: MERGE PACS IS DESIGNED AND MARKETED FOR SOFT COPY READING, COMMUNICATION AND STORAGE OF STUDIES PRODUCED BY DIGITIAL MODALITIES. CUSTOMER REPORTED ((B)(4)) THAT THE INCORRECT PATIENT'S IMAGES WERE DISPLAYED IN THE HANGING PROTOCOL COMPARISON WINDOW WHEN USING MERGE PACS. THE USER CLOSED THE VIEWER AND RELAUNCHED THE STUDY. ON THE SECOND LOAD, THE CORRECT COMPARISON WAS DISPLAYED. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE INCORRECT THE PATIENT'S IMAGES WERE BEING DISPLAYED FOR COMPARISON. | -incorrect product classification -inappropriate package design -incorrect package labeling | -data back-up problem -device contamination with biological material -device maintenance issue | Viewer
| FDA | Event Description: MERGE PACS PROVIDES A CENTRAL HUB FOR MANAGING AND VIEWING STUDIES, WHILE PROVIDING ACCESS TO REMOTE USERS VIA THE WEB. ITS COMMON WORKLISTS ENABLE EFFICIENT SUB-SPECIALTY READING FROM REMOTE LOCATIONS. ON (B)(6) 2015, A CUSTOMER REPORTED THAT THE STUDY WAS MISSING IMAGES WHEN VIEWED IN MERGE PACS. HOWEVER, THE IMAGES WERE STILL AVAILABLE ON THE MODALITY DEVICE. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE PATIENT IMAGES WERE NOT ACCESSIBLE TO MAKE A COMPREHENSIVE DIAGNOSIS. MERGE REMOVED THE STUDY FROM MERGE PACS AND HAD IT RESENT FROM THE MODALITY TO MAKE THE PATIENT IMAGES AVAILABLE FOR DIAGNOSTIC REVIEW. THE CUSTOMER VERIFIED THE ISSUE WAS RESOLVED. | -unable to view display -fail to access -unable to read | -human-device interface issue -unauthorized access to computer system -date-related software issue | Viewer
| FDA | Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016 A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT INDICATING THAT THEY WERE UNABLE TO FIND A STUDY. MERGE INVESTIGATED THE ISSUE AND DETERMINED THAT THE ALLEGED MISSING STUDY WAS STORED IN 2007 AND WATERMARKED OFF IN 2008. MERGE HEALTHCARE HAS NOT BEEN ABLE TO LOCATE THE IMAGE DATA. WITH HISTORIC IMAGES NOT BEING AVAILABLE FOR VIEWING, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS AND/OR TREATMENT. HOWEVER, THERE WERE NO REPORTS OF DELAY OR HARM TO PATIENTS AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THE ISSUE TO DETERMINE WHAT MAY HAVE HAPPENED TO THE ARCHIVED IMAGES. | -fail to store medical imaging -fail to process medical imaging -lack of display | -use of device issue -data issue -patient data issue | Viewer
| FDA | Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT X-RAY IMAGES WERE BLURRY. DUE TO IMAGES NOT PRESENTING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM. THERE IS NO INDICATION OF A DELAY IN TREATMENT OR DIAGNOSIS OR HARM THAT HAS OCCURRED AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THE CUSTOMER'S ALLEGATION. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT SHALL BE SUBMITTED. | -fail to store medical imaging -fail to process medical imaging -lack of display | -use of device issue -data issue -patient data issue | Viewer
| FDA | Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT THEY "CANT INVERT IMAGE WHEN USING WINDOW LEVEL TOOL ON ONE STUDY". WITH IMAGES NOT BEING PRESENTED AS EXPECTED TO A USER, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH MAY LEAD TO HARM. THERE HAVE BEEN NO REPORTS OF DELAY IN TREATMENT OR DIAGNOSIS TO A PATIENT OR OF HARM AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS INVESTIGATING THIS ALLEGATION BY THE CUSTOMER. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | -fail to store medical imaging -fail to process medical imaging -lack of display | -use of device issue -data issue -patient data issue | Viewer
| FDA | Event Description: MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. ON (B)(6) 2016, MERGE HEALTHCARE SUPPORT RECEIVED A CUSTOMER ALLEGATION INDICATING THAT X-RAY IMAGES WERE DARK AND THAT THESE IMAGES ARE UNABLE TO BE READ. FURTHER DISCUSSIONS WITH THE CUSTOMER REVEALED THAT A RADIOLOGIST OVERLOOKED AN IMAGE DUE TO THE IMAGE BEING BLACK WITH NO STUDY INFORMATION ON THE VIEWPORT. WITH RADSUITE NOT PRESENTING IMAGES IN A MANNER THAT ALLOWS FOR VIEWING, THERE IS A POTENTIAL FOR A DELAY IN A PATIENT'S TREATMENT OR DIAGNOSIS IF AN IMAGE IS OVERLOOKED. THE CUSTOMER HAS INDICATED THAT NO HARM HAS RESULTED TO A PATIENT AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: THE CUSTOMER'S ALLEGATION IS UNDER FURTHER INVESTIGATION BY MERGE HEALTHCARE. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE FILED. | -fail to store medical imaging -fail to process medical imaging -lack of display | -use of device issue -data issue -patient data issue | Viewer
| FDA | Event Description: MERGE RADSUITE IS AN APPLICATION THAT PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. THE APPLICATION PROVIDES WORKFLOW INTEGRATION CAPABILITIES FOR HEALTH CARE ENTERPRISES, WHEREIN: RADIOLOGISTS CAN VIEW, ANNOTATE, AND TAG STUDIES AS DIAGNOSTICALLY READ; REFERRING PHYSICIANS CAN VIEW PATIENT IMAGES AND RADIOLOGISTS' ANNOTATIONS AND; TERTIARY CARE PHYSICIANS, MEDICAL TECHNOLOGISTS, AND INFORMATION TECHNOLOGY PROFESSIONALS CAN RECEIVE PATIENT RECORDS. ON (B)(6) 2015, THE CUSTOMER REPORTED ((B)(6)) THAT CASES BEING SENT FROM RADSUITE TO AN OFFSITE DESTINATION WAS SLOW. THIS WAS A THIRD PARTY PACS VENDOR THAT DOES OFFSITE READING FOR THE CUSTOMER. THE CUSTOMER SENDS FROM RADSUITE AT NIGHT FOR THEIR STROKE PATIENTS. THE SITE REQUIRES 3 MINUTE AVERAGE TRANSFER IN ORDER TO MEET 15 MINUTES COMPLETE TO FINAL TURNAROUND TIME FOR STROKE/PERFUSION PATIENTS. EXPORTING AT NIGHT COULD TAKE 20 - 40 MINUTES. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION IS NOT AVAILABLE FOR REVIEW, LEADING TO A DELAY IN TREATMENT FOR STROKE VICTIMS. THE ROOT CAUSE OF THIS ISSUE WAS NOT LINKED TO RADSUITE OR ANY OTHER MERGE PRODUCT. IT WAS TIED TO A 3RD PARTY VENDOR THAT ST. JOHN'S IS USING TO READ THEIR STUDIES AT NIGHT. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE ROOT CAUSE WAS NOT RELATED TO RADSUITE BUT TO AN OUTSIDE SERVICE USED BY THE CUSTOMER. | -fail to process medical imaging -lack of display -inadequate message display | -use of device issue -data issue -inadequate user interface | Viewer
| FDA | Event Description: MERGE RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER REPORTED THAT THE PATIENT LOCATION OF ER WAS NOT COMING ACROSS THE SYSTEM. THE ER WORKLISTS ARE FILTERED BASED ON PATIENT LOCATION. A PRIOR STUDY WAS RETRIEVED AND TRIGGERED AN UPDATE TO THE LOCATION. AS SUCH, THE ER STUDY DROPPED OFF THE ER EXAMLIST. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT IMAGES ARE NOT AVAILABLE FOR A DIAGNOSIS. IT WAS DISCOVERED THAT THE CUSTOMER HAD AN INCORRECT UNDERSTANDING OF THE MERGE SOFTWARE. THE CUSTOMER HAD ASSUMED THAT THE PATIENT LOCATION WAS TRACKED AT AN ENCOUNTER LEVEL, BUT THE PRODUCT DOES NOT HAVE ENCOUNTER DRIVEN WORKFLOW. NO CHANGES BEING MADE TO THE PRODUCT. THIS ISSUE IS NOT A RESULT OF A DEFECT OR PRODUCT LABELING PROBLEMS. (B)(4). Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -fail to process medical imaging -fail to provide patient diagnosis info -fail to display data | -data issue -incorrect or inadequate result -defective item | Viewer
| FDA | Event Description: MERGE UNITY (B)(4) A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. INFORMATION WAS RECEIVED FROM A (B)(4) ADMINISTRATOR ON (B)(6) 2016 REPORTING A MISSING REPORT. THE UNITY PACS AUDIT TRAIL STATES THAT THE REPORT WAS SAVED (B)(6) 2016. MERGE SUPPORT INVESTIGATED AND CONFIRMED THAT A REPORT WAS SAVED ON THE STATION. HOWEVER, THERE WAS NO DOCUMENT IN THE REPORT DIRECTORY. A BLANK WORD DOCUMENT WAS PLACED IN THE DIRECTORY AND THE PHYSICIAN RE-READ THE EXAM. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -fail to manage data | -contamination during use -use of device issue -image display error | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS A CUSTOMER CONTACTED MERGE HEALTHCARE ON (B)(6) 2016 AND REPORTED SEEING A MESSAGE THAT A REPORT DID NOT CROSS TO THEIR RADIOLOGY INFORMATION SYSTEM (RIS) SYSTEM (EPIC). AFTER NOTICING THE VISIT NUMBER DID NOT MATCH, THE CUSTOMER CORRECTED IT IN MERGE UNITY AND TRIED TO RESEND. WHEN THE REPORT OPENED, IT WAS BLANK. THE CUSTOMER WAS UNABLE TO VIEW THE REPORT. THE MERGE HEALTHCARE UNITY SUPPORT TECHNOLOGIST PLACED A BLANK WORD DOCUMENT IN THE REPORT DIRECTORY ON THE IMAGE SERVER. THIS ALLOWED THE EXAM TO BE RE-READ BY THE PHYSICIAN AND THE REPORT SAVED AND WAS ACCESSIBLE. (B)(4). Manufacturer Narrative: EXAM WAS AVAILABLE FOR REVIEW. THE ISSUE WAS IDENTIFIED BY THE PACS ADMIN DURING INTERNAL QC WORK VIEWING FAILURES OF HL7 INTERFACES. LOGS FOR THIS STATION/INCIDENT WERE UNAVAILABLE FOR REVIEW. THERE WAS A "REPORT SAVE" EVENT IN THE AUDIT TRAIL, HOWEVER, NO EXAM REPORT UPLOADED TO THE IMAGE SERVER. IT WAS UNABLE TO BE DETERMINED IF THIS INCIDENT WAS RELATED TO SOFTWARE OR TO SITE NETWORK FAILURE. THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER ALLEGED ON (B)(6) 2016 THAT HE WAS UNABLE TO OPEN AN APPROVED REPORT FROM THE ONLINE TAB. MERGE HEALTHCARE SUPPORT WAS UNABLE TO LOCATE AN ASSOCIATED DOCUMENT IN THE REPORT DIRECTORY ON THE CUSTOMER'S IMAGE SERVER. A BLANK WORD DOCUMENT WAS PLACED IN THE DIRECTORY WHICH ALLOWED THE PHYSICIAN TO RE-READ THE EXAM. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NOT ASSOCIATING TO AN EXAM AND NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: THIS INCIDENT WAS A RESULT OF THE UNIVERSAL MANAGER CRASHING AT THE POINT THE RADIOLOGIST WAS CLOSING THE CASE. SITE SOFTWARE WAS UPDATED ON (B)(6) 2017 FROM (B)(6). Manufacturer Narrative: MERGE HEALTHCARE INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THERE WAS A "REPORTSAVE" EVENT IN THE AUDIT TRAIL BUT THERE WAS NO REPORT AVAILABLE FOR THE PATIENT. MERGE HEALTHCARE'S UNITY PACS SUPPORT WAS UNABLE TO REVIEW THE STATION SPECIFIC LOGS BECAUSE THEY HAD ALREADY BEEN TRUNCATED. MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT ON (B)(6) 2016 AND REPORTED A MISSING REPORT. AN AUDIT OF THE EXAM SHOWED THE CASE DID HAVE A 'REPORT SAVE' ENTRY IN THE TRAIL. THERE IS NO REPORT LISTED FOR THE EXAM. THE STATION SPECIFIC LOGS WERE TRUNCATED PRIOR TO UNITY TECHS REVIEW. THE REPORT WAS NOT SALVAGEABLE AND NEEDED TO BE RE-DICTATED BY THE PHYSICIAN. WITH A PATIENT'S STUDY REQUIRING RE-DICTATION THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT. HOWEVER, THERE HAS BEEN NO INDICATION OF ANY HARM AS A RESULT OF THIS ISSUE. (B)(6). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT ON (B)(6) 2017. THEY ALLEGED THAT THE REPORT FOR AN EXAM READ ON (B)(6) 2017 WAS NOT FOUND. THIS WAS NOTICED BY THE MANAGER WHILE RUNNING A REPORT ON (B)(6) 2017 IN THE MORNING. THE REPORT IS RUN TO MAKE SURE THE EXAMS WERE COMPLETE AND THEIR RESULTS CROSSED FROM UNITY PACS TO THEIR EMR SUCCESSFULLY. A MERGE UNITY TECH SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION, REVIEWED THE LOGS AND DETERMINED THAT THE REPORT WAS NOT SAVED. THE PHYSICIAN RE-READ THE EXAM ON (B)(6) 2017 FOR RESOLUTION. WHILE PATIENT IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER 100615 Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ALLEGATION TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A MERGE UNITY PACS ADMINISTRATOR REPORTS THAT AN EXAM REPORT DONE ON (B)(6) 2016 SHOWS AS NOT AVAILABLE WHEN THEY ATTEMPT TO VIEW THE REPORT. THERE IS A 'REPORT SAVE' EVENT IN THE AUDIT TRAIL. THE EXAM IN QUESTION WAS PART OF AN EXAM THAT WAS SPLIT AFTER ACQUISITION. THERE WAS NO EVIDENCE IN THE LOGS OF THE READING STATION THAT SHOW THE REPORT WAS ACTUALLY SAVED. WITH MERGE UNITY PACS NOT OPERATING AS EXPECTED THERE IS A POTENTIAL FOR A DELAY IN TREATMENT IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM HOWEVER, THE CUSTOMER HAS NOT ALLEGED ANY HARM. (B)(4). Manufacturer Narrative: EXAM WAS OPENED ON THE DOMINATOR READING STATION AND NO EVIDENCE EXISTED THAT A REPORT WAS SAVED. THE EXAM WAS SPLIT OUT FROM ANOTHER, THERE IS A POSSIBILITY THIS WAS A RESULT OF WORK FLOW FROM THE READING PHYSICIAN. PACS ADMIN DID GO OVER THE WORK FLOW WITH THE PHYSICIAN AS A PRECAUTION WHEN READING EXAMS THAT ARE SPLIT INTO TWO SEPARATE EXAMS. SITE SOFTWARE WAS UPGRADED FROM (B)(4) ON 1/21/2017. THE EXAM WAS AVAILABLE FOR REVIEW AT ALL TIMES, AND THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. INFORMATION WAS RECEIVED FROM A PACS ADMINISTRATOR ON (B)(6) 2017 ALLEGING THAT A PHYSICIAN DICTATED AND SAVED A PATIENT'S STUDY. THE PHYSICIAN THAT COMPLETED THE STUDY WAS CONTACTED THE NEXT DAY BY A REFERRING PHYSICIAN WITH PROBLEMS VIEWING THE PATIENT'S EXAM REPORT. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND FOUND THAT THE AUDIT TRAIL SHOWED A REPORT SAVE EVENT, HOWEVER, THE REPORT DISPLAYS WITH A MESSAGE THAT THE REPORT IS NOT YET FINALIZED. THERE WAS NO INFORMATION REGARDING THE EXAM FINDINGS FOR THIS PATIENT/EXAM DISPLAYED. THE READING PHYSICIAN WENT OVER FINDINGS WITH REFERRING PHYSICIAN OVER THE PHONE AND RE-READ THE EXAM. WITH MERGE UNITY PACS NOT FUNCTIONING AS EXPECTED WHERE A REPORT DOES NOT INCLUDE EXAM FINDINGS AS DICTATED BY A PHYSICIAN THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS/TREATMENT OR A MISDIAGNOSIS THAT COULD POTENTIALLY LEAD TO HARM. HOWEVER THERE IS NO INDICATION OF ANY HARM TO THE PATIENT AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: INVESTIGATION FOUND A SPECIFIC ERROR IN THE LOGS. THERE WAS A "READ" EVENT THAT OCCURRED BUT NOT A "SAVE" EVENT. IF A SAVE EVENT IS NOT FOUND FOLLOWING THE READ EVENT THE REPORT HAD NOT BE PROPERLY SAVED. THE EXAM WAS AVAILABLE FOR REVIEW AT ALL TIMES. THE READING PHYSICIAN COMMUNICATED INFORMATION TO THE REFERRING PHYSICIAN OF FINDINGS AND THEN RE-READ THE EXAM. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ISSUE AS REPORTED BY THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. INFORMATION WAS RECEIVED FROM THE PACS ADMINISTRATOR ON (B)(6) 2016 REPORTING AN EXAM REPORT WAS NOT AVAILABLE. AN INVESTIGATION SHOWED THERE IS A 'REPORT SAVE' EVENT IN THE AUDIT TRAIL, BUT NO EVIDENCE OF THE REPORT IN THE 'REPORTS' DIRECTORY. NO LOGS ARE AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THERE IS NO EVIDENCE THE REPORT WAS UPLOADED TO THE IMAGE SERVER AND ARCHIVED. THE REPORT WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT RESULTS IN HARM. HOWEVER, THERE IS NO INDICATION OF A SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THIS CASE/COMPLAINT WAS LINKED TO JIRA (B)(4). ATTEMPTS TO SAVE A REPORT TOO QUICKLY CAUSES A "MISSING REPORT" PROBLEM. THE JIRA WAS CLOSED AND VERIFIED AS BEING RESOLVED IN SOFTWARE BUILD R11.1.2. THE SITE WAS UPGRADED TO BUILD R11.1.2.2 ON 04/07/2017. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE WAS NOTIFIED THAT THERE WAS A DISCREPANCY BETWEEN THE OBSTETRIC RISK ANALYSIS PERCENTILES FROM THE CUSTOMER'S EXAMFORM AND MERGE'S STRUCTURED REPORT. THE CUSTOMER REPORTED THAT BECAUSE OF THE INCONSISTENCY IN THE NUMBERS, THE PATIENT WAS CONCERNED ABOUT THE WELLBEING OF THE BABY. THE PATIENT WAS REASSURED THAT THE BABY WAS WITHIN THE NORMAL RANGE. HOWEVER, AN INCONSISTENCY IN THE OBSTETRIC RISK ANALYSIS PERCENTILES CAN CAUSE OR CONTRIBUTE TO POTENTIAL HARM. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE HEALTHCARE WAS NOTIFIED BY A CUSTOMER ON (B)(6) 2017 THAT ALL USERS WERE NOT ABLE TO CONNECT TO THE SERVER. TESTING AND INVESTIGATION BY MERGE HEALTHCARE TECHNICAL SUPPORT DETERMINED THAT THE SITE HAD BEEN EXPERIENCING ONGOING STORAGE ISSUES WITH THEIR DRIVES. IN ORDER TO RESOLVE THIS HARDWARE STORAGE ISSUE, MERGE HEALHCARE SENT THE CUSTOMER SIX (6) NEW ONE TERABYTE (TB) DRIVES. THE NEW CONFIGURATION WOULD ALLOW FOR DISTRIBUTING THE PACS STORAGE ACROSS MULTIPLE DRIVES ENSURING THAT FREE SPACE AND REGULAR MAINTENANCE (DEFRAGGING) WOULD OCCUR. HOWEVER, DURING THE ON-SITE REPLACEMENT INSTALLATION, THE ARRAY DRIVE IN THE IMAGE SERVER FAILED. THE FAILED ARRAY WAS REPLACED. IN THE EFFORTS OF RESTORING THE DATA, FOUR EXAMS WERE UNABLE TO BE RECOVERED. THEREFORE, EACH PATIENT HAD TO BE RE-SCHEDULED FOR ANOTHER MAMMOGRAM EXAM AT A LATER DATE. THERE IS A POTENTIAL FOR PATIENTS THAT REQUIRE RE-IMAGING DUE TO DATA LOSS, TO EXPERIENCE A DELAY IN TREATMENT AND/OR DIAGNOSIS THAT COULD LEAD TO HARM. AT THIS TIME, THERE IS NO INDICATION OF AN ADVERSE EVENT FOR THE AFFECTED PATIENTS. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ISSUE REPORTED BY THE CUSTOMER. A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING DATABASE/DISK ARRAY GUARDIAN ISSUES. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE ARE THREE SIMILAR COMPLAINTS FOR THIS ISSUE. TRENDING WILL CONTINUE TO BE MONITORED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A REPORT WAS PROVIDED WITH A REQUEST FOR HELP WITH STRUCTURED REPORTS (SR) BY A CUSTOMER USING MERGE UNITY PACS WHO WAS ATTEMPTING TO CREATE A TEMPLATE FOR MULTIPLE FETUS EXAMS. THE MERGE UNITY STRUCTURED REPORT (SR) IS NOT SEPARATING FIELDS FOR ULTRASOUND (US) OBSTETRIC (OB) EXAMS WITH TWINS(AKA MULTIPLE FETUS EXAMS). WITH MERGE UNITY PACS NOT REPORTING PATIENT INFORMATION AS EXPECTED, THERE IS A POTENTIAL FOR MISTREATMENT OR MISDIAGNOSIS. THE ISSUE WAS IDENTIFIED DURING CREATION OF THE SR TEMPLATE AND NO PATIENT INVOLVEMENT WAS REPORTED. (B)(4). Manufacturer Narrative: FURTHER REVIEW BY OUR DEVELOPMENT TEAM UNDER (B)(4) FOUND. THERE WAS NO EXPECTATION THAT THE TWIN FETUS SR WOULD ONLY CONTAIN A SINGLE FETUS. THE SOFTWARE APPROPRIATELY PROVIDED AN ERROR MESSAGE WITH ONLY A SINGLE FETUS WAS PROVIDED. THE SOFTWARE WORKED AS EXPECTED WHEN USING STANDARD SR EXAM FORMS WITH A TWINS SR. THE SOFTWARE WORKED AS EXPECTED WITH USING OB SEND EXAM FORMS WITH AN SR WITH TWINS. THIS WAS CAUSED BY THE TECHNOLOGIST NOT PERFORMING THE STUDY CORRECTLY. THEY WERE PERFORMING ALL OF THE MEASUREMENTS FOR ONE FETUS AND SENDING THAT STRUCTURED REPORT TO THE PACS. THEY THEN PERFORMED ALL OF THE MEASUREMENT FOR THE SECOND FETUS AND SET THAT SR. THIS RESULTED IN AN ERROR AT TIME OF PROCESSING THE SR INFORMATION. THIS RESULTED IN NO OB CALCULATIONS BEING PERFORMED. THERE WAS A LITTLE CONFUSION IN THE INVESTIGATION AS WE MIXED THESE RESULTS WITH THE QA TESTING RESULTS AND INCORRECTLY THOUGHT THAT AFTER THE ERROR THE VALUES OF ONE FETUS WAS BEING PLACED IN BOTH LOCATIONS. THIS PROVED TO BE FALSE. AFTER THE ERROR MESSAGE OF NO VALUES FILLED OUT. THIS IS EXPECTED BEHAVIOR. THERE IS NO ATTEMPT BY THE PACS TO RECONCILE MULTIPLE SRS OF THE SAME TYPE. IT IS ASSUMED THAT THE LAST SR IS THE MOST ACCURATE AND UP TO DATE. IT HAS BEEN DETERMINED THIS IS MOST LIKELY RELATED TO A USE ERROR. NO FURTHER ACTION AT THIS TIME. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING PAGES. THE EXAM WAS CONDUCTED ON (B)(6) 2016; REPORT WAS APPROVED AND DICTATED ON (B)(6) 2016. THE ISSUE WAS REPORTED TO MERGE ON (B)(6) 2016 THAT ONLY 3 OF THE 4 PAGES WERE DISPLAYING. THEREFORE, THE REPORT NEEDED TO BE RE-DICTATED BY THE DOCTOR. THE REPORT WAS RE-READ ON (B)(6) 2016 AND APPROVED ON (B)(6) 2016. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: TRENDING OF MISSING REPORT IS CONTINUED TO BE MONITORED AND ADDITIONAL INVESTIGATION ACTIVITIES ARE BEING PERFORMED TO DETERMINE IF CORRECTIVE ACTIONS ARE REQUIRED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A PATIENT DEATH OCCURRED. THE NOTIFICATION WAS PROVIDED VIA EMAIL WHILE INQUIRING ABOUT WHETHER THERE IS A NOTIFICATION OR WARNING FOR AN EXAM WHEN A TECHNOLOGIST DOES NOT REVIEW ALL IMAGES. ON (B)(6) 2016, FOLLOW UP CALLS WERE MADE TO THE CUSTOMER IN ORDER TO GATHER ADDITIONAL INFORMATION. THE FACILITY PROVIDED FURTHER INFORMATION REGARDING THE INCIDENT. THE PATIENT WAS SEEN ON A FRIDAY AFTERNOON AND IMAGES WERE OBTAINED. AFTER THE IMAGES WERE CAPTURED, THE PATIENT WAS INSTRUCTED THAT THEY WERE NOT TO OPERATE A MOTOR VEHICLE BASED ON A KNOWN SEIZURE ISSUE. THE FACILITY FURTHER CLARIFIED THAT THEY HAVE NO MALFUNCTION TO REPORT REGARDING MERGE HEALTHCARE PRODUCTS. THE E-MAIL FROM THE CUSTOMER APPEARS TO HAVE ONLY BEEN AN INQUIRY SO THAT THE CUSTOMER COULD FURTHER REVIEW THEIR WORKFLOW OF SCANNING PATIENTS. (B)(4). Manufacturer Narrative: THE INITIAL REPORT THAT WAS SUBMITTED FOR 2183926-2016-00797 WAS ERRONEOUSLY SUBMITTED AS A 5 DAY REPORT. THIS WAS AN ERROR AS MERGE HEALTHCARE WAS NOT REQUIRED TO INITIATE ANY ACTIONS TO PREVENT AN UNREASONABLE RISK OF SUBSTANTIAL HARM. THE REPORT SHOULD HAVE BEEN SUBMITTED AS A 30 DAY REPORT. Manufacturer Narrative: A REVIEW OF THE COMPLAINT HISTORY FOUND NO OCCURRENCES RELATED TO THE REPORTED EVENT THAT COULD CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. MERGE WILL CONTINUE TO MONITOR COMPLAINTS FOR THIS ISSUE. SHOULD NEW INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE PROVIDED. BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER, THE CAUSE OF THIS EVENT IS MOST LIKELY NOT RELATED TO A MERGE PRODUCT AND APPEARS TO BE THE RESULT OF A PATIENT NOT ADHERING TO A CLINICIAN'S INSTRUCTIONS. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT AN EXAM WAS MARKED READ BY THE DOCTOR, HOWEVER NO REPORT WAS FOUND. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT A REPORT ID WAS CREATED BUT NO WORD DOCUMENTS COULD BE FOUND IN THE AUDIT TRAIL PROPERTIES FOR THE AUDIT TRAIL SAVE EVENT. THE REPORT COULD NOT BE RECOVERED AND THE EXAM NEEDED TO BE RE-READ. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING MISSING REPORTS OR REPORTS NOT SAVING. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, WHEN THE PRIMARY MONTAGE WAS TOGGLED OFF, THE USER EXPECTED THE MONTAGE TO TOGGLE OFF AND THE PRIMARY IMAGES WOULD BE RETAINED IN THE CURRENT LOCATION. HOWEVER, THE PRIMARY IMAGES SHIFTED ENTIRELY FROM THE SECOND MONITOR TO THE FIRST MONITOR. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE PRIMARY EXAM SHIFTING TO ANOTHER MONITOR WITHOUT ANY INDICATION OF DOING SO HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT IF THE USER IS UNAWARE THAT THE IMAGES HAVE MOVED. (B)(4). Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE FILED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A 3D MAMMOGRAM EXAM WAS STUCK ON INCORPORATING AND WAS NOT ABLE TO BE MOVED ONLINE TO BE READ. AN INVESTIGATION DONE BY MERGE DETERMINED THAT THE PATIENT'S MAMMOGRAM EXAM HAD 10 VIEWS, WHICH WAS TOO LARGE TO CROSS OVER TO THE PACS SYSTEM. THE EXAM WAS DELETED AND RESENT, THE RECOMMENDED COMPRESSION WAS APPLIED TO THE IMAGES ON THE MODALITY, THE IMAGES WERE SENT ONE AT A TIME, AND IN ADDITION, THE THIRD PARTY MODALITY VENDOR ASSISTED IN ATTEMPTING TO RESEND THE IMAGES ONE AT A TIME WITH NO SUCCESS. MERGE DETERMINED THAT THE IMAGES WERE TOO LARGE, EVEN AFTER THE COMPRESSION AND SENDING THEM ONE AT A TIME. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE RADIOLOGIST DECIDED TO RE-SCHEDULE THE PATIENT FOR A DIFFERENT EXAM, ONE VIEW, 2D MEDIOLATERAL OBLIQUE (MLO) VIEW MAMMOGRAM WITHOUT TOMOGRAPHY. THEREFORE, IMAGES THAT ARE TOO LARGE TO BE SENT FROM THE MODALITY TO THE PACS SYSTEM HAS THE POTENTIAL TO DELAY TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. A REVIEW OF THE MERGE UNITY USER GUIDES PROVIDED INFORMATION ON HOW TO MOVE EXAMS ONLINE AND SIZING MAMMOGRAMS. HOWEVER, DID NOT REVEAL ANY INFORMATION REGARDING TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES ABOUT MOVING EXAMS ONLINE. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT A REPORT WAS CREATED AND ATTEMPTED TO BE SAVED ON (B)(6) 2016 AT 23:31, HOWEVER, THE REPORT DID NOT UPLOAD SUCCESSFULLY TO THE SERVER. IT WAS NOTED THAT THIS REPORT WAS FOR A STAT CT CHEST/ABDOMEN/PELVIS. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO RECOVER THE REPORT OR FIND THE WORD DOCUMENT FILE FOR THE REPORT THAT WAS READ ON (B)(6) 2016. FURTHER INVESTIGATION DETERMINED THAT THE EXAM WAS RE-READ AND THE REPORT WAS REVISED ON (B)(6) 2016. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING MISSING REPORTS OR REPORTS NOT SAVING. MERGE HEALTHCARE IS EVALUATING THE COMPLAINT TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING. AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT A REPORT WAS CREATED & SAVED BY THE RADIOLOGIST. HOWEVER THE REPORT WAS MISSING. THE SYSTEM CRASHED BEFORE THE REPORT WAS SAVED TO THE SERVER. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO RECOVER THE REPORT OR FIND THE WORD DOCUMENT FILE FOR THE REPORT. THEREFORE, THE REPORT WAS LOST AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: AN INVESTIGATION DONE BY MERGE HEALTHCARE DETERMINED THAT IF A REPORT WAS SAVED TOO QUICKLY AFTER OPENING THE EXAM, THE REPORT WOULD NOT SAVE SUCCESSFULLY AS EXPECTED. THE USER MUST SAVE THE REPORT AFTER THE REPORT TEMPLATE HAS COMPLETED BUILDING IN ORDER TO ENSURE THAT THE REPORT SAVES SUCCESSFULLY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER ALLEGED THAT A REPORT FOR A STAT EXAM WAS MISSING. AFTER FURTHER INVESTIGATION BY MERGE HEALTHCARE SUPPORT, A REPORT TAB AND A REPORT ID WAS GENERATED WITHIN THE EXAM PROPERTIES. THIS INDICATED THAT A REPORT EXISTED. AFTER REVIEWING THE EVENT LOGS, IT WAS DETERMINED THAT THE REPORT WAS CORRUPTED. HOWEVER, THE MOST PROBABLE CAUSE FOR THE CORRUPTION COULD NOT BE DETERMINED. MERGE SUPPORT WAS UNABLE TO RECOVER THE REPORT. THEREFORE, THE REPORT NEEDED TO BE RE-DICTATED BY THE CUSTOMER. WHEN A CUSTOMER NEEDS TO RE-DICTATE A STAT EXAM IN ORDER TO OBTAIN A REPORT, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT. HOWEVER, THERE WAS NO REPORTED ADVERSE EVENT OR HARM TO A PATIENT. (B)(4). Manufacturer Narrative: WHILE MERGE HEALTHCARE INVESTIGATED THE CUSTOMER'S ALLEGATION, THE ASSIGNABLE CAUSE OF THE CORRUPTED REPORT COULD NOT BE DETERMINED. THE ISSUE WILL CONTINUE TO BE MONITORED TO DETERMINE IF ADDITIONAL DATA BECOMES AVAILABLE THAT WILL FURTHER THE INVESTIGATION. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS CONTACTED BY A CUSTOMER INDICATING THAT THEY WERE UNABLE TO FIND A REPORT FOR AN EXAM THAT WAS SAVED UNDER PENDING APPROVAL STATUS AFTER DICTATION. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO FIND THE REPORT OR THE WORD DOCUMENT FILE FOR THE REPORT. DUE TO NOT BEING ABLE TO FIND THE REPORT, RE-DICTATION WAS NECESSARY. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, NOT BEING ABLE TO FIND A REPORT THAT THEN NEEDS TO BE RE-DICTATED, THERE IS A POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, THE CAUSE OF THE MISSING REPORT COULD NOT BE DETERMINED AS THE LOG FILES WERE TRUNCATED, HOWEVER THE MOST PROBABLE CAUSE WAS AN ISSUE WITH UPLOADING THE REPORT TO THE SERVER. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 A CUSTOMER REPORTED THAT AN EXAM READ BY A RADIOLOGIST IN THE MORNING WAS MISSING THE REPORT HOWEVER, THE EXAM SHOWS THAT IT IS IN AN "APPROVED STATUS." DUE TO A USER NOT BEING ABLE TO LOCATE A REPORT FOR A PATIENT, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. WHILE THE EXAM/IMAGES WERE STILL AVAILABLE, THE PHYSICIAN HAD TO RE-READ THE EXAM TO CREATE A REPORT. THERE WAS NO INDICATION OF ANY HARM OR ADVERSE EVENT TO A PATIENT DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: SITE HAS A QC POLICY IN PLACE CHECKING EXAM REPORTS EACH DAY TO MAKE SURE ALL FAXES WENT OUT, REPORTS PRINTED WHEN NECESSARY. THIS INCIDENT WAS A RESULT OF THE UNIVERSAL MANAGER CRASHING AT THE POINT THE RADIOLOGIST WAS CLOSING THE CASE. SITE SOFTWARE WAS UPDATED ON 2/19/2017 FROM R11.0.1 TO R11.0.4. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 AT APPROXIMATELY 1:44AM PST, AFTER HOURS SUPPORT WAS CONTACTED BY MERGE FIELD SERVICE ENGINEER. THE ENGINEER WAS ONSITE PERFORMING A PLANNED AFTER HOURS SERVICE AND UPDATING THE FIRMWARE ON THE IMAGE SERVER. AFTER FIRMWARE UPDATE WAS COMPLETED AND THE IMAGE SERVER WAS RESTARTED, THE NETWORK DRIVE WAS NO LONGER AVAILABLE. ON CALL SPECIALIST WORKED THROUGH THE NIGHT WITH THE FIELD SERVICE ENGINEER AND THEY WERE ABLE TO HAVE THE IMAGE SERVER BACK ONLINE AND AVAILABLE THE NEXT MORNING PRIOR TO THE SITE OPENING. WHILE THE IMAGE SERVER WAS ONLINE AND THEY WERE ABLE TO ACQUIRE EXAMS, THERE WERE SEVERAL MOUNT POINTS (STORAGE LOCATIONS) THAT WERE UNAVAILABLE TO ACCESS. THESE MOUNT POINT WHICH CONTAINED EXAMS ACQUIRED PREVIOUSLY. WITH A COMBINATION OF EXAM RESTORING FROM LONG TERM ARCHIVE AND REPAIR, ALL EXAMS THAT ORIGINATED AT THIS SITE WERE RECOVERED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED DURING FOLLOW UP CALL WITH CUSTOMER 7/8/2016: PER THE PACS ADMIN AT THE CUSTOMER'S SITE, NO FURTHER ISSUES HAVE BEEN IDENTIFIED WITH THE IMAGE SERVER SINCE THE EVENT (B)(6) 2016 AND THE SUBSEQUENT REPAIR. HE ALSO REPORTED THAT NO INSTANCES OF MISSING EXAMS FROM THAT TIME FRAME HAVE BEEN IDENTIFIED. THE ASSIGNABLE CAUSE OF FAILURE WAS UNABLE TO BE CONFIRMED. THE CORRECTION TO THE STORAGE POINTS ON THE IMAGE SERVER HARD DRIVE REMOVED THE ABILITY TO IDENTIFY CAUSE. THIS ISSUE WILL CONTINUE TO BE MONITORED. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE IS NO SIGNIFICANT TREND. TRENDING WILL CONTINUE TO BE MONITORED. A REVIEW OF THE LABELING UNDER UPAX-2588, R11 IMPLEMENTATION GUIDE FOUND: "WHETHER YOU USE AN AUTOMATED OR MANUAL ARCHIVE, VERIFY THE FOLLOWING ITEMS DAILY." REFER TO PAGES 74-82 FOR SPECIFIC INFORMATION RELATED TO ARCHIVE. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED BY A PACS ADMINISTRATOR THAT SERVER MAINTENANCE HAD BEEN PERFORMED AND THE SITE WAS NOT ABLE TO RESTART THE DATABASE. MERGE HEALTHCARE WORKED WITH THE CUSTOMER TO FURTHER INVESTIGATE THE ISSUE. IT HAS BEEN DETERMINED THAT PRIOR TO THE SCHEDULED SERVER MAINTENANCE, THE DATABASE WAS UNEXPECTEDLY POWERED DOWN WHILE TASKS WERE STILL PENDING. THIS UNEXPECTED SHUTDOWN RESULTED SOME LOST DATA. MERGE HEALTHCARE ASSISTED THE SITE AND UTILIZING BACKUPS AND OTHER RESTORATION PROCEDURES THE DATABASE WAS ABLE TO BE RESTORED AND WAS FUNCTIONAL PRIOR TO THE NEXT BUSINESS DAY. THE PRIMARY IMPACT TO THE CLINICAL WORKFLOW WAS A RADIOLOGIST NEEDING TO RE-READ ABOUT 70 EXAMS. THERE ARE 4 REMAINING REPORTS THAT MERGE IS ASSISTING THE SITE WITH RECONSTRUCTING. THE INABILITY TO VIEW REPORTS AND IMAGES MAY RESULT IN A POTENTIAL TO DELAY IN TREATMENT AND/OR DIAGNOSIS A PATIENT. THE CUSTOMER DID NOT INDICATE ANY HARM TO PATIENTS WHOSE DATA MAY HAVE BEEN IMPACTED BY THE UNEXPECTED SHUTDOWN OF THE DATABASE. (B)(4). Manufacturer Narrative: BASED ON THE INVESTIGATION INTO THE DATABASE AND SCHEDULED SYSTEM MAINTENANCE AT THE SITE, THE DATABASE WAS SHUT DOWN A FEW HOURS PRIOR TO THE SCHEDULED SHUT DOWN. THIS WAS NOT SOMETHING THAT MERGE HEALTHCARE HAD ANY CONTROL OVER. THE DEVICE WAS OPERATING AS EXPECTED AND AS INTENDED UNTIL THE POWER WAS UNEXPECTEDLY SHUT DOWN. THIS IS LIKELY THE RESULT OF A USER ERROR. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT A REPORT WAS MISSING. THE CUSTOMER STATED THAT THE PHYSICIAN READ THE REPORT AND SAVED IT ON (B)(6). HOWEVER, WHILE DOING AN AUDIT, IT WAS NOTED THAT THIS REPORT WAS MISSING. IT WAS CONFIRMED THAT THE REPORT DID NOT UPLOAD AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A MISSING REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE REPORT IS MISSING. THIS INVESTIGATION IS IN PROGRESS AND NOT YET COMPLETE. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED BY A PHYSICIAN THAT AN IMAGE WAS MISSING FROM AN EXAM. THE TECHNOLOGIST WAS ABLE TO DISPLAY THE EXAM FROM THE ONLINE LIST AND ALL IMAGES WERE AVAILABLE. MERGE UNITY PACS IS A SERVER BASED SYSTEM THEREFORE DATA ACCESSED FROM ANY TERMINAL OR WORKSTATION IS EXPECTED TO BE EXACTLY THE SAME, AS IT IS COMING FROM ONE SOURCE. HOWEVER, THE INABILITY TO VIEW ALL IMAGES IN AN EXAM HAS THE POTENTIAL TO DELAY OR INADEQUATELY TREAT AND/OR DIAGNOSIS A PATIENT. AN INVESTIGATION INTO THIS ALLEGATION IS IN PROGRESS. (B)(4). | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE WHEN THE PHYSICIAN OPENED THE COMPARISON STUDY, THE STUDY INITIALLY WAS DISPLAYED ON THE CORRECT MONITOR. HOWEVER, WITHIN A SECOND, THE STUDY INDEPENDENTLY MOVED TO A DIFFERENT MONITOR. TO GET THE MONITOR TO DISPLAY THE CORRECT STUDIES, THE PHYSICIAN CREATED A WORK AROUND. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, STUDIES BEING MOVED TO A DIFFERENT MONITOR WITHOUT THE PHYSICIAN BEING AWARE HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT WHEN A USER DID NOT LOG OUT OF A STATION AND ANOTHER USER LOGGED IN, THE ORIGINAL USER'S NAME WAS STILL SHOWN AS LOGGED IN AT THE LOWER RIGHT HAND CORNER OF THE UNIVERSAL MANAGEMENT. THE CUSTOMER DID NOT KNOW WHAT THEIR AUTO LOG OUT FEATURE WAS SET TO BUT REPORTED THAT WHEN A REPORT IS READ, THE REPORT INDICATED THAT THE ORIGINAL USER READ THE REPORT INSTEAD OF THE INTENDED CURRENT USER. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THIS IS A POTENTIAL SAFETY ISSUE SINCE UNAUTHORIZED USERS CAN ACCESS SENSITIVE PATIENT INFORMATION AND MAKE UNNECESSARY CHANGES OR UPDATES. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, WHEN THE RADIOLOGIST TOGGLED THE MONTAGE FUNCTION ON THE FAR RIGHT OF THE MONITOR, THE POSITIONS OF THE PRIMARY AND PRIOR EXAMS WERE SWITCHED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, SWITCHING VIEWS FROM THE PRIMARY EXAM AND THE PRIOR EXAM WITHOUT ANY INDICATION OF DOING SO HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. Manufacturer Narrative: (B)(4). CONCLUSIONS CODE: - CONCLUSION NOT YET AVAILABLE-EVALUATION IN PROGRESS. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, AFTER RESENDING A REPORT WITH THE SAME ORDER NUMBER, THE REPORT UPDATED EVERYWHERE EXCEPT FOR THE ZERO-DOWNLOAD AMBASSADOR (ZDA). THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A REPORT NOT UPDATING WITH THE CORRECT INFORMATION HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT.Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, AFTER RESENDING A REPORT WITH THE SAME ORDER NUMBER, THE REPORT UPDATED EVERYWHERE EXCEPT FOR THE ZERO-DOWNLOAD AMBASSADOR (ZDA). THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A REPORT NOT UPDATING WITH THE CORRECT INFORMATION HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. (B)(4). Manufacturer Narrative: INVESTIGATION FOUND THAT THE CUSTOMER HAD THEIR REPORT MANAGER CONFIGURED NOT TO OVERWRITE EXISTING REPORTS IF THE MODALITY TYPE WAS DIFFERENT. SINCE THE CUSTOMER WAS CHANGING MODALITY TYPES WITH UPDATED ORDERS, THE SUPPORT ANALYST RECONFIGURED THE CUSTOMER'S REPORT MANAGER TO OVERWRITE ANY REPORTS REGARDLESS OF THE MODALITY TYPE IF THE ORDER NUMBER MATCHED. THE CUSTOMER CONFIRMED THAT THIS WAS THE DESIRED WORKFLOW AND TESTING CONFIRMED THAT THE REPORTS UPDATED AS EXPECTED AFTER THE RECONFIGURATION. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS USER INFORMED MERGE UNITY SUPPORT THAT A RADIOLOGIST WAS READING AN EXAM AND WENT TO PLAY THE REPORT. THE MERGE UNITY SYSTEM REPORTED THERE WAS NO EXISTING VOICE CLIP. MERGE UNITY TECHNICAL SUPPORT INVESTIGATED AND DETERMINED THAT THERE WAS A "REPORT SAVE" EVENT IN THE AUDIT TRAIL, BUT NO REPORT EXISTED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ AND SAVED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, A MISSING REPORT HAS THE POTENTIAL TO LEAD TO A DELAY PATIENT CARE THAT MAY LEAD TO HARM. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. WHEN PROMPTED TO LISTEN TO THE VOICE REPORT, THE PHYSICIAN RECEIVED A MESSAGE THAT NO VOICE CLIP EXISTED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(4) 2017, MERGE HEALTHCARE WAS NOTIFIED FROM THE CUSTOMER THAT AN INCOMPLETE REPORT WAS SENT OUT TO THE REFERRING PHYSICIAN WITHOUT THE MENTION OF A LUNG NODULE ON (B)(6) 2017. A RADIOLOGIST FROM THE CUSTOMER SITE READ A PATIENT'S CASE ON (B)(6) 2017 AND FOUND AN ABNORMALITY IN THE LUNG. THE RADIOLOGIST CALLED THE REFERRING DOCTOR TO INFORM THEM OF THE ABNORMALITY AND RECOMMENDED THAT THE PATIENT HAVE A FOLLOW UP CT. HOWEVER, THE REPORT THAT WAS SENT OUT TO REFERRING DOCTOR HAD NO MENTION OF THE LUNG NODULE IN THE REPORT BODY. ADDITIONALLY, THERE WAS NO INFORMATION REGARDING THE NODULE IN THE CONCLUSION SECTION OF THE REPORT. THE REFERRING DOCTOR'S OFFICE CALLED BACK FOR CLARIFICATION BECAUSE THEY WERE CONFUSED BY THE REPORT WHICH HAD NO MENTION OF THE LUNG NODULE. THE RADIOLOGIST AT THE SITE THEN SPOKE TO THE REFERRING DOCTOR AND ISSUED AN ADDENDUM REPORT WITH COMPLETE INFORMATION. THERE WERE NO KNOWN ADVERSE EVENTS TO THE PATIENT. HOWEVER, SENDING AN INCOMPLETE REPORT TO A REFERRING PHYSICIAN WITH MISSING INFORMATION COULD POTENTIALLY LEAD TO A DELAY IN TREATMENT AND/OR DIAGNOSIS THAT MAY LEAD TO HARM. (B)(4). Manufacturer Narrative: THE REPORT WAS ACCIDENTALLY APPROVED WITH MISSING INFORMATION AND WAS SENT WITHOUT CONCLUSION. WHEN APPROVING REPORTS, MERGE UNITY PACS IS INTERACTING WITH THIRD-PARTY SOFTWARE (IN THIS CASE, MICROSOFT WORD), AND THAT THE CONTROL OVER THE DOCUMENT ITSELF IS LIMITED. MERGE UNITY PACS HAS PRECAUTIONS IN PLACE WHEN THE RADIOLOGIST FOLLOWS A NORMAL RECOMMENDED WORKFLOW. THE WORKFLOW EXPECTATION IS FOR THE CAUTION STATEMENT IN THE REPORT TEMPLATE TO BE REPLACED WITH NO ROOM IN BETWEEN (WHEN USING SPEECH RECOGNITION, THE WHOLE CAUTION STATEMENT IS AUTOMATICALLY HIGHLIGHTED, AND THE INTENTION IS FOR THE RADIOLOGIST TO OVERWRITE IT WITH THE ACTUAL CONCLUSION TEXT). A POP-UP WARNING OCCURS IF THE CAUTION STATEMENT IN THE REPORT IS STILL PRESENT. THERE IS STILL ROOM FOR USER ERROR IF THE CAUTION STATEMENT IS REMOVED BUT NOT REPLACED BY A CONCLUSION, IT IS POSSIBLE TO APPROVE THE REPORT WITH TEXT MISSING. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE IS NO SIGNIFICANT TREND. TRENDING WILL CONTINUE TO BE MONITORED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT THE TRANSCRIPTIONIST WAS UNABLE TO OPEN THE DICTATION TO TYPE UP THE REPORT. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE UNITY EXAM AUDIT HAS INFORMATION RELATED TO THE REPORT'S SAVE EVENT. AS A RESULT OF NOT BEING ABLE TO FIND THE DICTATION, THE READING PHYSICIAN RE-READ THE EXAM TO RESOLVE THE ISSUE. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS WORK FLOW, THE DICTATION WAS CREATED BY THE PHYSICIAN. THE TRANSCRIPTIONIST ATTEMPTED TO OPEN THE DICTATION TO TYPE THE REPORT, BUT NO DICTATION WAS AVAILABLE. THE PHYSICIAN CONFIRMED THAT HE WAS ALSO UNABLE TO OPEN HIS DICTATION. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A UNITY PACS ADMINISTRATOR CONTACTED MERGE TECHNICAL SUPPORT ALLEGING THAT DURING THE PROCESS OF ADDING IMAGES TO AN EXAM, THE STATUS OF THAT EXAM DID NOT FUNCTION AS EXPECTED. THE EXAM WAS REVIEWED BY A RADIOLOGIST, AND ADDITIONAL IMAGES WERE ORDERED. THE TECHNOLOGIST REVERTED THE EXAM FROM THE VIEWING/REVIEWING "ONLINE" LOCATION BACK TO THE DICOM GATEWAY (ACQUISITION STATION) WHERE ADDITIONAL IMAGES WERE ACQUIRED AND SUBSEQUENTLY ADDED TO THE EXAM. WITH THE ADDITIONAL IMAGES ADDED, THE EXAM WAS THEN MOVED BACK ONLINE FOR VIEWING/REVIEWING. HOWEVER, THE EXAM WAS IN A READ STATUS (R) INSTEAD OF BEING IN THE EXPECTED SCANNED STATUS (S). A SCANNED STATUS INDICATES TO USERS THAT AN EXAM EITHER HAS NOT BEEN READ/REVIEWED OR THAT ADDITIONAL INFORMATION HAS BEEN ADDED TO THE EXAM AND REQUIRES RE-READING. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. AN EXAM WITH ADDITIONAL IMAGES ADDED, AND MOVED TO THE ONLINE LIST AS A "READ" STUDY HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THERE ARE ADDITIONAL IMAGES TO REVIEW. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED FOR EXAMS THAT HAVE BEEN REVERTED TO ADD ADDITIONAL IMAGES TO BE MOVED BACK TO THE ONLINE LIST IN A SCANNED (S) STATUS. EXAMS IN SCANNED STATUS SIGNAL TO THE READING PHYSICIAN THAT THE EXAM IS READY TO BE INTERPRETED AND REPORTS CREATED. IN THIS INSTANCE THE LOG FILES WERE NOT AVAILABLE FOR REVIEW. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT THE TRANSCRIPTIONIST WAS UNABLE TO OPEN THE DICTATION TO TYPE UP THE REPORT. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE UNITY EXAM AUDIT HAS INFORMATION RELATED TO THE REPORT'S SAVE EVENT. AS A RESULT OF NOT BEING ABLE TO FIND THE DICTATION, THE READING PHYSICIAN RE-READ THE EXAM TO RESOLVE THE ISSUE. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS WORK FLOW, THE DICTATION WAS CREATED BY THE PHYSICIAN. THE TRANSCRIPTIONIST ATTEMPTED TO OPEN THE DICTATION TO TYPE THE REPORT, NO DICTATION WAS AVAILABLE. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A UNITY PACS ADMINISTRATOR CONTACTED MERGE TECHNICAL SUPPORT ALLEGING THAT DURING THE PROCESS OF ADDING IMAGES TO AN EXAM, THE STATUS OF THAT EXAM DID NOT FUNCTION AS EXPECTED. THE EXAM WAS REVIEWED BY A RADIOLOGIST, AND ADDITIONAL IMAGES WERE ORDERED. THE TECHNOLOGIST REVERTED THE EXAM FROM THE VIEWING/REVIEWING "ONLINE" LOCATION BACK TO THE DICOM GATEWAY (ACQUISITION STATION) WHERE ADDITIONAL IMAGES WERE ACQUIRED AND SUBSEQUENTLY ADDED TO THE EXAM. WITH THE ADDITIONAL IMAGES ADDED, THE EXAM WAS THEN MOVED BACK ONLINE FOR VIEWING/REVIEWING. HOWEVER, THE EXAM WAS IN A READ STATUS (R) INSTEAD OF BEING IN THE EXPECTED SCANNED STATUS (S). A SCANNED STATUS INDICATES TO USERS THAT AN EXAM EITHER HAS NOT BEEN READ/REVIEWED OR THAT ADDITIONAL INFORMATION HAS BEEN ADDED TO THE EXAM AND REQUIRES RE-READING. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. AN EXAM WITH ADDITIONAL IMAGES ADDED, AND MOVED TO THE ONLINE LIST AS A "READ" STUDY HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THERE ARE ADDITIONAL IMAGES TO REVIEW. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED FOR EXAMS THAT HAVE BEEN REVERTED TO ADD ADDITIONAL IMAGES TO BE MOVED BACK TO THE ONLINE LIST IN A SCANNED (S) STATUS. EXAMS IN SCANNED STATUS SIGNAL TO THE READING PHYSICIAN THAT THE EXAM IS READY TO BE INTERPRETED AND REPORTS CREATED. IN THIS INSTANCE THE LOG FILES WERE NOT AVAILABLE FOR REVIEW. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT WHEN ATTEMPTING TO OPEN THE EXAM A MESSAGE IS DISPLAYED ON THE SCREEN THAT SHOWS "NO REPORT". MERGE SUPPORT INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT THERE WAS A SAVE EVENT FOR THE CUSTOMERS' ALLEGED REPORT. HOWEVER, MERGE UNITY TECHNICAL SUPPORT WAS UNABLE TO FIND THE SAVED REPORT FOR THE CUSTOMER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THIS CASE/SITE WAS LINKED TO (B)(4). MISSING REPORTS ISSUE: REPORTS ARE FAILING TO UPLOAD AFTER A POP-UP INDICATES AN ISSUE COMMUNICATING WITH THE SERVER. RETRYING ALSO FAILS. THE JIRA WAS CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION (B)(4). THE SITE WAS UPGRADED TO (B)(4) ON 08/22/2017. THE EXAM WAS AVAILABLE FOR REVIEW AT ALL TIMES, AND THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE LOG FILES FOR THE WORK STATION WERE TRUNCATED DUE TO THE AMOUNT OF TIME FROM THE EXAM READ TO THE DISCOVERY OF THE REPORT MISSING. AS A RESULT OF THE TRUNCATION, IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT WHEN ATTEMPTING TO OPEN THE EXAM A MESSAGE IS DISPLAYED ON THE SCREEN THAT SHOWS "NO REPORT". MERGE SUPPORT INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT THERE WAS A SAVE EVENT FOR THE CUSTOMERS' ALLEGED REPORT. HOWEVER, MERGE UNITY TECHNICAL SUPPORT WAS UNABLE TO FIND THE SAVED REPORT FOR THE CUSTOMER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THIS CASE/SITE WAS LINKED TO JIRA (B)(4). MISSING REPORTS ISSUE: REPORTS ARE FAILING TO UPLOAD AFTER A POP-UP INDICATES AN ISSUE COMMUNICATING WITH THE SERVER. RETRYING ALSO FAILS. THE JIRA WAS CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION R11.1.2. THE SITE WAS UPGRADED TO R11.1.2.5 ON 08/22/2017. THE EXAM WAS AVAILABLE FOR REVIEW AT ALL TIMES, AND THE REPORT WAS RE-READ BY THE PHYSICIAN. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE LOG FILES FOR THE WORK STATION WERE TRUNCATED DUE TO THE AMOUNT OF TIME FROM THE EXAM READ TO THE DISCOVERY OF THE REPORT MISSING. AS A RESULT OF THE TRUNCATION, IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A MERGE UNITY CUSTOMER (PACS ADMINISTRATOR) ALLEGED THAT A PATIENT'S EXAM, WHEN VIEWED ON THE ZERO DOWNLOAD AMBASSADOR (ZDA), SHOWS A DIFFERENT EXAM AND PATIENT'S REPORT. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THIS CUSTOMER'S ALLEGATION AND DETERMINED THAT THIS ISSUE WAS SEEN WITH TWO DIFFERENT EXAMS FROM THE SAME PATIENT. ON UNITY PACS SYSTEM, THE USER WILL GET A "NO REPORTS FOUND" MESSAGE, BUT ON THE ZDA A REPORT FROM A DIFFERENT PATIENT IS DISPLAYED. WITH THE ZDA NOT FUNCTIONING AS EXPECTED THERE IS A POTENTIAL THAT A CLINICIAN MAY NOT RECOGNIZE THAT THE EXAM AND REPORT THEY REQUESTED TO BE DISPLAYED VIA THE ZDA IS FOR A DIFFERENT PATIENT. THIS MAY LEAD TO A DELAY AND/OR INCORRECT DIAGNOSIS OR TREATMENT FOR A PATIENT WHICH COULD LEAD TO HARM. HOWEVER, THE CUSTOMER HAS NOT INDICATED THAT ANY HARM HAS RESULTED DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS HAS FUNCTIONALITY THAT WHEN ENABLED AND CONFIGURED, ALLOWS REFERRING PHYSICIANS TO VIEW EXAMS FROM OUTSIDE THE NETWORK VIA THE ZERO DOWNLOAD AMBASSADOR (ZDA). MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMERS ALLEGATION. IT WAS DETERMINED THAT A SINGLE PATIENT HAD 2 EXAMS ENTERED INTO THE SYSTEM UNDER THE SAME PLACER ORDER NUMBER. THEREFORE, WHEN THE REFERRING PHYSICIAN REQUESTED A PATIENT'S EXAM BY ENTERING THE PLACER ORDER NUMBER, AN INCORRECT REPORT WAS SHOWN. FURTHER INVESTIGATION FOUND INDICATIONS THAT THE EXAMS WERE IMPORTED FROM OUTSIDE OF THE MERGE UNITY PACS SYSTEM AS THERE WAS NO EVIDENCE OF DICOM DATA ORIGINATING IN THE CUSTOMER'S UNITY PACS SYSTEM. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT REPORTS HAVE NOT BEEN GENERATED. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORTS WERE BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. (B)(4). Manufacturer Narrative: THE INVESTIGATION INTO THIS CUSTOMER'S ALLEGATION DETERMINED THAT THE DELAY IN THE REPORT BEING FULLY APPROVED WAS A DEFECT INTRODUCED IN MERGE UNITY PACS RELEASE R11.1.2. THIS RELEASE WENT OUT TO A LIMITED NUMBER OF CUSTOMERS PRIOR TO BEING DETECTED. MERGE HEALTHCARE HAS SINCE DETERMINED THE ROOT CAUSE OF THE ISSUE AND HAS PERFORMED CORRECTIONS AND CORRECTIVE ACTIONS. A NEW RELEASE OF SOFTWARE HAS PROGRESSED THROUGH REVIEW AND VERIFICATION PHASES AND WILL BE AVAILABLE TO CUSTOMERS UPON RELEASE. A REVIEW OF THE MERGE UNITY PACS USER GUIDES, READING PHYSICIAN CORE CURRICULUM: ALL LEVELS DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING EXAMS THAT REMAIN STUCK IN A PENDING/HOLD STATUS. THE WORKAROUND PROVIDED TO THE CUSTOMER IS DEMONSTRATING EFFECTIVENESS FOR REPORTING RELATED ACTIVITIES. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT REPORTS HAVE NOT BEEN GENERATED. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORTS WERE BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. REFERENCE COMPLAINT NUMBER 83703. Manufacturer Narrative: INITIAL ONSET OF THIS WAS RESOLVED AT DEVELOPMENT'S RECOMMENDATION, PATCHED AUTOMATION BACK TO 11.1, WHICH RESULTED IN THE REPORTS GETTING COMPLETED SUCCESSFULLY. NEW REPORTS WERE ALSO BEING PROCESSED AS EXPECTED. THE JIRA OPENED FOR THIS ISSUE HAS BEEN CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION (B)(6) URIS-3378. AUTOMATION WILL NOT PROCESS REPORTS AFTER UPGRADE TO (B)(6). THE SITE WAS UPDATED TO SOFTWARE VERSION (B)(6). THE CAPA OPENED FOR THIS ISSUE HAS BEEN COMPLETED/CLOSED. QUALITY SYSTEM, QS-128273 [UNITY] AUTOMATION WILL NOT PROCESS REPORTS. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: THE INVESTIGATION INTO THIS CUSTOMER'S ALLEGATION DETERMINED THAT THE DELAY IN THE REPORT BEING FULLY APPROVED WAS A DEFECT INTRODUCED IN MERGE UNITY PACS RELEASE R11.1.2. THIS RELEASE WENT OUT TO A LIMITED NUMBER OF CUSTOMERS PRIOR TO BEING DETECTED. MERGE HEALTHCARE HAS SINCE DETERMINED THE ROOT CAUSE OF THE ISSUE AND HAS PERFORMED CORRECTIONS AND CORRECTIVE ACTIONS. A NEW RELEASE OF SOFTWARE HAS PROGRESSED THROUGH REVIEW AND VERIFICATION PHASES AND WILL BE AVAILABLE TO CUSTOMERS UPON RELEASE. A REVIEW OF THE MERGE UNITY PACS USER GUIDES, READING PHYSICIAN CORE CURRICULUM: ALL LEVELS DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING EXAMS THAT REMAIN STUCK IN A PENDING/HOLD STATUS. THE WORKAROUND PROVIDED TO THE CUSTOMER IS DEMONSTRATING EFFECTIVENESS FOR REPORTING RELATED ACTIVITIES. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED ISSUES WITH REPORTS NOT APPROVING. THE REPORT STATUS IS STUCK AFTER SIGNING OFF ON A REPORT. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORT WAS BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: INITIAL ONSET OF THIS WAS RESOLVED AT DEVELOPMENT'S RECOMMENDATION, PATCHED AUTOMATION BACK TO (B)(4), WHICH RESULTED IN THE REPORTS GETTING COMPLETED SUCCESSFULLY. NEW REPORTS WERE ALSO BEING PROCESSED AS EXPECTED. THE JIRA OPENED FOR THIS ISSUE HAS BEEN CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION (B)(4). URIS-3378 AUTOMATION WILL NOT PROCESS REPORTS AFTER UPGRADE TO (B)(4). THE SITE WAS UPDATED TO SOFTWARE VERSION (B)(4) ON (B)(6) 2017. THE CAPA OPENED FOR THIS ISSUE HAS BEEN COMPLETED/CLOSED. QUALITY SYSTEM. QS-128273 [UNITY] AUTOMATION WILL NOT PROCESS REPORTS. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: THE INVESTIGATION INTO THIS CUSTOMER'S ALLEGATION DETERMINED THAT THE DELAY IN THE REPORT BEING FULLY APPROVED WAS A DEFECT INTRODUCED IN MERGE UNITY PACS RELEASE R11.1.2. THIS RELEASE WENT OUT TO A LIMITED NUMBER OF CUSTOMERS PRIOR TO BEING DETECTED. MERGE HEALTHCARE HAS SINCE DETERMINED THE ROOT CAUSE OF THE ISSUE AND HAS PERFORMED CORRECTIONS AND CORRECTIVE ACTIONS. A NEW RELEASE OF SOFTWARE HAS PROGRESSED THROUGH REVIEW AND VERIFICATION PHASES AND WILL BE AVAILABLE TO CUSTOMERS UPON RELEASE. A REVIEW OF THE MERGE UNITY PACS USER GUIDES, READING PHYSICIAN CORE CURRICULUM: ALL LEVELS) DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING EXAMS THAT REMAIN STUCK IN A PENDING/HOLD STATUS. THE WORKAROUND PROVIDED TO THE CUSTOMER IS DEMONSTRATING EFFECTIVENESS FOR REPORTING RELATED ACTIVITIES | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT A REPORT THAT HAD BEEN DICTATED WAS UNABLE TO BE FOUND. MERGE SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND DETERMINED THAT THE REPORT WAS BEING HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. NORMALLY A REPORT TRANSITIONS FROM THE PAUSED OR HOLD STATUS TO A FULLY APPROVED STATUS THAT THEN TRIGGERS POST-APPROVAL REPORTING ACTIVITIES SUCH AS VIEWING OR FAXING. DUE TO MERGE UNITY REPORTS NOT PROGRESSING THROUGH THE APPROVALS PROCESS IN A TIMELY MANNER THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. HOWEVER, THE CUSTOMER HAS REPORTED NO SERIOUS INJURIES OR DEATHS AS A RESULT OF THIS ISSUE. THE SITE HAS BEEN PROVIDED WITH A WORKAROUND THAT ENSURES EXAMS ARE NOT HELD IN A PAUSED OR PENDING STATUS FOR AN EXTENDED PERIOD OF TIME. (B)(4). Manufacturer Narrative: INITIAL ONSET OF THIS WAS RESOLVED AT DEVELOPMENT'S RECOMMENDATION, PATCHED AUTOMATION BACK TO 11.1, WHICH RESULTED IN THE REPORTS GETTING COMPLETED SUCCESSFULLY. NEW REPORTS WERE ALSO BEING PROCESSED AS EXPECTED THE JIRA OPENED FOR THIS ISSUE HAS BEEN CLOSED AND VERIFIED AS RESOLVED IN SOFTWARE VERSION (B)(4). AUTOMATION WILL NOT PROCESS REPORTS AFTER UPGRADE TO (B)(4). THE SITE WAS UPDATED TO SOFTWARE VERSION (B)(4) ON 05/17/2017. THE CAPA OPENED FOR THIS ISSUE HAS BEEN COMPLETED/CLOSED. QUALITY SYSTEM (B)(4) [UNITY] AUTOMATION WILL NOT PROCESS REPORTS. NO FURTHER EVALUATION/INVESTIGATION WILL BE PERFORMED REGARDING THIS INCIDENT. Manufacturer Narrative: THE INVESTIGATION INTO THIS CUSTOMER'S ALLEGATION DETERMINED THAT THE DELAY IN THE REPORT BEING FULLY APPROVED WAS A DEFECT INTRODUCED IN MERGE UNITY PACS RELEASE R11.1.2. THIS RELEASE WENT OUT TO A LIMITED NUMBER OF CUSTOMERS PRIOR TO BEING DETECTED. MERGE HEALTHCARE HAS SINCE DETERMINED THE ROOT CAUSE OF THE ISSUE AND HAS PERFORMED CORRECTIONS AND CORRECTIVE ACTIONS. A NEW RELEASE OF SOFTWARE HAS PROGRESSED THROUGH REVIEW AND VERIFICATION PHASES AND WILL BE AVAILABLE TO CUSTOMERS UPON RELEASE. A REVIEW OF THE MERGE UNITY PACS USER GUIDES, READING PHYSICIAN CORE CURRICULUM: ALL LEVELS DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING EXAMS THAT REMAIN STUCK IN A PENDING/HOLD STATUS. THE WORKAROUND PROVIDED TO THE CUSTOMER IS DEMONSTRATING EFFECTIVENESS FOR REPORTING RELATED ACTIVITIES. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 08/18/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, UPLOADING A REPORT WOULD CAUSE THE DOMINATOR STATIONS TO CRASH. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE STATION CRASHING WHILE UPLOADING A REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, IF THE REPORT DOES NOT FULLY UPLOAD TO THE SERVER BEFORE THE STATION CRASHES. THE REPORT MAY NEED TO BE RE-DICTATED IF THE REPORT IS NOT RECOVERABLE. (B)(4). Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 08/02/2016, MERGE WAS NOTIFIED THAT A REPORT FOR AN EXAM WAS MISSING. AFTER FURTHER INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT A REPORT WAS CREATED & SAVED ON JUNE 26TH, 2016, HOWEVER, THERE WAS NO TRACE OF THE REPORT ON THE I DRIVE. IN ADDITION, AN AUDIT BY THE SUPPORT TECHNICIAN SHOWED NO TRAIL OF THE REPORT BEING COMBINED OR THE EXAM BEING MERGED. THE MERGE SUPPORT TECHNICIAN WAS NOT ABLE TO RECOVER THE REPORT OR FIND THE WORD DOCUMENT FILE FOR THE REPORT. THEREFORE, THE REPORT WAS UNABLE TO BE FOUND AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: AN INVESTIGATION DONE BY MERGE HEALTHCARE DETERMINED THAT IF A REPORT WAS SAVED TOO QUICKLY AFTER OPENING THE EXAM, THE REPORT WOULD NOT SAVE SUCCESSFULLY AS EXPECTED. THE USER MUST SAVE THE REPORT AFTER THE REPORT TEMPLATE HAS COMPLETED BUILDING IN ORDER FOR THE REPORT TO SAVE SUCCESSFULLY. THE CANDIDATE FIX FOR THIS ISSUE IS SCHEDULED TO BE IN RELEASE 11.1.2. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 07/27/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, UPLOADING A REPORT WOULD INTERMITTENTLY CAUSE THE DOMINATOR STATION TO CRASH. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, THE STATION CRASHING WHILE UPLOADING A REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, IF THE REPORT DOES NOT UPLOAD TO THE SERVER BEFORE THE STATION CRASHES. THE REPORT MAY NEED TO BE RE-DICTATED IF THE REPORT IS NOT RECOVERABLE. (B)(4). Manufacturer Narrative: TESTING AND INVESTIGATION DETERMINED THAT WHEN THE APPLICATION WAS WAITING FOR A LONG-RUNNING OPERATION (SUCH AS A DATABASE OPERATION) IN THE FOREGROUND AND THE BACKGROUND REPORT UPLOAD TASKS WERE COMPLETED BEFORE THE UNIVERSAL MANAGER MESSAGE PUMP RETURNED TO PROCESSING, THIS CAUSED THE STATON TO CRASH. THIS ISSUE WAS RESOLVED IN RELEASE 11.1.1.2, WHERE THE REPORT UPLOAD SYSTEM NO LONGER ATTEMPTS TO READ FROM THE TASK MEMORY IF THE TASK IS NOT ASSOCIATED WITH AN ACTIVE JOB. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER. THIS INVESTIGATION IS ONGOING AND NOT YET COMPLETE. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 06/23/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, SOME PATIENT EXAMS WERE NOT VISIBLE FROM THE ONLINE TAB. THROUGH AN INVESTIGATION BY MERGE SUPPORT, IT WAS DETERMINED THAT EXAMS WERE PROCESSED THROUGH THE DICOM GATEWAY BUT THE DATE OF BIRTH (DOB) WAS NOT FILLED IN ON AN EXAM. ONCE THE DOB WAS ADDED, THE EXAM WAS THEN PULLED FORWARD AND AVAILABLE ON THE ONLINE LIST OF EXAMS. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, PATIENT EXAMS THAT ARE NOT VISIBLE FROM THE ONLINE TAB HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: THIS ISSUE WAS INVESTIGATED BY UNITY SUPPORT AND DEVELOPMENT UNDER (B)(4). ACCORDING TO THE INVESTIGATION, EXAMS CAN BE VIEWED BY GOING TO THE PATIENTS TAB AND LOOKING AT PRIORS FOR THE ASSOCIATED PATIENT. IF AN EDIT IS MADE TO THE EXAM USING THIS PROCESS, THE EXAM SUBSEQUENTLY SHOWS UP WHEN SEARCHING FOR IT IN THE ONLINE TAB. PER THE INVESTIGATION, THIS ISSUE WAS FIXED IN A LATER VERSION 11.1.2. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 06/24/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, IT WAS NOTED THAT REPORTS WERE MISSING FROM EXAMS AFTER DICTATION. A REVIEW OF THE CUSTOMER'S AUDIT TRAIL INDICATED THAT THE REPORTS WERE SAVED. HOWEVER, FURTHER INVESTIGATION BY MERGE SUPPORT DETERMINED THAT THE CUSTOMER'S NETWORK CONNECTION WAS TERMINATED PRIOR TO THE REPORTS SAVING. THEREFORE, THE REPORTS WERE LOST. IN ADDITION, THE REPORT WAS NOT SALVAGEABLE AND NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NOT ASSOCIATING TO AN EXAM AND NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE REPORT IS NOT ATTACHED. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 04/1/2016, MERGE WAS NOTIFIED THAT ON AN UNKNOWN DATE, STUDIES WERE MARKED READ AND THE REPORTS WERE APPROVED, HOWEVER, NO DICTATION OR REPORTS WERE FOUND. MERGE SUPPORT WAS ABLE TO RECOVER ALL BUT 2 REPORTS, WHICH HAD TO RE-READ AND RE-DICTATED. HOWEVER, A CAUSE OF THE ISSUE COULD NOT BE DETERMINED SINCE THE REPORTS WERE OLDER, AND NO LOGS COULD BE REVIEWED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, A MISSING REPORT HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE REPORT IS MISSING. FURTHER INFORMATION IS NEEDED TO CLARIFY THIS EVENT. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER CONFIRMED THAT THERE WERE NO KNOWN ADVERSE EVENTS OR HARM TO A PATIENT. THE ASSIGNABLE CAUSE OF THE FAILURE WAS NOT CONFIRMED DUE TO THE LOGS FROM THE TIME PERIOD INDICATED BEING NO LONGER AVAILABLE. THIS ISSUE WILL CONTINUE TO BE MONITORED. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE IS NO SIGNIFICANT TREND. TRENDING WILL CONTINUE TO BE MONITORED. REVISED INFORMATION CONTAINED IN THE SUPPLEMENTAL REPORT INCLUDES THE FOLLOWING: ON 6/3/2016: NEW INFORMATION RECEIVED BY MANUFACTURER. Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO EVALUATE THIS ALLEGATION TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. DEPENDING ON THE RESULTS, A SUPPLEMENTAL REPORT MAY BE SUBMITTED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 10/19/2016, MERGE WAS NOTIFIED THAT A REPORT WAS APPROVED, HOWEVER NO REPORT WAS FOUND. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT A REPORT ID WAS CREATED, HOWEVER THE WORD DOCUMENT WAS NOT CREATED WHEN READ. THE MERGE SUPPORT TECHNICIAN RECREATED A BLANK WORD DOCUMENT IN THE REPORTS FOLDER, AND THE REPORT NEEDED TO BE RE-DICTATED. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. (B)(4). Manufacturer Narrative: A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING MISSING REPORTS OR REPORTS NOT SAVING. A REVIEW OF THE COMPLAINT HISTORY FOUND NO INCREASING TREND RELATED TO MISSING REPORTS WHERE THE WORD DOCUMENT WAS NOT CREATED. TRENDING WILL CONTINUE TO BE MONITORED. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 10/18/2016, MERGE WAS NOTIFIED THAT TWO ELECTROCARDIOGRAM (EKG) REPORTS WERE MARKED AS READ, BUT NO REPORTS WERE FOUND DURING AN AUDIT BY THE PACS ADMINISTRATOR. AN INVESTIGATION BY MERGE SUPPORT DETERMINED THAT THERE WERE NO INDICATIONS THAT THE REPORTS WERE SAVED. FURTHER INVESTIGATION FOUND THAT THE USER CLICKED THE "DONE" BUTTON INSTEAD OF CLICKING "READ/NEXT" TO SAVE THE REPORT. TO SAVE AND APPROVE A REPORT THE USER MUST CLICK THE READ/NEXT BUTTON. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, REPORTS NEEDING TO BE RE-DICTATED HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS.. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: A REVIEW OF THE COMPLAINT HISTORY FOUND NO OTHER POSSIBLE SIMILAR COMPLAINTS RELATED TO ELECTROCARDIOGRAM (EKG) REPORTS NOT SAVING IF THE USER PRESSES THE DONE BUTTON INSTEAD OF THE READ/NEXT BUTTON. TRENDING WILL CONTINUE TO BE MONITORED. TO SAVE AND APPROVE A REPORT THE USER MUST CLICK THE READ/NEXT BUTTON AND PRESSING SAVE WOULD NOT SAVE THE REPORT AS EXPECTED. THIS WAS A KNOWN ISSUE AND ALREADY ADDRESSED AND RESOLVED IN RELEASE 11.1.1. ONCE THE PATCH WAS APPLIED TO THE CUSTOMER'S WORK STATION, THE REPORTS SAVED AS EXPECTED. THIS RESOLVED THE CUSTOMER'S ISSUE. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 2/3/2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, WHEN TOGGLING THE MONTAGE TO VIEW IT IN THE FAR RIGHT MONITOR, THE TISSUE VIEW CHANGED TO BONE VIEW. MERGE WAS ABLE TO DUPLICATE THIS EVENT. HOWEVER, AFTER UPDATING THE READING PREFERENCES TO 'PRIMARY EXAM ON ONE PANEL (3 PANELS)' OR 'RIGHT-MOST PANEL FOR MONTAGE ONLY' THE IMAGES DID NOT CHANGE INDEPENDENTLY. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. HOWEVER, STUDIES CHANGING VIEWS WITHOUT THE RADIOLOGIST BEING AWARE HAS THE POTENTIAL TO LEAD TO AN INCORRECT DIAGNOSIS AND/OR TREATMENT. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: (B)(4). | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON JUNE 1, 2016 CUSTOMER REPORTED THE DISTANCE INDICATOR BETWEEN IMAGE SLICES ARE NOT REFLECTING WHAT IS DISPLAYED ON THE MODALITY. IT IS INTERMITTENT AND IS NOT PRESENTLY KNOWN IF IT IS LIMITED TO ONE MODALITY OR VENDOR. MERGE HEALTHCARE WAS NOTIFIED BY THE CUSTOMER THAT THEY WERE INFORMED BY AMERICAN COLLEGE OF RADIOLOGY (ACR) THAT THEIR FACILITY DID NOT PASS ACCREDITATION DUE TO SPACING LISTING ON IMAGES NOT CONSISTENT THROUGHOUT THE SERIES. INITIAL EVALUATION OF IMAGES SHOWS THE SAGITTAL IMAGES WITHIN A SPECIFIC SERIES SHOULD BE SPACED 3.9MM APART. THE MODALITY SHOWS THE CONSISTENT SPACING, HOWEVER, THE UNITY PACS VIEWER SHOWS A 4.0MM SPACING BETWEEN 2 OF THE IMAGES. DUE TO THE POTENTIAL FOR THE INCONSISTENCY IN SPACING MEASUREMENTS THAT IS DISPLAYED ON SOME IMAGES, THERE MAY BE CONFUSION WHICH COULD LEAD TO A POTENTIAL DELAY IN DIAGNOSIS OR TREATMENT FOR A PATIENT. THERE IS NO INDICATION THAT ANY PATIENTS WERE HARMED AS A RESULT OF THE DISPLAYED INCONSISTENCY IN SLICE MEASUREMENTS. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON SATURDAY (B)(6) 2016 A CUSTOMER ENTERED AN ALLEGATION REGARDING MERGE UNITY PACS VIA THE CUSTOMER PORTAL. MERGE HEALTHCARE'S UNITY SUPPORT BECAME AWARE OF THE CASE ON MONDAY MAY 9, 2016 DURING REGULAR BUSINESS HOURS. THE CUSTOMER REPORTED A RECENTLY CREATED TEST VERSION OF A MERGE UNITY STRUCTURED REPORT (SR) IS NOT SEPARATING FIELDS FOR ULTRASOUND (US) OBSTETRIC (OB) EXAMS WITH TWINS. WITH MERGE UNITY PACS NOT REPORTING PATIENT INFORMATION AS EXPECTED, THERE IS A POTENTIAL FOR MISTREATMENT OR MISDIAGNOSIS. THE ISSUE WAS IDENTIFIED DURING CREATION OF THE SR TEMPLATE AND NO PATIENT INVOLVEMENT WAS REPORTED. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to pan image | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. (B)(6) HOSPITAL REPORTED THAT IF 2 IMAGES ARE CAPTURED FOR THE SAME VIEW AND EACH VIEW IS ASSOCIATED TO A TOMO IMAGE, BOTH TOMOS ARE ASSOCIATED TO THE FIRST OF THOSE TWO IMAGES. THE USER MAY MISS SEEING THE ADDITIONAL VIEW WHICH WOULD HELP WITH DIAGNOSIS. WHAT IS EXPECTED IS THE ABILITY TO VIEW TOMO ON FIRST RCC IMAGE AND ABILITY TO VIEW TOMO ON THE SECOND RCC IMAGE. WHAT IS HAPPENING IS THAT BOTH TOMOS ARE ASSOCIATED TO THE FIRST RCC IMAGE. THE SECOND RCC IMAGE HAS NO TOMO OPTION AVAILABLE. RELEASE 11.0 OF UNITY HAS A DEFECT WHICH MISSED THE ASSOCIATE OF THE SECOND TOMO. DESIGN AND TESTING FAILED TO DETECT THERE COULD BE MORE THAN 2 TOMO IMAGES FOR SAME VIEW. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A MERGE UNITY CUSTOMER CONTACTED MERGE ((B)(4)) TO REPORT THAT IF THE LAB CAPTURES 2 IMAGES FOR THE SAME VIEW (EX. 2 RMLO IMAGES, 2 LCC IMAGES) AND EACH VIEW IS ASSOCIATED TO A TOMO IMAGE, BOTH TOMOS ARE ASSOCIATED TO THE FIRST OF THOSE TWO IMAGES. THIS IS A POTENTIAL SAFETY ISSUE BECAUSE CLINICIANS ARE UNABLE TO VIEW MULTIPLE IMAGES FROM THE SAME VIEW IN UNITY. THIS PLACES LIMITS ON THE IMAGES AVAILABLE FOR THE CLINICIAN TO VIEW AT THE TIME OF DIAGNOSIS. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A MERGE UNITY CUSTOMER CONTACTED MERGE (CASE (B)(4)) AND REPORTED A PHYSICIAN WAS READING TOMO AND THE REFERENCE LINES FOR LATERAL TO MEDIAL ARE NOT CORRECT. IT WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE BECAUSE THE INCORRECT DISPLAY OF THE REFERENCE LINES COULD CONTRIBUTE TO DIAGNOSIS ERRORS. AN INVESTIGATION DETERMINED THAT THE ISSUE WAS NOT CAUSED BY MERGE UNITY BUT WAS AN ERROR BY THE IMAGE MODALITY VENDOR, (B)(4) HAD INCORRECT IMAGE POSITION INFORMATION FOR ALL ANGULATED VIEWS IN THE DICOM MESSAGE. Manufacturer Narrative: A CORRECTION REPORT (B)(4) WAS ERRONEOUSLY REFERENCED IN THE INITIAL MDR. THERE WAS NO RECALL ASSOCIATED WITH THIS MDR. THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO REMOVE THE REFERENCE TO THE RECALL REPORT. Manufacturer Narrative: AN INVESTIGATION DETERMINED THAT THE ISSUE WAS NOT CAUSED BY MERGE UNITY BUT WAS AN ERROR BY THE IMAGE MODALITY VENDOR, (B)(4) HAD INCORRECT IMAGE POSITION INFORMATION FOR ALL ANGULATED VIEWS IN THE DICOM MESSAGE. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A TELERADIOLOGY CUSTOMER REPORTED ON (B)(6) 2016 THAT THEY WERE NOT ABLE TO FIND EXAMS THAT WERE SENT. THROUGH AN INVESTIGATION BY MERGE SUPPORT IT WAS DETERMINED THAT EXAMS WERE PROCESSED THROUGH THE DICOM GATEWAY BUT THE DATE OF BIRTH (DOB) WAS NOT FILLED IN ON AN EXAM. ONCE THE DOB WAS ADDED, THE EXAM WAS THEN PULLED FORWARD AND AVAILABLE ON THE ONLINE LIST OF EXAMS. DUE TO THE POTENTIAL FOR A DELAY IN DIAGNOSIS THIS ISSUE IS BEING REPORTED. (B)(4). Manufacturer Narrative: CUSTOMER IS RUNNING SOFTWARE RELEASE R11.0.3. ORIGINAL ISSUE WAS IDENTIFIED AS A SOFTWARE DEFECT AND WAS FIXED IN SOFTWARE RELEASE R11.1.2. LEVEL 3 TECH SUPPORT ENGINEER CONTACTED CUSTOMER TO ADVISE AND TALK ABOUT PATH TO UPGRADE THE SITE TO THE NEW RELEASE. CUSTOMER ADVISED TECH SUPPORT ENGINEER THAT THEY HAVE MOVED TO ANOTHER PACS SYSTEM AND NO LONGER ACQUIRE EXAMS TO MERGE UNITY PACS. CURRENTLY ONLY HAVE MERGE UNITY PACS POWERED FOR ACCESS TO PRIOR EXAMS FOR MIGRATION TO NEW PACS. CUSTOMER DOES NOT WANT TO UPGRADE THE MERGE UNITY PACS SYSTEM. NO FURTHER ACTION IS REQUIRED. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. CUSTOMER REPORTED THAT A GOOD NUMBERS OF REPORTS HAVE BEEN MISSING HEADERS. REPORTS ARE FOR DISTRIBUTION TO REFERRING PHYSICIANS. IF THE PATIENT DEMOGRAPHICS IS MISSING AND THE REPORT IS DISTRIBUTED VIA FAX OR PRINTER, THE REPORT COULD NOT BE ASSOCIATED TO ANY PATIENT SINCE NO PATIENT DEMOGRAPHICS ARE AVAILABLE. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO THE PATIENT. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. CUSTOMER REPORTED THAT A GOOD NUMBERS OF REPORTS HAVE BEEN MISSING HEADERS. REPORTS ARE FOR DISTRIBUTION TO REFERRING PHYSICIANS. IF THE PATIENT DEMOGRAPHICS IS MISSING AND THE REPORT IS DISTRIBUTED VIA FAX OR PRINTER, THE REPORT COULD NOT BE ASSOCIATED WITH ANY PATIENT SINCE NO PATIENT DEMOGRAPHICS ARE AVAILABLE. THIS EVENT DID NOT RESULT IN A SERIOUS INJURY TO THE PATIENT. (B)(4) Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. DIAGNOSTIC IMAGES SERVICES REPORTED THAT MONTAGE WAS DISPLAYING AN INCORRECT TOMO SLICE. THE FIRST TOMO IMAGE ADDED TO THE MONTAGE IS CHANGED WHEN ADDING A SECOND SLICE. (B)(6) RADIOLOGY ALSO REPORTED WHEN ADDING MULTIPLE TOMO IMAGES TO MONTAGE, THE IMAGES SELECTED ARE NOT THE ONES THAT INSERT TO THE MONTAGE. A PHYSICIAN ENCOUNTERED THIS ISSUE WHEN PRESENTING TO OTHER PHYSICIANS. THE MONTAGE THAT WAS CREATED WITH TOMO IMAGES DID NOT HAVE THE INTENDED TOMO IMAGES. DEFECT WAS FOUND IN UNITY PACS 11.0 CODE. CODE WAS NOT HANDLING MONTAGES CORRECTLY WHEN CLOSING EXAMS. (B)(4) Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER ALLEGED THAT STRUCTURED REPORTS WERE DISPLAYING MEASUREMENT UNITS IN CM VERSUS MM AS SENT FROM THE MODALITY. DUE TO THE ALLEGED LENGTH UNIT MEASUREMENT ISSUE, THIS ISSUE IS BEING REPORTED AS THERE IS A POTENTIAL FOR A MISDIAGNOSIS OR MISTREATMENT OF A PATIENT. THE CUSTOMER HAS NOT REPORTED ANY KNOWN IMPACTS TO A PATIENT. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE, IMAGES WERE SKIPPED WHILE SCROLLING WITH THE MOUSE WHEEL. THE CUSTOMER REPORTED THAT INITIALLY, A FRACTURE WAS MISSED, HOWEVER WAS CAUGHT AFTER FURTHER INVESTIGATION. IMAGES BEING SKIPPED OVER HAS THE POTENTIAL TO LEAD TO AN INCORRECT AND INCOMPLETE DIAGNOSIS AND/OR TREATMENT, ESPECIALLY IF THE PHYSICIAN OR HEALTH CARE PROFESSIONAL IS NOT AWARE THAT THE IMAGE ARE BEING SKIPPED. A FIELD SERVICE ENGINEER CONDUCTED AN INVESTIGATION ON SITE AND NOTED THAT THERE WAS A BROKEN WIRE IN THE CORD, CAUSING THE MOUSE TO CONNECT AND DISCONNECT FROM WINDOWS. THE MOUSE WAS REPLACED. (B)(4). | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 10/10/2015, A CUSTOMER REPORTED ((B)(4)) THAT READING PREFERENCES UNEXPECTEDLY CHANGED, RESULTING WITH "DISPLAY IMAGES ON PRIMARY MONITOR" AND "DISPLAY PRIMARY MONTAGE" ENABLED. THIS CAUSED THE PRIMARY EXAM ON THE 1+3 MONITOR SETUP TO PLACE THE PRIMARY STUDY ON THE COLOR MONITOR WHICH IS HIDDEN FROM VIEW BECAUSE THE ELECTRONIC REQUISITION WAS MAXIMIZED AND COVERED THE PRIMARY EXAM FROM VIEW. THE PHYSICIAN STOPPED WHEN HE NOTICED THAT THE EXAM HE WAS VIEWING AS THE PRIMARY WAS REALLY THE COMPARISON. HAD HE CONTINUED, THE SYSTEM WOULD HAVE ISSUED A POPUP WARNING MESSAGE WHICH LETS THEM KNOW THAT NOT ALL IMAGES HAVE BEEN VIEWED. ALSO, THE NON-PRIMARY IMAGES ARE HIGHLIGHTED "OTHER" BY THE SYSTEM. THIS PROBLEM WAS IDENTIFIED AS A POTENTIAL SAFETY COMPLAINT BECAUSE IT CAUSED USER CONFUSION. NO KNOWN PATIENTS WERE HARMED DUE TO THIS COMPLAINT. THE PHYSICAN NOTICED THE ISSUE BEFORE MAKING THE MISTAKE OF READING THE COMPARISON EXAM THINKING IT WAS THE PRIMARY. Manufacturer Narrative: THIS IS BEING REPORTED AFTER CONDUCTING A RETROSPECTIVE REVIEW. NO DEFECTS WERE FOUND WITH SOFTWARE. THE DIFFERENCE BETWEEN GLOBAL SETTINGS VS USER PREFERENCE SETTINGS CAUSED CONFUSION AND IMPACTED THE PROPER SET-UP OF THE PREFERENCES. SYSTEM WORKS AS DESIGNED. ADDTIONAL TRAINING WAS PROVIDED TO THE REPORTING USER. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER REPORTED THAT WHEN THEY LOAD MULTIPLE IMAGES TO THE MONTAGE THE PREVIOUS IMAGE GETS CHANGED AND HAS THE INCORRECT SLIDE NUMBER. THE FIRST TOMO IMAGE ADDED TO THE MONTAGE IS CHANGED WHEN ADDING A SECOND SLICE. IF THE MONTAGE IS MISSING IMAGES AND IF THE REFERRING PHYSICIAN OR SPECIALIST DO NOT READ THE FULL REPORT OR VIEW ALL THE IMAGES, THE DIAGNOSIS MAY BE INACCURATE OR THE PLAN OF CARE FOR THE PATIENT MAY NOT BE APPROPRIATE. REFERENCE: (B)(4). Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. NOTE THAT THE FDA TERMINATED THIS RECALL ON MARCH 6, 2017. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -fail to process medical imaging -fail to interpret images -unable to view display | -filling problem -data back-up problem -contamination during use | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A MERGE EMPLOYEE ALLEGED THAT THE 3D CALCIUM SCORING REPORT MAY BE INCORRECT OR INACCURATE FOR PATIENTS. SPECIFICALLY, WHEN A GENDER OF OTHER (O) IS SELECTED, THEY HAVE A CALCIUM SCORE GREATER THAN 75 AND/OR NO AGE IS PROVIDED IN THE STUDY. THIS COULD LEAD TO AN INCORRECT CALCULATION OR DELAY IN PATIENT CARE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION TO DETERMINE IF ANY CORRECTION OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 THE CUSTOMER REPORTED THAT AN EXAM REPORT WAS MISSING. THE MERGE HEALTHCARE UNITY TECHNICIAN'S INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT ON THE DAY EXAM WAS READ (B)(6) 2018. THERE WERE NO LOGS AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THE EXAM WAS RE-READ BY THE RADIOLOGIST. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. NO LOGS WERE AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THERE WAS A REPORT SAVE EVENT IN THE AUDIT TRAIL. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 A CUSTOMER REPORTED THAT A CASE WAS READ, HOWEVER, IS INCOMPLETE AND THE EXAM REPORT DID NOT FILE INTO HOSPITAL INFORMATION SYSTEM (HIS) SYSTEM. THE MERGE HEALTHCARE UNITY TECHNICIAN INVESTIGATION SHOWED THAT THE REPORT DID SAVE SUCCESSFULLY, HOWEVER, THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. THE EXAM REPORT WAS SUCCESSFULLY SAVED IN UNITY PACS. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 THE CUSTOMER REPORTED THAT THEY WERE UNABLE TO OPEN AN EXAM REPORT. THE UNITY TECHNICIAN'S INVESTIGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. THERE WERE NO LOGS AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. NO LOGS WERE AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THERE WAS A REPORT SAVE EVENT IN THE AUDIT TRAIL. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 THE CUSTOMER REPORTED THAT AN EXAM WAS COMPLETED, HOWEVER, THERE IS NO EXAM REPORT. THE UNITY TECHNICIAN'S INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE EXAM WAS READ, THERE WAS AN UPLOAD ERROR--THE LOCALLY CACHED REPORT WAS CORRUPT OR MISSING WHEN UNIVERSAL MANAGER TRIED TO UPLOAD IT: THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE EXAM WAS READ, THERE WAS AN UPLOAD ERROR. THE LOCALLY CACHED REPORT WAS CORRUPT OR MISSING WHEN UNIVERSAL MANAGER TRIED TO UPLOAD IT. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM WAS COMPLETED, HOWEVER, THERE IS NO EXAM REPORT. UNITY TECH INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. THE EXAM WAS READ, HOWEVER STATION ENCOUNTERED HTTP EXCEPTION ERRORS BEFORE A STATION CRASH (UNKNOWN IF CRASH WAS A FORCE QUIT) AND THE RESULTS WERE NEVER UPLOADED TO THE SITE IMAGE SERVER. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. THE EXAM WAS READ, HOWEVER THE USER'S WORK STATION ENCOUNTERED HTTP EXCEPTION ERRORS BEFORE A STATION CRASH (UNKNOWN IF CRASH WAS A FORCE QUIT) AND THE RESULTS WERE NEVER UPLOADED TO THE SITE'S IMAGE SERVER. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 A CUSTOMER REPORTED THAT AN EXAM REPORT WAS MISSING. THE MERGE UNITY SUPPORT TECHNOLOGIST'S INVESTIGATION INTO THE CUSTOMER'S ALLEGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. WHEN THE MERGE HEALTHCARE INVESTIGATION STARTED, THE CUSTOMER'S STATION LOGS HAD TRUNCATED, AND THEREFORE, WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 A CUSTOMER REPORTED THAT AN EXAM REPORT WAS MISSING. THE MERGE UNITY SUPPORT TECHNOLOGIST'S INVESTIGATION INTO THE CUSTOMER'S ALLEGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. WHEN THE MERGE HEALTHCARE INVESTIGATION STARTED, THE CUSTOMER'S STATION LOGS HAD TRUNCATED, AND THEREFORE, WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM REPORT WAS MISSING. THE MERGE UNITY SUPPORT TECHNOLOGIST'S INVESTIGATION INTO THE CUSTOMER'S ALLEGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. WHEN THE MERGE HEALTHCARE INVESTIGATION STARTED, THE CUSTOMER'S STATION LOGS HAD TRUNCATED, AND THEREFORE, WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2017, A CUSTOMER REPORTED THAT A PHYSICIAN DICTATED AN EXAM THE PREVIOUS NIGHT BUT WAS NOT ABLE TO FIND THE REPORT ON THE EXAM. MERGE UNITY TECHNICAL SUPPORT'S INVESTIGATION SHOWED THE CUSTOMER'S DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE PHYSICIAN'S WORK STATION LOGS SHOWED THE REPORT SAVE EVENT, WHICH WAS INTERRUPTED BY A WORK STATION CRASH, WHICH CAUSED THE ISSUE. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4) Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE READING STATION CRASHED WHILE UPLOADING THE REPORT. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT COMPLETE THE UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. LOGS REVEALED THE STATION CRASHED WHILE UPLOADING THE EXAM REPORT TO THE IMAGE SERVER. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2017, A PACS ADMIN REPORTED TO MERGE HEALTHCARE TECHNICAL SUPPORT THAT ONE EXAM REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. THE MERGE UNITY SUPPORT INVESTIGATION SHOWED THAT THE CUSTOMER'S DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE REPORT'S AUDIT TRAIL SHOWED THAT THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4) Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED. WHEN THE USER ATTEMPTED TO OPEN THE REPORT, WHILE THE REPORT HEADER AND PATIENT DEMOGRAPHICS DISPLAYED, THE BODY OF THE REPORT WAS MISSING. MERGE TECHNICAL SUPPORT'S INVESTIGATION SHOWED THAT THE EXAM WAS READ AND APPROVED, HOWEVER, THE BODY OF THE REPORT DID NOT UPLOAD TO THE IMAGE SERVER. TO RESOLVE THE ISSUE THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE PATIENT IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. WHEN PROMPTED TO VIEW THE REPORT, THE SYSTEM DISPLAYS THE REPORT HEADER AND PATIENT DEMOGRAPHICS, HOWEVER THE BODY OF THE REPORT IS MISSING. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. A UNITY PACS ADMINISTRATOR REPORTED TO MERGE TECHNICAL SUPPORT ON (B)(6) 2018 THAT AN INVALID IMAGE FORMAT WAS FOUND ON AN EXAM THAT WAS PREVIOUSLY READ BY A RADIOLOGIST. UPON INITIAL REVIEW ON APRIL 27, 2018 IT WAS BELIEVED THIS ISSUE WAS RELATED TO THE SITE'S SUPPLIED STORAGE VENDOR. THE COMPLAINT REMAINED OPEN FOR ADDITIONAL INFORMATION BUT WAS MARKED AS "NOT REPORTABLE". ON MAY 18, 2018 NEW INFORMATION BECAME AVAILABLE. THE ISSUE WITH THE INVALID IMAGE FORMAT WAS DETERMINED TO BE UNITY SOFTWARE RELATED. WHILE THE EXAM REPORT IS AVAILABLE, THE INABILITY TO RESTORE AN EXAM FROM LONG TERM STORAGE FOR VIEWING COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: THE INVESTIGATION INTO THIS ISSUE IS ACTIVE AND ONGOING. UNTY PACS HAS DISCOVERED HD ARCHIVE AND BURN TO CD ARE CORRUPTING IMAGES WITH LARGE LOOK UP TABLES (LUTS). IN THE CASE OF HD ARCHIVE, THE CORRUPTION GOES UNNOTICED UNTIL THE IMAGES ARE RESTORED. SOME IMAGES CAN BE RESTORED WHILE OTHER IMAGES CANNOT BE VIEWED (INVALID IMAGE). THERE ARE EXAM REPORTS AVAILABLE FOR REVIEW FOR ALL EXAMS. MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. ON MAY 23, 2018 THE LEAD UNITY SUPPORT TECHNICIAN CONTACTED THE DESIGNATED COMPLAINT HANDLING UNIT (DCHU) WITH ADDITIONAL INFORMATION LEADING DCHU TO SUBMIT THIS ISSUE TO FDA. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016, A CUSTOMER REPORTED THAT NON-ARCHIVED EXAMS HAD ACCIDENTALLY BEEN DELETED. THE CUSTOMER INDICATED THAT THE SYSTEM DID NOT GIVE ANY WARNING WHILE EXAMS THAT WERE NOT ARCHIVED WERE DELETED. WITH MERGE UNITY PACS NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR PATIENT DATA TO BE UNINTENTIONALLY DELETED WHICH MAY RESULT IN DATA USED FOR DIAGNOSIS OR TREATMENT DECISIONS BEING UNAVAILABLE. HOWEVER, THERE WAS NO SPECIFIC REPORTED PATIENT INVOLVEMENT OR HARM AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 A CUSTOMER REPORTED THAT THERE ARE MISSING STANDARD UPTAKE VALUES (SUVS) IN THE DICOM HEADER FROM A PET MODALITY. THE MISSING VALUES ARE CAUSING PROBLEMS WHEN THE RADIOLOGIST REVIEWS THE SUV ON PET EXAMS. DUE TO MERGE UNITY PACS NOT PRESENTING VALUES AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS THAT MAY LEAD TO HARM. HOWEVER, THERE WAS NO REPORTED ADVERSE EVENT OR HARM TO A PATIENT DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: DCHU SPOKE WITH PACS ADMINISTRATOR ON (B)(6) 2018. THE JIRA FOR THIS ISSUE HAS BEEN CLOSED AND VERIFIED AS BEGING RESOLVED IN R11. PACS ADMINISTRATOR REPORTS THIS SITE IS UPGRADING TO R11 ON (B)(6) 2018. UPGRADE TAB IN SALES FORCE CONFIRMS THIS SITE IS SCHEDULED TO GO DOWN AT 8:00 AM ON (B)(6) 2018 FOR UPGRADE TO R11.1.2.6. Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2016 A CUSTOMER ALLEGED THAT THEY WERE UNABLE TO FIND A PATIENT'S REPORT HOWEVER, THE PATIENT'S EXAM WAS FOUND. MERGE HEALTHCARE ASSISTED THE CUSTOMER AND DETERMINED THAT THE CUSTOMER'S EXAM WAS DONE ON (B)(6) 2016. WHILE THE CUSTOMER INDICATED THAT THE EXAM HAD BEEN REVIEWED AND A REPORT CREATED, NO REPORT WAS LOCATED. THEREFORE THE PHYSICIAN HAD TO RE-READ THE EXAM. WITH A REPORT NOT BEING FOUND WITHIN MERGE UNITY PACS AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN TREATMENT OR DIAGNOSIS WHICH CAN LEAD TO HARM. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT DUE TO THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 5/19/2021, A CUSTOMER CONTACTED MERGE HEALTHCARE SUPPORT TO REPORT THAT A PATIENT REPORT WAS MISSING. MERGE HEALTHCARE DETERMINED UPON INVESTIGATION THAT THE REPORT SAVED BUT FAILED TO UPLOAD. LOGS WERE NOT AVAILABLE FOR REVIEW. A BLANK REPORT DOCUMENT WAS MANUALLY PLACED IN THE FOLDER ON (B)(6) 2021 AND THE REPORT WAS RE-DICTATED AND SUBSEQUENTLY VIEWABLE AS EXPECTED. THIS SITE HAS REGULAR AND PERSISTENT NETWORK ISSUES THAT AFFECT REPORT UPLOADS. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEIR MERGE UNITY PACS SYSTEM WAS COMPLETELY DOWN. TROUBLESHOOTING AND TECHNICAL SUPPORT IS ONGOING BETWEEN MERGE HEALTHCARE AND THE CUSTOMER. WHEN MORE INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. MERGE HEALTHCARE DETERMINED UPON INVESTIGATION THAT THE REPORT SAVED BUT FAILED TO UPLOAD. LOGS WERE REVIEWED AND DID NOT INDICATE ANY ISSUE. A BLANK REPORT DOCUMENT WAS MANUALLY PLACED IN THE FOLDER AND THE REPORT WAS RE-DICTATED AND SUBSEQUENTLY VIEWABLE AS EXPECTED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. MERGE HEALTHCARE DETERMINED UPON INVESTIGATION THAT THE REPORT SAVED BUT FAILED TO UPLOAD. LOGS WERE NOT AVAILABLE FOR REVIEW. A BLANK REPORT DOCUMENT WAS MANUALLY PLACED IN THE FOLDER AND THE REPORT WAS RE-DICTATED AND SUBSEQUENTLY VIEWABLE AS EXPECTED. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED THAT THE REPORT ID AND EXAM ID WERE NOT LINKED. THE ROOT CAUSE IS BELIEVED TO BE 2 USERS WRITING TO THE DATABASE TABLE AT THE SAME MOMENT. THE CUSTOMER RE-READ THE EXAM AND WAS ABLE TO SUCCESSFULLY VIEW THE REPORT. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED THAT THE REPORT DID NOT SAVE DUE TO A WORKSTATION CRASH. A BLANK REPORT DOCUMENT WAS MANUALLY PLACED IN THE FOLDER AND THE REPORT WAS RE-DICTATED AND SUBSEQUENTLY VIEWABLE AS EXPECTED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED VIA APPLICATION LOGS THAT THE REPORT NEVER UPLOADED BECAUSE THE C DRIVE ON THE LOCAL STATION WAS FULL, WHICH PREVENTED THE REPORT FROM SAVING. SUPPORT ASSISTED THE CUSTOMER IN RE-GENERATING THE REPORT AND RECOMMENDED REVIEWING THE C DRIVE OF THE LOCAL STATION. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT AN EXAM REPORT WAS MISSING FROM A RECENT EXAM. CLICKING ON THE REPORT BUTTON RESULTS IN AN ERROR STATING "THE REPORT IS NOT FOUND". ATTEMPTING TO OPEN THE REPORT WITH TRANSCRIPTION RESULTS IN AN ERROR STATING "OBJECT NOT FOUND." UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED THAT HE REPORT WAS SAVED, BUT INVESTIGATING THE CONTAINING FOLDER INDICATED THAT NO DATA WAS UPLOADED AT THE TIME. THERE WAS A REPORT SAVE EVENT IN AUDIT TRAIL. USING THE REVISION BUTTON FROM THE DICTATION DIALOG BOX, THE REPORT TEMPLATE WAS RE-GENERATED AND THE REPORT WAS THEN VIEWABLE AS EXPECTED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. UPON INVESTIGATION AND TROUBLESHOOTING BY MERGE HEALTHCARE SUPPORT, IT WAS DETERMINED THAT THE REPORT NEVER UPLOADED. ONCE A BLANK WORD DOC WAS CREATED, AND THE REPORT TEMPLATE WAS REGENERATED, THE CUSTOMER WAS ABLE TO READ THE REPORT SUCCESSFULLY. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, THE CUSTOMER REPORTED RECEIVING AN EXAM REPORT WITH NO REPORT BODY TO (B)(6) HEALTH IMAGING ON THE SAME DAY. TROUBLESHOOTING FROM (B)(6) HEALTH IMAGING SUPPORT FOUND THAT THERE WERE NO PATIENT DEMOGRAPHICS, BODY, OR CONCLUSION DATA INCLUDED IN THE REPORT FILE. THERE WAS A REPORT SAVE EVENT IN THE AUDIT TRAIL. USING THE REVISION BUTTON FROM THE DICTATION DIALOG BOX, THE REPORT TEMPLATE WAS RE-GENERATED AND THE REPORT WAS THEN VIEWABLE AS EXPECTED. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2020, THE CUSTOMER REPORTED RECEIVING A "NO OBJECT FOUND" ERROR WHEN TRYING TO OPEN A REPORT FOR DICTATION TO IBM WATSON HEALTH IMAGING ON (B)(6) 2020. TROUBLESHOOTING FROM IBM WATSON HEALTH IMAGING SUPPORT FOUND THAT THE REPORT FOLDER WAS EMPTY, CAUSING THE "NO OBJECT FOUND" ERROR, AND THAT THERE WAS A REPORT SAVE EVENT PRESENT IN THE DEVICE AUDIT TRAIL. IBM WATSON HEALTH IMAGING SUPPORT REGENERATED THE REPORT TEMPLATE AND SAVED IT IN DICTATED (DID) STATUS. THE REPORT WAS SAVED WITH THE NEW TEMPLATE. THE REPORT WAS THEN READABLE. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED THAT AN EXAM WAS READ ON (B)(6) 2019 BUT THE REPORT WAS MISSING. TECH SUPPORT CONNECTED TO THE SITE AND FOUND THAT THERE WAS A REPORT SAVE EVENT BUT THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE WORD REPORT TEMPLATE WAS REGENERATED AND THE DOCTOR RE-READ THE EXAM ON (B)(6) 2019. THE REPORTED ISSUE WAS RESOLVED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE DOCTOR. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS READ ON (B)(6) 2019, HOWEVER, THE WORKSTATION HAD AN OPERATING SYSTEM FAILURE AND THE EXAM REPORT COULD NOT BE FOUND. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE REPORT WORD TEMPLATE WAS CREATED AND THE EXAM WAS RE-READ BY THE DOCTOR ON THE SAME DAY (B)(6) 2019. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE DOCTOR. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT ASSOCIATED CORRECTLY TO THE EXAM. THE EXAM WAS GENERATED AGAIN BY THE SITE AND THE THEY CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT NUMBER: (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT ASSOCIATED CORRECTLY TO THE EXAM. THE EXAM WAS GENERATED AGAIN BY THE SITE AND THE THEY CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE: COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS A WARNING MESSAGE WITH A "SAVE AS" ERROR INDICATING THE MICROSOFT WORD DOCUMENT IS IN PROTECTED MODE. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO THE IMAGE SERVER. A NEW MICROSOFT WORD TEMPLATE DOCUMENT WAS CREATED AND SUPPORT CREATED A NEW REPORTED THAT WAS LINKED TO THE EXAM. THE EXAM WAS SENT BACK TO DICTATED (DID) STATUS AND THE SITE CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)6) 2019 THE CUSTOMER REPORTED AN EXAM ( EXAM DATE 3/5/19) WAS MARKED READ ON 3/5/19, HOWEVER, EXAM REPORT COULD NOT BE FOUND. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE REPORT WORD TEMPLATE WAS CREATED AND THE EXAM WAS RE-READ BY THE DOCTOR ON (B)(6) 2019. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE DOCTOR. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE SALES FORCE (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORT ID GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO THE IMAGE SERVER. A NEW MICROSOFT WORD TEMPLATE DOCUMENT WAS CREATED AND SUPPORT CREATED A NEW REPORT ID THAT WAS LINKED TO THE EXAM. THE EXAM WAS SENT BACK TO DICTATION ON HOLD (DIH) STATUS AND THE SITE CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT ASSOCIATED CORRECTLY TO THE EXAM. THE EXAM WAS GENERATED AGAIN BY THE SITE AND THE THEY CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM REPORT WAS UNABLE TO OPEN. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS A WARNING MESSAGE WITH A "SAVE AS" ERROR INDICATING THE (B)(6) DOCUMENT IS IN PROTECTED MODE. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORT-ID GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO THE IMAGE SERVER. A NEW (B)(6) TEMPLATE DOCUMENT WAS CREATED AND SUPPORT CREATED A NEW REPORT-ID THAT WAS LINKED TO THE EXAM. THE EXAM WAS SENT BACK TO DICTATED (DID) STATUS AND THE SITE CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT. THE REPORT WAS SAVED, BUT NOT ASSOCIATED PROPERLY WITH THE EXAM. A CAUSE OF THE FAILED ASSOCIATION WAS NOT ABLE TO BE DETERMINED BECAUSE THE READING STATION LOGS WERE TRUNCATED. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT ASSOCIATED CORRECTLY TO THE EXAM. THE EXAM WAS SENT BACK TO "SCANNED" STATUS AND THE SITE CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2019 THE CUSTOMER REPORTED AN EXAM WAS MARKED READ, HOWEVER, THERE IS NO EXAM REPORT OR SAVE EVENT IN THE AUDIT TRAIL. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I:\ DRIVE. THE EXAM WAS SENT, A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED. THE EXAM WAS SENT BACK TO "SCANNED" STATUS AND THE SITE CONFIRMED THAT THE EXAM WAS ABLE TO BE RE-READ WITHOUT INCIDENT. NO FURTHER ACTION REQUIRED AT THIS TIME. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM WAS COMPLETED, HOWEVER, THERE IS NO EXAM REPORT. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTID GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I:\ DRIVE. A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTID. THE EXAM WAS RE-READ AND CONFIRMED WITH THE CUSTOMER THAT THE REPORT WAS RE-READ WITHOUT FURTHER ISSUES. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON DECEMBER 13, 2018 THE CUSTOMER REPORTED AN EXAM WAS COMPLETED, HOWEVER, THERE IS NO EXAM REPORT. AN INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY AND HAD TO BE READ AGAIN BY THE DOCTOR. THE EXAM WAS READ, THERE WAS AN UPLOAD ERROR AT THE TIME OF THE SAVE EVENT. THERE WAS NO REPORTED ADVERSE EVENT TO A PATIENT. A MISSING REPORT DOES HAVE THE POTENTIAL TO DELAY PATIENT CARE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I:\ DRIVE. A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. PENDING FURTHER REVIEW. A SUPPLEMENTAL REPORT WILL BE PROVIDED. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 A CUSTOMER REPORTED AN EXAM REPORT FAILED TO UPLOAD. THE MERGE HEALTHCARE, UNITY TECHNICIAN'S INVESTIGATION SHOWED THAT A SAVE EVENT DID OCCUR IN THE AUDIT TRAIL. HOWEVER, THE READ REPORT ITSELF FAILED TO UPLOAD PROPERLY. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. THE MERGE HEALTHCARE UNITY TECHNICIAN'S INVESTIGATION SHOWED THAT A SAVE EVENT DID OCCUR IN THE AUDIT TRAIL. HOWEVER, THE READ REPORT ITSELF FAILED TO UPLOAD PROPERLY. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. (B)(4). | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM REPORT IS CORRUPT. THE MERGE HEALTHCARE UNITY TECH'S INVESTIGATION SHOWED THAT THE REPORT DID SAVE SUCCESSFULLY, HOWEVER, THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. THE EXAM REPORT WAS SUCCESSFULLY SAVED IN UNITY PACS. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE MS WORD DOCUMENT (EXAM REPORT) WAS CORRUPT. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED THEY WERE UNABLE TO OPEN AN EXAM REPORT. THE MERGE HEALTHCARE UNITY SUPPORT TECH'S INVESTIGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT ON THE DAY EXAM WAS READ (B)(6) 2018. THERE WERE NO LOGS AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THE EXAM WAS RE-READ BY THE RADIOLOGIST ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. NO LOGS WERE AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THERE WAS A REPORT SAVE EVENT IN THE AUDIT TRAIL. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED THERE WERE NO REPORTS FOR 6 MAMMOGRAPHY EXAMS. THE REPORTS WERE READ BY ONE PHYSICIAN ON A HOME READING STATION IN ONE SESSION. THE MERGE HEALTHCARE UNITY SUPPORT TECHNICIAN INVESTIGATED THE EXAMS, AND ALL OF THEM DID HAVE REPORT IDS WITH FOLDERS IN I:\REPORTS. TWO OF THE SIX EXAM REPORTS WERE RECOVERED, THE OTHER 4 EXAMS WERE RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN, WHO WAS REVIEWING THE EXAMS FROM A HOME COMPUTER, RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION ALLOWED THE RECOVERY OF TWO OF THE REPORTS. THE CAUSE OF THE OTHER FOUR REPORTS FAILING TO LOAD WAS DUE TO THE PHYSICIAN'S HOME WORKSTATION CRASHING DURING THE UPLOAD OF THE SIX REPORTS. THIS CAUSED FOUR OF THE SIX REPORT DOCUMENTS TO BE LOST. TO RESOLVE THIS ISSUE, THE OTHER 4 EXAMS WERE RE-READ BY THE PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 THE CUSTOMER REPORTED AN EXAM WAS COMPLETED BUT THERE IS NO EXAM REPORT. UNITY SUPPORT INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISSING REPORT DOES HAVE THE POTENTIAL TO DELAY PATIENT CARE. REFERENCE SALES FORCE COMPLAINT-(B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. PENDING FURTHER REVIEW. A SUPPLEMENTAL REPORT WILL BE PROVIDED. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE CUSTOMER REPORTED THAT THE EXAM WAS STUCK AND THE REPORT WAS NOT AVAILABLE. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY REVIEWING THE MERGE UNITY PACS WORKSTATION. THE AUDIT TRAIL OF THE EXAM SHOWED THE ORIGINAL EXAM WAS READ ON THE SAME DAY AS THE EXAM DATE (B)(6) 2018 HOWEVER THE USER DID NOT SAVE THE REPORT AND THAT THERE WAS NO SAVE REPORT EVENT ON (B)(6) 2018 IN THE LOG. THERE WAS NO DATA UPLOADED AND THEREFORE THE REPORT WAS NOT AVAILABLE. THE EXAM WAS REQUIRED TO BE RE-READ BY THE PHYSICIAN. THE EXAM WAS RE-READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER: (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. THE EXAM AUDIT TRAIL DID NOT SHOW A SAVED REPORT EVENT. IT WAS NOT POSSIBLE TO DETERMINE IF THE ERROR MESSAGE DISPLAYED ON THE SCREEN FOR THE DOCTOR WHEN THE REPORT WAS FIRST READ AND SAVED NOTIFIED THE FAILURE TO UPLOAD. THERE WAS NO EVIDENCE THAT REPORT WAS SAVED. THE UNITY SUPPORT TECHNICIAN'S INVESTIGATION SHOWED THAT THE REPORT WAS NOT UPLOADED AND SAVED. AS A RESULT OF FINALIZED REPORT BEING UNAVAILABLE, THE EXAM WAS RE-READ BY A PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 A CUSTOMER REPORTED THAT AN EXAM REPORT WAS MISSING. THE MERGE UNITY SUPPORT TECHNOLOGIST'S INVESTIGATION INTO THE CUSTOMER'S ALLEGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. WHEN THE MERGE HEALTHCARE INVESTIGATION STARTED, THE CUSTOMER'S STATION LOGS HAD TRUNCATED, AND THEREFORE, THE WORKSTATION LOGS WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. TO RESOLVE THE ISSUE, THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 A CUSTOMER REPORTED TO MERGE UNITY TECHNICAL SUPPORT THAN AN EXAM REPORT WAS MISSING. THE MERGE UNITY SUPPORT TECHNOLOGIST'S INVESTIGATION INTO THE CUSTOMER'S ALLEGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 4/13/2018, THE CUSTOMER REPORTED THAT THE REPORT WAS MISSING FOR AN EXAM (DATED (B)(6) 2018) THAT WAS DICTATED BY THE DOCTOR ON (B)(6) 2018. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY REVIEWING THE MERGE UNITY PACS WORKSTATION. TECHNICAL SUPPORT DID IDENTIFY A STATION CRASH AND THERE WAS NO DATA UPLOADED AND THEREFORE THE REPORT WAS NOT AVAILABLE. THE EXAM WAS REQUIRED TO BE RE-READ BY THE PHYSICIAN. THE EXAM WAS RE-READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. STATION LOGS WERE UNAVAILABLE FOR REVIEW. IT WAS NOT POSSIBLE TO DETERMINE IF THE ERROR MESSAGE DISPLAYED ON THE SCREEN FOR THE DOCTOR WHEN THE REPORT WAS FIRST READ AND SAVED NOTIFIED THE FAILURE TO UPLOAD. THERE WAS NO EVIDENCE THAT REPORT WAS SAVED. THE UNITY SUPPORT TECHNICIAN'S INVESTIGATION SHOWED THAT THE UNITY PACS STATION CRASHED BEFORE THE REPORT WAS UPLOADED AND SAVED. AS A RESULT OF FINALIZED REPORT BEING UNAVAILABLE, THE EXAM WAS RE-READ BY A PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE CUSTOMER REPORTED THAT THE REPORT WAS MISSING FOR AN EXAM (DATED (B)(6) 2018) THAT WAS DICTATED BY THE DOCTOR ON (B)(6) 2018. MERGE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY REVIEWING THE WORKSTATION LOGS FOR MERGE UNITY PACS. TECHNICAL SUPPORT DID IDENTIFY A SAVE EVENT IN THE LOGS BUT THERE WAS NO DATA UPLOADED AND THEREFORE THE REPORT WAS NOT AVAILABLE. THE EXAM WAS REQUIRED TO BE RE-READ BY THE PHYSICIAN. THE EXAM WAS RE-READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. STATION LOGS WERE AVAILABLE FOR REVIEW. IT WAS NOT POSSIBLE TO DETERMINE IF THE ERROR MESSAGE DISPLAYED ON THE SCREEN FOR THE DOCTOR WHEN THE REPORT WAS FIRST READ AND SAVED NOTIFIED THE FAILURE TO UPLOAD. THERE WAS NO EVIDENCE THAT REPORT WAS SAVED. THE UNITY SUPPORT TECHNICIAN'S INVESTIGATION SHOWED THAT THE UNITY DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD. AS A RESULT OF THE FINALIZED REPORT BEING UNAVAILABLE, THE EXAM WAS RE-READ BY A PHYSICIAN. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, A CUSTOMER REPORTED THAT AN EXAM WAS MISSING A REPORT AND WOULD LIKE ASSISTANCE IN RETRIEVING IT. THE EXAM WAS READ ON (B)(6) 2018 PER THE CUSTOMER BUT NO REPORT WAS GENERATED. THE EXAM WAS RE-READ ON (B)(6) 2018 AND A REPORT WAS GENERATED. MERGE HEALTHCARE TECHNICAL SUPPORT PULLED THE AUDIT LOGS THAT COULD BE ANALYZED FROM RELATED WORKSTATIONS FOR FURTHER ANALYSIS. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE TECHNICAL SUPPORT PULLED THE AUDIT TRAIL FROM THE UM ON STATION 51 AND LOGS FOR 704 ON STATION 9. THE 704 LOGS ARE TRUNCATED AS OF (B)(6) 2018. THE AUDIT TRAIL WAS REVIEWED BY UNITY TECHNICAL SUPPORT. THE AUDIT TRAIL SHOWS FOUR EVENTS FOR (B)(6) 2018 WITH A REPORT SAVE EVENT, BUT DOES NOT EXPLAIN WHAT HAPPENED. THERE IS NO FURTHER INFORMATION AVAILABLE FOR ANALYSIS. THE LOGS FOR 704 ARE TRUNCATED AS OF (B)(6) 2018 AS ARE THE AUTOMATION LOGS. NO ADDITIONAL INFORMATION REGARDING THE ALLEGED REPORT SAVING MALFUNCTION WAS AVAILABLE FOR REVIEW. THE STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. DUE TO THE LACK OF INFORMATION FOR FURTHER ANALYSIS, NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON MARCH 15, 2018, THE CUSTOMER REPORTED AN EXAM REPORT THAT WAS DICTATED ON (B)(6) 2018 WAS NOT SAVED AND IS MISSING. UNITY TECH INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT BUT HEADER, BODY, AND REPORT FILES WERE NOT PRESENT. STATION LOGS HAD TRUNCATED, AND WERE UNABLE TO BE REVIEWED. THE EXAM WAS READ ON (B)(6) 2018. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY TECH INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER; THE REPORT HEADER, BODY, AND REPORT FILES WERE NOT PRESENT. THE EXAM WAS RE-READ BY A PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW AND FURTHER INVESTIGATION. AS A RESULT, NO OTHER REVIEW/INVESTIGATION CAN BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE UNITY CUSTOMER REPORTED THEY ARE UNABLE TO VIEW AN EXAM FROM 2013 FOR USE AS A COMPARISON STUDY. THEY ARE MET WITH A MESSAGE WHEN TRYING TO OPEN THIS PARTICULAR STUDY. ERROR READING EXAM DATA FIELD, EXAM CANNOT BE VIEWED. UNITY TECH INVESTIGATION SHOWED THERE WAS AN ONLINE RECORD FOR THIS EXAM AND ALTHOUGH THE DATABASE SHOWED AN OWNING LOCATION, THERE WAS NO IMAGE FOLDER THERE NOR WAS THERE ANY PATIENT FILE. AUDIT TRAIL DOES NOT EXPLAIN WHAT COULD HAVE HAPPENED. EXAM WAS NEVER ARCHIVED AND NOW APPEARS THAT THE IMAGE DATA WAS LOST ALONG THE WAY. THE READING PHYSICIAN DOES HAVE A PHYSICAL REPORT FROM THIS 2013 EXAM FOR REFERENCE. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISSING EXAM DOES HAVE THE POTENTIAL TO DELAY PATIENT CARE. (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS FURTHER INVESTIGATING THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE PACS ADMINISTRATOR REPORTED THAT NO REPORT WAS DISPLAYING FOR 2 EXAMS THAT HAD BEEN READ. THE UNITY TECH'S INVESTIGATION SHOWED ONE REPORT IS MISSING SINCE THE EXAM WAS USING A COMBINED REPORT. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. THE UNITY TECH'S INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. THE WORK STATION'S LOGS HAD TRUNCATED, AND WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT ALLEGED HE HAD A EXAM ONLINE THAT WAS STUCK IN APD* STATUS. THE MERGE SUPPORT TECHNICIAN TESTED DICTATION FOR THE NOTED EXAM. THERE WAS AN "OBJECT NOT FOUND" ERROR WHEN CLICKING THE DICTATE BUTTON. TECHNICAL SUPPORT DETERMINED THAT WHEN THE EXAM WAS READ AND APPROVED, THE REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. THE EXAM WAS RE-DICTATED, SAVED AND WAS STORED TO THE SERVER. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. STATION LOGS WERE NOT AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THE EXAM WAS RE-READ BY THE PHYSICIAN. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, A PACS ADMINISTRATOR REPORTED THEY WERE ATTEMPTING TO FINISH A STAT REPORT HOWEVER, THE SYSTEM WILL NOT ALLOW THEM TO FINISH READING. THE CUSTOMER REPORTED THAT THE SYSTEM DISPLAYED AN ERROR THAT 'OBJECT NOT FOUND'. MERGE HEALTHCARE TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION AND THE SYSTEM LOGS SHOWED THAT THE EXAM WAS READ AND APPROVED ON (B)(6) 2018, HOWEVER, THE EXAM REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. STATION LOGS WERE NOT AVAILABLE FOR REVIEW AS THEY HAD TRUNCATED. THE EXAM WAS RE-READ BY THE PHYSICIAN. DUE TO THE STATION LOG INFORMATION BEING UNAVAILABLE, NO FURTHER INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2018, THE CUSTOMER REPORTED AN EXAM REPORT WAS NOT SAVED AND IS MISSING. THE UNITY TECH INVESTIGATION SHOWED THAT THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WOULD HAVE TO BE RECREATED. THE AUDIT TRAIL DOES SHOW A "REPORT SAVE" EVENT. AT THE TIME OF THE INVESTIGATION, STATION LOGS HAD TRUNCATED, AND WERE UNABLE TO BE REVIEWED. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION LOGS HAD TRUNCATED, THEREFORE WERE UNAVAILABLE FOR REVIEW. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, THE PACS ADMIN REPORTED AN EXAM REPORT WAS STUCK IN APD* (APPROVED/ERROR) STATUS. USERS WERE UNABLE TO OPEN THE REPORT. MERGE HEALTHCARE UNITY TECHNICAL SUPPORT PERFORMED AND INVESTIGATION AND IT SHOWED THE CUSTOMER'S WORKSTATION HAD CRASHED DURING A REPORT SAVE. THIS WORKSTATION CRASH WAS CAUSED BY A MODAL CONFLICT WITH (B)(6). THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. STATION CRASH CAUSED BY MODAL CONFLICT WITH (B)(6) WHICH PREVENTED ORIGINAL REPORT UPLOAD. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS ADMIN REPORTED THEY HAVE 4 EXAMS MARKED AS "READ" BY THE RADIOLOGIST BUT NO REPORT WAS SAVED. THE UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE RADIOLOGIST. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS ADMIN REPORTED THEY HAVE 4 EXAMS MARKED AS "READ" BY THE RADIOLOGIST BUT NO REPORT WAS SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE RADIOLOGIST. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, THE PACS ADMIN REPORTED THEY HAVE 4 EXAMS MARKED AS "READ" BY THE RADIOLOGIST BUT NO REPORT WAS SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE RADIOLOGIST. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON DECEMBER 7, 2017, THE PACS ADMIN REPORTED A RADIOLOGIST IS UNABLE TO VIEW A REPORT. RECEIVING ERROR WHEN OPENING THE REPORT FOR DICTATION. REPORTING NO OBJECT FOUND. THE MERGE UNITY SUPPORT INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER REPORTED THAT A RADIOLOGIST WAS UNABLE TO VIEW A REPORT. THE RADIOLOGIST WAS RECEIVING AN ERROR WHEN OPENING THE REPORT FOR DICTATION. THE CUSTOMER REPORTED THAT NO OBJECT COULD BE FOUND. MERGE UNITY TECHNICAL SUPPORT INVESTIGATED THE CUSTOMER'S ALLEGATION BY CONNECTING TO THE WORKSTATION AND SEARCHING IN THE REPORTS FOLDER. THE INVESTIGATION SHOWED THAT THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ AND SAVED SUCCESSFULLY BY THE PHYSICIAN WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT (B)(6). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WAS DISCOVERED AFTER THEY TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS ADMINISTRATOR REPORTED THAT THEY HAVE AN EXAM ONLINE. THE EXAM IS MARKED AS READ BUT DOESN'T LIST A REPORT STATUS OR WHO DICTATED THE REPORT. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THAT THE CUSTOMER'S THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. AN AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE READING STATION CRASHED WHILE UPLOADING THE REPORT. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE EXAM WAS RE-READ BY THE PHYSICIAN. UNITY INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT COMPLETE THE UPLOAD TO THE SERVER. AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. LOGS REVEALED THE STATION CRASHED WHILE UPLOADING THE EXAM REPORT TO THE IMAGE SERVER. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6)2017 A CUSTOMER CONTACTED MERGE HEALTHCARE TECHNICAL SUPPORT AND ALLEGED THAT THEY HAD AN EXAM THAT HAD NO EXAM REPORT. MERGE UNITY TECHNICAL SUPPORT'S INVESTIGATION SHOWED THAT THE CUSTOMER'S DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. TO ENSURE THAT AN EXAM REPORT WAS COMPLETED, THE EXAM WAS RE-READ BY THE PHYSICIAN. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE'S INVESTIGATION SHOWED THE DATABASE WAS UPDATED BUT THE REPORT DID NOT UPLOAD TO THE SERVER. THE AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. THE EXAM WAS THEN RE-READ BY THE PHYSICIAN. LOGS WERE UNAVAILABLE FOR REVIEW AS THIS ISSUE WASDISCOVERED AFTER THE LOGS HAD BEEN TRUNCATED ON THE STATION. NO OTHER REVIEW/INVESTIGATION IS AVAILABLE FOR DETERMINING THE ROOT CAUSE OF THIS CUSTOMER'S ALLEGATION. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, MERGE HEALTHCARE INTERNALLY FOUND THAT WHEN ADDING AN ADDITIONAL BIOPSY WITHOUT ATYPICAL HYPERPLASIA TO A PATIENT THAT PREVIOUSLY HAD A BIOPSY WITH ATYPICAL HYPERPLASIA, IT WILL OVERWRITE THE CONDITION AND REMOVE THE ATYPICAL HYPERPLASIA COMPLETELY. THIS ISSURE COULD RESULT IN AN INCORRECT GAIL RISK CALCULATION, POTENTIALLY RESULTING IN A DELAY IN TREATMENT AND/OR INCORRECT TREATMENT FOR A PATIENT. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ISSUE FOUND INTERNALLY AND RECENTLY SUBMITTED A CORRECTION TO THE FDA ON 25JUL2017. (2183926-07/25/2017-011-C). ONLY CUSTOMERS WITH A LICENSE FOR MAMMO TRACKING ARE IMPACTED IN RELEASES 10.0.6 AND 10.0.7. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE TECHNICAL SUPPORT WITH QUESTIONS ABOUT HOW THE MERGE UNITY BREAST RISK ASSESSMENT CALCULATES THE RISK. THE CUSTOMER REPORTED THAT A PATIENT HAD TWO PREVIOUS NEGATIVE EXAMS THAT AND THE PATIENT'S FAMILY HAS NO HISTORY OF BREAST CANCER. HOWEVER, WHEN PERFORMING THE RISK ASSESSMENT IN MERGE UNITY FOR THIS PARTICULAR PATIENT, THE LIFETIME RISK LEVEL WAS NOT WHAT THE CUSTOMER EXPECTED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. A MISCALCULATION REGARDING THE POTENTIAL RISK OF BREAST CANCER COULD RESULT IN A PATIENT NOT RECEIVING THE RECOMMENDED TREATMENT AND/OR RECEIVING TREATMENT THAT IS NOT NEEDED. (B)(4). Manufacturer Narrative: UPON RESEARCHING THE CUSTOMER'S QUESTION REGARDING THE CALCULATIONS OF RISK, MERGE HEALTHCARE IDENTIFIED AN ISSUE WHERE ATYPICAL HYPERPLASIA MAY NOT BE FACTORED INTO THE BREAST RISK CALCULATION. A MISCALCULATION REGARDING POTENTIAL RISK COULD RESULT IN A PATIENT NOT RECEIVING THE RECOMMENDED TREATMENT AND/OR RECEIVING TREATMENT THAT IS NOT NEEDED. MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE ALLEGATION FROM THE CUSTOMER AND IS PLANNING A CORRECTION (2182926-07/21/2017-010-C). . | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN OLDER EXAM WAS FOUND WITH A MISSING REPORT. THERE WAS NO PDF, AND IT WOULD NOT OPEN WITH DICTATION. THERE ARE NO ERRORS, THE OPEN OPTION ON THE DICTATION BOX IS SIMPLY GRAYED OUT. THE REPORT'S AUDIT TRAIL SHOWED THE EXAM WAS READ, APPROVED AND SAVED. UNITY TECHNICAL SUPPORT OBTAINED THE REPORT ID FROM THE EXAM PROPERTIES, AND SEARCHED FOR IT IN THE REPORTS FOLDER. THE FOLDER WAS THERE, BUT THERE WERE NO CONTENTS IN IT (DICTATION, DOCUMENT, ETC). UNITY TECHNICAL SUPPORT CREATED A BLANK DOCUMENT AND NAMED IT TO MATCH THE REPORT ID. IT WAS CONFIRMED BY THE CUSTOMER THAT THE BLANK REPORT CAN BE OPENED WITH DICTATION. THE TRANSCRIPTIONIST WAS ABLE TO RE-CREATE THE REPORT UTILIZING THE ORIGINAL VOICE FILE OF THE REPORT. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT NUMBER (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. AT THIS FACILITY, HOWEVER, THE PHYSICIAN READS THE EXAM AND AN AUDIO FILE IS CREATED FOR A TRANSCRIPTIONIST TO TYPE. DUE TO THE TIME LAPSE BETWEEN THE EXAM DATE AND THE DATE THAT THE REPORT WAS FOUND TO BE NOT PRESENT, DOES NOT ALLOW FOR LOGS TO BE REVIEWED FOR CAUSE. THE ORIGINAL AUDIO FILE WAS AVAILABLE AND THE REPORT WAS RE-CREATED BY THE TRANSCRIPTIONIST. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A CUSTOMER CONTACTED MERGE UNITY TECHNICAL SUPPORT AND ALLEGED THAT AN EXAM REPORT WAS SAVED BUT WHEN ATTEMPTING TO OPEN THE REPORT, A MESSAGE IS DISPLAYED ON THE SCREEN THAT SHOWS AN ERROR. ERROR IS 'OBJECT NOT FOUND'. MERGE SUPPORT INVESTIGATED THE CUSTOMERS ALLEGATION AND DETERMINED THAT THERE WAS A SAVE EVENT FOR THE CUSTOMERS' ALLEGED REPORT. MERGE UNITY TECHNICAL SUPPORT WAS UNABLE TO FIND THE SAVED REPORT FOR THE CUSTOMER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. WHEN PROMPTED TO VIEW THE REPORT, THE SYSTEM DISPLAYS A MESSAGE OF "NO OBJECT FOUND." IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(6) 2017, A PACS USER INFORMED MERGE UNITY SUPPORT THAT A RADIOLOGIST WAS COMPLETING A REPORT IN MERGE REPORTING AND RECEIVED A SERIES OF ERROR MESSAGES. THE ERRORS WERE OBSERVED BY THE PHYSICIAN. THE SITE ATTEMPTED TO RESOLVE THIS ISSUE INTERNALLY. AFTER THE CUSTOMER CONTACTED MERGE SUPPORT AN INVESTIGATION SHOWED THAT THE REPORT FAILED TO UPLOAD TO THE IMAGE SERVER. THIS REQUIRED THE PHYSICIAN TO RE-READ THE EXAM FOR RESOLUTION. WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. HOWEVER, THE CUSTOMER DID NOT ALLEGE ANY HARM, SERIOUS INJURY OR DEATH AS A RESULT OF THIS ISSUE. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. MERGE HEALTHCARE IS CONTINUING TO FURTHER INVESTIGATE THE ALLEGATION FROM THE CUSTOMER TO DETERMINE IF ANY CORRECTIONS OR CORRECTIVE ACTIONS ARE NECESSARY. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER CONTACT REPORTED ON (B)(6) 2019 THERE WERE FIVE EXAMS WITH CORRUPTED REPORTS FROM (B)(6) 2019 FROM TWO DIFFERENT READING STATIONS. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES WERE AVAILABLE, A CORRUPTED REPORT COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY LOGS FOUND NO ERRORS INDICATING THE FAILURE OF THE REPORTS TO UPLOAD, HOWEVER IT WAS NOTED THAT UNIVERSAL MANAGER (UM) CRASHED. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THE EXAMS IN THE FAILED RESULTS OUT QUEUE. THE EXAMS WERE ALL READ AND APPROVED IN APPROVED (APD) STATUS WITH REPORT ID'S. THE REPORTS FOLDER WAS MANUALLY CHECKED, HOWEVER, NONE OF THE REPORTS HAD REPORT FOLDERS. THE REPORT STATUSES WERE UPDATED TO DICTATED (DID) AND RESENT TO THE USER RE-DICTATION. UNITY SUPPORT CONFIRMED THAT ALL FIVE REPORTS WERE RE-READ SUCCESSFULLY. NO FURTHER INVESTIGATION WILL BE PERFORMED. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER ALLEGED AN EXAM WAS READ ON (B)(6) 2019 BUT THE REPORT WAS NOT VIEWABLE. MERGE HEALTHCARE TECHNICAL SUPPORT CONNECTED TO THE SITE AND FOUND THAT THERE WAS A REPORT SAVE EVENT BUT THE REPORT DID NOT UPLOAD. THE REPORT WAS LISTED IN THE AUDIT TRAIL, HOWEVER, NO WORD DOCUMENT/REPORT WAS LISTED IN THE REPORT DIRECTORY OR THE EXAM PROPERTIES. TECH SUPPORT REGENERATED THE REPORT TEMPLATE FOR READING AGAIN. THE EXAM WAS RE-READ AND THE REPORT WAS GENERATED AND APPROVED BY THE DOCTOR ON THE SAME DAY (B)(6) 2019. ONLY THIS EXAM WAS AFFECTED; OTHER EXAMS READ FROM THE SAME STATION HAD NO ISSUES. THE REPORTED ISSUE WAS RESOLVED WITH THE RE-READING OF THE EXAM AND GENERATION OF THE REPORT. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES WERE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. REFERENCE COMPLAINT (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. IT WAS NOT POSSIBLE TO DETERMINE IF THE MESSAGE DISPLAYED ON THE SCREEN FOR THE PHYSICIAN WHEN THE REPORT WAS FIRST READ AND SAVED. THE MERGE HEALTHCARE UNITY INVESTIGATION SHOWED THE REPORT DID NOT UPLOAD SUCCESSFULLY, AND WILL HAVE TO BE RECREATED. TO RESOLVE THIS ISSUE, THE EXAM WAS RE-READ BY THE DOCTOR. NO OTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. THE CUSTOMER CONTACT REPORTED ON (B)(6) 2018 THAT A "NO OBJECT FOUND" ERROR WAS NOTED WHEN TRYING TO VIEW A REPORT FOR A SPECIFIC MAMMOGRAM EXAM, SUBSEQUENTLY, THE STATION KEPT HAVING TO BE RESTARTED. THERE WAS NO REPORTED ADVERSE EVENT TO THE PATIENT. HOWEVER, WHILE IMAGES ARE AVAILABLE, A FINALIZED REPORT NOT BEING AVAILABLE FOR SUBSEQUENT REVIEW BY THE PATIENT'S PHYSICIAN COULD POTENTIALLY RESULT IN A DELAY IN CARE THAT COULD LEAD TO HARM. (B)(4). Manufacturer Narrative: MERGE UNITY PACS IS DESIGNED TO DISPLAY A MESSAGE TO THE USER WHEN AN EXAM REPORT FAILS TO UPLOAD TO THE IMAGE SERVER. IN THIS INSTANCE, THERE WAS NO EVIDENCE THAT THE PHYSICIAN RECEIVED THE MESSAGE WARNING THAT THE REPORT DID NOT UPLOAD. TROUBLESHOOTING FROM UNITY SUPPORT FOUND THAT THERE WAS A REPORTED GENERATED, HOWEVER THE DOCUMENT WAS NOT UPLOADED TO I:\ DRIVE. A NEW WORD DOCUMENT WAS CREATED AND RENAMED REPORTED. THE EXAM WAS RE-READ AND CONFIRMED WITH THE CUSTOMER THAT THE REPORT WAS PRESENT. NO FURTHER REVIEW/INVESTIGATION WILL BE PERFORMED FOR CAUSE. | -fail to process medical imaging -fail to interpret images -image orientation incorrect | -use of device issue -instruction for use issue -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS IS A MEDICAL IMAGE, AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON 11/05/2020, A CUSTOMER CONTACTED MERGE HEALTHCARE TO INFORM US THAT THEY WERE UNABLE TO VIEW AN EXAM REPORT FROM A RECENT EXAM. MERGE HEALTHCARE DETERMINED UPON INVESTIGATION THAT THE REPORT SAVED BUT FAILED TO UPLOAD. LOGS WERE NOT AVAILABLE FOR REVIEW. A BLANK REPORT DOCUMENT WAS MANUALLY PLACED IN THE FOLDER AND THE REPORT WAS RE-DICTATED, AND SUBSEQUENTLY VIEWABLE AS EXPECTED. THIS SITE HAS REGULAR AND PERSISTENT NETWORK ISSUES THAT AFFECT REPORT UPLOADS. NO FURTHER ACTION IS NECESSARY AT THIS TIME. THIS HAS THE POTENTIAL TO DELAY PATIENT TREATMENT, AND/OR DIAGNOSIS. THERE HAVE BEEN NO REPORTS OF PATIENT INJURY, OR HARM AS A RESULT OF THIS ISSUE. REFERENCE COMPLAINT: (B)(4). | -fail to interpret images -image orientation incorrect -inadequate image types | -use of device issue -contamination during use -filling problem | Viewer
| FDA | Event Description: MERGE UNITY PACS, A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT IS USED FOR VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. ON (B)(4) 2017, MERGE WAS NOTIFIED THAT ON AN UNSPECIFIED DATE OVER THE WEEKEND OF (B)(6), AN ERROR WAS RECEIVED WHEN ATTEMPTING TO BURN IMAGES ONTO A COMPACT DISC (CD). THE CUSTOMER WAS NOT ABLE TO BURN IMAGES OF A COMPUTED TOMOGRAPHY (CT) OF THE HEAD/NECK ONTO A CD FROM THEIR RIMAGE BURNER, WHICH WAS NEEDED FOR A PATIENT BEING AIR-LIFTED TO ANOTHER FACILITY. A PATIENT WAS BEING EVALUATED IN THE EMERGENCY DEPARTMENT FOR AN ATTEMPTED SUICIDE. IT WAS REPORTED THAT THE PATIENT WAS UNRESPONSIVE AND INTUBATED. THE CUSTOMER REPORTED THAT THE PATIENT WAS AIR LIFTED TO A PEDIATRIC TRAUMA CENTER IN LUBBOCK, TX IN GRAVE CONDITION. AS A RESULT OF THE CUSTOMER NOT BEING ABLE TO BURN IMAGES ONTO THE CD, THE CUSTOMER REPORTED THAT ANOTHER CT HAD TO BE PERFORMED AT THE HOSPITAL RECEIVING THE PATIENT. THERE WERE NO KNOWN ADVERSE EVENTS TO THE PATIENT. HOWEVER, HAVING TO RE-DO A CT OF THE HEAD/NECK CAN DELAY PATIENT CARE AND/OR DIAGNOSIS. IN ADDITION, THIS CAN EXPOSE THE PATIENT TO ADDITIONAL UNNECESSARY RADIATION. (B)(4). Manufacturer Narrative: TESTING AND INVESTIGATION BY MERGE SUPPORT DETERMINED THAT THE ERROR "DEVICE: AUTOLOADER 1, HID-0 ERROR CODE: 363 FAILED TO START LABEL RE-TRANSFER ON EVEREST PRINTER; PRINTER SENSE CODE = 254 (UNDEFINED SENSE CODE)" WAS LIKELY RELATED TO THE LABEL/RIBBON ASSEMBLY. THE CUSTOMER FOUND THAT THE COLOR RIBBON (A PART OF THE LABEL/RIBBON ASSEMBLY) WAS OUT. ONCE THE COLOR RIBBON WAS CHANGED, THE ISSUE WAS RESOLVED. A REVIEW OF THE MERGE UNITY USER GUIDES DID NOT REVEAL ANY TROUBLESHOOTING OR OTHER ANOMALIES/ISSUES REGARDING BURNING IMAGES ONTO A CD. A REVIEW OF THE COMPLAINT HISTORY INDICATES THERE IS NO SIGNIFICANT TREND. TRENDING WILL CONTINUE TO BE MONITORED. | -fail to process medical imaging -fail to interpret images -inadequate information | -use of device issue -filling problem -contamination during use | Viewer
| FDA | Event Description: NO KNOWN EVENT OCCURRED. PROBLEM: THE IMPORT OF ENHANCED ULTRASOUND FORMATTED DATA (NEMA DICOM SUPPLEMENT 43: STORAGE OF 3D ULTRASOUND DATA, VERSION: FINAL TEXT - (B)(6) 2009), CAN RESULT IN INCORRECT DISPLAY OF IMAGE ORIENTATION. THE REASON WAS AN INCORRECT UNDERSTANDING OF THE DATA FORMAT. THIS FAULT DOES NOT RESIDE IN THE TRUE 3D RENDER ENGINE, BUT RATHER IS LOCALIZED TO THE DICOM INPUT TASK OF THE APPLICATION PROBLEM: THE IMPORT OF PHILIPS 3DDCM FORMATTED ULTRASOUND IMAGES CAN RESULT IN INCORRECT DISPLAY OF IMAGE ORIENTATION. THE REASON WAS AN INCORRECT UNDERSTANDING OF THE DATA FORMAT. THIS FAULT DOES NOT RESIDE IN THE TRUE 3D RENDER ENGINE, BUT RATHER IS LOCALIZED TO THE DICOM INPUT TASK OF THE APPLICATION. ISSUE WAS IDENTIFIED ON RELEASE VERSIONS: 1.6.1 AND 1.6.2. | -illegible data -incorrect device orientation -inadequate message display | -data issue -data back-up problem -medical gas supply problem | Viewer
| FDA | Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. CONTEXT: INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM WITH THE MAGNIFICATION GLASS TOOL WHICH, IN SOME INSTANCES, DISPLAYS A DIFFERENT IMAGE THAN THE ONE CURRENTLY BEING SHOWN IN THE SELECTED VIEWPORT. IT APPEARS THAT THE TOOL IS DISPLAYING A MAGNIFIED IMAGE FOR A DIFFERENT SLICE OF THE SERIES. THIS COULD POTENTIALLY DELAY A DIAGNOSIS. THIS PROBLEM WAS REPORTED BY AN INTELERAD EMPLOYEE WHILE AT A CLIENT SITE. INVESTIGATION: (B)(4). THE SORTING ORDER OF AN IMAGE STACK IS NOT TAKEN INTO ACCOUNT BY THE MAGNIFICATION GLASS TOOL. IF THE USER IS USING REVERSE ORDERING (EITHER MANUALLY SELECTED OR THROUGH A PREFERENCE IN THE CASE OF TOMOSYNTHESIS) AND THEN USES THE MAGNIFICATION GLASS TOOL, THE IMAGE SHOWN IN THE TOOL WILL NOT BE THE ONE IN THE CURRENT VIEWPORT BUT RATHER THE CORRESPONDING POSITION FROM THE OPPOSITE END OF THE STACK. THIS DEFECT APPLIES TO ANY MODALITY WITH A STACK OF IMAGES, EX. TOMOSYNTHESIS, CT. THE DISCREPANCY IS HIGHLY DETECTABLE BY USER AND IS EXTREMELY UNLIKELY TO LEAD TO HARM TO THE PATIENT. DESCRIPTION OF HAZARDOUS SITUATION IN USE OF THE DEVICE: INCORRECT IMAGE BEING SHOWN WHILE USING THE MAGNIFICATION TOOL, WHICH COULD CAUSE A LOW POSSIBILITY OF DELAY IN TREATMENT. THE INCORRECT IMAGES DISPLAYED ARE HIGHLY EVIDENT TO RADIOLOGISTS. THIS PROBLEM IS NOT CURRENTLY DOCUMENTED AS A KNOWN RISK. A MITIGATION IS CURRENTLY IN PLACE TO REDUCE THIS RISK AND TO PREVENT RISK TO PATIENT. THE RISK CAN BE FURTHER REDUCED THROUGH ADDITIONAL INTERNAL SOFTWARE CHECKS TO PREVENT INCORRECT IMAGE DISPLAY WHEN INCONSISTENCIES ARE DETECTED. PROBABILITY OF RISK: REMOTE (2) SEVERITY OF RISK: IMPROBABLE (1) DETECTABILITY (USER AWARENESS. IF DEVICE FAILURE OCCURS, IS IT EASILY RECOGNIZED BY USER): HIGHLY DETECTABLE BY USER. RISK VS. BENEFIT COMPARISON: BENEFIT OF AN INCREASE IN AVERAGE TIME SAVINGS, DIAGNOSTIC ACCURACY, DATA INTEGRITY AND IMAGE ACCESSIBILITY OUTWEIGHS THE RISK OF LOW PROBABILITY EVENTS LEADING TO DELAYED DIAGNOSIS. THE PROBLEM IS ONLY CAUSED IN MULTI IMAGE SITUATIONS WHERE THE STACK ORDER IS REVERSED AND THE MAGNIFICATION GLASS TOOL IS USED. THE SOFTWARE FAILURE IS VERY EVIDENT AND EXTREMELY UNLIKELY TO LEAD TO HARM TO PATIENT. THE FEATURE HAS BEEN IN USE FOR EIGHT YEARS WITH ONLY A SINGLE REPORTED PROBLEM, WITH NO INDICATION OF REAL HARM. CORRECTIVE ACTIONS: INTELERAD WILL: 1) IDENTIFY ALL CLIENTS AFFECTED (19 SEPTEMBER 2017). 2) INFORM ALL CONCERNED CLIENTS OF THE ISSUE (SEPTEMBER 2017). 3) PROVIDE ALL CLIENTS WITH A SOFTWARE UPGRADE (INSTALLATION DATE AS AGREED WITH CLIENTS). ESTIMATED COMPLETION DATE OF THE PROPOSED RECALL ACTION IS JUNE 2018. | -shattered glass -fail to resolve problem -device malfunction | -use of device issue -device operates differently than expected -improper or incorrect procedure or method | Viewer
| FDA | Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. FIVE RELATED CUSTOMER COMPLAINTS ((B)(4)) WERE REGISTERED AND INVESTIGATED FROM OCT 2014 TO JUNE 2015 FOR RELATED ISSUES. THE INVESTIGATION OF THESE COMPLAINTS LED TO THE DISCOVERY OF ISSUES WITH IMAGE INGESTION. BUGS IDENTIFIED: (B)(4) LDS: LDS: FAILED TO INGEST IMAGE WITH OVERLAYS BELONGING TO 60XX GROUP HAVING ALPHANUMERIC ID (IE, 600A) (B)(4) LDS: ROWS STUCK IN INCOMING_FILES DUE TO FILES MISSING IN THE INCOMING FILES FOLDER. INVESTIGATION: INVESTIGATION OF THE LOGS IN (B)(4) REVEALED THAT SOME OF THE EXPECTED OVERLAY FILES WERE NOT FOUND ON THE FILE SYSTEM; THIS LED TO DISCOVERY OF THE MISSING FILES LEFT BEHIND AT AN EARLIER STAGE OF IMAGE INGESTION. ROOT CAUSE OF THE PROBLEM: DUE TO THE ASYNCHRONOUS NATURE OF IMAGE INGESTION PROCESSING, A FAILURE AT THE EARLY STAGES OF THE INGESTION PROCESS CAN LEAD TO IMAGES NOT BEING VISIBLE TO THE END USER. THERE IS NO INDICATION TO THE USER THAT THIS HAS OCCURRED. REASON FOR POSSIBLE PATIENT SAFETY: THE RADIOLOGIST MAY PROVIDE A DIAGNOSIS WITHOUT BEING AWARE THAT IMAGES ARE MISSING. DESIGN MITIGATION TO REDUCE THE RISK: PROPOSED NEW MITIGATION: THE PACS WILL REPORT PROBLEMS INGESTING IMAGES AND INDICATE SUCH ISSUES TO THE USER ((B)(4)). THIS MITIGATION WILL BE INCLUDED IN A COMING SOFTWARE RELEASE. (B)(4): LDS: EXPOSE BROKEN SERIES TO USER (WARN ABOUT ADDITIONAL IMAGES PENDING INGESTION WHEN VIEWING IMAGES). DETECTABILITY OF PROBLEM: UNTIL THE SOFTWARE UPDATE THAT WILL INCLUDE THE PROPOSED MITIGATION IS INSTALLED ON THE CLIENT'S SYSTEM, THE USER MAY NOT BE AWARE THAT SOME IMAGES MAY BE MISSING. HOWEVER, IF THE CLIENT'S SYSTEM HAS OTHER MECHANISMS TO VERIFY IMAGE COUNT BETWEEN SOURCE AND DESTINATION, THEN THE USER MIGHT BE ABLE TO KNOW ABOUT THE DISCREPANCY. Manufacturer Narrative: IN ORDER TO CORRECT THIS ISSUE, INTELERAD WILL: PROVIDE THE AFFECTED CLIENTS WITH A LIST OF STUDIES THAT WERE AFFECTED AND POSSIBLY RECOVER THE MISSING IMAGES. DEPLOY A SOFTWARE UPDATE THAT INCLUDES A FIX AND MITIGATION TO ALL CLIENTS WHERE THE ISSUE COULD POTENTIALLY OCCUR. THIS WILL BE DONE TENTATIVELY. PERIODICALLY AUDIT ALL AFFECTED CLIENT, ANY NEW OCCURRENCES OF THIS ISSUE UNTIL THE MITIGATION (SOFTWARE UPDATE) IS IN PLACE. THE SOFTWARE UPDATE WILL INCLUDE ONLY BUG FIXES - YOUR WORKFLOW WILL NOT BE AFFECTED. THE UPDATE WILL RESULT IN ABOUT 5 MINUTES OF DOWNTIME PER SERVER AS WE PERFORM A FULL SERVICE RESTART. IMPACT ON USERS WILL BE MINIMAL, IF ANY. CORRECTIVE ACTION CALENDAR: FOR ALL CLIENTS WITH PACS VERSION 4-5-1 OR HIGHER, THE SOFTWARE UPDATE WILL BE DEPLOYED WITHIN A TIMEFRAME TO BE AGREED WITH CLIENTS (COMPLETION ESTIMATED BY END OF MARCH 2016). | -patient incident -patient experienced adverse event -improper device position | -use of device issue -Protective Measure Issue -user used incorrect product for intended use | Viewer
| FDA | Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. ONE CUSTOMER (B)(4) WAS REGISTERED AND INVESTIGATED. BUGS IDENTIFIED: (B)(4): (REVERT OF (B)(4)) IV: RADIOGRAPHIC SCALING FACTOR DERIVED FROM DISTANCE TO PATIENT AND DETECTOR IS BEING APPLIED TO NON-PROJECTION RADIOGRAPHY MODALITIES (B)(4): IV (MEASUREMENT): MEASUREMENTS ARE NOT CALIBRATED WHEN DICOM IS MISSING MAG FACTOR, BUT HAS DISTANCE TO PATIENT AND DETECTOR PLATE. INVESTIGATION: THERE IS INDEED A PROBLEM WITH THE MEASUREMENTS AFTER THE FIX INTRODUCED IN (B)(4). IV NOW USES THE RATIO OF THE DISTANCE FROM THE SOURCE TO PATIENT OVER DISTANCE FROM SOURCE TO DETECTOR TO COMPUTE A SCALING FACTOR TO APPLY TO MEASUREMENTS, BUT THAT SCALING SHOULD ONLY BE APPLIED WHEN DEALING WITH A PROJECTION RADIOGRAPHY MODALITY (CR, DX, MG, PX, RF, OR XA). UNFORTUNATELY, IT IS NOW BEING APPLIED TO ALL MODALITIES. THERE IS A GUARD IN THE CODE THAT PREVENTS COMMENTS ABOUT THE SCALING FACTOR TO BE DISPLAYED ON MODALITIES THAT ARE NOT PROJECTION RADIOGRAPHY, BUT THE SCALING FACTOR ITSELF IS APPLIED IN ALL CASES, WHICH IS WRONG. ROOT CAUSE OF THE PROBLEM: SINCE THE INTRODUCTION OF A FIX TO THE MEASUREMENTS ALGORITHM ((B)(4)), INTELEVIEWER NOW USES THE RATIO OF THE DISTANCE FROM THE SOURCE TO PATIENT OVER DISTANCE FROM SOURCE TO DETECTOR TO COMPUTE A SCALING FACTOR TO APPLY TO MEASUREMENTS WHEN THE ESTIMATED RADIOGRAPHIC MAGNIFICATION FACTOR (DICOM (0018,1114)) IS NOT PRESENT. THAT SCALING SHOULD ONLY BE APPLIED WHEN DEALING WITH A PROJECTION RADIOGRAPHY MODALITY (CR, DX, MG, PX, RF, OR XA), HOWEVER, IT ENDED UP BEING APPLIED TO OTHER MODALITIES, SUCH AS CT, GENERATING ERRONEOUS MEASUREMENTS. REASON FOR POSSIBLE PATIENT SAFETY: THE RADIOLOGIST COULD BASE THEIR DIAGNOSIS ON INCORRECT ANATOMICAL MEASUREMENTS. RISK ANALYSIS: THIS PROBLEM IS DOCUMENTED AS A KNOWN RISK AND IS MITIGATED BY DESIGN: CONSISTENT USE OF MAPPING FUNCTIONS TO PATIENT ANATOMICAL SPACE AND CALIBRATED UNITS BEFORE PERFORMING COORDINATE AND INTENSITY CALCULATIONS. FOR CLIENTS USING THIS SPECIFIC SOFTWARE VERSION, THIS PROBLEM COULD WILL OCCUR WITH APPROXIMATELY 50% OF ALL CT STUDIES. RISK TO THE PATIENT COULD OCCUR ONLY IF THE MODALITY IS CT AND IF THE DISCREPANCY OF THE MEASUREMENT IS SMALL. IN ORDER TO CORRECT THIS ISSUE, INTELERAD WILL: IMMEDIATELY: NOTIFY TO CLIENTS TO STOP USING THE AFFECTED VERSION AND TO REVERT TO AN ANTERIOR OR ULTERIOR SOFTWARE VERSION. IMMEDIATELY: REMOVE THE AFFECTED SOFTWARE VERSION FROM USE - MAKE IT UNAVAILABLE ON THE CLIENTS' SERVERS. IMMEDIATELY: FLAG THE VERSION INTERNALLY AS A NO DEPLOY VERSION. FIX THE AFFECTED VERSION. RUN SERVER SIDE SCRIPTS ON CLIENTS' SYSTEMS TO IDENTIFY STUDIES READ WITH THE AFFECTED VERSION. SEND RECALL NOTICE TO AFFECTED CLIENTS ( 4 ). COMMUNICATE TO CLIENTS THE LIST OF EXAMS THAT WERE READ USING THE DEFECTIVE VERSION. RUN ADDITIONAL SERVER-SIDE SCRIPTS TO ENSURE THAT NO STUDIES WERE READ WITH THE AFFECTED VERSION AFTER CLIENTS WERE NOTIFIED. ESTIMATED DATE TO COMPLETE CORRECTIVE ACTION: END OF FEBRUARY 2016 | -patient experienced adverse event -device alarmed -device malfunction | -Device Operates Differently than Expected -no audible alarm -Mechanical Issue | Viewer
| FDA | Event Description: NO PATIENT INCIDENT REPORTED AT THIS MOMENT, THIS ISSUE IS DOCUMENTED AS A PRECAUTION. THE ISSUE WAS REPORTED BY OUR CLIENT. THEY PROVIDE TELERADIOLOGY SERVICES TO VARIOUS HOSPITALS. THEIR PROTOCOL FOR ACCEPTING STAT CASES FROM THESE HOSPITALS REQUIRES THAT THE SENDING FACILITY TRANSMIT A SEPARATE TEXT MESSAGE THAT IDENTIFIES EACH STAT CASE BY ACCESSION NO, TOGETHER WITH THE STUDY DESCRIPTION AND THE EXPECTED NUMBER OF IMAGES TO BE TRANSMITTED. FURTHER, BASED ON THIS PROTOCOL, THE CLIENT'S ADMINISTRATIVE STAFF WHO CONTROL THE WORKFLOW, CONFIRMS THE IMAGE COUNT TRANSMITTED IS THE SAME AS THE IMAGE COUNT AT THE READING CTR (IN INTELEPACS) BEFORE ACCEPTING THE CASE AND ASSIGNING TO A RADIOLOGIST FOR READING. DESCRIPTION OF PROBLEM: IMAGES SUCCESSFULLY TRANSMITTED TO INTELEPACS ARE NOT AVAILABLE FOR READING IN ALL STUDIES, AND THE READER IS UNAWARE OF THE POSSIBILITY OF MISSING IMAGES. THE PROBLEM MANIFESTS UNDER VERY SPECIFIC CIRCUMSTANCES, LIMITED TO A FEW CLIENT SITES AS A POSSIBLE PATIENT SAFETY ISSUE. THE POSSIBLE PATIENT SAFETY ARISES WHEN RECEIVING A SERIES AND CAN MANIFEST IN TWO SCENARIOS: FIRST: FOR THE 4 CLIENTS THAT EMPLOY INGEST SERVERS TO PROVIDE TEMPORARY STORE-AND-FORWARD CAPABILITIES IN THEIR ARCHITECTURE, IN SOME CONDITIONS, THESE SERVERS FAIL TO DELIVER ALL RECEIVED IMAGES TO THE PERMANENT STORAGE SERVERS AND NO RECORD OF THESE IMAGES PERSIST IN THE MAIN ARCHIVE. AS THESE INGEST SERVERS ARE NOT CONFIGURED AS VALID IMAGE LOCATIONS FOR ACCESS BY DIAGNOSTIC VIEWERS, THE STUDY CAN BE DISPLAYED IN INTELEPACS WITH LESS THAN FULL SET OF TRANSMITTED IMAGES. SECOND: FOR THE REMAINING CLIENT ARCHITECTURES THAT DO NOT EMPLOY TEMPORARY INGEST SERVERS, THERE IS STILL A VERY LOW PROBABILITY OF FAILURE (FAR LESS LIKELY THAN THE FIRST SCENARIO) FOR THE DEVICE TO PROPERLY REPLICATE AND REGISTER IMAGES WHEN TRANSFERRING THEM TO ANOTHER SERVER. IF THE SECOND COPY IS A VALID IMAGE LOCATION, IT IS POSSIBLE FOR A SOME IMAGES TO BE MISSING WHEN VIEWING. WE WILL: IMMEDIATELY VERIFY AND APPLY A BUG FIX TO ALL THE CLIENTS WHERE THIS ISSUE CAN OCCUR; APPLY AUDIT SCRIPTS TO IDENTIFY ANY OCCURRENCE OF THE ISSUE AND ALSO VALIDATE THE BUG FIX. CORRECTIVE ACTION CALENDAR: FOR THE THREE CLIENTS USING A CONFIGURATION WITH INGEST SERVERS AS TEMPORARY STORAGE, AN IMMEDIATE UPDATE PROCESS WAS UNDERTAKEN VIA THE CLIENT SPECIFIC RELEASE PROCESS (COMPLETED AS OF 04/23/2015) TO AVOID THE POSSIBILITY OF DATA LOSS. FOR THE REMAINING CLIENTS, THE NORMAL FULL RELEASE UPDATE (FRU) WILL BE EXPEDITED BUT UNDERGO THE USUAL TESTING AND RELEASE PROCESS. Manufacturer Narrative: AFFECTED SOFTWARE VERSIONS: MODELS: 4.7.1 R27 AND UP, 4.8.1 R20 AND UP, 4.9.1 R17 AND UP, 4.10.1 R9 AND UP, 4.11.1 R3 AND UP. | -patient experienced adverse event -patient incident -fail to provide indicator | -Failure of device to Self-Test -Device Operates Differently than Expected -Protective Measure Issue | Viewer
| FDA | Event Description: OCTOBER 5, 2016 SIEMENS BECAME AWARE OF DATA LOSS AT THE CUSTOMER'S SITE IN (B)(6). THE SHORT TERM STORAGE (STS) ATTACHED TO THE SIEMENS SYNGO.PLAZA SYSTEM GOT CORRUPTED DUE TO A HANDLING ERROR DURING ROUTINE MAINTENANCE OF THE HARD DISC UNIT (RAID) ON THE STS. DURING EXCHANGE OF THE HARD DISC, THE CONFIGURATION WAS LOST; THEREFORE DATA STORED ON THE RAID BECAME NO LONGER AVAILABLE ON THE STS. HOWEVER, NOT ALL OF THE DATA WAS COMPLETELY LOST DUE TO REDUNDANCY IN THE SETUP OF THE SYNGO.PLAZA ENVIRONMENT AT THIS SITE: ú A LARGE AMOUNT OF DATA LOST ON THE STS HAS BEEN ARCHIVED ON THE LONG TERM STORAGE (LTS) ATTACHED TO THE SYNGO.PLAZA UNIT. THUS THIS DATA IS STILL AVAILABLE TO THE USER. ú A SEPARATE SUBSET OF THE DATA, WHICH WAS PREVIOUSLY STORED ON THE STS PRIOR TO THE INCIDENT, WAS FOUND ON THE SCANNERS WHERE IT HAD BEEN INITIALLY CREATED. THIS DATA WAS RE-SENT TO THE STS AND IS AVAILABLE TO THE USER AS WELL. ú IN TOTAL 64 SERIES CONTAINING 43 STUDIES WERE PERMANENTLY LOST. THIS DATA WAS OBTAINED BETWEEN APRIL AND AUGUST 2016. SIEMENS SYNGO.PLAZA SYSTEM DID NOT CONTRIBUTE TO THE DATA LOSS. THERE ARE NO INJURIES AND NO NEGATIVE IMPACT HAS BEEN REPORTED BY THE CUSTOMER. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. NO PATIENT RESCANS WERE REQUIRED AS THE LOST STUDIES WERE OBTAINED PRIOR TO AUGUST 31, 2016, THEREFORE THE DATA HAD BEEN ALREADY ASSESSED AND USED FOR CLINICAL DIAGNOSIS PRIOR TO THE DATE OF THE INCIDENT ((B)(6) 2016). HOWEVER, SIEMENS CONSIDERS THIS INCIDENT AN ADVERSE EVENT AS THE DATA OLDER THAN 4 WEEKS MAY STILL HAVE CLINICAL RELEVANCE AND BE USED BY RADIOLOGISTS FOR COMPARISON TO IMAGES RECEIVED DURING NEWER PATIENT EXAMINATIONS. THE CLINICAL IMPACT OF THIS INHIBITED COMPARISON IS ASSESSED AS LOW. (B)(6). | -patient experienced adverse event -device damaged -device misoriented | -device maintenance issue -power conditioning issue -loss of power | Viewer
| FDA | Event Description: OFFICEPACS IS USED FOR THE MANIPULATION AND DISPLAYING OF MEDICAL IMAGES. OFFICEPACS IS DESIGNED AND MARKETED FOR SOFT COPY READING AND STORAGE OF STUDIES PRODUCED BY DIGITAL MODALITIES. ON (B)(6) 2013 A CUSTOMER REPORTED MISSING IMAGES FROM AN OLD STUDY. MERGE CUSTOMER SUPPORT CONDUCTED A SEARCH OF THE SERVER AND WAS UNABLE TO LOCATED THE MISSING STUDIES. THIS MAY POTENTIALLY RESULT IN DELAY IN PATIENT CARE. THERE WAS NO REPORT OF ADVERSE PATIENT IMPACT. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -fail to process medical imaging -inadequate image manipulation -inadequate graphic display | -use of device issue -data back-up problem -capturing issue | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT A MONITOR IMAGE BECAME PURPLISH DURING THE LAPAROSCOPIC CHOLECYSTECTOMY WITH THE SUBJECT DEVICE AND THE LAPARO-THORACO VIDEOSCOPE LTF-VH. THE FACILITY CONTINUED TO USE THE SUBJECT DEVICE AND THE LTF-VH, AND THE PROCEDURE WAS COMPLETED. AFTER THE PROCEDURE, BILE SPILLAGE OCCURRED. THE FACILITY PERFORMED AN OPEN SURGERY BECAUSE THERE WAS NO ALTERNATE DEVICE AT THAT TIME. THE FACILITY COMMENTED THE IMAGE DIFFICULTY MIGHT CAUSE THE BILE SPILLAGE. THE LTF-VH WAS SUSPECTED TO HAVE MALFUNCTION AT FIRST, SO OLYMPUS CHECKED THE LTF-VH ON (B)(6) 2016. BUT THERE WAS NO ABNORMALITY OF THE LTF-VH. AFTER THAT, THE PHENOMENON OCCURRED AGAIN WHEN THE FACILITY USED THE SUBJECT DEVICE CONNECTED WITH ANOTHER SCOPE. THEREFORE, THE SUBJECT DEVICE MIGHT BE ATTRIBUTED TO THE CAUSE OF THIS PHENOMENON. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2016-00237 TO PROVIDE DEVICE EVALUATION RESULTS. THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CHECKED THE SUBJECT DEVICE BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED. OLYMPUS WILL CONTINUE TO INVESTIGATE THE CAUSE OF THIS PHENOMENON. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00106. Manufacturer Narrative: OMSC TRIED TO REPRODUCE THE PHENOMENON WITH THE SUBJECT DEVICE AND THE ENDOSCOPE THAT THE FACILITY OWN, AND THE PHENOMENON WAS REPRODUCED. OMSC INVESTIGATED THE BEHAVIOR OF THE VIDEO SIGNAL PROCESSING ABOUT THE SUBJECT DEVICE, BUT OMSC COULD NOT FIND ANY ABNORMALITY. OMSC IS INVESTIGATING ABOUT THIS ISSUE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT #: 8010047-2016-00106. Manufacturer Narrative: OMSC INVESTIGATED THE BEHAVIOR OF THE VIDEO SIGNAL PROCESSING. THE SUBJECT DEVICE DID NOT HAVE ANY ABNORMALITY. OMSC EVALUATED THE ENDOSCOPE THAT THE FACILITY OWN, AND FOUND THAT THE TIMING OF THE VIDEO SIGNAL PROCESSING OF THIS ENDOSCOPE DEVIATED FROM THE STANDARD. AS A RESULT OF THESE INVESTIGATIONS, OMSC CONCLUDED THAT THIS PHENOMENON DID NOT CAUSE BY THE SUBJECT DEVICE, AND SURMISED THAT THE MALFUNCTION OF THE ENDOSCOPE CAUSED THIS PHENOMENON. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT #: 8010047-2016-00106. Manufacturer Narrative: THE SUBJECT DEVICE IS NOT RETURNED TO OLYMPUS YET. OLYMPUS WILL INVESTIGATE THE SUBJECT DEVICE TO DETERMINE THE CAUSE OF THIS PHENOMENON WHEN OLYMPUS RECEIVES IT. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00106. | -loss of image -image orientation incorrect -image reversal | -fluid leak -fitting problem -No Known Device Problem | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT AN ENDOSCOPE IMAGE IN THE MONITOR BECAME REDDISH SUDDENLY DURING LAPAROSCOPIC CHOLECYSTECTOMY. THE FACILITY COMPLETED THE PROCEDURE BY EXCHANGING THE CH-S190-XZ-E TO ANOTHER DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT EXCEPT EXCHANGING THE DEVICE. Manufacturer Narrative: OLYMPUS CONFIRMED THAT THE PHENOMENON WAS NOT REPRODUCED IN THE FACILITY AND THE DEVICE WORKED PROPERLY FOR 6 HOURS SINCE THE DEVICE WAS TURNED ON. OLYMPUS COULD NOT DETERMINE THE CAUSE OF THIS PHENOMENON BECAUSE THE PHENOMENON WAS NOT REPRODUCED. THIS PHENOMENON MIGHT BE ATTRIBUTED TO INADEQUATE CONNECTION TO THE VIDEO PROCESSOR DEVICE OR ADHESION OF FOREIGN SUBSTANCES ON ELECTRICAL CONTACTS OF THE DEVICE CAUSED CONTACT FAILURE TEMPORARILY. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. Manufacturer Narrative: OLYMPUS MEDICAL SYSTEMS CORPORATION (OMSC) PERFORMED A MDR RETROSPECTIVE REVIEW AND OMSC FOUND THAT THIS SUPPLEMENTAL REPORT IS REQUIRED ON DECEMBER 15, 2015. THE SUBJECT VIDEO SYSTEM CENTER (OTV-S190) WHICH THE FACILITY USED IN THE PROCEDURE WAS RETURNED TO OLYMPUS. OLYMPUS CONNECTED THE SUBJECT CAMERA HEAD (CH-S190-XZ-E) TO THE SUBJECT VIDEO SYSTEM CENTER TO CHECK OPERATION, BUT THE PHENOMENON WAS NOT REPRODUCED. HOWEVER, THERE WAS GAUZE FLUFF AROUND THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET IN THE SUBJECT VIDEO SYSTEM CENTER. OLYMPUS COULD REPRODUCE THE PHENOMENON TO CONNECT THE SUBJECT CAMERA HEAD TO THE SUBJECT VIDEO SYSTEM CENTER WITH THE ELECTRICAL CONTACT OF THE SUBJECT CAMERA HEAD ISOLATED. THEREFORE, CONTACT FAILURE BY THE GAUZE FLUFF ON THE VIDEO CONNECTOR SOCKET MIGHT CAUSE THE PHENOMENON. OLYMPUS ATTRIBUTED THIS PHENOMENON MIGHT BE CAUSED BY THE SUBJECT VIDEO SYSTEM CENTER. THE INSTRUCTION MANUAL OF OTV-S190 MENTIONS NOTIFICATIONS FOR DEVICE HANDLING. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT OTV-S190, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -inappropriate image rendered -loss of image -system migrated | -no pressure -Air Leak -leak | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT DURING CHOLECYSTECTOMY, A MONITOR IMAGE DISAPPEARED WHEN THE FACILITY OUTPUTTED THE ELECTRIC KNIFE. THE FACILITY COMPLETED THE PROCEDURE TO REPLACE A DEVICE TO THEIR SPARE DEVICE BUT THERE WAS NO INFORMATION WHAT DEVICE THE FACILITY REPLACED. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE IS NOT RETURNED TO OLYMPUS YET. OLYMPUS WILL INVESTIGATE THE SUBJECT DEVICE TO DETERMINE THE CAUSE OF THIS PHENOMENON AFTER OLYMPUS RECEIVES IT. OLYMPUS IS UNDER INVESTIGATION FOR THE CAUSE OF THIS PHENOMENON. OLYMPUS WILL SUBMIT A SUPPLEMENTAL MDR REPORT AFTER THE CAUSE OF THIS PHENOMENON IS DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -loss of image -image display artifact -device malfunction | -device maintenance issue -fluid leak -component missing | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT DURING LAPAROSCOPIC CHOLECYSTECTOMY, A MONITOR IMAGE DISAPPEARED SUDDENLY IN THE COMBINATION WITH THE SUBJECT DEVICE AND THE LTF-VH. THE FACILITY CONSIDERED THE CAUSE OF THIS PHENOMENON AS THE MALFUNCTION OF THE SUBJECT DEVICE OR THE LTF-VH. THE FACILITY TURNED OFF THE SUBJECT DEVICE IMMEDIATELY, AND THEN, THEY REMOVED AND INSERTED THE POWER SUPPLY CABLE. AFTER THAT, THE FACILITY TURNED ON THE SUBJECT DEVICE AGAIN AND THE DEVICES WORKED PROPERLY. THE FACILITY COMPLETED THE PROCEDURE WITHOUT REPLACING THE DEVICES. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2015-00925 TO PROVIDE DEVICE EVALUATION RESULTS. THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR INVESTIGATION. OLYMPUS CHECKED THE SUBJECT DEVICE WITH THE COMBINATION SCOPE WHICH THE FACILITY USED DURING THE PROCEDURE, AND THE PHENOMENON WAS SOMETIMES REPRODUCED WHEN THE VIDEO CONNECTER OF THE SCOPE WAS PULLED. CONTACT FAILURE BETWEEN THE SCOPE AND THE SUBJECT DEVICE MIGHT OCCUR TEMPORARILY BECAUSE THE FACILITY ADDED FORCE TO THE VIDEO CONNECTER OF THE SCOPE TOWARD PULLING DIRECTION FOLLOWING THE CABLE WAS PULLED WHILE USING THE SCOPE. SO THE SUBJECT DEVICE MIGHT NOT BE ABLE TO SUPPLY POWER TO THE SCOPE OR TO PROCESS IMAGE SIGNAL PROPERLY AT THAT TIME. CONSEQUENTLY, A MONITOR IMAGE MIGHT DISAPPEAR. IT WAS ALREADY PASSED OVER 7 YEARS SINCE THE FACILITY BEGAN TO USE THE SUBJECT DEVICE. THE FACILITY HAD INSERTED AND REMOVED A SCOPE FROM THE VIDEO CONNECTER SOCKET REPEATEDLY. SO CONTACT FAILURE BETWEEN THE SCOPE CONNECTER AND THE VIDEO CONNECTER SOCKET MIGHT OCCUR TEMPORARILY BECAUSE THE SOCKET WAS EASILY RATTLED BY WEAR. THERE WAS ALSO A POSSIBILITY THAT DETERIORATING CONTACT PINS OF THE VIDEO CONNECTER SOCKET MIGHT AFFECT THIS PHENOMENON. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS CAUTIONS FOR THE DEVICE HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2015-00928. Manufacturer Narrative: THE SUBJECT DEVICE IS NOT RETURNED TO OLYMPUS YET. OLYMPUS WILL INVESTIGATE THE SUBJECT DEVICE TO DETERMINE THE CAUSE OF THIS PHENOMENON AFTER OLYMPUS RECEIVES IT. OLYMPUS IS UNDER INVESTIGATION FOR THE CAUSE OF THIS PHENOMENON. OLYMPUS WILL SUBMIT A SUPPLEMENTAL MDR REPORT AFTER THE CAUSE OF THIS PHENOMENON IS DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2015-00928. | -loss of image -image display artifact -device detached | -fluid leak -fitting problem -electrical issue | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT DURING THE TUR (TRANSURETHRAL RESECTION) PROCEDURE THE HALATION HAD OCCURRED ON THE ENDOSCOPIC IMAGE REPEATEDLY. THE FACILITY HAD COMPLETED THE PROCEDURE WITH USE OF THE OTHER DEVICE. THERE WAS NO REPORT OF THE PT'S INJURY REGARDING THIS EVENT. Manufacturer Narrative: THE REFERENCED OTV-S7V AND CONCOMITANTLY-USED LIGHT SOURCE AND CAMERA HEAD WERE RETURNED TO THE OLYMPUS MEDICAL SYSTEMS CORP (OMSC) FOR EVAL. OMSC EVALUATED THE DEVICES AND FOUND THAT THEY HAD NO ABNORMALITY AND OPERATED CORRECTLY. THE EXACT CAUSE OF THIS PHENOMENON CANNOT BE CONCLUSIVELY DETERMINED, HOWEVER THERE IS THE POSSIBILITY OF THIS PHENOMENON IS ATTRIBUTED TO THE USER HANDLING OF THE DEVICES. OMSC STATED THE APPROPRIATE HANDLING OF THE OTV-S7V IN THE INSTRUCTION MANUAL WHEN THE OTV-S7V HAD ABNORMALITIES. ALSO, OMSC CHECKED THE MFR HISTORY OF THE SUBJECT DEVICE, THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADD'L INFO IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -patient experienced adverse event -image orientation incorrect -loss of image | -crack -electromagnetic interference (emi) -leak | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT DURING UNKNOWN PROCEDURE, A MONITOR IMAGE DISAPPEARED SUDDENLY IN THE COMBINATION WITH THE SUBJECT DEVICE AND THE VIDEOSCOPE (B)(4). THE FACILITY COMPLETED THE PROCEDURE TO REPLACE THE SCOPE TO ANOTHER ONE. THERE WAS NO PATIENT'S INJURY IN THIS EVENT. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2015-00974 TO PROVIDE DEVICE EVALUATION RESULTS. OLYMPUS COULD REPRODUCE THE PHENOMENON WHILE PERFORMING AN OPERATIONAL CHECK OF THE SUBJECT DEVICE WITH THE VIDEOSCOPE LTF-S190-10 WHICH USED DURING THE PROCEDURE. BUT THE PHENOMENON COULD NOT BE REPRODUCED AGAIN AFTER THE DEVICE WAS TURNED ON AND OFF. OLYMPUS ALSO CONFIRMED THE ELECTRICAL CONTACTS OF THE VIDEO CONNECTER SOCKET OF THE SUBJECT DEVICE GOT CORRODED. THE FACILITY MIGHT CONNECT THE SUBJECT DEVICE WITH THE SCOPE WHOSE CONNECTOR WAS WET AND WATER CONGESTION OCCURRED INSIDE THE VIDEO CONNECTER SOCKET. THEREFORE, THE ELECTRICAL CONTACTS GOT CORRODED, RESULTING IN CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE SCOPE. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS CAUTIONS FOR VIDEO CONNECTOR HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2015-00975. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR INVESTIGATION, BUT THE PHENOMENON WAS NOT REPRODUCED SO FAR. OLYMPUS WILL PERFORM FURTHER INVESTIGATION TO DETERMINE THE CAUSE OF THIS PHENOMENON. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. OLYMPUS WILL SUBMIT A SUPPLEMENTAL REPORT AFTER THE CAUSE OF THIS PHENOMENON IS DETERMINED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2015-00975. | -loss of image -unable to scroll image -device disconnected | -fluid leak -particulates -device inoperable | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT NOISE ON THE MONITOR SCREEN OCCURRED DURING LAPAROSCOPIC HERNIA RESTORATION. THE FACILITY REPLACED THE SUBJECT OTV-S190 TO ANOTHER SYSTEM TO COMPLETE THE PROCEDURE. THERE WAS NO INJURY REPORT OF THIS EVENT EXCEPT EXCHANGING THE DEVICE. Manufacturer Narrative: THE SUBJECT OTV-S190 WAS RETURNED TO OLYMPUS FOR EVAL. OLYMPUS COULD NOT CONFIRM THE PHENOMENON BUT THERE WERE FOREIGN SUBSTANCES NEAR THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET. OLYMPUS COULD REPRODUCE THIS PHENOMENON BY CONNECTING A SCOPE TO THE OTV-S190, PLACING FOREIGN SUBSTANCES ON THE SCOPE CONNECTOR PINS FOR INSULATION. THIS PHENOMENON MIGHT BE ATTRIBUTED TO CONDUCTION FAILURE BETWEEN THE SCOPE AND THE OTV-S190 BECAUSE ADHESION OF FOREIGN SUBSTANCES INSIDE THE VIDEO CONNECTOR SOCKET OR WATER ON THE SCOPE CONNECTOR MIGHT INTERRUPT CONNECTION BETWEEN THE TWO DEVICES TEMPORARILY. OLYMPUS ALSO CHECKED THE MFG HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTIONS MANUAL ALREADY STATED NOTIFICATIONS OF ELECTRICAL CONTACT HANDLING. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -noise emitted -incorrect screen display -device alarmed | -break -circuit failure -leak | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT THE MONITOR IMAGE DISAPPEARED DURING TRANSURETHRAL URETEROLITHOTRIPSY. THE FACILITY COMPLETED THE PROCEDURE TO REPLACE THE SUBJECT OTV-S7PRO TO ANOTHER DEVICE. THERE WAS NO REPORT OF PATIENT'S INJURY IN THIS EVENT. Manufacturer Narrative: THIS IS A SUPPLEMENTAL REPORT FOR MFR REPORT #8010047-2015-00731 TO PROVIDE DEVICE EVALUATION RESULTS. OLYMPUS PERFORMED FURTHER INVESTIGATION FOR THE SUBJECT DEVICE, BUT THE PHENOMENON WAS NOT REPRODUCED. OLYMPUS COULD NOT DETERMINE THE CAUSE OF THIS PHENOMENON CONCLUSIVELY. OLYMPUS ATTRIBUTED THE CAUSE OF THIS ISSUE TO THE ASSUMPTION THAT THE PARTS ON THE CIRCUIT BOARD IN THE SUBJECT DEVICE, WHICH PROCESSED IMAGE OUTPUT, MIGHT NOT WORK WELL TEMPORARILY FOR SOME REASON. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR INVESTIGATION. OLYMPUS INVESTIGATED THE SUBJECT DEVICE BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED SO FAR. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE ISSUE IS UNDER INVESTIGATION, SO OLYMPUS WILL SUBMIT A SUPPLEMENTAL MDR REPORT AFTER THE CAUSE OF THIS PHENOMENON IS DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -loss of image -image display artifact -device malfunction | -device maintenance issue -leak -data back-up problem | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THAT THE MONITOR SHOWED THE ERROR E315 (UNSUPPORTED SCOPE) DURING UNSPECIFIED PROCEDURE WITH THE SUBJECT DEVICE AND THE LAPARO-THORACO VIDEOSCOPE LTF-VH. THE FACILITY REPLACED THE LTF-VH TO ANOTHER LAPARO-THORACO VIDEOSCOPE LTF-V3, BUT THE SAME PHENOMENON STILL OCCURRED. AFTER THAT, THE LTF-VH WHICH THE FACILITY USED AT FIRST WAS CONNECTED TO THE SUBJECT DEVICE AND THERE WAS NO ERROR. THE FACILITY COMPLETED THE PROCEDURE WITH THE SUBJECT DEVICE AND THE LTF-VH. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION AND OLYMPUS COULD REPRODUCE THE PHENOMENON. OLYMPUS CONFIRMED THERE WAS FOREIGN SUBSTANCE IN THE ELECTRICAL CONTACT OF THE VIDEO CONNECTER SOCKET OF THE SUBJECT DEVICE. OLYMPUS ALSO CONFIRMED AT THE FACILITY THAT THERE WERE FOREIGN SUBSTANCES IN THE ELECTRICAL CONTACT OF THE VIDEO PLUG OF BOTH SCOPES WHICH THE FACILITY USED AT THAT TIME. THERE WAS A POSSIBILITY THAT THE SCOPES WHICH HAD FOREIGN SUBSTANCES IN THE ELECTRICAL CONTACT OF THE VIDEO PLUG WERE CONNECTED TO THE SUBJECT DEVICE, RESULT IN CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE SCOPES. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF OTV-S190, LTF-VH AND LTF-V3 ALREADY MENTIONS CAUTIONS FOR VIDEO CONNECTOR HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT #: 8010047-2016-00245, 8010047-2016-00246. | -improper device position -device malfunction -error message | -fluid leak -leak -electromagnetic interference (emi) | Viewer
| FDA | Event Description: OLYMPUS WAS INFORMED THE FOLLOWING. DURING THE LAPAROSCOPIC CHOLECYSTECTOMY, THE OPERATOR PERFORMED THE CUT PREPARATION. AT THAT TIME, THE ENDOSCOPIC IMAGE WAS DISAPPEARED. THE FACILITY REPLACED THIS OTV-S190 WITH ANOTHER SIMILAR DEVICE MADE BY STRYKER AND THE FACILITY COMPLETED THE PROCEDURE. THERE WAS NO REPORT OF THE PATIENT'S INJURY REGARDING THIS EVENT. ALTHOUGH A CAUSAL RELATIONSHIP WITH THIS EVENT IS UNKNOWN, THE HOSPITALIZATION WAS PROLONGED. Manufacturer Narrative: THE SUBJECTED DEVICE WAS NOT RETURNED TO OLYMPUS MEDICAL SYSTEMS CORP.(OMSC). OMSC CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE OTV-S190 INSTRUCTION MANUALS STATE THE CORRESPONDING METHOD WHEN THERE IS AN ABNORMALITY IN THE ENDOSCOPIC IMAGE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. | -loss of image -image orientation incorrect -unable to edit | -fluid leak -Air Leak -Bleed Back | Viewer
| FDA | Event Description: ON (B)(6) 2012 A MERGE CARDIO CUSTOMER REPORTED REPORTS DONE UNDER WRONG PATIENT. REPORT FROM 2ND PATIENT IS ALSO SHOWING UP AS A 2ND REPORT THREAD FOR PATIENT #1. THE PDF GENERATION FUNCTIONALITY WAS NOT DESIGNED TO BE EXCLUSIVE PER PATIENT. IF TWO OR MORE USERS WERE REPORTING ON DIFFERENT PATIENTS AND REQUESTED THE GENERATION OF A PDF CONCURRENT, THE SYSTEM WOULD INCORRECTLY ATTACH THE FIRST PDF REPORT FILE GENERATED TO THE OTHER PATIENT(S). Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -incorrect patient information -fail to associate patient -fail to provide patient diagnosis info | -disconnection -leak -Connection Issue | Viewer
| FDA | Event Description: ON (B)(6) 2015 MCKESSON SUPPORT RECEIVED A REQUEST FROM THE CUSTOMER TO PROVIDE ASSISTANCE WITH AN INVESTIGATION INTO THE CHANGE HISTORY FOR A PRELIMINARY REPORT ATTACHED TO A CT STUDY FOR A PATIENT. DURING A F/U DISCUSSION WITH THE REPORTING FACILITY ON 04/23/2015, MCKESSON BECAME AWARE THE PATIENT EXPIRED ON (B)(6) 2015 DUE TO COMPLICATIONS FOLLOWING ABDOMINAL SURGERY. Manufacturer Narrative: IN RESPONSE TO THE REQUEST FROM THE REPORTING FACILITY, MCKESSON INITIATED AN INVESTIGATION IN ORDER TO DETERMINE WHEN THE PRELIMINARY REPORT HAD BEEN CREATED, MODIFIED, OR ACCESSED BASED ON INFO CONTAINED IN APPLICATION LOGS. DURING FURTHER DISCUSSIONS WITH THE REPORTING FACILITY TO CLARIFY THE SPECIFIC INFO BEING REQUESTED, AND ANALYSIS OF APPLICATION LOGS, MCKESSON'S INVESTIGATION CONCLUDED THE FOLLOWING: A PRELIMINARY REPORT WAS CREATED ON (B)(6) 2015 BY A RADIOLOGIST. UPON FURTHER REVIEW OF THE PATIENT'S CONDITION BY ANOTHER RADIOLOGIST ON-CALL THE FOLLOWING DAY, THE PATIENT UNDERWENT ABDOMINAL SURGERY ON (B)(6) 2015. THE REPORTING FACILITY INDICATED THAT THE PATIENT EXPIRED DUE TO COMPLICATIONS POST-SURGERY AND EXPIRED ON (B)(6) 2015. BASED ON THE INVESTIGATION PERFORMED AND ANALYSIS OF THE APPLICATION LOGS, THERE IS NO EVIDENCE OF A DEVICE MALFUNCTION. | -fail to provide indicator -difficult to maintain -missing part | -detachment of device component -fitting problem -disconnection | Viewer
| FDA | Event Description: ON (B)(6) 2015, VITAL CUSTOMER SUPPORT RECEIVED A CALL FROM (B)(6) REGARDING DISPLAY OF INCORRECTLY INVERTED BRAIN PERFUSION COLOR MAPS (CBF AND MTT) ON OUR VITREA CT BRAIN PERFUSION 4D SOFTWARE APPLICATION. THE REPORTER INDICATED THAT A DELAY IN PATIENT TREATMENT HAD OCCURRED BECAUSE OF THE ISSUE. AGAIN ON (B)(6) 2015, THE SAME REPORTER INFORMED OUR CUSTOMER SUPPORT TEAM REGARDING ANOTHER PATIENT ISSUE WITH THE SOFTWARE IN WHICH THE CT BRAIN PERFUSION RESULT FROM OUR SOFTWARE WAS NOT IN AGREEMENT WITH THE PERFORMED MR RESULTS. THE MRI SHOWED AN OBVIOUS STROKE (INFARCT), WHILE THE CT PERFUSION RESULTS ON OUR SOFTWARE DID NOT. THE PATIENT WAS TREATED BASED ON MRI RESULTS.Event Description: ON 08/25/2015, VITAL CUSTOMER SUPPORT RECEIVED A CALL FROM (B)(6) REGARDING DISPLAY OF INCORRECTLY INVERTED BRAIN PERFUSION COLOR MAPS (CBF AND MTT) ON OUR VITREA CT BRAIN PERFUSION 4D SOFTWARE APPLICATION. THE REPORTER INDICATED THAT A DELAY IN PATIENT TREATMENT HAD OCCURRED BECAUSE OF THE ISSUE. AGAIN ON 08/26/2015, THE SAME REPORTER INFORMED OUR CUSTOMER SUPPORT TEAM REGARDING ANOTHER PATIENT ISSUE WITH THE SOFTWARE IN WHICH THE CT BRAIN PERFUSION RESULT FROM OUR SOFTWARE WAS NOT IN AGREEMENT WITH THE PERFORMED MR RESULTS. THE MRI SHOWED AN OBVIOUS STROKE (INFARCT), WHILE THE CT PERFUSION RESULTS ON OUR SOFTWARE DID NOT. THE PATIENT WAS TREATED BASED ON MRI RESULTS. Manufacturer Narrative: THE INVESTIGATION IS ONGOING WITH USER FACILITY. Manufacturer Narrative: THE CUSTOMER HAS PROVIDED US TWO DATASETS TO INVESTIGATE. THE RESULTS OF THE INVESTIGATION IDENTIFIED TWO SEPARATE ISSUES. ISSUE #1: UNEXPECTED MTT/CBF MAPS. THIS CUSTOMER STATED THE MTT/CBF MAPS APPEAR TO BE INVERTED. THE INVESTIGATION SHOWS THE DATASET HAS A TRUNCATED SCAN, THE TIME-DENSITY GRAPH SHOWS THAT THERE IS NOT ENOUGH WASHOUT FOR OUR SOFTWARE TO PRODUCE ACCURATE RESULTS. FROM OUR NEURO EDUCATION AND REFERENCE GUIDE (VPMC-13505 A, PAGE 122), WE STATE "CAUTION: FACTORS THAT CAN AFFECT THE ACCURACY OF THE RESULTS ARE: úNOT FULLY CAPTURING THE CONTRAST BOLUS (INCLUDING AT LEAST 2 SECONDS OF PRE-ENHANCEMENT IMAGES AND POST-ENHANCEMENT)." ALSO, FROM OUR NEURO EDUCATION AND REFERENCE GUIDE (VPMC-13505 A, PAGE 128), WE INSTRUCT THE USER TO REVIEW THE TIME-DENSITY GRAPH. ADDITIONALLY, FROM OUR NEURO EDUCATION AND REFERENCE GUIDE (VPMC-13505 A, PAGE 152), WE PROVIDE EXAMPLES OF OPTIMAL AND SUB-OPTIMAL TIME-DENSITY GRAPHS. IN CONCLUSION: THE SOFTWARE HAS BEHAVED AS IT IS DESIGNED. DUE TO INCOMPLETE DATA, THE PERFUSION RESULTS COULD NOT HAVE BEEN CALCULATED ACCURATELY, ALTHOUGH THE SOFTWARE DID PRODUCE ACCURATE RESULTS WITH RESPECT TO THE INCOMPLETE DATA. THE SOFTWARE IS DESIGNED TO ASSIST THE USER IN THIS SITUATION BY PROMINENTLY DISPLAYING THE TIME-DENSITY GRAPH AND BY PROVIDING ACCURATE INSTRUCTIONS FOR USE ON HOW TO INTERPRET THE TIME-DENSITY GRAPH. ISSUE #2: CT PERFUSION RESULTS DO NOT AGREE WITH MR RESULTS FOR THIS ISSUE, THE USER STATED THEY WERE NOT CONFIDENT IN THE RESULTS PROVIDED BY OUR SOFTWARE AND PROCEEDED TO PERFORM AN MR SCAN TO OBTAIN ADDITIONAL INFORMATION. WHEN THE USER REVIEWED THE MR SCAN DATA, A CLEAR INFARCT WAS SHOWN, WHICH WAS NOT SHOWN BY THE CT PERFUSION RESULTS. A DETAILED INVESTIGATION WAS DONE, THE SUMMARY RESULTS ARE: THE CT PERFUSION RESULTS FROM OUR SOFTWARE ARE IN AGREEMENT WITH ANOTHER VENDOR'S ALGORITHM, WHICH SUGGESTS OUR SOFTWARE GENERATED THE RESULTS ACCURATELY BASED ON THE DATA. ADDITIONALLY, A VISUAL REVIEW OF THE CTA DATA SHOWS THAT THE CT PERFUSION RESULTS ARE IN AGREEMENT WITH THE CTA DATA, WHICH SUGGESTS THE SOFTWARE GENERATED RESULTS ACCURATELY. | -incorrect display value -inappropriate color scheme -device marked incorrectly | -use of device issue -infusion or flow issue -inadequate user interface | Viewer
| FDA | Event Description: ON (B)(6) 2016 PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT ORIENTATIONS MARKED IN SOME IMAGES LOOKS WRONG, THE ISSUE IS STILL UNDER INVESTIGATION. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. Manufacturer Narrative: AN INVESTIGATION WAS PERFORMED BY THE ENGINEERING TEAM AND CLINICAL APPLICATION EXPERTS. FURTHER INVESTIGATION FOUND THAT THIS ISSUE IS SPECIFIC TO SAVING SAGITTAL VOLUME FILES. THIS ISSUE WILL ONLY OCCUR IF THE USER SAVES THE RECONSTRUCTED DATA IN THE SAGITTAL ORIENTATION AND THEN LOADS THE SAGITTAL VOLUME INTO THE VIEWING APPLICATION. CLINICAL APPLICATIONS TRAINS USERS TO ONLY SAVE AND USE THE TRANSVERSE IMAGE VOLUME FILES AFTER EXECUTING THE RECONSTRUCTION APPLICATION (AUTOSPECT PRO). THE ENGINEERING TEAM DETERMINED THAT THE RISK IS ACCEPTABLE REGARDING THIS ISSUE. THE CUSTOMER IS CURRENTLY USING THE SYSTEM AND THEY HAVE BEEN INSTRUCTED BY CLINICAL APPLICATIONS ON HOW TO AVOID THE ISSUE. Manufacturer Narrative: (B)(4). THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. | -wrong print -device marked incorrectly -inappropriate marking symbol | -break -disconnection -air leak | Viewer
| FDA | Event Description: ON (B)(6) 2017 A CUSTOMER SUBMITTED A COMPLAINT INDICATING THAT MAMMOGRAM IMAGES APPEARED DARK WHEN LOADED IN CLEARCANVAS RIS PACS. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: THE CUSTOMER DEMONSTRATED THAT THE ANATOMY WAS NOT VISIBLE THROUGH REMOTE ACCESS. THE PROBE TOOL WAS USED TO CHECK THE BRIGHTNESS VALUES. BRIGHTEST AREAS HAD VALUES AROUND 190-200. DICOM HEADERS DISPLAYED A WINDOW CENTER (0028,1050) VALUE OF 2,700. CENTER VALUE SHOULD BE ROUGHLY BETWEEN THE BRIGHTEST AND DARKEST PIXEL VALUES. IT WAS DETERMINED THAT THE MODALITY IS WRITING THE CENTER VALUE DIFFERING FROM THE BRIGHTEST SPOT. IT WAS DETERMINED AT THIS TIME THAT THE PRODUCT FUNCTIONED AS INTENDED AND THIS ISSUE WAS DUE TO INCORRECT DICOM DATA. A WORKAROUND WAS PROVIDED TO THE CUSTOMER IF THE ISSUE IS ENCOUNTERED AGAIN. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME (B)(4) CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. THIS DEFECT IS ONLY ENCOUNTERED IN SELECT SCENARIOS, WHERE A MODALITY PRODUCES IMAGES THAT ARE COMPRESSED USING (B)(4), AND THE IMAGE PIXEL DATA IS LESS THAN 16-BIT, AND THE "BITNESS" OF THE COMPRESSED DATA STREAM DOES NOT MATCH THAT OF THE IMAGE PIXEL DATA. THE OUTCOME OF THE ABOVE IS A LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA, WHICH CAUSES THE SYMPTOMS THAT WERE REPORTED BY THE CUSTOMER. THE ROOT CAUSE IS THAT A SOFTWARE DEFECT FOUND IN THE DEVICE THAT IS ENCOUNTERED WHEN IT IS USED WITH NON-DICOM COMPLIANT (B)(4) COMPRESSED IMAGES. THE SOFTWARE IMPROPERLY USES THE DICOM BIT DEPTH (I.E. THE BITS STORED TAG) TO DECOMPRESS THE COMPRESSED PIXEL DATA STREAM FOR DISPLAY, INSTEAD OF THE BIT DEPTH THAT IS ENCODED IN THE COMPRESSED PIXEL DATA STREAM ITSELF. IN NON-DICOM COMPLIANT IMAGES WHERE THE DICOM AND COMPRESSED PIXEL DATA STREAM BIT DEPTHS DO NOT MATCH, THE SOFTWARE OUTPUTS AN IMAGE WITH SOME LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA. IN THESE SCENARIOS, THE MODALITY PRODUCING THE IMAGE IS ITSELF ALSO NOT COMPLIANT TO DICOM (B)(4). MDR REPORTS: 3012075008-2020-00002, 3012075008-2020-00003 AND 3012075008-2020-00004 HAVE THE SAME ROOT CAUSE. A CAPA HAS BEEN OPENED AND RECALL INITIATED TO ADDRESS THE ISSUE IN AFFECTED DEVICES. | -dark patches -loss of image -image brightness out of specification | -device operates differently than expected -no display or display failure -break | Viewer
| FDA | Event Description: ON (B)(6) 2017 AGFA BECAME AWARE OF A SECOND OCCURRENCE REPORTED BY THE CUSTOMER IN WHICH REPORTS CROSSING FROM CARDIO SYSTEM (CV12.1SU1) INTO EPIC ARE SHOWING UP WITH INCORRECT DEMOGRAPHICS. THE FIRST OCCURRENCE HAS BEEN REPORTED IN FDA MDR 1225058-2015-14119. PREVIOUSLY THE ROOT CAUSE HAS BEEN IDENTIFIED AS PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. AGFA REPORTED TO FDA VIA 21 CFR 806 REPORTING (Z-0283-2016) A CORRECTION FOR THIS ISSUE. THE CUSTOMER DECLINED THE CORRECTION, HOWEVER THE CUSTOMER DID ACKNOWLEDGE AND ACCEPT THE RISK ASSOCIATED WITH NOT ACCEPTING THE CORRECTED VERSION OF THE SOFTWARE AND NOT REMOVING THE AFFECTED VERSION OF THE SOFTWARE. AGFA IS WORKING WITH THE CUSTOMER TO RESOLVE THIS ISSUE. THERE HAS BEEN NO REPORTED HARM TO PATIENT OR USER DURING THIS EVENT. A SUPPLEMENTAL REPORT WILL BE PROVIDED.Event Description: THIS SUPPLEMENT REPORT #1 IS BEING SUBMITTED TO PROVIDE CORRECTIVE ACTIONS TAKEN. | -cross flow -incorrect message display -incorrect device orientation | -no display or display failure -application program issue -break | Viewer
| FDA | Event Description: ON (B)(6) 2017, PHILIPS RECEIVED FEEDBACK THAT INTELLISPACE PORTAL 7.0.2.20700 MR FIBERTRAK APPLICATION IS NO LONGER ABLE TO POST PROCESS CERTAIN DATA. THE ISSUE IS STILL UNDER INVESTIGATION. Manufacturer Narrative: WE HAVE RECEIVED NEW INFORMATION AND THEREFORE MODEL # BRAND NAME: AND SERIAL # WERE CORRECTED IN THIS REPORT THE RISK ASSESSMENT CONCLUDED THAT THE OVERALL RISK IS ACCEPTABLE. A PHILIPS APPLICATIONS SPECIALIST WENT ONSITE AND TRAINED THE CUSTOMER ON HOW TO PROPERLY CREATE STUDIES TO ANALYZE ON THE SYSTEM. Manufacturer Narrative: WE HAVE NOT COMPLETED OUR INVESTIGATION OF THIS EVENT. WE WILL FILE A FOLLOW-UP EMDR AT THE COMPLETION OF THE INVESTIGATION. (B)(4). | -fail to store data -fail to display data -fail to manage data | -failure to read input signal -filtration issue -failure of device to self-test | Viewer
| FDA | Event Description: ON (B)(6) 2019 A CUSTOMER REPORTED THAT IMAGES RETRIEVED FROM VRAD (PACS) IS DISPLAYED VERY DARK WHEN LOADED INTO CLEARCANVAS WORKSTATION PERSONAL EDITION 13.1. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: THE CUSTOMER ALSO NOTED THE IMAGES APPEAR FINE WHEN VIEWED IN THE VRAD WEB VIEWER. HE ALSO SAID THAT THEY APPEARED NORMAL IN AN OLDER VERSION OF CC WORKSTATION (2.0 SP1 - OPEN-SOURCE). THE SYNAPTIVE REP INITIATED SCREEN SHARE TO OBSERVE. SAMPLE WAS A DIGITAL X-RAY FROM A KONICA MINOLTA PORTABLE SCANNER. W/L WAS CHANGED TO AROUND 491/177 BEFORE IT WAS VIEWABLE. INITIAL W/L (AUTO) = 4096/2047. OPENED A FEW MORE IMAGES THAT WERE VERY SIMILAR. THE CUSTOMER EXPLAINED THAT THE PREVIOUS VERSION WAS CC WS 2.0 SP1 (OPEN SOURCE), BUT WAS UNABLE TO CONFIRM WHETHER THESE PARTICULAR IMAGES OPENED CORRECTLY WITH THAT VERSION. THE SYNAPTIVE REP GAVE THE CUSTOMER A WORKAROUND VIA THE WINDOW/LEVEL HOTKEY CONFIGURATION. ON 18-MAR-2019 AN INVESTIGATION FOUND THE ISSUE TO BE IDENTICAL TO FEED-3544 (REPORTED AS 3012075008-2020-00001): THE BITS STORED (0028,0101) VALUE IS SET TO 12-BIT IN THE DICOM HEADER HOWEVER, IT APPEARS THE JPEG 2000 COMPRESSED PIXEL DATA IS ACTUALLY 16-BIT. WHEN THE TAG VALUE IS EDITED TO 16, THE IMAGE APPEARS CORRECT (I.E. NOT DARK) ON LOAD. THEREFORE, THE DICOM DATA IS INCONGRUENT. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME JPEG 2000 CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. THIS DEFECT IS ONLY ENCOUNTERED IN SELECT SCENARIOS, WHERE A MODALITY PRODUCES IMAGES THAT ARE COMPRESSED USING JPEG 2000, AND THE IMAGE PIXEL DATA IS LESS THAN 16-BIT, AND THE "BITNESS" OF THE COMPRESSED DATA STREAM DOES NOT MATCH THAT OF THE IMAGE PIXEL DATA. THE OUTCOME OF THE ABOVE IS A LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA, WHICH CAUSES THE SYMPTOMS THAT WERE REPORTED BY THE CUSTOMER. THE ROOT CAUSE IS THAT A SOFTWARE DEFECT FOUND IN THE DEVICE THAT IS ENCOUNTERED WHEN IT IS USED WITH NON-DICOM COMPLIANT JPEG 2000 COMPRESSED IMAGES. THE SOFTWARE IMPROPERLY USES THE DICOM BIT DEPTH (I.E. THE BITS STORED TAG) TO DECOMPRESS THE COMPRESSED PIXEL DATA STREAM FOR DISPLAY, INSTEAD OF THE BIT DEPTH THAT IS ENCODED IN THE COMPRESSED PIXEL DATA STREAM ITSELF. IN NON-DICOM COMPLIANT IMAGES WHERE THE DICOM AND COMPRESSED PIXEL DATA STREAM BIT DEPTHS DO NOT MATCH, THE SOFTWARE OUTPUTS AN IMAGE WITH SOME LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA. IN THESE SCENARIOS, THE MODALITY PRODUCING THE IMAGE IS ITSELF ALSO NOT COMPLIANT TO DICOM PS3.5 8.2.4. MDR REPORTS: 3012075008-2020-00001, 3012075008-2020-00002, AND 3012075008-2020-00003 HAVE THE SAME ROOT CAUSE. A CAPA HAS BEEN OPENED AND RECALL INITIATED TO ADDRESS THE ISSUE IN AFFECTED DEVICES. | -image display error -dark patches -loss of image | -no display or display failure -detachment of device component -leak | Viewer
| FDA | Event Description: ON (B)(6) 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. THE POWERJACKET WINDOW CAN BECOME OUT OF SYNC AND DISPLAY THE DETAILS FOR THE PREVIOUSLY LOADED PATIENT. ON DECEMBER 16, 2020 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER COMP- (B)(4). | -inappropriate window level -incorrect product classification -device malfunction | -failure of device to self-test -detachment of device component -component missing | Viewer
| FDA | Event Description: ON (B)(6) 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH THE SYNAPSE PACS POWERJACKET. THE POWERJACKET WAS SHOWING INCORRECT PATIENT STUDIES/IMAGES; ALL THE INFORMATION WAS CORRECT EXCEPT FOR THE IMAGES. THE CUSTOMER INDICATED THIS OCCURS A FEW TIMES A DAY. THE ISSUE WAS REVIEWED AND ESCALATED TO THE ENGINEERING DEPARTMENT. ON (B)(6) 2020 A RISK ASSESSMENT WAS PERFORMED BY THE ENGINEERING DEPARTMENT TO INVESTIGATE RISK TO PATIENT SAFETY AND FMSU QUALITY AND REGULATORY AFFAIRS DEPARTMENTS WERE INFORMED OF THE SEVERITY. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A JAVA SCRIPT ERROR. A SOFTWARE FIX FOR THE ISSUE IS AVAILABLE AND WILL BE PROVIDED TO THE CUSTOMER. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -illumination power out of specification -power out of range -loss of power | -failure of device to self-test -component missing -break | Viewer
| FDA | Event Description: ON (B)(6) 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) WAS NOTIFIED THAT THE SYNAPSE PACS SHOWS INCORRECT PATIENT STUDY/IMAGE IN THE POWERJACKET. THE ISSUE WAS REVIEWED AND ESCALATED TO THE ENGINEERING DEPARTMENT. ON (B)(6) 2020 A RISK ASSESSMENT WAS PERFORMED BY THE ENGINEERING DEPARTMENT TO INVESTIGATE RISK TO PATIENT SAFETY AND FMSU QUALITY AND REGULATORY AFFAIRS DEPARTMENTS WERE INFORMED OF THE SEVERITY. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A JAVA SCRIPT ERROR. THE ISSUE HAS BEEN RESOLVED; THE CUSTOMER'S SOFTWARE WAS UPGRADED. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER COMP- (B)(4). | -incorrect patient information -incorrect display -image orientation incorrect | -label -failure to read input signal -leak | Viewer
| FDA | Event Description: ON (B)(6) 2020 THE FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) ENGINEERING DEPARTMENT INITIATED A COMPLAINT FOR AN INTERNAL ISSUE WHERE THE SYNAPSE PACS WAS SHOWING INCORRECT 3D SPHERE MAX VALUES FOR SPECIFIC CT STUDIES. THE 3D SPHERE MAX SHOULD GIVE A MAX DENSITY VALUE WITHIN THE SPHERE'S RANGE; HOWEVER THE VALUE SHOWN WAS MUCH HIGHER THAN EXPECTED. THE ISSUE IS NOT PRESENT WITH ANY OTHER DENSITY TOOLS. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THIS ISSUE WAS DISCOVERED INTERNALLY; THIS SOFTWARE VERSION HAS NOT BEEN RELEASED TO ANY CUSTOMERS. THE CAUSE WAS TRACED TO INSUFFICIENT TESTING DURING DEVELOPMENT OF SOFTWARE VERSION. THE INCORRECT VALUE IS HIGHLY DETECTIBLE AND WOULD BE NOTICED BY THE CLINICIAN. THE ISSUE WILL BE RESOLVED PRIOR TO RELEASE TO CUSTOMERS. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -incorrect 3d model -incorrect display format -incorrect message display | -device alarm system issue -failure to read input signal -Device displays error message | Viewer
| FDA | Event Description: ON (B)(6) 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR AN ISSUE WITH THE EJECTION FRACTION (EF) AND LEFT VENTRICULAR (LV) VOLUME METHOD OF DISK BI-PLANE (BP MOD) CALCULATION. THE EF OR LV VOLUME BP MOD DID NOT UPDATE IN THE CLINICAL REPORTING APPLICATION (CRA) REPORT WHEN MEASUREMENTS WERE CHANGED BY A DIFFERENT USER. AFTER THE SONOGRAPHER PERFORMS LV VOLUME BP MOD AND SAVES THE REPORT, AND ANOTHER USER TRIES TO EDIT THE "TRACE", THE UPDATED/NEW VOLUME AND EJECTION FRACTION VALUES DO NOT UPDATE IN THE CRA. IF THE SONOGRAPHER WHO INITIALLY PERFORMED THE BP MOD EDITS THE "TRACE", THE VALUES UPDATE CORRECTLY. THE ISSUE WAS REVIEWED, AND ESCALATED TO THE ENGINEERING DEPARTMENT. ON OCTOBER 30, 2020, A RISK ASSESSMENT WAS PERFORMED BY THE ENGINEERING DEPARTMENT TO INVESTIGATE RISK TO PATIENT SAFETY, FMSU QUALITY, AND REGULATORY AFFAIRS DEPARTMENTS WERE INFORMED OF THE SEVERITY. THERE WAS NO SERIOUS INJURY, OR DEATH ASSOCIATED WITH THIS EVENT. THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO INSUFFICIENT TESTING DURING DEVELOPMENT. THE ISSUE HAS BEEN RESOLVED IN THE NEXT VERSION, AND THIS CUSTOMER WAS ALREADY PROVIDED WITH A FIX. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF.#: INTERNAL COMPLAINT NUMBER: (B)(4). | -fail to eject -part ejected -component issue | -Device displays error message -communication or transmission issue -device alarm system issue | Viewer
| FDA | Event Description: ON (B)(6) 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS STANDARD UPTAKE VALUES (SUV). ON (B)(6) 2020, FMSU ENGINEERING DEPARTMENT INITIATED A BUG FOR THE ISSUE TO INVESTIGATE THE RISK TO PATIENT SAFETY. ON (B)(6) 2020, ENGINEERING DEPARTMENT FOUND THAT THE MAX SUV VALUES WERE DIFFERENT WHEN USING 2D REGION OF INTEREST (ROI) AS COMPARED TO WHEN USING 3D SPHERE ROI. VALUES ARE USED TO DETERMINE THE SIZE OF THE LESIONS AND CONTRIBUTE TO A DIAGNOSIS. THE ISSUE WAS FOUND TO ONLY IMPACT THE VALUES WHEN USED WITH A SPECIFIC NON-FUJIFILM DEVICE FOR PET CALCULATIONS. THERE WAS NO PATIENT SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO INSUFFICIENT TESTING DURING DEVELOPMENT. A BUG FIX IS BEING DEVELOPED AND AFFECTED CUSTOMERS WILL BE NOTIFIED AND PROVIDED THE FIX. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. INTERNAL COMPLAINT NUMBER (B)(4). | -incorrect display value -inadequate message display -incorrect measurement | -break -failure of device to self-test -no flow | Viewer
| FDA | Event Description: ON (B)(6) 2021 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS. THE POWERJACKET WINDOW CAN BECOME OUT OF SYNC AND DISPLAY THE DETAILS FOR THE PREVIOUSLY LOADED PATIENT. ON FEBRUARY 05, 2021 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. ON FEBRUARY 11, 2021, THE RISK ASSESSMENT WAS REVISED BASED ON AVAILABILITY OF NEW INFORMATION. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -inappropriate window level -unable to access -device malfunction | -failure of device to self-test -detachment of device component -no flow | Viewer
| FDA | Event Description: ON (B)(6) 2021 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER NOTIFICATION REGARDING AN ADVERSE EVENT RELATED TO SYNAPSE PACS. TWO REPORTS CHANGED FROM A QUALITATIVE EJECTION FRACTION (EF) VALUE OF 55-60% TO ONE REPORTED AS <20%. THE ISSUE LED TO ONE PATIENT RECEIVING AN UNNECESSARY INVASIVE CORANARY ANGIO PROCEDURE. A SECOND PATIENT WAS UNAFFECTED BECAUSE THE ERROR WAS CAUGHT PRIOR TO ASSESSMENT. THE PATIENT HAS BEEN REPORTED TO US AS BEING HEALTHY WITH NO ISSUES AND WAS DISCHARGED. Manufacturer Narrative: UNKNOWN. THE EF VALUE IS SELECTED BY THE USER VIA A DROP-DOWN MENU. ONCE THE DROP-DOWN MENU HAS BEEN SELECTED, THE EF VALUE CAN BE CHANGED BY THE UP AND DOWN ARROW KEYS. THE USER DID NOT CHANGE FOCUS FROM THE DROP-DOWN MENU IN THE ADVANCED REPORTING WINDOW TO THE VIEWER WINDOW BEFORE USING THE ARROW KEYS IN AN ATTEMPT TO MOVE TO THE NEXT IMAGE. THIS CAUSED THE SELECTED EF VALUE TO UNKNOWINGLY CHANGED. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -patient experienced adverse event -lead to adverse trend -device malfunction | -device alarm system issue -break -missing test results | Viewer
| FDA | Event Description: ON 09-MAY-2016 PHILIPS RECEIVED SERVICE CALL FROM A CUSTOMER REPORTING THAT ORIENTATIONS MARKED IN SOME IMAGES LOOKS WRONG IN TWO EBW-NM STATIONS. THIS ISSUE IS STILL UNDER INVESTIGATION. THERE WAS NO REPORT ON PATIENT INVOLVEMENT WITH THIS EVENTS, AND NO REPORT OF PATIENT HARM. ON THE 03-MAY-2016 WE ALREADY GOT A FEEDBACK FROM THIS CUSTOMER ABOUT THE SAME ISSUE IN ONE OF THESE SYSTEMS AND WE ALREADY REPORTED THE INCIDENT AS MFR. REPORT #:1525965-2016-00027. THIS REPORT AIMED TO COVER THE SECOND SYSTEM. Manufacturer Narrative: AN INVESTIGATION WAS PERFORMED BY THE ENGINEERING TEAM AND CLINICAL APPLICATION EXPERTS. FURTHER INVESTIGATION FOUND THAT THIS ISSUE IS SPECIFIC TO SAVING SAGITTAL VOLUME FILES. THIS ISSUE WILL ONLY OCCUR IF THE USER SAVES THE RECONSTRUCTED DATA IN THE SAGITTAL ORIENTATION AND THEN LOADS THE SAGITTAL VOLUME INTO THE VIEWING APPLICATION. CLINICAL APPLICATIONS TRAINS USERS TO ONLY SAVE AND USE THE TRANSVERSE IMAGE VOLUME FILES AFTER EXECUTING THE RECONSTRUCTION APPLICATION (AUTOSPECT PRO). THE ENGINEERING TEAM DETERMINED THAT THE RISK IS ACCEPTABLE REGARDING THIS ISSUE. THE CUSTOMER IS CURRENTLY USING THE SYSTEM AND THEY HAVE BEEN INSTRUCTED BY CLINICAL APPLICATIONS ON HOW TO AVOID THE ISSUE. Manufacturer Narrative: (B)(4). | -device marked incorrectly -wrong print -configuration confused | -no display or display failure -device markings issue -break | Viewer
| FDA | Event Description: ON 13-JUN-2017 A CUSTOMER REPORTED THAT IMAGES THAT APPEARED VERY BRIGHT (WASHED OUT) WHEN LOADED IN CLEARCANVAS WORKSTATION PERSONAL EDITION 13.1. SHE RECEIVED THE IMAGES VIA A DVD. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: A SYNAPTIVE REP USED "GOTOASSIST" TO VIEW IMAGES DIRECTLY. THE REP FOUND THE WINDOW CENTER AND WINDOW WIDTH DICOM TAG VALUES WERE SET TO 2047 AND 4096 RESPECTIVELY. THE PROBE TOOL SHOWED THAT THE BRIGHTEST WHITE HAD VALUE OF 60-70, BLACKS WERE 256. MANUAL WINDOW/LEVEL APPLIED TO MAKE IMAGE VISIBLE WAS IN 250/150 RANGE. THE REP FOUND: BITS STORED (0028,0101) VALUE SET TO 12. AFTER CORRECTING THE VALUE TO 16 AND SAVING IMAGE IT DISPLAYED PROPERLY UPON RELOADING. THE ISSUE WAS FOUND TO BE SIMILAR TO FEED-3544 (REPORTED AS 3012075008-2020-00001). IMAGE IS ACTUALLY 16-BIT, BUT DICOM HEADER TAG (0028,0101) INCORRECTLY STATES IT IS 12-BIT. LEFT CUSTOMER TO CONTACT THE SITE THAT PRODUCED THE DVD TO SEE WHY THE DICOM TAG VALUE DOES NOT MATCH WHAT IS IN THE PIXEL DATA. THIS WAS LOGGED AS AND ENHANCEMENT REQUEST FOR IMPROVED INTEROPERABILITY WITH OTHER DICOM VENDORS. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME JPEG 2000 CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. THIS DEFECT IS ONLY ENCOUNTERED IN SELECT SCENARIOS, WHERE A MODALITY PRODUCES IMAGES THAT ARE COMPRESSED USING JPEG 2000, AND THE IMAGE PIXEL DATA IS LESS THAN 16-BIT, AND THE "BITNESS" OF THE COMPRESSED DATA STREAM DOES NOT MATCH THAT OF THE IMAGE PIXEL DATA. THE OUTCOME OF THE ABOVE IS A LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA, WHICH CAUSES THE SYMPTOMS THAT WERE REPORTED BY THE CUSTOMER. THE ROOT CAUSE IS THAT A SOFTWARE DEFECT FOUND IN THE DEVICE THAT IS ENCOUNTERED WHEN IT IS USED WITH NON-DICOM COMPLIANT JPEG 2000 COMPRESSED IMAGES. THE SOFTWARE IMPROPERLY USES THE DICOM BIT DEPTH (I.E. THE BITS STORED TAG) TO DECOMPRESS THE COMPRESSED PIXEL DATA STREAM FOR DISPLAY, INSTEAD OF THE BIT DEPTH THAT IS ENCODED IN THE COMPRESSED PIXEL DATA STREAM ITSELF. IN NON-DICOM COMPLIANT IMAGES WHERE THE DICOM AND COMPRESSED PIXEL DATA STREAM BIT DEPTHS DO NOT MATCH, THE SOFTWARE OUTPUTS AN IMAGE WITH SOME LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA. IN THESE SCENARIOS, THE MODALITY PRODUCING THE IMAGE IS ITSELF ALSO NOT COMPLIANT TO DICOM PS3.5 8.2.4. MDR REPORTS: 3012075008-2020-00001, 3012075008-2020-00002, AND 3012075008-2020-00004 HAVE THE SAME ROOT CAUSE. A CAPA HAS BEEN OPENED AND RECALL INITIATED TO ADDRESS THE ISSUE IN AFFECTED DEVICES. | -image brightness out of specification -distorted images -loss of image | -failure of device to self-test -human-device interface issue -break | Viewer
| FDA | Event Description: ON DECEMBER 02, 2020 FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE CARDIOVASCULAR. STUDY WAS SENT IN FOR A NEW PATIENT WITH THE CORRECT MRN. INCORRECT MRN IS SHOWING IN CV CLIENT. ON DECEMBER 16, 2020, A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY, OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. THE NUMERIC ID ASSIGNED TO A PATIENT RECORD COULD BE REUSED ONCE A NEWER PATIENT RECORD WAS MERGED TO AN OLDER PATIENT RECORD. WHEN THIS ACTION IS DONE, THE NEWER ID IS MADE AVAILABLE AGAIN. THIS WAS NOT AN ISSUE PRIOR TO ADVANCED REPORTING. WITH THE INTRODUCTION OF ADVANCED REPORTING THE NUMERIC ID IS NOT RELEASED WITHIN THE ADVANCED REPORTING DATABASE SCHEMA AND THAT IS HOW THE MISMATCH BETWEEN THE PATIENT DEMOGRAPHIC DATA CAN OCCUR. REF: INTERNAL COMPLAINT NUMBER COMP (B)(4). | -loss of function -inadequate device information -device malfunction | -failure of device to self-test -component missing -Torn Material | Viewer
| FDA | Event Description: ON DECEMBER 15, 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. THE POWERJACKET WINDOW CAN BECOME OUT OF SYNC AND DISPLAY THE DETAILS FOR THE PREVIOUSLY LOADED PATIENT. ON JANUARY 06, 2021 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS A ONE TIME OCCURRENCE THAT COULD NOT BE REPLICATED IN THE FMSU LAB OR BY THE CUSTOMER. HOWEVER, IT IS SIMILAR IN NATURE TO OTHER PATIENT MISMATCH ISSUES FOUND IN THIS PRODUCT. FUJIFILM BELIEVES THAT THIS ISSUE IS LIKELY CAUSED BY THE SAME DEFECT. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -inappropriate window level -unable to connect -device malfunction | -communication or transmission issue -failure of device to self-test -component missing | Viewer
| FDA | Event Description: ON FEBRUARY 7, 2017, AGFA BECAME AWARE OF AN EVENT IN WHICH A CUSTOMER HAS LOGGED AN FDA MEDWATCH FDA 3500 FORM STATING THE AGFA IMPAX SOFTWARE IS AT FAULT FOR AN IMPROPERLY PERFORMED MEDICAL PROCEDURE. AT THIS TIME, THE ONLY INFORMATION WE WERE PROVIDED WITH WAS THAT THE FDA MEDWATCH 3500 FORM WAS SUBMITTED TO THE FDA AND THAT AN UNNECESSARY MEDICAL PROCEDURE HAD BEEN PERFORMED. ON FEBRUARY 24, 2017, THE CUSTOMER PROVIDED AGFA WITH THE MEDWATCH FDA 3500 FORM WHICH PROVIDED THE DETAILS OF THE SITUATION. THE CUSTOMER REPORTED A SINGLE IMAGE WITHIN A CT SERIES WAS FLIPPED AND SAVED BY A RADIOLOGIST DURING INTERPRETATION. THE CUSTOMER SUBSEQUENTLY INDICATED THERE WAS NOT A SIGNIFICANT VISUAL INDICATION THAT THE IMAGE HAD BEEN SAVED IN THE OPPOSITE DIRECTION OF ALL OTHER IMAGES IN THE SERIES WHICH THEN LED TO A PATIENT HAVING SURGERY ON THE WRONG SIDE. THE PRELIMINARY INVESTIGATION FROM AGFA SHOWS THIS TO BE A USER ERROR AND THAT THE PRODUCT IS WORKING AS DESIGNED. INVESTIGATION IS UNDER WAY TO DETERMINE THE ROOT CAUSE. A SUPPLEMENTAL REPORT WILL BE PROVIDED.Event Description: THIS SUPPLEMENTAL REPORT # 1 IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION RELATED TO THE DESCRIPTION OF THE OVERALL EVENT AND TO PROVIDE THE ROOT CAUSE. Manufacturer Narrative: ON FEBRUARY 7, 2017, AGFA HEALTHCARE BECAME AWARE OF AN EVENT FOR WHICH A CUSTOMER LOGGED AN FDA MEDWATCH FDA 3500 FORM STATING THE AGFA IMPAX SOFTWARE WAS INVOLVED IN AN IMPROPERLY PERFORMED MEDICAL PROCEDURE. AT THAT TIME, THE ONLY INFORMATION PROVIDED TO AGFA WAS THAT THE FDA MEDWATCH 3500 FORM WAS SUBMITTED TO THE FDA AND THAT AN UNNECESSARY MEDICAL PROCEDURE HAD BEEN PERFORMED. ON FEBRUARY 24, 2017, THE CUSTOMER PROVIDED AGFA WITH THE MEDWATCH FDA 3500 FORM, WHICH CONTAINED DETAILS OF THE SITUATION. AGFA HAS SINCE CONFIRMED THAT THERE HAVE BEEN NO OTHER REPORTS OF SIMILAR COMPLAINTS FROM THIS CUSTOMER OR ANY OTHER CUSTOMER. ADDITIONALLY, AGFA HAS CONFIRMED THAT THIS WAS THE FIRST REPORT OF THIS SPECIFIC INCIDENT FROM THIS CUSTOMER. THE CUSTOMER REPORTED THAT A SINGLE IMAGE WITHIN A CT SERIES WAS FLIPPED AND SAVED BY A RADIOLOGIST DURING INTERPRETATION. THE CUSTOMER SUBSEQUENTLY INDICATED THERE WAS NOT A SIGNIFICANT VISUAL INDICATION THAT THE IMAGE HAD BEEN SAVED IN THE OPPOSITE DIRECTION OF ALL THE OTHER IMAGES IN THE SERIES. IT SHOULD BE NOTED THAT THERE ARE, IN FACT, MANY OBVIOUS VISUAL INDICATORS OF IMAGE MANIPULATION THAT ARE DESCRIBED LATER IN THIS DOCUMENT. UNFORTUNATELY, THE USER OVERLOOKED THESE INDICATORS WHICH LED TO A PATIENT HAVING SURGERY ON THE WRONG SIDE OF THE BODY. BEYOND THE UNNECESSARY TREATMENT ITSELF, THE CUSTOMER HAS CONFIRMED THAT THE PATIENT DID NOT EXPERIENCE ANY INJURY OR ADVERSE OUTCOME AS A RESULT OF THE SURGERY. INVESTIGATION BY AGFA HEALTHCARE CONFIRMS THAT THIS WAS A USER ERROR AND WAS NOT CAUSED BY THE IMPAX SOFTWARE, WHICH FUNCTIONED AS DESIGNED. THE MANIPULATIONS PERFORMED BY THE CUSTOMER ARE WITHIN THE EXPECTED BEHAVIOR OF THE IMPAX PRODUCT AND ARE CLEARLY DESCRIBED IN THE IMPAX KNOWLEDGE BASE (IMPAX USER MANUAL). THERE ARE MANY TOOLS AVAILABLE TO USERS THAT CAN BE USED TO MANIPULATE INDIVIDUAL IMAGES IN ORDER TO FACILITATE VARIOUS WORKFLOWS. HOWEVER, IT IS IMPOSSIBLE FOR AGFA HEALTHCARE TO ENFORCE HOW USERS APPLY THESE TOOLS. ALTHOUGH THE CUSTOMER HAS REQUESTED THAT AGFA UPDATE THE IMPAX SOFTWARE SUCH THAT IMAGE MANIPULATIONS APPLY ONLY TO A STACK/SERIES OF IMAGES AND NEVER TO A SINGLE IMAGE, THE SINGLE-IMAGE MANIPULATION FUNCTIONALITY IS REQUIRED FOR SOME USE CASES (SUCH AS MAMMOGRAPHY). AS SUCH, REMOVING THE FUNCTIONALITY (AS PROPOSED BY THE CUSTOMER) WOULD REDUCE THE EFFECTIVENESS OF THE PRODUCT. IN FACT, THE PROGRAM OFFICE OF THE U.S. MILITARY HAS LISTED THE SINGLE-IMAGE MANIPULATION FUNCTIONALITY AS A REQUIREMENT VIA THE GSPS (GREYSCALE SOFTCOPY PRESENTATION STATE STORAGE) STANDARD FOR IMAGE DISPLAY. WHEN USING THE TOOL, THERE ARE MULTIPLE VISUAL INDICATORS TO ALERT THE USER TO IMAGE MANIPULATION, INCLUDING CURSOR ICONS, MENU BUTTONS, LEFT/RIGHT LABELS, AND IMAGE MANIPULATION LABELS. ONCE THE IMAGE(S) ARE SAVED, THE LEFT/RIGHT LABELS AND IMAGE MANIPULATION LABELS REMAIN, AND ARE DEEMED TO BE SUFFICIENT VISUAL INDICATORS. ADDING ADDITIONAL VISUAL INDICATORS WOULD START TO CROWD THE IMAGES, ULTIMATELY REDUCING THE EFFECTIVENESS OF THE PRODUCT. ALTHOUGH NO CHANGES WILL BE MADE TO THE IMPAX PRODUCT IN THIS REGARD, AGFA HEALTHCARE HAS CREATED A QUICK REFERENCE GUIDE FOR THE (B)(6), WHICH WILL BE PRESENTED TO THE CUSTOMER IN ORDER TO RAISE THEIR AWARENESS OF THIS PARTICULAR FUNCTIONALITY. NO FURTHER ISSUES HAVE BEEN REPORTED. | -log error -fail to log out -software defect | -device inoperable -fail-safe design failure -Contamination During Use | Viewer
| FDA | Event Description: ON FRIDAY (B)(6) 2016 (B)(4) CALLED TO REPORT THAT THE XG-1 (SMARTCR) AT THEIR CUSTOMER SITE, (B)(6) ALLEGEDLY EXPERIENCED A SMOKING SMART CR ( SERIAL NUMBER:(B)(4)) INCIDENT ON (B)(6). AT THE POINT THAT THIS REPORTED SMOKING INCIDENT OCCURRED THE SITE HAD NOT HAD CAR-102 PERFORMED. THE DEALER HAS BEEN ON SITE AND REPLACED THE POWER SUPPLY. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. Manufacturer Narrative: THIS REPORT IS BEING MADE OUT OF AN ABUNDANCE OF CAUTION. DEALER SERVICE ENGINEER VISITED SITE AND EVALUATED DEVICE TO DETERMINE CAUSE OF EVENT. POWER SUPPLY WAS REPLACED AND THE DEVICE WAS RETURNED TO USE. THE SITE INFORMATION IS: (B)(6). A CR READER UNIT IS USED FOR READING X-RAY IMAGES THAT HAVE BEEN EXPOSED ON A CR IMAGE PLATE. THE SMARTCR IS ALSO KNOWN AS AN XG-1 OR A CR-IR 346. THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. FUJIFILM SUBMITTED THE CORRECTION REPORT IDENTIFIED IN RECALL ON MARCH 22, 2016. THIS ADDRESSES A CORRECTIVE ACTION THAT WILL REPLACE THE POWER SUPPLY IN ALL SMART CR SYSTEMS. NOTIFICATION TO USERS WAS MADE. | -patient experienced adverse event -fuse blown -device malfunction | -Defective Component -Device displays error message -break | Viewer
| FDA | Event Description: ON JUNE 26, 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. WHEN OPENING THE POWERJACKET AND SELECTING A PATIENT OFF THE WORKLIST, THE POWERJACKET DISPLAYED INCORRECT PATIENT INFORMATION. THE DOCTOR INDICATED THAT THE ISSUE OCCURRED AGAIN WHEN SELECTING ANOTHER PATIENT AND THE POWERJACKET AGAIN DISPLAYED A DIFFERENT PATIENT. THE FUJIFILM CUSTOMER SERVICE REPRESENTATIVE DID NOT IDENTIFY ANY ISSUES WITHIN THE IIS LOGS OF THE CUSTOMER AND THE ISSUE COULD NOT BE REPRODUCED. ON JULY 08, 2020 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT INVOLVEMENT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. A BUG FIX WAS DEVELOPED AND PROVIDED TO FIX THE ISSUE. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: (B)(4). | -fail to supply power -incorrect product classification -unable to connect | -failure of device to self-test -communication or transmission issue -sticking | Viewer
| FDA | Event Description: ON MARCH 15, 2021 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) WAS MADE AWARE OF AN INCIDENT WITH SYNAPSE PACS. WHEN A USER OPENS A STUDY WITH AN URL-CALL FROM THEIR HIS (SAP), IMAGES ARE SHOWN FROM PATIENT A BUT THE INFORMATION PANE SHOWS STUDIES FROM PATIENT B. ON (B)(6) 2021 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL; HOWEVER THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. FUJIFILM INITIATED A RECALL ON 3/2/2021 TO ALERT CUSTOMERS OF SEVERAL PATIENT MISMATCH ISSUES EXISTING IN ALL SYNAPSE PACS 5 VERSIONS UP TO 5.7.210. FUJIFILM SUBMITTED C&R REPORT (1000513161-03/11/2021-001-C) TO FDA, WHICH HAS BEEN CLASSIFIED AS CLASS II AND ASSIGNED RECALL NUMBER Z-1348-2021. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -inappropriate user -stored inappropriately -unable to edit | -fitting problem -human-device interface issue -compliance issue | Viewer
| FDA | Event Description: ON NOVEMBER 05, 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. WHEN OPENING THE POWERJACKET AND SELECTING A PATIENT OFF THE WORKLIST, THE POWERJACKET INTERMITTENTLY DISPLAYED INCORRECT PATIENT INFORMATION. ON NOVEMBER 06, 2020, A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY, OR DEATH ASSOCIATED WITH THIS EVENT. THE ISSUE IS CONSIDERED HIGHLY DETECTABLE BY A HEALTHCARE PROFESSIONAL, HOWEVER, THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS EVALUATED, AND REPLICATED IN THE FMSU LAB. THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS IS A KNOWN ISSUE WITH AN ONGOING FIELD ACTION, AND THE CUSTOMER WAS PROVIDED THE FIELD ACTION VERIFICATION FORM, AND WILL BE SCHEDULED FOR AN UPGRADE. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. INTERNAL COMPLAINT NUMBER: (B)(4). | -fail to generate indication -incorrect product classification -fail to supply power | -failure of device to self-test -communication or transmission issue -sticking | Viewer
| FDA | Event Description: ON NOVEMBER 09, 2020 FUJIFILM MEDICAL SYSTEMS USA, INC. (FMSU) SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS POWERJACKET. THE POWERJACKET WINDOW CAN BECOME OUT OF SYNC AND DISPLAY THE DETAILS FOR THE PREVIOUSLY LOADED PATIENT. ON NOVEMBER 25, 2020 A RISK ASSESSMENT WAS PERFORMED TO INVESTIGATE THE RISK TO PATIENT SAFETY. THERE WAS NO PATIENT IMPACT, SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. THE ISSUE WAS EVALUATED AND REPLICATED IN THE FMSU LAB; THE CAUSE WAS TRACED TO A SOFTWARE CONFIGURATION. THIS ISSUE IS VERY RARE AND REQUIRES MANY TRIES WITH LARGE STUDIES TO REPRODUCE; IT IS CONSIDERED TO BE A VERY RARE OCCURRENCE IN A CLINICAL ENVIRONMENT. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REF: INTERNAL COMPLAINT NUMBER (B)(4). | -inappropriate window level -component missing -incorrect product classification | -communication or transmission issue -failure of device to self-test -component missing | Viewer
| FDA | Event Description: ON NOVEMBER 14, 2017, AGFA WAS CONTACTED BY (B)(6) FROM (B)(6) ABOUT AN EVENT THAT HAD OCCURRED (B)(6) 2017 WHILE VIEWING STUDIES IN THE CLINICAL SIDE BAR OF AGFA'S ENTERPRISE IMAGING CLIENT. THE CUSTOMER HAD REPORTED TO NHS THAT THEY REALIZED ON (B)(6) 2017, THAT A PATIENT'S IMAGES HAD BEEN MISINTERPRETED BY THE RADIOLOGIST APPROXIMATELY FIVE MONTHS EARLIER. INVESTIGATION BY AGFA AND THE CUSTOMER REVEALED ON (B)(6) 2017, A PATIENT UNDERWENT A CT PULMONARY ANGIOGRAM (CTPA) DURING A HOSPITALIZATION FOR A LOWER RESPIRATORY TRACT INFECTION. THE CTPA SHOWED NO EVIDENCE OF PULMONARY EMBOLISM. NOTE THAT THIS PATIENT WAS ALSO DIAGNOSED WITH PLEURAL ADENOCARCINOMA OF THE LUNG IN (B)(6) 2017. THE CANCER WAS TREATABLE, BUT NOT CURABLE, WITH A LIFE EXPECTANCY OF 9 TO 20 MONTHS. ON (B)(6) 2017, THE SAME PATIENT WAS ADMITTED TO THE HOSPITAL WITH INCREASING SHORTNESS OF BREATH. ANOTHER CPTA WAS PERFORMED AND REPORTED ON (B)(6) 2017 AS BEING SIMILAR TO THAT FROM (B)(6) 2017, WITH NO EVIDENCE OF ACUTE PULMONARY THROMBOEMBOLISM. THE PATIENT WAS DIAGNOSED WITH A WORSENING OF THEIR EXISTING COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE) AND WAS TREATED ACCORDINGLY. THE PATIENT WAS NOT TAKING ANTI-COAGULANTS AT THIS TIME. ON (B)(6) 2017, THE PATIENT HAD A THORAX, ABDOMEN, AND PELVIS CT SCAN TAKEN TO ASSESS THE EFFECTS OF CHEMOTHERAPY. IT WAS REPORTED ON SEPTEMBER 12, 2017 AND ON SEPTEMBER 13, 2017, AN ADDENDUM WAS ADDED TO THE (B)(6) 2017 REPORT, INDICATING THE IMAGES HAD BEEN MISINTERPRETED AND THE SCAN DID, IN FACT, SHOW MULTIPLE PULMONARY EMBOLI. THE RADIOLOGIST WHO REPORTED THE (B)(6) 2017 SCAN MISTAKENLY LOOKED AT THE IMAGES FROM THE (B)(6) 2017, SCAN AND REPORTED BASED ON THOSE INSTEAD OF THE CORRECT IMAGES. THIS WAS NOT REPORTED TO AGFA HEALTHCARE AT THIS TIME. ON APRIL 20, 2018, AGFA WAS CONTACTED BY THE CUSTOMER REQUESTING INPUT INTO THE ROOT CAUSE INVESTIGATION OF THE ABOVE DESCRIBED EVENT. AFTER REVIEWING THE SEQUENCE OF THIS OVERALL EVENT, AGFA OPENED COMPLAINT (B)(4) TO ASSESS FOR CLINICAL RISK AND ALSO OPENED NONCONFORMITY (B)(4) TO ADDRESS INSUFFICIENT COMPLAINT DOCUMENTATION. THE PRIMARY ROOT CAUSE WAS IDENTIFIED AS CUSTOMER MISUSE. THE CUSTOMER HAD RECEIVED TRAINING AND WAS MADE AWARE OF THIS ISSUE AND THAT THEY MAY NEED TO SORT OUT PRIOR STUDIES IN PROPER DATE SEQUENCE. IN MAY 2017, THE ISSUE WAS ASSESSED BY AN AGFA HEALTHCARE CLINICAL ANALYST AS POSING NO RISK TO PATIENT HEALTH. ADDITIONAL ROOT CAUSE WAS A KNOWN ISSUE WHICH TRIGGERS A CHANGE OF THE SORT ORDER OF THE STUDIES IN THE CLINICAL SIDE BAR (CSB). THE ISSUE WAS CORRECTED IN ENTERPRISE IMAGING 8.1.1. THE CUSTOMER REPORTED IN THIS EVENT RECEIVED THIS CORRECTION ON JULY 27, 2017. AFTER REVIEWING THIS EVENT, AGFA OPENED COMPLAINT (B)(4) TO ASSESS THE CLINICAL RISK FOR THE EVENT AND ALSO OPENED NONCONFORMITY (B)(4) TO ADDRESS INSUFFICIENT COMPLAINT DOCUMENTATION. THERE HAS BEEN NO REPORTED HARM TO PATIENT OR USER DURING THIS EVENT. | -incorrect product classification -device malfunction -inadequate labeling | -Device displays error message -leak -application program issue | Viewer
| FDA | Event Description: ON SEPTEMBER 27, 2019, FUJIFILM MEDICAL SYSTEMS USA, INC. SERVICE DEPARTMENT RECEIVED A CUSTOMER INQUIRY FOR ASSISTANCE WITH SYNAPSE PACS TOMO MARKERS AND WAS DETERMINED NOT TO BE A COMPLAINT BASED ON THE SERVICE TICKET. ON JANUARY 24, 2020, FUJIFILM MEDICAL SYSTEMS USA, INC. ENGINEERING DEPARTMENT WAS NOTIFIED OF CUSTOMER CONCERNS REGARDING SYNAPSE PACS TOMO ISSUES. INITIAL INVESTIGATION INDICATED THAT THE ORIENTATION MARKER WAS BEING DISPLAYED ON THE WRONG SIDE OF THE TOMO IMAGE. THE RESOLUTION WAS PROVIDED THROUGH A BUG FIX. BASED ON RISK AND SEVERITY OF THE ISSUE AND FIX, THE COMPLAINT WAS DETERMINED NOT TO BE REPORTABLE. ON APRIL 1, 2020, FUJIFILM MEDICAL SYSTEMS USA, INC QUALITY AND REGULATORY DEPARTMENT WAS NOTIFIED OF CUSTOMER CONCERNS REGARDING ADDITIONAL SYNAPSE PACS TOMO ISSUES. UPON FURTHER INVESTIGATION, ON APRIL 28, 2020 IT WAS DETERMINED THAT THERE WERE THREE (3) ISSUES WITH THE LATERALITY. ISSUE 1: INDICATORS WERE MISSING ON LXCCL, RXCCL, LLMO, AND RLMO. THIS ISSUE WAS FIXED IN VERSION (B)(4), HOWEVER, HAVING THE REFERENCE POINTS ON THE SLIDE BAR BROUGHT ANOTHER ISSUE TO LIGHT. ISSUE 2: THE LLMO AND RLMO ARE DISPLAYING LATERALITY INCORRECTLY BECAUSE THE LATERAL AND MEDIAL REFERENCE POINTS WERE SWITCHED, THE SLIDE BAR WAS IN THE INCORRECT DIRECTION, ISSUE 3: THE IMAGE COUNT WAS DISPLAYING IN REVERSE ORDER. THERE WAS NO PATIENT SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. THIS REPORT IS BEING SUBMITTED IN ABUNDANCE OF CAUTION. Manufacturer Narrative: THE ISSUE WAS REPLICATED IN THE FUJIFILM MEDICAL SYSTEMS USA LAB. THE CAUSE WAS TRACED TO THE SOFTWARE NOT READING THE CORRECT DICOM TAB FOR BREAST TOMO IMAGES. A BUG FIX HAS BEEN DEVELOPED AND ALL AFFECTED CUSTOMERS WILL BE NOTIFIED OF THE ISSUE AND PROVIDED THE FIX. IF ANY ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | -fail to service -fail to deliver therapy -loss of function | -no flow -component missing -failure of device to self-test | Viewer
| FDA | Event Description: ON THE (B)(6) 2015 PHILIPS RECEIVED FEEDBACK FROM A CUSTOMER REPORTING THAT WHEN RECONSTRUCTING DATA BY THE INTELLISPACE PORTAL AUTOSPECT_PRO APPLICATION SOME OF THE RESULTS MAY BE DIFFERENT THAN EXPECTED. THE FEEDBACK INDICATE THAT THE AFFECTED DATA IS CREATED ON LEGACY AXIS/IRIX GAMMA CAMERAS MANUFACTURED BY MARCONI MEDICAL SYSTEMS USING 102 DEGREE ACQUISITION PROTOCOL. THERE WAS NO REPORT OF PATIENT INVOLVEMENT WITH THIS EVENT, AND NO REPORT OF PATIENT HARM. Manufacturer Narrative: DATA RECEIVED FROM A PREVIOUS SIMILAR COMPLAINT WAS INVESTIGATED BY ENGINEERING AND REVEALED THAT THE AUTOSPECT PRO APPLICATION WAS ONLY DESIGNED TO RECONSTRUCT CARDIAC SPECT DATA OBTAINED WITH DETECTORS POSITIONED AT 90ø OR 180ø RELATIVE TO ONE ANOTHER. HOWEVER, CERTAIN GAMMA CAMERAS, E.G., THE MARCONI AXIS AND LRIX CAMERAS, PERMIT ACQUISITIONS AT OTHER RELATIVE DETECTOR ANGLES. AFFECTED CUSTOMERS WERE NOTIFIED BY FIELD SAFETY NOTICE (FSN 88100036) SENT ON 09-MAR-2016. FIELD CHANGE ORDER ((B)(4)) WAS RELEASED ON (B)(6) 2016 TO IMPLEMENT A SOFTWARE UPDATE TO ELIMINATE THE PROBLEM. Manufacturer Narrative: (B)(4). | -fail to retrieve data -fail to configure patient information -error message | -unauthorized access to computer system -break -failure of device to self-test | Viewer
| FDA | Event Description: ONE SITE HAS REPORTED MISSING 400-500 STUDIES THAT NEVER MADE IT TO THE ARCHIVE FOR LONG TERM STORAGE BEFORE THE WATERMARKING PROCESS DELETED THEM. THE CODE PURGES SERIES FROM THE RADSUITE (AV) IMAGE CACHE PRIOR TO THE SERIES BEING STORED IN ENTERPRISE CONTENT MANAGER (ECM) LONG TERM IMAGE ARCHIVE, AND STUDIES ARE NOT STORED IN THE LONG TERM ARCHIVE (LTA) OR THE AV IMAGE CACHE FOR FUTURE REFERENCE. Manufacturer Narrative: SUBMITTED THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS TO SECTION. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -missing maintenance -fail to follow process -missing data | -device maintenance issue -improper or incorrect procedure or method -medical gas supply problem | Viewer
| FDA | Event Description: ORIGINAL ISSUE WAS ESCALATED AS A SIMPLE REQUEST FOR SERVICE SUPPORT IN DETERMINING WHY THE SYNGO DYNAMICS PRODUCT COULD NOT RETRIEVE STUDY IMAGES FROM THE SITE'S ARCHIVE SYSTEM. SERVICE INVESTIGATION FOUND THAT THE SITE'S LONG TERM ARCHIVE SYSTEM (NON-SIEMENS PRODUCT) HAD FAILED. THE SITE HAS TURNED OVER THE ARCHIVE DISKS TO A DATA RECOVERY COMPANY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S HISTORICAL DATA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES WILL STILL BE AVAILABLE IN THE SD SYSTEM UNLESS DELETED FROM THE SD SYSTEM BY SITE IT POLICY OR USER ACTION. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. | -fail to upload images -fail to retrieve data -fail to select | -Device Handling Issue -therapeutic or diagnostic output failure -device unsafe to use in environment | Viewer
| FDA | Event Description: ORIGINAL ISSUE WAS ESCALATED TO SIEMENS BY THE CUSTOMER AS A SIMPLE REQUEST FOR SERVICE SUPPORT TO HELP DETERMINE REASON WHY THE SYNGO DYNAMICS PRODUCT COULD NOT RECEIVE EXAMINATIONS. DURING THE INVESTIGATION SIEMENS FOUND THAT THE SITE'S LONG TERM ARCHIVE SYSTEM HAD FAILED. SIEMENS SERVICE ASSISTED THE SITE AND (B)(4) (MANUFACTURER OF THE HARD DRIVE) IN ESTABLISHING A WORKING ARCHIVE DRIVE. HOWEVER, THIS SITE HAS NO CONTRACT WITH SIEMENS AND HAS TURNED OVER THE FAILED DISKS TO A DATA RECOVERY COMPANY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S HISTORICAL DATA. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES WILL STILL BE AVAILABLE IN THE SYNGO DYNAMICS SYSTEM UNLESS DELETED FROM THE SYSTEM BY SITE IT POLICY OR USER ACTIONS. NO INJURY IS ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH WOULD RESULT IN THE NEED FOR A PATIENT RESCAN. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO REASONABLY FORESEEABLE PATH TO OCCURRENCE OF HARM CAN BE DETERMINED. CUSTOMER'S ADDRESS: (B)(6). | -inaccurate syringe marking -unable to obtain readings -fail to provide indicator | -inaccurate delivery -device unsafe to use in environment -environmental control or utility issue | Viewer
| FDA | Event Description: OUR FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITHIN THE ACCEPTABLE RISK THRESHOLD FOR THE PRODUCT, HOWEVER, RELEVANT RISK FOR THE SITE WAS IDENTIFIED. THE SYNGO.PLAZA HAS BEEN CONFIGURED TO USE A PRIMARY (NAS1) AND BACKUP (NAS2) NETWORK ATTACHED STORAGE (NAS) MOUNT POINT FOR ARCHIVING PURPOSES. DUE TO AN UNRECOVERABLE HARDWARE FAILURE ON THE NAS1 A RESTORATION OF ALL DATA WAS PERFORMED FROM THE NAS2. HOWEVER, THE NAS2 DID NOT CONTAIN ALL ARCHIVED DATA AS ON NAS1. 8523 SERIES DATED WITHIN THE TIMEFRAME 2011-2016 WERE NOT IN THE ARCHIVE AND NO LONGER AVAILABLE. THIS IS CONSIDERED A DATA LOSS OF CLINICAL RELEVANT DATA. THE ROOT CAUSE ANALYSIS WAS ABLE TO LIMIT DOWN THE ISSUE TO A MANUAL INTERVENTION ON THE NAS2'S FILESYSTEM, WHICH IS NOT NO LONGER TRACEABLE. THE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE MISSING DATA IN THE NAS2 ARCHIVE. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURED IN THE (B)(6). Manufacturer Narrative: SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITHIN THE ACCEPTABLE RISK THRESHOLD FOR THE PRODUCT, HOWEVER, RELEVANT RISK FOR THE SITE DUE TO DATA LOSS OF CLINICAL RELEVANT DATA WAS IDENTIFIED. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. | -fail to identify risk -incorrect product classification -inappropriate variable value | -fluid leak -device inoperable -protective measure issue | Viewer
| FDA | Event Description: OUR HOSPITAL WAS REVIEWING A PATIENT CONCERN. A PATIENT REPORTED THAT WHEN THEY SQUEEZED THE BALL, THERE WAS NO RESPONSE FROM THE CONTROL ROOM. A PNEUMATIC CALL-BALL ATTACHED TO AN AIR FILLED CORD IS A CONNECTION TO AN EXTERNAL AUDIO SYSTEM AND LIGHT SYSTEM. WHEN THE BALL IS SQUEEZED, AN AUDIBLE SOUND/ ALERT IS SOUNDED, HEARD AND VISUALIZED AS A LIGHT IN THE MRI CONTROL ROOM. THE CALL-BALL IS AN AIR FILLED BULB-BALL THAT IS CONNECTED TO A FLEXIBLE TUBING WHICH UTILIZES A MALE CONNECTOR TO A FEMALE PORT. THIS AIR FILLED TUBING IS CONNECTED TO THE FOOT OF THE MRI PATIENT TABLE. THE PATIENT ACTIVATES THE SYSTEM BY SQUEEZING THE CALL-BALL WHICH ALERTS THE TECHNOLOGIST IN THE CONTROL AREA WITH BOTH VISUAL AND AUDIBLE ALARM AND THEN THIS HAS TO BE DEACTIVATED BY THE MRI TECHNOLOGIST. WE FOUND THAT IF THE CORD IS CRIMPED IN ANYWAY OR CLOSURE OF THE AIR FILLED TUBING OR FAILURE TO SQUEEZE THE BALL WITH SUFFICIENT FORCE DEACTIVATES THE FEATURE AND NO SOUND OR LIGHT IS SEEN. WE ARE REPORTING THIS AS A SAFETY CONCERN. | -delayed response -unable to compress -pin detached | -Device Handling Issue -fluid leak -Contamination During Use | Viewer
| FDA | Event Description: PACS APPLICATION FAILS TO IDENTIFY THAT AN EXTERNAL RADIOLOGY REPORT HAS BEEN RECEIVED AND PERMITS A RADIOLOGIST TO CONTINUE DICTATING A CASE THAT HAS BEEN ALREADY FINALIZED. THIS ALLOWS THE DISCREPANT REPORTS TO EXIST (1 INTERNAL TO OUR PACS AND 1 FROM THE LOCAL RADIOLOGIST WHO HAS THEIR REPORT ENTERED INTO THE PERMANENT MEDICAL RECORD). THE LOCALLY GENERATED REPORT IS NOT ARCHIVED IN PACS AND ANY SUBSEQUENT REVIEW OF THE CASE AS A PRIOR OR FOR QA PURPOSES, WILL HAVE CONCLUSIONS BASED ON THE INTERNALLY GENERATED REPORT WHICH IS NOT THE REPORT OF RECORD. | -fail to analyze medical imaging -fail to retrieve medical imaging -fail to generate notification | -improper or incorrect procedure or method -device inoperable -failure of device to self-test | Viewer
| FDA | Event Description: PATIENT HAD A PROCEDURE DONE IN CATHETER LABORATORY. UNABLE TO FIND THE FILMS ON THE MCKESSON ARCHIVAL SYSTEM. THERE WAS AN ERROR MESSAGE ON THE COMPUTER DURING THIS TIME. THE FILMS COULD NOT BE TRANSFERRED OVER UNTIL THE ERROR MESSAGE WAS CLEARED. ONCE THE ERROR MESSAGE WAS CORRECTED THE FILMS WERE AVAILABLE. ALL FILMS FOR THIS PATIENT ARE PRESENT. THERE IS A PROCESS IN PLACE, SO THAT THIS WILL NOT HAPPEN. FILMS ARE TO BE CHECKED DAILY TO BE SURE THEY HAVE TRANSFERRED OVER TO MCKESSON. ANY MISSING FILMS DUE TO ERROR MESSAGES SHOULD BE ADDRESSED IN REAL TIME. CATH LAB ASSISTANT NURSE MANAGER HUDDLED WITH TEAM MEMBERS REGARDING THIS PROCESS AND TO BE SURE THAT THEY ARE FOLLOWING IT DAILY AND WITH ALL CASES. | -error message -inadequate film processing -device malfunction | -crack -Device displays error message -improper or incorrect procedure or method | Viewer
| FDA | Event Description: PATIENT PRESENTED FOR A MAMMOGRAM AND ULTRASOUND. HE INDICATED THAT HE COULDN'T STAND FOR THE EXAM, SO HE WAS POSITIONED SITTING DOWN. WHILE SETTING UP FOR THE IMAGE THE PATIENT WAS HEARD YELLING AND IT WAS NOTED THE C-ARM OF THE UNIT HAD COME DOWN ON THE PATIENT'S LEFT FEMUR. THE UNIT WAS RAISED AND PATIENT COMPLAINED OF LEFT HEEL PAIN. HE WAS SUBSEQUENTLY TAKEN TO THE EMERGENCY ROOM WHERE IMAGING STUDIES CONFIRMED A COMMINUTED CALCANEAL FRACTURE. | -fail to enable mri setting -fail to interpret images -image orientation incorrect | -Device Handling Issue -fluid leak -poor quality image | Viewer
| FDA | Event Description: PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. DUE TO AN IDENTIFIED HARDWARE FAILURE IN NETWORK ATTACHED STORAGE (NAS), THERE ARE MISSING FOLDERS CONTAINING PATIENT IMAGES IN THE CONFIGURED NAS (NAS1). A BACKUP NAS IS CONFIGURED, HOWEVER, IT IS MANAGED BY A 3RD PARTY APPLICATION AND NOT CONFIGURED IN THE SYNGO.PLAZA SYSTEM. CONTENTS OF NAS1 AND NAS2 ARE IDENTICAL. BOTH NAS SYSTEMS HAVE A LARGE AMOUNT OF EMPTY FOLDERS FOR DATES RANGING BETWEEN JANUARY 6, 2017 AND APRIL 19, 2017. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED DATA; HOWEVER, THE AFFECTED DATA IS CONSIDERED AS CLINICALLY RELEVANT. THE PRODUCT SYNGO.PLAZA DID NOT CONTRIBUTE TO THE DATA LOSS, AS THE ISSUE IS RELATED TO A HARDWARE FAILURE AND THE SYNGO.PLAZA IS A SOFTWARE SOLUTION ONLY (B)(6). | -accessory missing -loss of data -loss of performance | -device inoperable -particulates -defective component | Viewer
| FDA | Event Description: PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO MIX-UP OCCURRED, HENCE, NO PATIENT RESCAN WAS REQUIRED. CUSTOMER REPORTED THAT CERTAIN STUDIES CANNOT BE FOUND IN SYNGO.PLAZA PATIENT BROWSER. THE SYSTEM WAS MIGRATED AND UPGRADED FROM AN OLD SYNGO.PLAZA VERSION TO A NEW SERVER AND NEW SW VERSION VB20A IN (B)(6) 2017. DURING THE ANALYSIS OF THIS ISSUE IT HAS BEEN IDENTIFIED THAT AN INCORRECT LICENSE FILE IS PRESENT ON THE SYSTEM. THUS STUDIES WERE NOT ARCHIVED TO SYNGO.PLAZA'S LONG TERM ARCHIVE. THE AFFECTED STUDIES WERE PRESENT IN THE ONLINE STORAGE FOR REPORTING OF CURRENT EXAMINATIONS. AFTER A CERTAIN FILL LEVEL WAS REACHED IN THE ONLINE STORAGE, THE STUDIES WERE REMOVED BY AUTODELETION FUNCTION. THE AFFECTED STUDIES RANGE FROM 01/2017 TO 07/2017 AND ARE NOT AVAILABLE FOR PRIOR COMPARISON. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. ACTIVITIES ARE ONGOING TO EVALUATE THE EXACT AMOUNT OF AFFECTED STUDIES AND THE POSSIBILITY OF RESTORING THEM. NO ISSUE FROM SYNGO.PLAZA SOFTWARE SINCE SYNGO.PLAZA DID NOT CONTRIBUTE TO THE DATA LOSS AT THE SITE. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. (B)(6). | -component missing -device malfunction -defective component | -device operates differently than expected -no audible alarm -no flow | Viewer
| FDA | Event Description: PERMANENT LOSS OF 15 STUDIES WAS DISCOVERED IN THE SYNGO.PLAZA'S DATABASE. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE DATA LOSS AS THE ISSUE IS RELATED TO A HARDWARE FAILURE AND THE SYNGO.PLAZA SYSTEM IS A SOFTWARE ONLY SOLUTION. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. THE LOST DATA IS CONSIDERED CLINICALLY RELEVANT AS THE STUDIES WERE BETWEEN 1 AND 6 MONTHS OLD AND MAY BE STILL REQUESTED FOR COMPARISON. THE REPORTED EVENT OCCURRED IN (B)(6. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. ON 2017-10-26 CUSTOMER REPORTED MISSING OR CORRUPTED IMAGES APPEARING IN SYNGO.PLAZA. DURING THE ANALYSIS IT WAS IDENTIFIED THAT ON THE (B)(6) 2017, ONE OF THE HARD DRIVES USED BY THE SYNGO.PLAZA'S SHORT TERM STORAGE HAD FILE SYSTEM PROBLEMS CAUSED BY HARDWARE FAILURES. (B)(6. | -loss of therapy -device damaged -device malfunction | -use of device issue -device alarm system issue -defective item | Viewer
| FDA | Event Description: POTENTIAL HARM WITH SOFTWARE; I WORK FOR THE DEPARTMENT OF (B)(6). OUR RADIOLOGY PACS IS SUPPLIED BY INTELERAD MEDICAL SYSTEMS. IN THE PAST, A CODE DEFECT WAS IDENTIFIED THAT ALLOWED OUR RADIOLOGISTS TO DICTATE A RADIOLOGY REPORT ON A STUDY THAT HAD ALREADY BEEN READ BY A LOCAL RADIOLOGIST. THE PROBLEM HAS RETURNED AND HAS RESULTED IN CONFLICTING REPORTS AND CALLBACKS TO ORDERING PROVIDERS FOR CRITICAL RESULTS, WITH RESULTS THAT DIFFERED FROM THE LOCAL RADIOLOGIST REPORT, LEADING TO CONFUSION ON THE PART OF THE ORDERING PROVIDER. THIS SOFTWARE DEFICIENCY COULD RESULT IN A DIFFERING DIAGNOSIS AND/OR CHANGE IN TREATMENT PLANS IF THE LOCAL RADIOLOGIST REPORT DIFFERS SIGNIFICANTLY FROM THE (B)(6) RADIOLOGY REPORT. SINCE THIS IS NOW THE SECOND TIME THAT THIS PROBLEM HAS OCCURRED AND IS A PATIENT SAFETY ISSUE, I AM ELEVATING THIS TO THE FDA. FDA SAFETY REPORT ID# (B)(4). | -software defect -defect introduced -error message | -failure to sense -device alarm system issue -use of device issue | Viewer
| FDA | Event Description: PROBLEM DESCRIPTION: FOLLOWING A CLIENT COMPLAINT, INTELERAD HAS IDENTIFIED A SOFTWARE PROBLEM IN THE IMAGE FUSION MODULE WHICH MAY DISPLAY AN INCORRECT CALCULATION OF STANDARD UPTAKE VALUE (SUV) WHEN LOADING THE PRESENTATION STATE OF IMAGE FOR A PREVIOUSLY CALCULATED (AND SAVED) REGION OF INTEREST (ROI). THIS PROBLEM COULD LEAD RADIOLOGISTS TO BASE THEIR DIAGNOSIS/REPORT ON INCORRECT ANATOMICAL MEASUREMENTS LEADING TO INCORRECT MEDICAL TREATMENT. NO PATIENT INCIDENT HAS BEEN REPORTED AT THIS MOMENT - THIS PROBLEM IS DOCUMENTED AS A PRECAUTION. INVESTIGATION DETAILS: ONE CUSTOMER COMPLAINT ((B)(4)) BUG IDENTIFIED: (B)(4). THIS IS A BUG DUE TO A MISSING COORDINATE TRANSFORMATION, WHICH MANIFESTS ITSELF ONLY WHEN THE CT AXIAL SPACING IS NOT 1MM. MOST IF NOT ALL OF OUR TESTING DATA TO THAT POINT HAD BEEN FROM GE WHICH HAD 1MM SPACING, SO THIS BUG WAS NOT SEEN, UNTIL WE HAD A PET/MR DATASET WITH AXIAL SPACING NOT EQUAL TO 1MM. EVEN THEN, THE BUG AS STATED WAS THAT THE MAX PIXEL WAS NOT IN THE CUBICAL VOI REGION. WE DID NOT NOTE THAT THE MAX PIXEL WAS INCORRECT FOR THE GIVEN REGION. THE BUG ONLY APPLIES TO RESTORED PRESENTATION STATES; ANY MODIFICATION OF A VOI READ FROM PRESENTATION STATES CAUSES MAX SUV TO BE CALCULATED CORRECTLY THE INCORRECT MAX SUV IS THAT OF AN ADJACENT REGION, SO IS RANDOM, AND MORE THAN LIKELY NOT TO CORRESPOND TO THE ACTUAL REGION OF INTEREST, THEREFORE THE ERROR SHOULD BE NOTICEABLE. THE DISPLAYED SUV AT THE CURSOR LOCATION IS CORRECT, CONTRADICTING THE INCORRECT MAX SUV. THE "GO TO MAX" ASSOCIATED WITH THE VOI WILL TRIANGULATE OUT OF THE VOI BOX, INDICATING THAT SOMETHING IS WRONG. ALL OF THE MAX SUVS IN THE PRESENTATION STATES WILL BE INCORRECT AS ABOVE, INCREASING THE LIKELIHOOD THAT THE ERROR WILL BE NOTICED THE FIX ((B)(4)) IS AVAILABLE IN INTELEVIEWER 4-11-1 AND LATER VERSIONS. THE INVESTIGATION WAS CONDUCTED ON THE CLIENT'S SYSTEMS AND WAS REPRODUCED ON INTELERAD'S SYSTEMS (INTELEVIEWER VERSION 4-101-1-P180). ASSESSMENT OF RISK TO USER: DESCRIPTION OF HAZARDOUS SITUATION IN USE OF THE DEVICE: INCORRECT CALCULATION OF SUV WHEN LOADING PRESENTATION STATE OF IMAGE FOR A PREVIOUSLY CALCULATED (AND SAVED) ROI. THIS MAY LEAD TO INCORRECT DIAGNOSIS. THIS PROBLEM IS CURRENTLY DOCUMENTED AS A KNOWN RISK: RISK #2.3 IN RISK MANAGEMENT SUMMARY (ENG-007 ISSUE 21). THE MITIGATION IN PLACE REDUCES THE RISK FOR THE GENERAL CASE OF SUV MEASUREMENTS, BUT NOT FOR DATASETS SUCH AS THE ONE THAT CAUSED THE INCIDENT, NOR THE SPECIFIC USE CASE OF DISPLAYING A MEASUREMENT WHEN RELOADING A PRESENTATION STATE. THE RISK CAN BE FURTHER REDUCED THROUGH ADDITIONAL DESIGN FACTORS TAKING INTO ACCOUNT SPECIFIC DATASETS AND THE CASE OF RELOADING THE PRESENTATION STATE. PROBABILITY OF RISK: BEFORE MITIGATION: OCCASIONAL (3). AFTER MITIGATION: REMOTE (2). SEVERITY OF RISK: BEFORE MITIGATION: INJURY REQUIRING TREATMENT (3) (IN THE WORST CASE). AFTER MITIGATION: INJURY REQUIRING TREATMENT (3) (IN THE WORST CASE) . DETECTABILITY (USER AWARENESS. IF DEVICE FAILURE OCCURS, IS IT EASILY RECOGNIZED BY USER): THE REPORTED INCORRECT SUV IS MORE OR LESS RANDOM, SO DOES NOT CORRESPOND TO THE ASSOCIATED VOLUME. DETECTABILITY INCREASES WITH THE NUMBER OF MEASUREMENTS, SINCE THE LACK OF CORRESPONDENCE BETWEEN INCORRECT VALUES AND ASSOCIATED VOLUMES BECOMES MORE OBVIOUS. CAN BE RECOGNIZED IN THE NORMAL CASE BY TRAINED USERS. RISK VS. BENEFIT COMPARISON (TO HELP DETERMINE THE RECALL STRATEGY AND TYPE OF THE RECALL): RELOADING A SAVED PRESENTATION STATE IS A BASIC FUNCTION OF THE DEVICE. THE RISK IS SMALL COMPARED TO THE BENEFIT AS IT IS A SIGNIFICANT TIME SAVER, AND NORMALLY REDUCES THE RISK OR ERRORS OF REPEATING MEASUREMENTS BY HANDS. HOW OFTEN DOES THE PROBLEM OCCUR? HOW OFTEN CAN THE PROBLEM CAUSE DIRECT OR INDIRECT HARM TO THE PATIENT? ONLY CAUSED BY DATA WHERE THE AXIAL SPACING IS NOT 1MM. IT MAY CAUSE HARM IF IT GOES UNDETECTED, AND THERE IS A MAJOR DISCREPANCY BETWEEN THE REAL VALUE AND WHAT IS DISPLAYED, AND THERE IS NO OTHER MEANS TO CONFIRM THE DIAGNOSIS BEFORE THE PATIENT IS TREATED. THE DEVICE HAS BEEN IN USE FOR TWO YEARS WITH ONLY A SINGLE REPORTED INCIDENT, WITH NO INDICATION OF REAL HARM. DESCRIBE THE FACTORS THAT MAY CONTRIBUTE TO PRODUCT RISK TO THE PATIENT PROBLEM OCCURS FOR PATIENT HIGH/LOW RESOLUTION CT OR MR SCANS WHERE AXIAL RESOLUTION IS != 1, AND WHERE PATIENT MEASUREMENTS ARE STORED FOR LATER RELOAD. FACTORS THAT MAY CONTRIBUTE TO PRODUCT RISK (I.E. DEVICE DESIGN, MANUFACTURING PROBLEMS, OR USE ERROR): AS ABOVE. RISK ASSESSMENT SUMMARY: FUNCTION: DISPLAY PREVIOUSLY SAVED SUV MEASUREMENT. FAILURE MODE/DEFECT: INCORRECT SUV DISPLAYED. LOCAL EFFECT: INCORRECT MEASUREMENT, LEADING TO INCORRECT DIAGNOSIS IF UNDETECTED. SYSTEM EFFECT: NONE. ADVERSE HEALTH CONSEQUENCES: INCORRECT DIAGNOSIS LEADING TO INCORRECT TREATMENT IF UNDETECTED, RESULTING IN HARM FROM INCORRECT TREATMENT OR MISSING/DELAYED TREATMENT. SEVERITY: 3 INJURY REQUIRING TREATMENT. PROBABILITY: 3 OCCASIONAL. RISK RATING: MEDIUM. CORRECTIVE ACTION: REMOVING THE AFFECTED VERSION WILL PREVENT USERS FROM ENCOUNTERING THE PROBLEM. PROPOSED RECALL ACTION: INTELERAD WILL: IDENTIFY ALL CLIENTS AFFECTED, I.E. ALL CLIENTS WITH ACCESS TO THE IMAGE FUSION MODULE OF INTELEPACS - 31 AUGUST 2016. INFORM ALL CONCERNED CLIENTS OF THE ISSUE (RECALL LETTER) - 31 AUGUST 2016. PROVIDE ALL CLIENTS THAT DO NOT ALREADY HAVE A FIXED VERSION OF THE DEVICE WITH A SOFTWARE UPDATE - 30 SEPTEMBER 2016. ESTIMATED COMPLETION DATE OF THE PROPOSED RECALL ACTION IS END OF SEPTEMBER 2016. | -incorrect message display -incorrect display -incorrect measurement | -use of device issue -improper or incorrect procedure or method -device markings issue | Viewer
| FDA | Event Description: PROBLEM: A FEW OF OUR CARDIOLOGISTS HAVE HAD ISSUES WITH FUJI CHANGING EJECTION FRACTION (EF) VALUES IN THE PAST TWO MONTHS. ONE WEEK AGO, THIS HAPPENED ON ONE OUTPATIENT STRESS ECHOCARDIOGRAMS AND ONE OF THE INPATIENT CASES. THIS INPATIENT UNDERWENT A CORONARY ANGIOGRAM AND MRI DUE TO THE ABERRANCY ON HIS ECHO REPORT. BOTH REPORTS CHANGED FROM A QUALITATIVE EF 55-60% TO ONE REPORTED AS <20%. PER THE MRI, EJECTION FRACTION = NORMAL. IMPACT ON THE INPATIENT: ADDITIONAL CARDIOLOGY FOLLOW UP (CONSULT, PROGRESS NOTES, ECHOCARDIOGRAM, CARDIAC MRI, ANGIOGRAM, AND EXTRA HOSPITAL DAY'S STAY). FOLLOW UP: THE EVENT RESULTING IN THE ERROR WAS IDENTIFIED (FOUND ONE OF OUR MDS AND TECHNICIANS) AND IS TIED TO THE PROVIDER WORKFLOW. WHEN THE READING MD SIGNS THE REPORT (I.E. THE READER CHOOSES A NORMAL EF AND SIGNS THE REPORT, BUT THE FINAL REPORT STATES EF 20%). THIS OCCURS WHEN YOU SELECT THE EF STATEMENT AND YOU SELECT ONE OF THE ARROW KEYS ON THE KEYBOARD INSTEAD OF CHANGING TO THE NEXT IMAGE, THE EF STATEMENT WILL CHANGE TO ANOTHER OPTION ON THE LIST. INTERIM ACTION: FACILITY CARDIOLOGY DEPARTMENT LEADS HAVE SPOKEN WITH THE MANUFACTURER AND INDICATED THAT THIS IS A TOP PRIORITY FOR COMPANY TO FIND A RESOLUTION. IN THE INTERIM, ALL USERS WERE NOTIFIED TO BE AWARE WHEN USING THE ARROWS ON THE KEYBOARD TO MOVE THROUGH THE REPORTS. ALSO INSTRUCTED "IF YOU SEE AN ECHO THAT STATES EF < 20%, (ESPECIALLY IF THIS IS AN UNEXPECTED RESULT OR IS ASSOCIATED WITH OTHER DISCREPANT STATEMENTS SUCH AS ¨NORMAL WALL MOTION¨ OR ¨NO CHANGE FROM PRIOR¨ WHEN THE PRIOR EF WAS NORMAL), PLEASE PULL UP THE ECHO IMAGES TO VERIFY THE LVEF. IF YOU FEEL THE <20% WAS AN ERROR PLEASE REACH OUT TO THE READING CARDIOLOGIST SO THE REPORT CAN BE CORRECTED..." | -fail to eject -device malfunction -part broken or damaged | -break -Sticking -device markings issue | Viewer
| FDA | Event Description: PROBLEM: NO PATIENT INVOLVEMENT, PROBLEM IDENTIFIED DURING INTERNAL TESTING. POTENTIAL HAZARD(S): WHEN USING ENHANCED CT FORMAT DICOM DATA, THERE MAY BE: INCORRECT ORIENTATION OF DISPLAYED IMAGE DATA INDICATING INCORRECT PATIENT SIDE (LEFT, RIGHT, SUPERIOR, INFERIOR, ANTERIOR OR POSTERIOR) (SIDEDNESS). INACCURATE MEASUREMENT. Manufacturer Narrative: NO SPECIFIC MEDICAL DEVICE FAILURE. NO PATIENTS INVOLVED. ISSUE IDENTIFIED DURING ECHOPIXEL INTERNAL TESTING. HEALTH HAZARD EVALUATION COMPLETED BY QUALITY MANAGEMENT.. ALL CURRENT USERS OF PRODUCT HAVE BEEN CONTACTED, PROVIDED INFORMATION ON THE ISSUE. ALL CURRENT USERS HAVE INDICATED THAT NONE USE THE STATED ENHANCED MULTI-FRAME CT FORMATTED DICOM DATA, THEREFORE NO RISK TO PATIENTS CORRECTIVE ACTION PROCESS IN PLACE. NO SPECIFIC DEVICE FAILURE INVOLVED. | -incorrect device orientation -incorrect operation -incorrect labeling | -use of device issue -incorrect measurement -insufficient flow or underinfusion | Viewer
| FDA | Event Description: Q-STATION MIXED INFORMATION.Event Description: THE CUSTOMER REPORTED THAT WHEN USING Q-STATION, DATA FROM A PREVIOUS PATIENT APPEARS ON THE CURRENT PATIENT'S DATA REPORT. THERE WAS NO REPORT OF INJURY TO PATIENT OR NEGATIVE IMPACT ON CLINICAL JUDGMENT. Manufacturer Narrative: THIS ISSUE IS CURRENTLY UNDER INVESTIGATION. FINAL RESULTS OF THE EVALUATION WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS RESULTING FROM THIS ISSUE. SOFTWARE. Manufacturer Narrative: IT WAS DETERMINED THAT THE ISSUE WAS A KNOWN SOFTWARE DEFECT. MANDATORY FIELD CHANGE ORDER,(B)(4), WAS DEVELOPED AND RELEASED TO CORRECT THIS ISSUE. THE FIX IS INCLUDED IN Q-STATION VERSION 3.3.2 AND HIGHER. THE AFFECTED CUSTOMERS WILL BE PROVIDED WITH THE NEWEST VERSION OF Q-STATION (3.3.2) ONCE THE NEW ORDERS HAVE BEEN CREATED AND SHIPPED. | -patient data breached -incorrect patient information -incorrect device orientation | -fluid leak -data issue -incorrect or inadequate result | Viewer
| FDA | Event Description: RADSUITE IS A MEDICAL IMAGE VIEWING SOFTWARE APPLICATION THAT ALLOWS IMAGE DATA TO BE VIEWED FROM A FACILITY IMAGE CACHE (FIC) OR FROM A LONG TERM ARCHIVE (LTA). A CUSTOMER NOTICED THAT PATIENT A'S STUDY WAS STORED UNDER THE CORRECT PATIENT IN THE FIC BUT WAS STORED AS A DIFFERENT, PATIENT B, IN THE LTA. WHEN PATIENT A'S CURRENT STUDY WAS OPENED IN THE RADSUITE VIEWER ON THE LTA, PATIENT B'S STUDY WAS OPENED AS A PRIOR INSTEAD OF PATIENT A'S STUDY, AND WAS LABELED WITH PATIENT A'S INFORMATION. | -fail to view image stack -fail to interpret images -fail to monitor loading images | -data issue -application program issue -problem with software installation | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER FILED A COMPLAINT (CASE (B)(4)) AGAINST RADSUITE STATING THAT THE PIXEL VALUE TOOL'S BODY WEIGHT AND IDEAL BODY WEIGHT VALUES ARE OFF IN RADSUITE COMPARED TO THE SCANNER. THE CUSTOMER CONTACT STATED THAT THE RADIOLOGIST'S USE THIS INFORMATION TO DETERMINE IF THE TUMOR IS MALIGNANT OR BENIGN AND NEEDS IT TO BE ACCURATE. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS. NOTE THAT FDA TERMINATED THIS RECALL ON FEBRUARY 13, 2017. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE INVESTIGATION BY MERGE HEALTHCARE SHOWED THAT THE VALUE IS OFF BY APPROXIMATELY A FACTOR OF 20. THIS SHOULD BE EASILY RECOGNIZED BY THE USER BECAUSE OF THE DRASTIC DIFFERENCE BETWEEN THE VALUE AND THE VISUAL (QUALITATIVE) APPEARANCE OF THE IMAGE. THE DEFECT WAS INTRODUCED INTO THE RADSUITE V8.30.7.8 RELEASE. (B)(6) WAS THE FIRST AND ONLY CUSTOMER TO RECEIVE THE V8.30.7.8 RELEASE THAT CONTAINED THE DEFECT. A FIX HAS BEEN APPLIED AT (B)(6) TO CORRECT THE CALCULATION ERROR. | -fail to display data -inadequate message display -fail to provide diagnostic test | -use of device issue -data issue -device operates differently than expected | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED ((B)(4)) THAT AN INCORRECT STUDY WAS ATTACHED TO A PATIENT AND THEY WERE TRYING TO DELETE THE IMAGES. THEY WERE UNABLE TO DELETE THE IMAGES FROM THE STUDY. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION COULD NOT BE CORRECTED ON THE STUDY, LEADING TO A POTENTIAL MISDIAGNOSIS. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -incorrect display format -fail to process medical imaging -inadequate message display | -communication or transmission issue -data issue -compliance issue | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED (CASE (B)(4)) THAT STUDIES ARE STORING BUT ARE ALSO LOCKING UP. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION COULD NOT BE AVAILABLE FOR REVIEW. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -fail to process medical imaging -fail to provide diagnostic test -inadequate image manipulation | -incorrect or inadequate result -data issue -improper or incorrect procedure or method | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. A CUSTOMER REPORTED (CASE (B)(4)) THAT THEY WERE UNABLE TO RETRIEVE A STUDY BACK FROM THE ARCHIVE TO THE RADSUITE VIEWER. THE CUSTOMER WAS TRYING TO BRING THE STUDY BACK FROM THE ARCHIVE TO CORRECT MRN INFORMATION. THIS WAS DETERMINED TO BE A POTENTIAL SAFETY ISSUE IN THE EVENT THE INFORMATION CANNOT BE CORRECTED ON THE STUDY. Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -fail to process medical imaging -unable to scroll image -fail to provide diagnostic test | -unauthorized access to computer system -human-device interface issue -incorrect or inadequate result | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER COMPLAINT (B)(4) FROM (B)(6) OPENED (B)(6) 2015 STATES PREFETCHER IS NOT WORKING AS EXPECTED (APPEARS DELAYED). CUSTOMER REPORTED THAT PREFETCHER IS NOT WORKING AND IS AFFECTED , POTENTIALLY, PATIENT CARE. THE PREFETCHER IS INTENDED TO PULL PATIENT PRIORS FOR RADIOLOGIST TO REVIEW DURING DIAGNOSIS. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to display data -inadequate image types | -use of device issue -problem with software installation -electronic property issue | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. CUSTOMER REPORTED ISSUES WITH THEIR PRIMARY PRODUCTION RADSUITE CAUSING "DO NOT ROUTE" EXCEPTIONS. ALL CASES INVESTIGATED BY CUSTOMER ROUTED AND THEN THE EXCEPTIONS POPPED UP AFTERWARDS. ALSO, AN ICON NEXT TO THE STUDY WILL SHOW A RED "X" NOTIFYING USERS OF AN EXCEPTION. PER THE COMPLAINT CLAIM, THE APPEARANCE OF THOSE ICONS CONFUSED A RADIOLOGIST AND DELAYED READING OF A STAT STUDY BY 14 HOURS ON A CRITICAL FINDING. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -fail to display data -inadequate image manipulation | -human-device interface issue -communication or transmission issue -incorrect or inadequate result | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. ON (B)(6) 2013 A CUSTOMER REPORTED A PROBLEM WHEN THE MOUSE CURSOR HOVERED OVER THE ROI IN THE SECONDARY SERIES IT DISPLAYED THE STATISTICAL DATA BELONGING TO THE FIRST IMAGE. THIS IS A POTENTIAL SAFETY ISSUE SINCE THE ROI DISPLAYS AND MEASURES HOUNSFIELD UNITS (HUS). HOUNSFIELD MEASUREMENTS ARE USED TO MEASURE TISSUE DENSITY AND MAY BE USED AS AN INPUT TO CLINICAL DECISION MAKING. THERE WAS NO REPORT OF ADVERSE PATIENT IMPACT. (B)(4). Manufacturer Narrative: UPON REVIEW OF THIS CUSTOMER ISSUE, IT WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -inadequate image manipulation -fail to display data -lack of display | -use of device issue -failure of device to self-test -human-device interface issue | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. STUDIES ARE WATERMARKING OFF IN AV WITHIN 5-6 DAYS. THIS IS CAUSING THE RADIOLOGISTS TO HAVE TO PRE-FETCH STUDIES, CAUSING DISSATISFACTION. PREVIOUSLY THERE WERE 27 DAYS BEFORE STUDIES WOULD WATERMARK OFF AV. STUDY TODAY (B)(6) DID NOT PRE-FETCH PRIOR STUDY FROM (B)(6) 2015. RADIOLOGIST FEELS THIS IS A PATIENT SAFETY ISSUE. (B)(4). Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE AN MDR. | -fail to process medical imaging -inadequate image types -inadequate message display | -use of device issue -data issue -user used incorrect product for intended use | Viewer
| FDA | Event Description: RADSUITE PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. THE APPLICATION PROVIDES WORKFLOW INTEGRATION CAPABILITIES FOR HEALTH CARE ENTERPRISES, WHEREIN: RADIOLOGISTS CAN VIEW, ANNOTATE, AND TAG STUDIES AS DIAGNOSTICALLY READ. REFERRING PHYSICIANS CAN VIEW PATIENT IMAGES AND RADIOLOGISTS' ANNOTATIONS. TERTIARY CARE PHYSICIANS, MEDICAL TECHNOLOGISTS, AND INFORMATION TECHNOLOGY PROFESSIONALS CAN RECEIVE PATIENT RECORDS. ON 5/25/2017, MERGE HEALTHCARE TECHNICAL SUPPORT WAS NOTIFIED BY A RADSUITE CUSTOMER THAT PROCEDURES WERE NOT ABLE TO BE PERFORMED CAUSING A DELAY IN TREATMENT FOR 2 PATIENTS. PER THE CUSTOMER SITE, EMERGENCY DEPARTMENT STUDIES ARE TYPICALLY READ WITHIN 30 MINUTES HOWEVER, THERE WERE TWO STUDIES THAT WERE NOT READ UNTIL THE NEXT MORNING. WITH MERGE RADSUITE NOT FUNCTIONING AS EXPECTED, THERE IS A POTENTIAL FOR A DELAY IN DIAGNOSIS OR TREATMENT THAT MAY LEAD TO HARM. THERE WAS NO REPORTED HARM TO EITHER PATIENT A OR PATIENT B AS A RESULT OF THE DELAY. (B)(4). Manufacturer Narrative: PER THE INVESTIGATION BY MERGE SUPPORT, BASED ON THE EXAMPLES PROVIDED BY THE CUSTOMER, IT APPEARS THAT THE STUDIES WERE ORDERED AND SENT IN WITHOUT A PATIENT LOCATION. THE PATIENT LOCATION WAS LATER UPDATED ON THE WORK LIST. HOWEVER, THE PATIENT LOCATION WASN'T UPDATED ON THE LONG TERM ARCHIVE BECAUSE THE STUDY HAD ALREADY BEEN ROUTED TO THE ARCHIVE PRIOR TO THE UPDATE TO THE PATIENT LOCATION. THE STUDIES ARE FOUND AND VIEWED FROM THE WORKLIST AND PATIENT LOCATION IS AN IMPORTANT FIELD TO THIS CUSTOMER. THE CUSTOMER'S WORKFLOW INCLUDES ADDING IN PATIENT LOCATION AFTER AN ORDER HAS ALREADY BEEN COMPLETED. MERGE TECHNICAL SUPPORT ASSISTED THE CUSTOMER BY ADDING IN A SHORT (1 MINUTE) DELAY BETWEEN CUSTOMER UPDATES TO THE WORKLIST AND THE RELATED STUDIES BEING SENT TO THE LONG TERM ARCHIVE. BASED ON INFORMATION RECEIVED FROM THE CUSTOMER, THE SHORT DELAY IN SENDING UPDATED REPORTS TO THE ARCHIVE HAS RESOLVED THE ISSUE. | -fail to process medical imaging -inadequate image manipulation -fail to manage 3d model | -use of device issue -data issue -incorrect or inadequate result | Viewer
| FDA | Event Description: REPORTEDLY, AN ATTENDING PHYSICIAN WAS INTERPRETING AN ECHOCARDIOGRAM IN SYNAPSE CARDIOVASCULAR AND WAS MEASURING DOPPLER. THE PATIENT'S HEART RHYTHM WAS IRREGULAR, SO THE PHYSICIAN HAD TO AVERAGE DOPPLER SIGNAL MEASUREMENTS FOR EACH VALVE. WHILE DOING DOPPLER SIGNAL TRACINGS OF THE MITRAL VALVE, HE NOTICED THAT VALUES THAT SHOULD HAVE BEEN CONSTANTS IN THE CALCULATIONS FROM THE INITIAL TRACINGS WOULD SPONTANEOUSLY CHANGE AS MORE SIGNALS WERE AVERAGED (SPECIFICALLY MEAN PRESSURE GRADIENTS). REPORTEDLY, THIS HAS BEEN AN ONGOING ISSUE THAT IS CURRENTLY BEING INVESTIGATED BY THE MANUFACTURER. UNKNOWN HOW LONG THIS HAS BEEN AN ISSUE AND HOW MANY PATIENTS HAVE BEEN AFFECTED. | -fail to drive artificial heart -incorrect readings -confused indication | -fluid leak -Contamination During Use -Air Leak | Viewer
| FDA | Event Description: SICAT IMPLANT IS A STANDALONE SOFTWARE APPLICATION, I.E. A RADIOLOGICAL VISUALIZATION SOFTWARE FOR DIAGNOSIS AND DENTAL IMPLANT PLANNING. A DENTIST FOUND THAT IMPLANT PLANNING DATA EXPORTED FROM SICAT IMPLANT ARE NOT CORRECT FOR ONE SPECIFIC EXPORT FORMAT. IF IT WERE TO RECUR, THE ERROR IN THE EXPORT FORMAT MIGHT CAUSE DENTISTS TO MANUFACTURE (OR HAVE MANUFACTURED) SURGICAL GUIDES WHICH ARE INCORRECT. USING THOSE GUIDES MIGHT CAUSE DRILLING INTO THE JAW TO BE EITHER NOT DEEP ENOUGH OR TO DEEP. A DENTIST MIGHT HAVE TO STOP A SURGICAL PROCEDURE. OR THE DRILLING WHEN TO DEEP MIGHT CAUSE DAMAGE TO STRUCTURES WITHIN THE JAW WHICH NEED INTERVENTION. Manufacturer Narrative: SICAT IMPLANT V2.0 IS A STANDALONE SOFTWARE APPLICATION. THE DEVICE PROBLEM WAS ANALYZED AND IT WAS FOUND THAT IMPLANT PLANNING DATA EXPORTED FROM SICAT IMPLANT ARE NOT CORRECT FOR ONE SPECIFIC EXPORT FORMAT. A CODING ERROR WAS DETECTED CAUSING THE DEVICE PROBLEM. IMMEDIATE ACTION HAS BEEN TAKEN WITHIN THE LICENSING SYSTEM OF SICAT IMPLANTV 2.0 TO PREVENT NEW CUSTOMERS/USERS FROM ACTIVATING THE APPLICATION. A FIELD SAFETY NOTICE HAS BE SEND OUT TO THE USERS AND CONSIGNEES OF SICAT IMPLANT V2.0. THE DEVICE WILL BE CORRECTED IN THE FIELD BY DISTRIBUTING AND APPLYING A SOFTWARE PATCH TO REMOVE THE PROBLEM. EVALUATION RESULTS AND EVALUATION REPORT WILL BE PROVIDED IN A FOLLOW-UP REPORT. | -inadequate software configuration -incorrect device orientation -incomplete specification | -problem with software installation -data issue -loose or intermittent connection | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF A CATASTROPHIC LOSS OF THE CUSTOMER'S STORAGE AREA NETWORK (SAN) AT THIS SITE. THE SAN IS PROVIDED AND MAINTAINED BY THE USER; IT IS THEIR CHOSEN LOCATION FOR THE DEPLOYMENT AND MANAGEMENT OF MULTIPLE MEDICAL SOFTWARE PRODUCTS, INCLUDING SIEMENS HEALTHCARE PRODUCTS. THE SPECIFIC CAUSE OF THE FAILURE OF THE SAN WAS NOT SHARED BY THE CUSTOMER WITH SIEMENS. THERE IS NO INDICATION THAT THE DEPLOYED SIEMENS PRODUCTS CAUSED THE FAILURE OF THE SAN, EITHER BY THE CUSTOMER OR AS A TECHNICAL POSSIBILITY. IT WAS SUBSEQUENTLY DETERMINED THAT THERE WERE NO RECENT DATA BACKUPS FROM WHICH DATA RECOVERY WOULD BE POSSIBLE AS THE CUSTOMER ALSO USED THEIR SAN AS THE SOLE BACKUP LOCATION. THE SITE HAS AN ONGOING RECOVERY EFFORT WITH A 3RD PARTY DATA RECOVERY COMPANY. IT HAS BEEN IDENTIFIED THAT ONLY SOME DRIVES FROM THE NETWORK HAVE A POTENTIAL FOR DATA RECOVERY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S DATA. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE LOSS OF POTENTIALLY CLINICALLY RELEVANT DATA MAY CONTRIBUTE TO A DELAY IN TREATMENT OR TO A PATH FOR INCORRECT TREATMENT. THE EXTENT OF THE PROBABILITY OF HARM OCCURRENCE FOR THE SITE IS INCALCULABLE AS THE EXTENT OF THE TOTAL RECOVERY AND INTERIM CLINICAL OPERATIONS AT THE CONCERNED SITE IS UNKNOWN TO SIEMENS. THE REPORTED EVENT OCCURED IN (B)(6). Manufacturer Narrative: THE USER, WHO IS RESPONSIBLE FOR MAINTAINING THE STORAGE AREA NETWORK, CHOSE THIS LOCATION FOR THE DEPLOYMENT AND MANAGEMENT OF MULTIPLE MEDICAL SOFTWARE PRODUCTS AT THE FACILITY. THE SAN WAS ALSO USED AS THE SOLE DATA BACKUP LOCATION. THE SPECIFIC CAUSE OF THE FAILURE OF THE SAN WAS NOT SHARED BY THE CUSTOMER WITH SIEMENS. THE SYNGO PLAZA PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED. NO FURTHER UPDATES WILL BE MADE FOR THIS REPORTED ISSUE. INTERNAL ID # (B)(4). | -loss of data -fail to maintain status -loss of performance | -device maintenance issue -fluid leak -protective measure issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF A CATASTROPHIC LOSS OF THE CUSTOMER'S STORAGE AREA NETWORK (SAN) AT THIS SITE. THE SAN IS PROVIDED AND MAINTAINED BY THE USER; IT IS THEIR CHOSEN LOCATION FOR THE DEPLOYMENT AND MANAGEMENT OF MULTIPLE MEDICAL SOFTWARE PRODUCTS, INCLUDING SIEMENS HEALTHCARE PRODUCTS. THE SPECIFIC CAUSE OF THE FAILURE OF THE SAN WAS NOT SHARED BY THE CUSTOMER WITH SIEMENS. THERE IS NO INDICATION THAT THE DEPLOYED SIEMENS PRODUCTS CAUSED THE FAILURE OF THE SAN, EITHER BY THE CUSTOMER OR AS A TECHNICAL POSSIBILITY. IT WAS SUBSEQUENTLY DETERMINED THAT THERE WERE NO RECENT DATA BACKUPS FROM WHICH DATA RECOVERY WOULD BE POSSIBLE AS THE CUSTOMER ALSO USED THEIR SAN AS THE SOLE BACKUP LOCATION. THE SITE HAS AN ONGOING RECOVERY EFFORT WITH A 3RD PARTY DATA RECOVERY COMPANY. IT HAS BEEN IDENTIFIED THAT ONLY SOME DRIVES FROM THE NETWORK HAVE A POTENTIAL FOR DATA RECOVERY. IT IS NOT POSSIBLE FOR SIEMENS TO DETERMINE THE EXTENT OR FUTURE AVAILABILITY OF THE SITE'S DATA. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE LOSS OF POTENTIALLY CLINICALLY RELEVANT DATA MAY CONTRIBUTE TO A DELAY IN TREATMENT OR TO A PATH FOR INCORRECT TREATMENT. THE EXTENT OF THE PROBABILITY OF HARM OCCURRENCE FOR THE SITE IS INCALCULABLE AS THE EXTENT OF THE TOTAL RECOVERY AND INTERIM CLINICAL OPERATIONS AT THE CONCERNED SITE IS UNKNOWN TO SIEMENS. THE REPORTED EVENT OCCURED IN (B)(6). Manufacturer Narrative: THE USER, WHO IS RESPONSIBLE FOR MAINTAINING THE STORAGE AREA NETWORK, CHOSE THIS LOCATION FOR THE DEPLOYMENT AND MANAGEMENT OF MULTIPLE MEDICAL SOFTWARE PRODUCTS AT THE FACILITY. THE SAN WAS ALSO USED AS THE SOLE DATA BACKUP LOCATION. THE SPECIFIC CAUSE OF THE FAILURE OF THE SAN WAS NOT SHARED BY THE CUSTOMER WITH SIEMENS. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED. NO FURTHER UPDATES WILL BE MADE FOR THIS REPORTED ISSUE. INTERNAL ID # (B)(4). | -loss of data -fail to maintain status -loss of performance | -device maintenance issue -fluid leak -protective measure issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF A DATA LOSS INCIDENT AT THE USER'S SITE. SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITHIN THE ACCEPTABLE RISK THRESHOLD. DUE TO A DELETION OF FOLDERS FROM THE SHORT TERM STORAGE (STS) PARTITION OF THE SYNGO.PLAZA SERVER, BELONGING TO 517 STUDIES. WITH CURRENT AVAILABLE INFORMATION, IT IS CONFIRMED THAT THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE DATA LOSS IN THE STS. THE STORAGE OF THE IMAGE FILES WAS SUCCESSFULLY COMPLETED UPON RECEIPT AND THE IMAGE FILES WERE PROPERLY STORED (E.G. THE SERIES STATUS IS MARKED AS COMPLETE). THE IMAGES WERE PRESENT AFTER STORAGE. THERE ARE NO INJURIES ATTRIBUTED TO THE EVENT. THE EVENT OCCURRED IN THE (B)(6). Manufacturer Narrative: THE PRODUCT SYNGO.PLAZA DID NOT CONTRIBUTE TO THE DATA LOSS IN THE STS. FOLDER DELETION FROM THE STS CANNOT BE ATTEMPTED FROM ANY OF THE SYNGO.PLAZA WORKFLOWS. IF ANY IMAGE FILE IS DELETED FROM SYNGO.PLAZA, IT IS LOGGED IN THE APPLICATION EVENT LOGS. IN THE REPORTED CASE THERE ARE NO ENTRIES IN THE APPLICATION EVENT LOGS THAT SUGGEST THAT THE FILES HAVE BEEN DELETED FROM THE SYNGO.PLAZA. STS IS INTENDED ONLY FOR TEMPORARY STORAGE OF PATIENT DATA UNTIL IT IS TRANSFERRED TO LONG TERM STORAGE (LTS). NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. WITH THIS INFORMATION, A DEATH OR SERIOUS INJURY IS HIGHLY IMPROBABLE. | -loss of data -loss of function -loss of traceability | -device alarm system issue -protective measure issue -inaudible voice prompts | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF A HARDWARE CRASH OF THE STORAGE DRIVES ATTACHED TO THE SYNGO IMAGING PRODUCT CAUSED BY POWER FAILURE. SEVERAL DRIVES HOSTING THE SYNGO IMAGING INTERNAL DATABASE AND DATA BECAME INOPERABLE. THE CONCERNED USER SITE WAS IN THE PROCESS OF TRANSITIONING ITS PRIMARY PACS ROLE FROM THE SYNGO IMAGING TO THE SYNGO.PLAZA. THEREFORE, A SYNCHRONIZATION OF DATA TRANSMISSION WAS ESTABLISHED; HOWEVER, IT WAS PLANNED TO OPERATE BOTH SYSTEMS IN PARALLEL UNTIL MIGRATION OF ALL DATA TO THE SYNGO.PLAZA WAS FINALIZED. NO IMPACT ON THE CLINICAL ROUTINE AT THE SITE HAS BEEN OBSERVED. THE USER CONTINUES TO READ AND SAVE CURRENT EXAMINATIONS USING THE SYNGO.PLAZA. HOWEVER, IN CASE PRIOR IMAGES THAT ARE LOCATED ON THE INOPERABLE DRIVES ARE NEEDED FOR A FOLLOW UP READING AND COMPARISON, THEY WOULD NOT BE AVAILABLE. THE ROOT CAUSE OF THE ISSUE IS A HARDWARE ERROR. THE SYNGO IMAGING PRODUCT DID NOT CONTRIBUTE TO THE UNAVAILABILITY OF IMAGES. THERE ARE NO INJURIES RELATED TO THIS EVENT. Manufacturer Narrative: (B)(4). ACTIVITIES TO REPAIR THE DEFECTIVE DRIVES AT THE USER SITE AS WELL AS FINALIZATION OF DATA MIGRATION TO THE SYNGO.PLAZA ARE ON-GOING. THE FINAL AMOUNT OF DATA LOSS (IF ANY) CAN BE DETERMINED ONLY AFTER THE RECOVERY ACTIVITIES AT SITE ARE COMPLETED. NO CONSEQUENCES FROM THE USER HAVE BEEN REPORTED SO FAR. | -device damaged -fail to access -device burned | -medical gas supply problem -device markings issue -fitting problem | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF A RANSOMWARE VIRUS ATTACK ON THE SYNGO.PLAZA SERVER. THE NAME OF THE VIRUS IS UNKNOWN; HOWEVER, IT COMES FROM THE FAMILY OF RANSOMWARE VIRUSES. THE SYNGO.PLAZA SERVER WAS NOT INFECTED, ONLY THE FILES ACCESSIBLE THROUGH SAMBA (SMB) SHARES WERE ENCRYPTED AS A RESULT OF THE ATTACK. THE VIRUS ATTACKED THE SYNGO.PLAZA THROUGH THE CLIENTS OR OTHER COMPUTERS IN THE NETWORK. AT THE TIME OF THE ATTACK THE ANTIVIRUS WAS NOT INSTALLED ON THE SYNGO.PLAZA SERVER. ACCESS TO SHARE WAS NOT LIMITED AS RECOMMENDED BY THE SYNGO.PLAZA PRODUCT DOCUMENTATION AS IT WAS SET WITH PERMISSION TO "EVERYONE" WITH "READ, WRITE, EXECUTE". THE ATTACK RESULTED IN 310,075 SOPS (SERVICE OBJECT PAIRS) IN THE STS (SHORT TERM STORAGE) TO BE ENCRYPTED. THESE SOPS BELONG TO 50,504 SERIES (FULLY OR PARTIALLY AFFECTED) OUT OF 665,628 STORED SERIES IN THE STS. MULTIPLE SEPARATE LTS (LONG TERM STORAGE) SYSTEMS WERE AFFECTED BY THIS ATTACK AS WELL. THIS INCIDENT OCCURRED IN (B)(6). Manufacturer Narrative: POSSIBILITY TO RECOVER DATA FROM THE ENCRYPTED LTS IS CURRENTLY BEING ASSESSED BY SIEMENS. HOWEVER, WITH CURRENTLY AVAILABLE INFORMATION, IT CAN BE CONFIRMED THAT THE SYNGO.PLAZA SYSTEM DID NOT CONTRIBUTE TO THE DATA LOSS. | -patient experienced adverse event -device damaged -support compromised | -no display or display failure -detachment of device component -human-device interface issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. ALL IMAGES IN THE SHORT TERM STORAGE (STS) WERE CORRUPTED AFTER A HARD DISK (RAID) FAILURE OCCURRED. AT THE USER SITE THE SYNGO.PLAZA SYSTEM IS OPERATED IN A STANDALONE MODE, I.E. WITHOUT ARCHIVING FUNCTIONALITY. FURTHERMORE ,CUSTOMER HAS NO OTHER ARCHIVING SOLUTION AVAILABLE. NONE OF THE PATIENT DATA COULD BE RECOVERED DUE TO ALL DATABASE DRIVES AND BACKUP DRIVES BEING CORRUPTED. TOTAL NUMBER AND TIME PERIOD OF LOST DATA ARE AS FOLLOWS: TOTAL NUMBER OF IMAGES LOST: 97868. TOTAL NUMBER OF STUDIES CORRUPTED: 19343. TIMELINE OF THE LOST DATA CANNOT BE DETERMINED AS ALL CORRUPTED STUDY RECORDS HAVE BEEN REMOVED FROM THE DATABASE AND THE DATA WAS PRESENT ONLY UNTIL DECEMBER 6, 2019. NONE OF THE STUDIES INCLUDING THE RECENT STUDIES (DATED BETWEEN DECEMBER 1 AND 6, 2019) WERE AVAILABLE IN THE MODALITY TO RESEND. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATION. THE REPORTED INCIDENT OCCURRED DUE TO A HARDWARE ISSUE WITHOUT CONTRIBUTION FROM THE SYNGO.PLAZA. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. INTERNAL ID # (B)(4). | -data corrupted -fail to interpret images -loss of image | -leak -device maintenance issue -device inoperable | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. THE APPLICATION LOG FILES SHOWED MULTIPLE INSTANCES OF ABNORMAL SHUTDOWN (POWER OUTAGE OR WINDOWS BLUE SCREEN/BSOD ISSUES) OF THE MACHINE. DUE TO THIS INTERRUPTION IN FILE TRANSFER PROCESS, THE IMAGE FILES WERE NOT SUCCESSFULLY STORED TO THE LOCAL FILE SYSTEM, RESULTING IN ARCHIVE FAILURE AND LOSS OF PHYSICAL FILES. TOTAL # OF IMAGES LOST - 2067. TOTAL # OF STUDIES LOST- 52. TOTAL # OF SERIES AFFECTED - 73. AFFECTED STUDIES ARE DATED 2014, 2015, 2016, 2017, 2018 AND 2019. LOADING, ARCHIVING AND AUTO-ROUTING OF SUCH IMAGES WILL FAIL AS SOME OF THE IMAGES ARE CORRUPTED AND UNREADABLE. PRIOR IMAGES ARE NO AVAILABLE FOR COMPARISON EITHER. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATIONS. THE REPORTED ISSUE AT THE CONCERNED SITE IS CAUSED BY OUTDATED BSOD - MICROSOFT SOFTWARE. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID #(B)(4). | -inadequate log file -power outage -incorrect shutdown procedure | -incorrect or inadequate result -device alarm system issue -data issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. THE LOG FILES SHOWED MULTIPLE INSTANCES OF ABNORMAL SHUTDOWN (POWER OUTAGE OR WINDOWS BLUE SCREEN/BSOD ISSUES) OF THE MACHINE, WHILE RECEIVING THE IMAGES FROM THE MODALITY. DUE TO THIS INTERRUPTION IN FILE TRANSFER PROCESS, THE IMAGE FILES WERE NOT SUCCESSFULLY STORED TO THE LOCAL FILE SYSTEM, RESULTING IN ARCHIVE FAILURE AND LOSS OF PHYSICAL FILES. AS OF DECEMBER 18, 2019, TOTAL NUMBER AND TIME PERIOD OF MISSING DATA ARE AS FOLLOWS: TOTAL # OF IMAGES LOST - 72, TOTAL # OF STUDIES AFFECTED - 14, AFFECTED STUDIES ARE DATED 2018 AND 2019. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATION AS THE SYNGO.PLAZA SENT AS ERROR STATUS TO SCU (MODALITY) FOR THE CORRUPTED IMAGES. THE SCU SHOULD HAVE RESENT THE IMAGES TO THE SYNGO.PLAZA. IN THIS CASE SYNGO.PLAZA APPLICATION WORKED AS SPECIFIED. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID # (B)(4). | -power outage -incorrect shutdown procedure -power interrupted | -data issue -incorrect or inadequate result -battery issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN INCIDENT ON THE SYNGO.PLAZA UNIT. THE LOG FILES SHOWED MULTIPLE INSTANCES OF ABNORMAL SHUTDOWN (POWER OUTAGE OR WINDOWS BLUE SCREEN/BSOD ISSUES) OF THE MACHINE. DUE TO THIS INTERRUPTION IN FILE TRANSFER PROCESS, THE IMAGE FILES WERE NOT SUCCESSFULLY STORED TO THE LOCAL FILE SYSTEM, RESULTING IN ARCHIVE FAILURE AND LOSS OF PHYSICAL FILES. TOTAL # OF IMAGES LOST - 2235. TOTAL # OF STUDIES AFFECTED - 8. AFFECTED STUDIES ARE DATED 2016, 2017 AND 2018. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA WORKS ACCORDING TO SPECIFICATIONS. THE REPORTED ISSUE AT THE USER SITE IS CAUSED BY OUTDATED BSOD - MICROSOFT SOFTWARE. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID #: (B)(4). | -power outage -incorrect shutdown procedure -power interrupted | -data issue -incorrect or inadequate result -battery issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN ISSUE WHEN SERVICE SUPPORT REQUEST WAS ESCALATED FOR THE SYNGO DYNAMICS PRODUCT. THE SYNGO DYNAMICS PRODUCT COULD NOT RETRIEVE HISTORICAL STUDY IMAGES FROM THE SITE'S ARCHIVE SYSTEM. IT WAS DETERMINED THAT 619 IMAGE STUDIES WERE MISSING FROM THE SITE'S ARCHIVE. THE CLINICAL RELEVANCE OF THE POTENTIAL DATA LOSS IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES WILL STILL BE AVAILABLE IN THE SYNGO DYNAMICS SYSTEM UNLESS DELETED BY THE SITE IT POLICY OR USER ACTIONS. THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH HAS ANY INDICATION OF RESULTING IN THE NEED FOR A PATIENT RESCAN. Manufacturer Narrative: NO EVIDENCE WAS FOUND TO INDICATE THAT A FAILURE OF THE SYNGO DYNAMICS PRODUCT WAS RESPONSIBLE FOR THE LOSS OF HISTORICAL STUDY IMAGES FROM THE SITES LONG TERM ARCHIVE. AS THE DATA IS NO LONGER AVAILABLE IN THE SYNGO DYNAMICS PRODUCT CACHE, THIS DATA IS CONSIDERED UNRECOVERABLE. IT IS NOT POSSIBLE FOR SIEMENS TO FURTHER IDENTIFY A ROOT CAUSE FOR THIS DATA LOSS. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO REASONABLY FORESEEABLE PATH TO OCCURRENCE OF HARM CAN BE DETERMINED. INTERNAL ID # (B)(4). | -fail to retrieve data -incorrect product classification -fail to analyze signal | -material disintegration -Device Handling Issue -Decreased Sensitivity | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH A STORAGE DEVICE ATTACHED TO THE SYNGO IMAGING PRODUCT. THE STORAGE DEVICE EXPERIENCES FAILURES DURING STARTUP PREVENTING LOGICAL VOLUMES FROM GOING ONLINE. THIS CAUSES INABILITY OF SYNGO IMAGING INTERNAL SERVICE TO START AND MAKES IMAGES INACCESSIBLE. IT IS PRESENTLY UNKNOWN WHETHER ANY DATA IS PERMANENTLY LOST. THE ROOT CAUSE OF THE ISSUE IS ASSUMED TO BE A HARDWARE ERROR. THERE IS NO CONTRIBUTION OF THE SYNGO IMAGING PRODUCT TO THE UNAVAILABILITY OF IMAGES. NO HARM HAS BEEN REPORTED FROM THIS SITE, HOWEVER, PRIOR DATA MAY NOT BE AVAILABLE IF NEEDED FOR COMPARISON, LEADING TO REDUCED BASE FOR DIAGNOSIS. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE TOTAL AMOUNT OF POTENTIALLY LOST DATA HAS NOT YET BEEN DETERMINED. SIEMENS IS CONDUCTING A THOROUGH INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | -device malfunction -insufficient storage space -device damaged | -leak -disconnection -fitting problem | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH DISK 5, WHERE THE STS (SHORT TERM STORAGE) IS LOCATED, ON THE SYNGO.PLAZA UNIT. DURING NORMAL OPERATION DISK 5 WAS MARKED AS "UNALLOCATED', RESULTING IN THE SYNGO.PLAZA BEING UNABLE TO READ IMAGES/PATIENT DATA. THE REPORTED ISSUE IS A HARDWARE ISSUE. POSSIBILITY TO RECOVER THE DISK 5 IS CURRENTLY BEING ASSESSED BY SIEMENS TECHNICAL EXPERTS. THE AMOUNT OF POTENTIALLY LOST DATA IS CURRENTLY UNKNOWN. THERE ARE NO INJURIES ATTRIBUTED TO THE REPORTED EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. STS IS INTENDED ONLY FOR TEMPORARY STORAGE OF PATIENT DATA UNTIL IT IS TRANSFERRED TO LTS (LONG TERM STORAGE). THE SYNGO.PLAZA IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE SYNGO.PLAZA PRODUCT. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. (B)(6). | -unable to operate -unsecured data -component unattached | -loss of power -leak -circuit failure | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH THE POTENTIAL OF DATA PATIENT MIX-UP FROM A SITE IN (B)(6). PER THE COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. HOWEVER, SIEMENS FACTORY EXPERTS EVALUATED THE REPORTED ISSUE AS LOW RISK FOR THE PRODUCT AND RISK FOR THE SITE. IT WAS REPORTED THAT PATIENT DATA GOT MIXED UP FOR ABOUT 3 DAYS DURING 2 OCCURRENCES. THE MIX-UP WAS CORRECTED AND NO IMPACT FOR PATIENTS WAS REPORTED. SOME OF THE DATA MIX-UP WAS DETECTED BY THE USER. THE COMPLETE AMOUNT OF DATA MIX-UP WAS DETERMINED AND ALL DATA WAS CORRECTED. THE SYNGO.PLAZA PRODUCT IS WORKING AS DESIGNED AND INTENDED; THE ESTABLISHED MITIGATION TO ROLL BACK A FAILED STORING ACTIVITY IS CURRENTLY IN PLACE AND EFFECTIVE UNDER NORMAL CIRCUMSTANCES. HOWEVER, THE IMPLEMENTATION DID NOT CONSIDER SUCH AN EXCEPTIONAL SCENARIO AS IN THIS CASE. THEREFORE AN INVESTIGATION HAS STARTED TO DETERMINE WHETHER THE MITIGATION CAN BE ENHANCED. Manufacturer Narrative: THIS SITE HAS A HIGH AVAILABILITY SETUP WITH TWO REDUNDANT SYNGO.PLAZA SERVERS, COMBINED IN A SO CALLED SERVER CLUSTER. FOR THE DICOM-PARTNERS IN THE CLINICAL NETWORK THIS CLUSTER WILL APPEAR AND WILL BE TREATED TRANSPARENTLY AS A SINGLE SYNGO.PLAZA SERVER. THE INTENTION OF THIS SETUP IS TO ENSURE AN OPERATION OF THIS SINGLE OUT-SIDE VISIBLE SYNGO.PLAZA SERVER WITHOUT INTERRUPTION. TO ACHIEVE THIS THE "CLUSTER MANAGER" SOFTWARE (PROVIDED BY MICROSOFT) WILL MONITOR THE STATE OF THE ACTIVATE "OPERATIVE" SERVER FOR PROPER OPERATION. IN CASE THIS "OPERATIVE" SERVER SHOULD FAIL, THE CLUSTER MANAGER WILL ACTIVATE THE SECOND - "SPARE" SERVER TO TAKE OVER. AT ANY TIME ONLY ONE OF THE SERVERS EXCLUSIVELY WILL TAKE CARE FOR THE TASKS TO BE DONE BY THE SERVER CLUSTER. AT THIS SITE THE CLUSTER MANAGER GOT INTO A DISORDERED STATE CAUSED BY HARDWARE PROBLEMS. IN THIS ERROR STATE HE FREQUENTLY AND RAPIDLY SWITCHED BETWEEN THE TWO SERVERS OF THE CLUSTER. AS A RESULT THE TWO SERVERS AT SOME POINTS IN TIME SPORADICALLY WERE RUNNING IN PARALLEL - NOT EXCLUSIVELY. DURING THESE TIMEPOINTS THE TASKS TO BE PERFORMED BY THE SERVER-CLUSTER WERE DONE IN PARALLEL BY BOTH OF THE SERVERS, WRITING CONCURRENTLY ONTO THE SAME DATA FILES AND SAME DATABASE RECORDS. DEPENDING ON THE NON-PREDICTABLE SEQUENCE OF THE ACTIONS IN THIS ERROR SCENARIO, THE CONTENT OF DATA FILES AND DATABASE REFERENCES HAD MIXED PATIENT INFORMATION. THE SYNGO.PLAZA IS WORKING AS DESIGNED AND INTENDED. WITH THIS INFORMATION, A DEATH OR SERIOUS INJURY IS HIGHLY IMPROBABLE. (B)(6). | -incorrect data -illegible data -incorrect data entered | -use of device issue -device alarm system issue -device operates differently than expected | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF AN ISSUE WITH TWO ORACLE DATABASE PARTITIONS ON THE SYNGO IMAGING SYSTEM. THE ORACLE DATABASE PARTITION IS A TABLE THAT CONTAINS DICOM INSTANCE DETAILS FOR THE SYNGO IMAGING ARCHIVE. THESE PARTITIONS ARE NOT ACCESSIBLE DUE TO I/O ERRORS IN THE FILESYSTEM. IMAGES MAINTAINED IN THESE PARTS OF THE DATABASE CANNOT BE OPENED. NO IMPACT ON CLINICAL ROUTINE AT THE SITE HAS BEEN OBSERVED. THE USER CONTINUES TO READ AND SAVE CURRENT EXAMINATIONS. HOWEVER, IF PRIOR IMAGES ARE NEEDED FOR A FOLLOW UP READING AND THEY ARE LOCATED IN THE INOPERABLE PARTS OF THE DATABASE, IT WOULD NOT BE POSSIBLE TO ACCESS THEM LEADING TO A REDUCED BASIS FOR CLINICAL DIAGNOSIS. THE ROOT CAUSE OF THE REPORTED ISSUE IS A HARDWARE ERROR. THE SYNGO IMAGING PRODUCT DID NOT CONTRIBUTE TO THE UNAVAILABILITY OF IMAGES. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/19. ACTIVITIES TO REPAIR THE DEFECTIVE DATABASE AT THE USER SITE AS WELL AS REVIEW FOR POTENTIALLY MISSING DATA AND ITS FURTHER RESTORATION ARE ON-GOING. THE FINAL AMOUNT OF DATA LOSS (IF ANY) CAN BE DETERMINED ONLY AFTER THE RECOVERY ACTIVITIES AT SITE ARE COMPLETED. NO CONSEQUENCES FROM THE USER HAVE BEEN REPORTED SO FAR. CUSTOMER'S ADDRESS: (B)(6). | -fail to generate indication -device damaged -device malfunction | -device alarm system issue -data issue -data back-up problem | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS CAUSED BY A HARDWARE MAINTENANCE ISSUE. OUR FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITH NO RISK FOR THE PRODUCT WITH POTENTIAL RISK FOR THE USER. THE DRIVES D:\ (INSTALLATION DRIVE -VIRTUAL), F:\ AND G:\ ON THE SYNGO.PLAZA SERVER WERE EITHER NOT ACCESSIBLE OR EMPTY. FIVE HDD'S WERE FOUND TO BE DEFECTIVE ON THE STORAGE/DATASTORE FOR THE PLAZA VM, WHICH IS CONFIGURED FROM RHDASSRV1. THESE DRIVES HOLD THE PARTITIONS OF THE SYNGO PLAZA SERVER. NAS1 SERVER/LONG TERM STORAGE (LTS) HAS 2 HDD'S NOT ACCESSIBLE/EMPTY AS WELL. TOTAL OF 7 HDD'S ARE EITHER DEFECTIVE/EMPTY OR NOT ACCESSIBLE BY SERVERS. ALL PRIOR INFORMATION IS LOST. CLINICAL DIAGNOSIS MAY BE POTENTIALLY AFFECTED AS OLD ARCHIVED DATA IS NOT AVAILABLE TO PRIOR COMPARISON. THE REPORTED DATA LOSS WAS NOT RELATED TO THE SYNGO.PLAZA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED INCIDENT OCCURRED IN (B)(6). Manufacturer Narrative: THE ADMINISTRATOR MANUAL HAS A CAUTION "PLAZA_RISK_DOCUPS", INFORMING USER TO CONNECT ALL SYSTEM COMPONENTS TO AN UNINTERRUPTIBLE POWER SUPPLY (UPS). ANOTHER CAUTION - "PLAZA_RISK_DOCDATABASENOTAVAILABLE" - EXISTS TO INDICATE THE NEED TO MAINTAIN BACKUP TO AVOID DATA LOSS DURING HARDWARE FAILURES. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS CUSTOMER. | -loss of product function -loss of electrical connection -loss of vacuum | -compliance issue -human-device interface issue -protective measure issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON NAS SERVER PLACED ON THE SYNGO.PLAZA PRODUCT. THE AFFECTED DATA IS DATED TO THE TIME FRAME (B)(6) 2017 AND (B)(6) 2017. OUR FACTORY EXPERTS EVALUATED THE REPORTED ISSUE WITH NO RISK FOR THE PRODUCT WITH POTENTIAL RISK FOR THE USER. THE USER SITE WAS A CENTRAL SERVER DEPLOYMENT WITH ONE NAS MOUNT POINT - NAS1 - CONFIGURED AS PRIMARY IN THE ADMIN PORTAL. USING AN EXTERNAL APPLICATION ( QNAP-QTS ), A SECOND NAS MOUNT POINT - NAS2 - WAS SETUP AS BACKUP FOR NAS1. SEQUENCE OF EVENTS WAS ANALYZED AND EXTRACTED FROM THE LOG FILES. DETAILS ARE BELOW: ARCHIVE AS WELL AS DE-ARCHIVE FUNCTIONS WERE WORKING AS INTENDED UNTIL (B)(6) 2017 - 17:31:41. ON (B)(6) 2017 THE NAS1 HAD SOME HARDWARE ISSUE. CENTRAL SERVER COULD NOT REACH THE NAS1. DUE TO THIS, BOTH ARCHIVING AS WELL AS DE-ARCHIVING FUNCTIONS STOPPED WORKING UNTIL (B)(6) 2017. ON (B)(6) 2017 THE BACKUP NAS2'S IP ADDRESS WAS CHANGED MANUALLY TO NAS1 IP ADDRESS, TO RESOLVE THE ARCHIVING AND DE-ARCHIVING ISSUES IN THE USER SITE. AFTER THIS CHANGE, BOTH FUNCTIONS STARTED WORKING FINE. FROM (B)(6) TO (B)(6) ARCHIVING WAS BEING PERFORMED TO NAS2. DURING THIS PERIOD THE ARCHIVED FILES WERE ONLY PRESENT ON THE NAS2. AFTER FIXING THE HARDWARE ISSUE, THE IP ADDRESS WAS SWITCHED BACK TO RE-START ARCHIVING TO THE NAS1. THE EXTERNAL APPLICATION (QNAP-QTS) WAS ACTIVATED TO RESTART BACKUP FROM THE NAS1 TO THE NAS2. AT THIS POINT, DUE TO THE REPLICA CONFIGURATION OF QNAP-QTS, THE NAS2 WAS BROUGHT TO CONTENT OF THE NAS1 (ADDING NEW DATA BUT REMOVING DATA NOT PRESENT EARLIER ON THE NAS1). ADDITIONAL FILES ARCHIVED IN THE NAS2 DURING THE INTERMEDIATE SERVICING TIME WERE LOST DUE TO THIS REPLICATION. OUR EXPERTS ASSESSED THAT DATA FOR 4811 PATIENTS RANGING FROM (B)(6) 2017 TO (B)(6) 2017 WAS LOST. CLINICAL DIAGNOSIS MAY BE POTENTIALLY AFFECTED AS OLD ARCHIVED DATA IS NOT AVAILABLE TO PRIOR COMPARISON. THE SYNGO. PLAZA PRODUCT IS WORKING AS DESIGNED AND INTENDED. THE REPORTED DATA LOSS WAS NOT RELATED TO THE SYNGO. PLAZA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA IS WORKING AS DESIGNED AND INTENDED. WITH THIS INFORMATION, A DEATH OR SERIOUS INJURY IS HIGHLY IMPROBABLE. (B)(6). | -loss of data -patient data breached -sensor data breached | -data issue -application security issue -communication or transmission issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE (B)(6) SYSTEM. HDDS AT THE CLINICAL FACILITY WERE DAMAGED OR BECAME UNAVAILABLE DUE TO POWER LOSS. THIS SCENARIO CORRUPTED USER'S ENTIRE RAID STORAGE VOLUME (HOSTING SYSTEM AND DATA). AFTER SYSTEM RECOVERY WAS PERFORMED AND DATA WAS RESTORED FROM ALL AVAILABLE SOURCES, THE TOTAL EXTENT OF DATA LOSS WAS DETERMINED AS FOLLOWS: (B)(4) STUDIES, (B)(4) SERIES, (B)(4) IMAGES. THESE (B)(4) STUDIES ARE DATED WITHIN THE TIME PERIOD OF (B)(6) 2002 TO (B)(6) 2019. THEY ARE ASSESSED AS PARTIALLY CLINICAL RELEVANT. THE (B)(6) IS WORKING AS DESIGNED AND INTENDED, AND THE REPORTED DATA LOSS WAS UNRELATED TO THE (B)(6) PRODUCT. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE (B)(6) IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE (B)(6) PRODUCT. THE (B)(6) PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS USER. (B)(6). | -data corrupted -loss of data -loss of power indication | -power conditioning issue -loss of data -device difficult to setup or prepare | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE SYNGO PLAZA SYSTEM. THE ONLINE CONSISTENCY CHECK (OCC) TOOL REPORTED DATA LOSS IN ONLINE PARTITION (FILES NOT AVAILABLE AND FILES CORRUPTED). THERE WERE MULTIPLE OCCURRENCES OF BLUE SCREEN OF DEATH (BSOD) 0X51 ERRORS CAUSING LOSS OR CORRUPTION OF IMAGE FILES. SOME OF THE MISSING/CORRUPT PATIENT DATA WAS SUCCESSFULLY RECOVERED. HOWEVER, THE TOTAL EXTENT OF REPORTED DATA LOSS IS AS FOLLOWS: 28 PATIENTS. 28 STUDIES: ALL STUDIES ARE DATED BETWEEN (B)(6) 2016 AND (B)(6) 2018. 53 SERIES. 1577 IMAGES. THE AFFECTED DATA WAS STORED ON THE SHORT TERM STORAGE (STS) AND WAS NOT YET ARCHIVED. THE STUDY REPORTS WERE NOT YET GENERATED IN THE SYNGO PLAZA SYSTEM. SINCE THE STUDIES ARE FROM RECENT TIMELINE, THEY ARE STILL ASSESSED AS CLINICALLY RELEVANT FOR FOLLOW-UP DIAGNOSIS. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. STS IS INTENDED ONLY FOR TEMPORARY STORAGE OF PATIENT DATA UNTIL IT IS TRANSFERRED TO LTS (LONG TERM STORAGE). THE SYNGO.PLAZA IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE SYNGO.PLAZA PRODUCT. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS USER. (B)(6). | -loss of data -data corrupted -configuration data corrupted | -data issue -application security issue -data back-up problem | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE SYNGO.PLAZA SYSTEM. IT WAS REPORTED THAT PATIENT IMAGES WERE MISSING FROM THE SHORT TERM STORAGE (STS). THE MISSING IMAGES WERE DATED BETWEEN JANUARY 2017 AND JANUARY 2019 AND THE TOTAL EXTENT OF DATA LOSS WAS DETERMINED AS FOLLOWS: 29 PATIENTS / 32 STUDIES / 38 SERIES / 4973 IMAGES. AT THE TIME OF THE ISSUE THE CONCERNED SITE WAS RUNNING ON THE SYNGO.PLAZA WITH SW VERSION VB20A_HF05. THE AVAILABLE LOG FILES SHOWED THAT ALL SYNGO.PLAZA SERVICE SHUT DOWN ABNORMALLY. THE TRIGGER FOR THE SHUT DOWN IS UNKNOWN AS NO EVENT LOG FILES ARE AVAILABLE FOR THAT DATE. IT IS RESPONSIBILITY OF THE USER TO MAINTAIN THE HARDWARE. THE SYSTEM SINCE THEN HAS BEEN UPDATE TO NEWER VERSION - SYNGO.PLAZA WITH SW VB20A_HF07. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: THE SYNGO.PLAZA IS WORKING AS DESIGNED AND ACCORDING TO THE SPECIFICATIONS; THE REPORTED DATA LOSS WAS UNRELATED TO THE SYNGO.PLAZA PRODUCT. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THE CONCERNED SITE. NO CONSEQUENCES HAVE BEEN REPORTED FROM THIS USER. | -missing review of pre-operative patient images -loss of data -loss of image | -component missing -data back-up problem -device markings issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF DATA LOSS ON THE SYNGO.PLAZA SYSTEM. THE INITIAL ISSUE OF HARDWARE RECOVERY AND OPERATIONAL SYSTEMS/DATABASE RESTORATION WAS REPORTED UNDER MDR # 2240869-2019-69334, HOWEVER, IT WAS LATER RECOGNIZED THAT APPROXIMATELY 900 STUDIES DATING TO THE TIMEFRAME 2016 - 2019 HAD NOT BEEN ARCHIVED. SEVERAL DATA FILES (PLAZA MTS FILES) WERE FOUND WITHOUT REFERENCE TO THE SYNGO.PLAZA DATABASE. AFTER THE MTS FILES WERE IMPORTED, THE DATA LOSS WAS EVALUATED AS 525 STUDIES WITH 4617 SERIES. THE AFFECTED DATA CONSISTED OF RECENTLY ACQUIRED IMAGES AND IT WAS STORED ON THE SHORT-TERM STORAGE (STS). THESE IMAGES ARE POTENTIALLY CLINICALLY RELEVANT FOR FOLLOW-UP DIAGNOSIS. THE USER IS AWARE OF THE UNAVAILABILITY OF THIS DATA. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. THE SYNGO.PLAZA IS WORKING AS DESIGNED AND INTENDED; THE REPORTED DATA LOSS WAS UNRELATED TO THE SYNGO.PLAZA PRODUCT. THE SYSTEM IS EQUIPPED WITH PREVENTATIVE CAUTION - PLAZA_RISK_DOCDATABASENOTAVAILABLE - THAT INDICATES THE NEED TO MAINTAIN BACKUP TO AVOID DATA LOSS DURING HARDWARE FAILURES. THE SYNGO.PLAZA PRODUCT DID NOT CONTRIBUTE TO THE PARTIAL DATA LOSS AT THIS SITE. NO CONSEQUENCES HAVE BEEN REPORTED THE USER. CUSTOMER'S ADDRESS: (B)(6). | -loss of data -loss of sensing data -loss of conduction | -Delivery System Failure -Device Handling Issue -data issue | Viewer
| FDA | Event Description: SIEMENS BECAME AWARE OF FILES BEING CORRUPTED ON THE SYNGO.PLAZA SYSTEM WITH SOFTWARE VERSION (B)(4). FROM THE AVAILABLE LOGS, IT WAS OBSERVED THAT ALL SYNGO.PLAZA SERVICES SHUT DOWN ABNORMALLY AT THE TIME OF FILE CREATION. THIS RESULTED IN CORRUPTED FILES FOR STUDIES, WHICH WERE RECEIVED DURING THE SHUTDOWN. IN TOTAL 125 FILES FROM 10 STUDIES WERE CORRUPTED - 100 FILES FROM 2019 AND 25 FILES FROM 2017. THERE WAS NO INJURY ASSOCIATED WITH THE ISSUE. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: WHEN THE USER LOADED THE STUDIES, AN ERROR SHOULD HAVE BEEN DISPLAYED TO THE USER INDICATING AN ISSUE. IT WAS CONFIRMED THAT THE SYNGO.PLAZA HAD SENT THE ERROR STATUS TO MODALITY FOR THE CORRUPTED IMAGES. IN THIS CASE SYNGO.PLAZA APPLICATION WORKED AS SPECIFIED. NO CONSEQUENCES HAVE BEEN REPORTED, NEITHER FROM THIS CUSTOMER NOR FROM THE INSTALLED BASE. INTERNAL ID # (B)(4). | -software corrupted -data corrupted -system abused | -device operates differently than expected -device inoperable -Burn of Device or Device Component | Viewer
| FDA | Event Description: SIEMENS HEALTHINEERS BECAME AWARE OF A DATA LOSS ON THE SYNGO IMAGING SYSTEM. INITIALLY 100 SERIES WERE REPORTED WITH NOTIFICATION "ARCHIVE FAILED". HOWEVER, THE ACTUAL INVESTIGATION SHOWED THAT ONLY 30 IMAGES OUT OF 2 SERIES WERE REALLY LOST DUE TO A HARDWARE FAILURE. THE LOST DATA WAS CLINICALLY RELEVANT; HOWEVER, NO PATIENT RESCANS WERE REQUIRED. THE INCOMPLETE DATA WAS INDICATED TO THE USER AT THE TIME OF IMAGE LOADING. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019. NO CONSEQUENCES FROM THIS INCIDENT HAVE BEEN REPORTED. WITH THIS INFORMATION, A DEATH OR SERIOUS INJURY IS HIGHLY IMPROBABLE. | -loss of data -patient experienced adverse event -loss of communication | -device alarm system issue -communication or transmission issue -human-device interface issue | Viewer
| FDA | Event Description: SIEMENS RECEIVED FEEDBACK FROM THE CUSTOMER REPORTING THAT THE DISPLAYED STANDARD UPTAKE VALUE (SUV) WAS INCORRECT DURING THE CASE AS IT WAS NOT RECONSTRUCTED FROM THE VERY BEGINNING. THIS OCCURRED WHEN THE CUSTOMER VIEWED THE LIST MODE RECONSTRUCTION DATA WITH MI APPLICATIONS ON THE SYNGO MULTIMODALITY WORKPLACE, WHICH WAS INSTALLED ON CUSTOMER'S BIOGRAPH 40 UNIT. ACCORDING TO THE CUSTOMER, THE DISPLAYED SUV WAS WRONG IN THE REPORTED CASE AS IT WAS NOT BEING RECONSTRUCTED FROM THE VERY BEGINNING. INCORRECT SUV MAY LEAD TO A POTENTIAL MISDIAGNOSIS AND INCORRECT PATIENT TREATMENT. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: **RESUBMISSION OF INITIAL REPORT AS PER FDA ON (B)(6) 2020** OUR EXPERTS WERE ABLE TO REPRODUCE THE REPORTED ISSUE. THE REPORTED ISSUE IS UNDER INVESTIGATION AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. CUSTOMER'S ADDRESS: (B)(6). | -incorrect display value -incorrect display format -incorrect message display | -no display or display failure -Failure to Sense -leak | Viewer
| FDA | Event Description: SIEMENS RECEIVED FEEDBACK FROM THE CUSTOMER REPORTING THAT THE DISPLAYED STANDARD UPTAKE VALUE (SUV) WAS INCORRECT DURING THE CASE AS IT WAS NOT RECONSTRUCTED FROM THE VERY BEGINNING. THIS OCCURRED WHEN THE CUSTOMER VIEWED THE LIST MODE RECONSTRUCTION DATA WITH MI APPLICATIONS ON THE SYNGO MULTIMODALITY WORKPLACE, WHICH WAS INSTALLED ON CUSTOMER'S BIOGRAPH 40 UNIT. ACCORDING TO THE CUSTOMER, THE DISPLAYED SUV WAS WRONG IN THE REPORTED CASE AS IT WAS NOT BEING RECONSTRUCTED FROM THE VERY BEGINNING. INCORRECT SUV MAY LEAD TO A POTENTIAL MISDIAGNOSIS AND INCORRECT PATIENT TREATMENT. THE REPORTED ISSUE OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE REPORTED ISSUE IS UNDER INVESTIGATION AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. (B)(6). | -incorrect display value -incorrect display format -incorrect message display | -no display or display failure -Failure to Sense -leak | Viewer
| FDA | Event Description: SIEMENS RECENTLY BECAME AWARE OF DATA LOSS THAT HAD OCCURRED ON THE SYNGO IMAGING SYSTEM. FEBRUARY 12, 2016, AFTER PERFORMING REPAIR ACTIONS ON THE HARD DISC UNIT (RAID) OF THE ARCHIVE THAT IS ATTACHED TO THE SYNGO IMAGING PRODUCT, IT WAS NOTICED THAT 6 STUDIES WERE MISSING AND 6 STUDIES WERE PARTIALLY MISSING. THESE STUDIES RANGED FROM JANUARY 31, 2016 TO FEBRUARY 5, 2016. THIS DATA LOSS WAS DISCOVERED BY A SIEMENS SERVICE ENGINEER WHO WAS PERFORMING THE REPAIR ACTION. THE ENGINEER COMPARED THE ARCHIVE STATUS BEFORE AND AFTER THE REPAIR ACTION AND DISCOVERED DATA LOSS. HOWEVER, THE CUSTOMER ALLOWED TO DELETE STUDIES CREATED PRIOR TO FEBRUARY 6, 2016 ON AFFECTED IMAGE PARTITION ON THE COURSE OF THE REPAIR ACTIONS. THE DATA LOSS WAS CAUSED BY A HARDWARE ERROR AS THE RAID UNIT PARTIALLY CRASHED. THE DEFECTIVE DISC WAS REPLACED BY A (B)(4) SERVICE ENGINEER. THE LOST DATA WAS ABOUT A WEEK OLD. IT IS UNKNOWN WHETHER IT WAS USED FOR CLINICAL DIAGNOSIS. HOWEVER, THE CUSTOMER CONFIRMED THAT DATA FROM THAT PERIOD OF THE TIME CAN BE DELETED. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE REPORTED EVENT WAS EVALUATED BY OUR FACTORY EXPERTS. THE SYNGO IMAGING SYSTEM IS FUNCTIONING ACCORDING TO SPECIFICATIONS AND DID NOT IN ANY WAY CONTRIBUTE TO THE DATA LOSS. THE REPORTED DATA LOSS WAS CAUSED BY A HARDWARE PROBLEM IN THE HARD DISC UNIT (RAID) AT THE CUSTOMER SITE. NO CONSEQUENCES FOR PATIENTS HAVE REPORTED. (B)(6). | -loss of data -loss of performance -loss of function | -battery issue -data back-up problem -electrical shorting | Viewer
| FDA | Event Description: SINCE 2010, THE IMAGE ARCHIVE IS DONE IN ONE NAS (NETWORK ATTACHED STORAGE), WHICH IS ATTACHED TO THE SYNGO IMAGING XS SERVER. ON (B)(6), THE NAS SYSTEM CRASHED SUDDENLY AND AFTER A RESTARTED, ONE OF THE LOGICAL DRIVES DISAPPEARED FROM THE OPERATING SYSTEM. THE DRIVE HAS BEEN RE-ENABLED AND DURING THE INVESTIGATIONS IT WAS IDENTIFIED THAT SEVERAL ZIP FILES CONTAINING THE PATIENT STUDIES ARE CORRUPTED. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 7/28/20. THERE ARE NO INJURIES ATTRIBUTED TO THE REPORTED EVENT. HOWEVER, THE INVESTIGATION WHETHER CLINICALLY RELEVANT DATA LOSS OCCURRED IS ON-GOING. THE AMOUNT OF DATA LOSS IS CURRENTLY UNKNOWN. ACTIVITIES ARE ONGOING AT SITE TO EVALUATE THE AMOUNT OF DATA LOST AND AFTER THOSE ACTIVITIES ARE FINISHED THE AMOUNT OF DATA LOST WILL BE KNOWN. THIS ACTIVITY WILL TAKE SEVERAL WEEKS. THIS REPORT WAS SUBMITTED JANUARY 19, 2017. CUSTOMER'S ADDRESS: (B)(6). | -loss of image -unable to scroll image -device crushed | -use of incorrect control settings -failure to read input signal -no display or display failure | Viewer
| FDA | Event Description: SINCE 2010, THE IMAGE ARCHIVE IS DONE IN ONE NAS (NETWORK ATTACHED STORAGE), WHICH IS ATTACHED TO THE SYNGO IMAGING XS SERVER. ON 15TH OF DECEMBER, THE NAS SYSTEM CRASHED SUDDENLY AND AFTER A RESTARTED, ONE OF THE LOGICAL DRIVES DISAPPEARED FROM THE OPERATING SYSTEM. THE DRIVE HAS BEEN RE-ENABLED AND DURING THE INVESTIGATIONS IT WAS IDENTIFIED THAT SEVERAL ZIP FILES CONTAINING THE PATIENT STUDIES ARE CORRUPTED. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT AS PER FDA ON 4/3/2019 THERE ARE NO INJURIES ATTRIBUTED TO THE REPORTED EVENT. HOWEVER, THE INVESTIGATION WHETHER CLINICALLY RELEVANT DATA LOSS OCCURRED IS ON-GOING. THE AMOUNT OF DATA LOSS IS CURRENTLY UNKNOWN. ACTIVITIES ARE ONGOING AT SITE TO EVALUATE THE AMOUNT OF DATA LOST AND AFTER THOSE ACTIVITIES ARE FINISHED THE AMOUNT OF DATA LOST WILL BE KNOWN. THIS ACTIVITY WILL TAKE SEVERAL WEEKS. (B)(6). | -loss of image -fail to upload images -unable to scroll image | -no display or display failure -defective item -use of incorrect control settings | Viewer
| FDA | Event Description: SITE BEGAN USING THE GE PURE FILTER ON CADSTREAM STUDIES AND AFTER SOME TIME REALIZED THEY WERE REFERRING LESS PATIENTS TO BIOPSY. SITE CONCERNED THAT THE PURE FILTER AND/OR CADSTREAM WERE CONTRIBUTING TO THEIR CHANGE IN FOLLOW-UP. THE OVERLAYS CALCULATED BY CADSTREAM MAY HAVE CONTRIBUTED TO THE INCORRECT PATIENT FOLLOW-UP FROM THE SITE. Manufacturer Narrative: SUBMITTING THIS SUPPLEMENTAL REPORT TO ADD FDA CORRECTION AND REMOVAL REFERENCE NUMBERS Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THE ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. | -filter plugged -incorrect labeling -flow blocked | -use of device issue -filtration issue -failure to read input signal | Viewer
| FDA | Event Description: TABLE VERTICAL BELT BROKE DURING PATIENT DIAGNOSTIC TEST. ARTIFACT RENDERED SCAN NOT READABLE. | -incorrect device orientation -device broken or damaged -part broken or damaged | -user used incorrect product for intended use -insufficient flow or underinfusion -patient data issue | Viewer
| FDA | Event Description: THE BUG WAS DISCOVERED DURING TESTING INTERNALLY IN THE MANUFACTURER'S FACILITY. THE BUG WAS EVALUATED TO HAVE PATIENT RISK. THE BUG CAUSES BOLD ACTIVATIONS MAPS TO BE VISUALIZED AS OVERLAYS WITHOUT TAKING THE COREGISTRATION INTO ACCOUNT. ANY OUTPUT CREATED FROM THESE ACTIVATION MAPS AS OVERLAYS WILL NOT BE ADJUSTED ACCORDING TO THE COREGISTRATION. THE BUG ONLY AFFECTS THE BOLD FMRI ANALYSIS AND NO OTHER MODULES OR ANALYSES. FURTHERMORE, THE PROBLEM WILL NOT OCCUR IF THE BOLD/DTI WIZARD IS USED, OR IF COREGISTRATION IS DONE BEFORE THE BOLD MODULE ANALYSIS IS STARTED. Manufacturer Narrative: DEVELOPMENT TEAM DURING INTERNAL TESTING DISCOVERED A BUG CONCERNING HANDLING OF COREGISTRATION FOR BOLD SERIES. THIS BUG WAS VERIFIED IN CURRENT VERSION 2.3.14. ALL USERS WILL BE NOTIFIED TO STOP USING THE BOLD MODULE IN NORDICICE AND CONTACT NORDICNEUROLAB FOR FURTHER INSTRUCTIONS FOR BLOCKING THE BOLD MODULE FOR FURTHER USE. AFFECTED USERS WILL BE OFFERED A TRANSITION TO NORDICBRAINEX ((B)(4)) WHICH IS OPTIMIZED FOR THE BOLD WORKFLOW. RISK EVALUATION: THE PROBLEM CAN POTENTIALLY CAUSE A MISALIGNMENT OF THE BOLD ACTIVATIONS IN RELATIONS TO THE STRUCTURAL UNDERLAY. GIVEN THAT THE HEAD MOTION CORRECTED BY THE COREGISTRATION IS USUALLY SMALL, THE EFFECT OF THIS BUG IS EQUALLY SMALL. THE MOST COMMON SIZE OF MISALIGNMENTS ARE ON THE SAME ORDER OF MAGNITUDE AS THE INHERENT UNCERTAINTIES OF THE ANALYSIS ITSELF, WHICH MAKES IT UNLIKELY THAT IS SHOULD INTRODUCE ANY ADDED RISK PATIENT HARM IN THE LAST STEP. FOR LARGER MISALIGNMENTS, IT IS UNLIKELY THAT THE ERROR IS NOT DETECTED BY THE USER. | -fail to identify risk -fail to identify hazard -fail to interpret images | -compliance issue -user used incorrect product for intended use -use of incorrect control settings | Viewer
| FDA | Event Description: THE CUSTOMER BELIEVES THAT THE STANDARDIZED UPTAKE VALUES (SUV) MEASUREMENTS ARE NOT AS ACCURATE WHEN THE MPR IS DONE USING A 2D IMAGE. AT THE TIME OF THIS REPORT, THE DEGREE OF INACCURACY AND ITS CLINICAL SIGNIFICANCE HAS NOT BEEN CONFIRMED. Manufacturer Narrative: AN INVESTIGATION INTO THE COMPLAINT FROM A CUSTOMER WHO ALLEGED THAT THE STANDARDIZED UPTAKE VALUE (SUV) MEASUREMENTS ARE NOT AS ACCURATE WHEN THE MPR IS DONE USING A 2D IMAGE, HAS FOUND THAT THE SUV MEASUREMENTS ARE ACCURATE IN 2D AND 3D . EVIDENCE TO SUPPORT THIS CONCLUSION INCLUDES THE FOLLOWING: A MERGE HEALTHCARE IMPLEMENTATION ENGINEER CONTACTED THE CUSTOMER TO OBTAIN ADDITIONAL INFORMATION AND TO SEE IF THE END USER WOULD HAVE ANY EXAMPLES OF THE ALLEGATION AVAILABLE. A WEB MEETING WAS HELD WITH AN ENGINEER AT MERGE AND A PHYSICIAN AT THE COMPLAINANT'S FACILITY. THE PHYSICIAN COMPLETED SUV MEASUREMENTS ON THE SAME PET/CT IMAGE USING THREE DIFFERENT VIEWERS: "NORMAL" PACS VIEWER, THE MPR SINGLE VIEWER AND THE MPR FOUR PANEL VIEWER. THE PHYSICIAN FOUND ALL OF THE SUV MEASUREMENTS WERE WITHIN APPROXIMATELY 1/10TH OF EACH OTHER. THE PHYSICIAN ACKNOWLEDGED THAT THESE MEASUREMENTS WERE ACCEPTABLE FOR HIM AND HIS PRACTICE. TO ENSURE THAT THE MERGE PACS PRODUCT FUNCTIONS AS EXPECTED, A MERGE SENIOR SOFTWARE DEVELOPER INVESTIGATED THE MERGE PACS SUV CALCULATION RELATED CODE. THE DEVELOPER CONFIRMED THAT THE CODE FUNCTIONS AS INTENDED. ADDITIONALLY, WHEN A REQUIRED SUV VARIABLE SUCH AS PATIENT WEIGHT, IS NOT AVAILABLE IN THE STUDY, A VALUE OF "N/A" IS DISPLAYED FOR THE MEASUREMENT. THE INVESTIGATION INTO THIS COMPLAINT ALSO DETERMINED THAT MERGE PACS WAS TESTED ACCORDING TO THE QUANTITATIVE IMAGING BIOMARKERS ALLIANCE (QIBA)/ RADIOLOGICAL SOCIETY OF NORTH AMERICA (RSNA) DIGITAL REFERENCE OBJECT (DRO) GUIDANCE. MERGE HEALTHCARE HAS ACKNOWLEDGED THIS REPORT AND COMPLETED THE SUV CALCULATIONS USING THE DRO. ALL TEST DATA WAS FOUND ACCEPTABLE. IN CONCLUSION, MERGE HEALTHCARE'S INVESTIGATION INTO THE ALLEGED INACCURATE SUV MEASUREMENTS WERE MOST LIKELY DUE TO A USER ERROR. MERGE PACS SATISFIES THE RSNA/QIBA GUIDELINES FOR SUV MEASUREMENTS. THE CUSTOMER WAS ABLE TO CONFIRM AFTER ADDITIONAL RESEARCH AND CALCULATIONS USING 3 DIFFERENT VIEWERS, THAT THE SUV MEASUREMENTS WERE WITHIN AN ACCEPTABLE RANGE. MERGE HEALTHCARE'S INVESTIGATION CONFIRMED THAT THIS IS NOT A DEVICE MALFUNCTION AND THERE WAS NO INJURY, POTENTIAL INJURY OR DEATH. Manufacturer Narrative: INVESTIGATION INTO THE ROOT CAUSE OF THE PROBLEM IS UNDERWAY AT THE TIME OF THIS REPORT. | -inaccurate measurement -inadequate measurement unit -inadequate image manipulation | -use of device issue -incorrect measurement -leak | Viewer
| FDA | Event Description: THE CUSTOMER CALLED IN A COMPLAINT ON (B)(6) 2015, STATING THEY HAD SENT TWO ACCESSION NUMBERS OVER THAT SAME MORNING AND ON THE AGILITY SIDE, THE PATIENT NAME IS INCORRECT. UPON INVESTIGATION, THE STUDY USER ID MATCHED SHOWING THE STUDY WAS IN IMPAX CORRECTLY AND IN AGILITY WITH THE SAME STUDY_USER ID, BUT IT HAD THE WRONG PATIENT DEMOGRAPHICS. AGFA DISCOVERED THERE WERE CASES WHERE SEPARATE PATIENTS EXISTED IN IMPAX WITH THE SAME PATIENT ID. AGFA R&D BECAME INVOLVED AND THE MAPPINGS WERE CHANGED IN AGILITY SO ALL STUDIES FROM CEDAR COUNTY HAD AN ASSIGNING AUTHORITY OF CCMH. THE CUSTOMER SITE IS AN IMPAX SITE THAT ROUTINELY SENDS STUDIES TO AN AFFILIATED SITE, NEVADA REGIONAL MEDICAL CENTER. NRMC JUST RECENTLY MIGRATED FROM AN IMPAX SYSTEM TO IMPAX AGILITY. BECAUSE THE CUSTOMER SITE HISTORICALLY REISSUES PATIENT IDS, MANY OF THEIR PATIENTS SHARE PATIENT IDS. IN IMPAX, THESE STUDIES ARE SUCCESSFULLY DIFFERENTIATED BY THE USE OF THE PUID (PATIENT USER ID) ATTRIBUTE. IMPAX AGILITY, HOWEVER, DOES NOT USE PUID; IT USES PATIENT ID AND ISSUER OF PATIENT ID TO DIFFERENTIATE/LINK PATIENTS. THIS MEANS WHEN STUDIES ARE SENT TO IMPAX AGILITY, THE PUID ATTRIBUTE IS DROPPED ALLOWING PATIENTS WITH SHARED PIDS TO BECOME LINKED (ESSENTIALLY MERGED). BECAUSE THE ISSUER OF PID HAS NOT CHANGED AND THE PIDS ARE THE SAME, IMPAX AGILITY TREATS THESE PATIENTS AS ONE AND ONLY ONE SET OF PATIENT DEMOGRAPHICS IS APPLIED TO ALL THOSE WHO WERE LINKED. WHEN PATIENTS WITH THE SAME PATIENT ID, BUT DIFFERENT PUID, ARE MIGRATED FROM WITHIN IMPAX TO IMPAX AGILITY, THESE PATIENTS GET MERGED AND THE PATIENT'S DEMOGRAPHICS FROM ONE PATIENT WILL BE LINKED TO THE OTHER PATIENT'S STUDIES. THE DATA SHOULD HAVE BEEN MIGRATED WITH A DIFFERENT ASSIGNING AUTHORITY DEPENDING ON THE IMPAX SYSTEM FROM WHERE THE DATA WAS PROVIDED (ISSUER). ON (B)(6) 2015, THE MIGRATION WAS HALTED. THE AFFECTED PATIENTS' RECORDS WERE IDENTIFIED AND A SOLUTION TO APPLY AN APPROPRIATE ASSIGNING AUTHORITY FOR NEW STUDIES WAS IMPLEMENTED. THE AFFECTED STUDIES WERE DELETED AND MIGRATED BACK INTO AGILITY WITH THE NEW SEPARATE ASSIGNING AUTHORITIES. THE ISSUE WAS CORRECTED REMOTELY AND FINISHED ON (B)(6) 2015. THE MIGRATION OF THE AFFILIATED AFFECTED SITE, NEVADA REGIONAL MEDICAL CENTER IS CURRENTLY IN PROGRESS AND IS BEING MONITORED. CORRECT MAPPINGS BETWEEN PUID & AA HAVE BEEN PUT IN PLACE. ROOT CAUSE IS IMPAX USES A DIFFERENT MECHANISM THAN AGILITY TO MAKE A DISTINCTION BETWEEN PATIENTS WITH THE SAME PATIENT ID FROM A DIFFERENT ISSUER. THIS WAS NOT STREAMLINED OR TAKEN INTO CONSIDERATION WHEN MIGRATING THE DATA FROM DIFFERENT IMPAX SYSTEMS TO AGILITY. TO PREVENT THIS FROM HAPPENING AT OTHER IMPAX AGILITY SITES IN THE FUTURE, THIS ISSUE AND THE WAY TO PREVENT IT FROM HAPPENING HAS BEEN UPDATED AND DESCRIBED IN THE IMPAX AGILITY - INTEGRATION GUIDE. A 21 CFR PART 806 REPORT WILL BE REPORTED TO THE FDA, TO INCLUDE FURTHER CORRECTIVE ACTIONS AND TRAINING TO IMPROVE THE KNOWLEDGE ON ASSIGNING AUTHORITY WITHIN AGFA SUPPORT ORGANIZATION. A LONG TERM FINAL SOLUTION WILL BE TO IMPLEMENT A PATIENT MATCHING RULE WHICH PROVIDES A SECONDARY CHECK TO ENSURE TWO PATIENT RECORDS WITH THE SAME PID AND ISSUER ARE INDEED THE SAME PERSON BY CHECKING OTHER PATIENT DEMOGRAPHICS IS BEING MADE TO THE PRODUCT. THERE ARE NO OTHER SITES AFFECTED. THERE WAS NO HARM TO PATIENTS OR USERS DURING THIS EVENT. ANY FURTHER ACTIONS WILL BE REPORTED VIA AGFA RRC (B)(4)AND THE FDA 21 CFR PART 806 REPORT. | -incorrect serial number -incorrect pin number -incorrect patient information | -break -hole in material -Device Damaged Prior to Use | Viewer
| FDA | Event Description: THE CUSTOMER COMPLAINED, THAT THE CR READER OVERHEATED CAUSING SMOKE TO FILL EMERGENCY DEPARTMENT, TWO PERSONS OF THE STAFF INHALED SMOKE. ACCORDING TO INFORMATION RECEIVED ONE PERSON WAS TREATED WITH 4LITER OF OXYGEN FOR 30 MINUTES. FURTHERMORE BLOOD WORK WAS DONE AND ONE PERSON APPARENTLY RECEIVED A CHEST X RAY. Manufacturer Narrative: THE COMPUTED RADIOGRAPHY (CR) READER PROCESSES DIGITAL IMAGE PLATES, AND SENDS THE IMAGE DATA TO THE IMAGE ACQUISITION SYSTEM WORKSTATION FOR REVIEWING, POST PROCESSING AND ARCHIVING THE CLINICAL IMAGE. VISUAL INSPECTION BY THE PHILIPS FIELD SERVICE ENGINEER SHOWED THAT THERE WAS NO VISIBLE EVIDENCE OF BURNING OF (A PART OF) THE READER. FOR TESTING REASON THE READER WAS TURNED ON BY FIELD SERVICE ENGINEER. NO SMOKE WAS DETECTED, BUT THE POWER SUPPLY EMITTED A BURNING SMELL WITHIN A FEW MINUTES OF OPERATION. THE ROOT CAUSE OF THE ISSUE POINTS TO A PREMATURE FAILURE OF AN ELECTRICAL COMPONENT INSIDE THE POWER SUPPLY. THE DEVICE IS COMPLIANT TO APPLICABLE STANDARDS LIKE EN 60601 1, EN 60950, AND UL/CSA, WHICH ADEQUATELY ADDRESS SUCH HAZARDS. THE OXYGEN TREATMENT TO THE TWO PERSONS IS CONSIDERED A PREVENTIVE MEASURE, NOT A THERAPEUTIC ONE. THEREFORE THIS INCIDENT IS FINALLY NOT CLASSIFIED AS A SERIOUS INJURY. CORRECTED DATA: ADVERSE EVENT OR PRODUCT PROBLEM (CHANGED FROM ADVERSE EVENT AND PRODUCT PROBLEM TO PRODUCT PROBLEM). TYPE OF REPORTABLE EVENT (CHANGED FROM SERIOUS INJURY TO MALFUNCTION). (B)(4). Manufacturer Narrative: (B)(4). THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. | -device overheating -patient experienced adverse event -part burned | -leak -disconnection -Protective Measure Issue | Viewer
| FDA | Event Description: THE CUSTOMER COMPLAINT, THAT BASED ON A STITCHING IMAGE PATIENT WAS TAKEN TO SURGERY UNNECESSARILY.Event Description: THE CUSTOMER COMPLAINT, THAT BASED ON A STITCHING IMAGE, WHEREON BROKEN RODS FOR THE SPINE WAS SHOWN, PATIENT WAS TAKEN TO SURGERY. CUSTOMER FOUND THAT SURGERY WAS NOT NEEDED AS RODS IN THE PATIENTS BACK WERE NOT BROKEN. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. (B)(4). Manufacturer Narrative: STITCHING OF X-RAY IMAGES IS OF INTEREST IN CASE OF DISEASE PATTERNS LIKE SCOLIOSIS OR ASYMMETRIES IN THE STRUCTURE OF LEG BONES. THE SYSTEM OPTIONALLY SUPPORTS THE AUTOMATIC AND MANUAL JOINING OF SEVERAL IMAGES TO PRODUCE ONE LARGE IMAGE THIS EVENT OCCURRED SINCE THE OPERATOR FLIPPED ONLY THE FIRST IMAGE MANUALLY OF THE TWO GENERATED IMAGES, THEN BOTH WERE STITCHED TOGETHER, SHOWING A WRONG COMPOSITE (MISALIGNED OR BROKEN RODS IN THE BACK.) BASED ON THIS STITCHING RESULT, A SURGERY WAS DONE, NO CONTROL X-RAY WAS PERFORMED FOR VERIFYING THE ISSUE. (B)(4). | -unraveled stitching -poor anatomy image -unclear image | -device operates differently than expected -break -failure of device to self-test | Viewer
| FDA | Event Description: THE CUSTOMER CONTACTED SYNAPTIVE MEDICAL INC ON 2017-03-08 AND REPORTED THAT SOME MAMMOGRAPHY IMAGES APPEAR VERY DARK ON IMAGE LOAD. NO ADVERSE EFFECTS TO THE PATIENT HAVE BEEN REPORTED AS OCCURRING. Manufacturer Narrative: ON (B)(6) 2017, THE CUSTOMER SENT SAMPLE ANONYMIZED IMAGES TO SYNPATIVE REP. OLDER IMAGES APPEARED CORRECT (JPEG LOSSLESS). THE PROBLEMATIC ONES COME FROM THIS YEAR AND ARE JPEG 2000 COMPRESSED. THE CUSTOMER FOUND THERE ARE NO DATA LUTS (LOOK UP TABLES) ENCODED INTO THE IMAGES. ALL IMAGES HAVE WINDOW CENTER ((B)(4)) AND WINDOW WIDTH ((B)(4)) VALUES SET TO 2047 AND 4096 RESPECTIVELY. THE PROBE TOOL REVEALED THE FOLLOWING VALUES: BLACK PIXELS HAVE VALUES OF 0.0 (AS EXPECTED), WHITE/GREY PIXELS HAVE VALUES THAT RANGE FROM: 1 TO ~250 ON STUDY #1 (PROBLEMATIC IMAGE), 1 TO ~3800 ON STUDY #2 ("GOOD" IMAGE) THE CUSTOMER DETERMINED THE WINDOW CENTER/WIDTH VALUES ARE IN AN APPROPRIATE RANGE FOR THE "GOOD" IMAGES (I.E. STUDY 2 AND 3), BUT ARE COMPLETELY OUT OF RANGE FOR THE "BAD" IMAGE (STUDY 1). THE CUSTOMER ALSO NOTED: "CLEARCANVAS'S ANALYSIS IS INCORRECT AS THE IMAGES HAVE DATA BEYOND 256. THE CUSTOMER CONCLUDED THERE IS A BUG DECOMPRESSING JPEG 2000 LOSSLESS IMAGES AND SO THEIR PROBE ANALYSIS IS INCORRECT" (NOTE: DATA BEYOND 256" IS REFERRING TO PIXEL DEPTH). AT THE TIME, FURTHER ANALYSIS FROM SYNAPTIVE FOUND THE FOLLOWING: THE BITS STORED ((B)(4)) VALUE IS SET TO 12-BIT IN THE DICOM HEADER; HOWEVER, IT APPEARS THE JPEG 2000 COMPRESSED PIXEL DATA IS ACTUALLY 16-BIT; WHEN THE TAG VALUE IS EDITED TO 16, THE IMAGE APPEARS CORRECT (I.E. NOT DARK) ON LOAD. THEREFORE, THE DICOM DATA IS INCONGRUENT. A WORKAROUND WAS PROVIDED TO THE CUSTOMER AT THE TIME: CORRECT THE DICOM HEADER TO REFLECT THE PIXEL DATA CORRECTLY. THIS SHOULD BE FIXED AT THE IMAGE SOURCE. OTHERWISE, MANUALLY CHANGE THE BITS STORED ((B)(4)) VALUE TO 16. ON 04-FEB-2020, A SOFTWARE DEFECT WAS DISCOVERED ON WHILE INVESTIGATING AN ISSUE REGARDING A RESEARCH DEVICE. THIS LED TO A RETROSPECTIVE ANALYSIS OF FEEDBACK TICKETS. THE RESEARCH DEVICE USES THE SAME JPEG 2000 CODEC AS THE DEVICE DESCRIBED IN THIS REPORT. WHILE THE LIKELIHOOD OF POTENTIAL SERIOUS HEALTH CONSEQUENCES IS REMOTE, THE USE OF THE DEFECTIVE SOFTWARE ASSOCIATED COULD RESULT IN MISDIAGNOSIS, POTENTIALLY CAUSING SIGNIFICANT INDIRECT HARM NECESSITATING MEDICAL INTERVENTION THAT IS SERIOUS BUT TEMPORARY. THIS DEFECT IS ONLY ENCOUNTERED IN SELECT SCENARIOS, WHERE A MODALITY PRODUCES IMAGES THAT ARE COMPRESSED USING JPEG 2000, AND THE IMAGE PIXEL DATA IS LESS THAN 16-BIT, AND THE "BITNESS" OF THE COMPRESSED DATA STREAM DOES NOT MATCH THAT OF THE IMAGE PIXEL DATA. THE OUTCOME OF THE ABOVE IS A LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA, WHICH CAUSES THE SYMPTOMS THAT WERE REPORTED BY THE CUSTOMER. THE ROOT CAUSE IS THAT A SOFTWARE DEFECT FOUND IN THE DEVICE THAT IS ENCOUNTERED WHEN IT IS USED WITH NON-DICOM COMPLIANT JPEG 2000 COMPRESSED IMAGES. THE SOFTWARE IMPROPERLY USES THE DICOM BIT DEPTH (I.E. THE BITS STORED TAG) TO DECOMPRESS THE COMPRESSED PIXEL DATA STREAM FOR DISPLAY, INSTEAD OF THE BIT DEPTH THAT IS ENCODED IN THE COMPRESSED PIXEL DATA STREAM ITSELF. IN NON-DICOM COMPLIANT IMAGES WHERE THE DICOM AND COMPRESSED PIXEL DATA STREAM BIT DEPTHS DO NOT MATCH, THE SOFTWARE OUTPUTS AN IMAGE WITH SOME LOSS OF PRECISION IN THE DECOMPRESSED PIXEL DATA. IN THESE SCENARIOS, THE MODALITY PRODUCING THE IMAGE IS ITSELF ALSO NOT COMPLIANT TO DICOM PS3.5 8.2.4. NOTE: THE CUSTOMER HAS TWO DEVICE LICENSES, (B)(4) THE COMPLAINT DESCRIBES ISSUES ENCOUNTERED WHILE USING ONE DEVICE BUT DID NOT DISCLOSE WHICH DEVICE WAS BEING USED. AS THE COMPLAINT IS REGARDING ONE DEVICE, A SINGLE MDR WILL BE REPORTED. MDR REPORTS: 3012075008-2020-00001, 3012075008-2020-00003 AND 3012075008-2020-00004 HAVE THE SAME ROOT CAUSE. A CAPA HAS BEEN OPENED AND RECALL INITIATED TO ADDRESS THE ISSUE IN AFFECTED DEVICES. | -dark patches -distorted images -loss of image | -no display or display failure -inadequate user interface -use of device issue | Viewer
| FDA | Event Description: THE CUSTOMER DISCOVERED A MISMATCH IN THE SYNGO.PLAZA PATIENT BROWSER FOR A CERTAIN PATIENT BETWEEN THE STUDY DESCRIPTION AND THE RELATED THUMBNAILS. THIS WAS NOTICED ON TIME; THEREFORE NO ACTUAL PATIENT MIX-UP OCCURRED. THE ROOT CAUSE OF THE PATIENT MISMATCH WAS THE MICROSOFT CLUSTER STARTING APPLICATION PROCESSES ON TWO SYNGO.PLAZA SYSTEMS SIMULTANEOUSLY. AFTER RESENDING THE AFFECTED STUDY, THE ISSUE COULD BE SOLVED BY THE SIEMENS HC CUSTOMER SERVICE. Manufacturer Narrative: THE ISSUE IS A REPORTABLE ADVERSE EVENT IN THE U.S. DUE TO PATIENT DATA MISMATCH, WHICH HAS A POTENTIAL TO BECOME PATIENT MIX-UP, WITHOUT ANY CONTRIBUTION OF THE SIEMENS SYNGO.PLAZA PRODUCT. (B)(6). | -part mismatched -adapter mismatched -incorrect patient data loaded | -fluid leak -High Readings -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER HAS CONFIGURED HANGING PROTOCOL (HP) CONFIGURATION OF "FIT TO WINDOW" AS THE ZOOM FACTOR. THE CUSTOMER STATED THAT THERE WAS AN INTERMITTENT ISSUE IN THAT SOME SERIES WERE ZOOMED AND SOME WERE NOT WHEN AXIAL IMAGES ARE LINKED, AND THEN ¨DRAGGED AND DROPPED¨ ONTO THE MONITOR. THE CUSTOMER INDICATED THAT A PARTICULAR REGION WAS NOT INCLUDED IN ONE OF THE IMAGES THAT WERE ZOOMED. THE CUSTOMER INDICATED THAT THE REGION IS CAPTURED IF USING HP WITH A MATRIX OF 1 PRESENTATION. AN ADVERSE EVENT DID NOT OCCURRED, NO MISDIAGNOSIS WAS MADE; NO TREATMENT RENDERED; SAME PHYSICIAN WAS MANUALLY CHANGING THE ZOOM ON THE IMAGES AND NOTICED THAT THE DEFAULT FIT TO WINDOW ZOOM DID NOT SHOW ALL THE PATHOLOGY. Manufacturer Narrative: A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. Manufacturer Narrative: INVESTIGATION IDENTIFIED A PRODUCT ANOMALY WAS FOUND RESULTING FROM CHANGES IN VALUES OF PIXEL SPACING (0028,0030) DICOM TAG BETWEEN IMAGES. WHENEVER A USER MOVES FROM AN IMAGE WITH ONE TYPE OF PIXEL SPACING TO ANOTHER IMAGE WHICH HAS A DIFFERENT PIXEL SPACING ON THE SAME WINDOW WHERE LINKING IS ENABLED, THEN FIT TO WINDOW ZOOM FACTOR MAY BE INCORRECT. INVESTIGATION OF ISSUE INDICATES THAT IT WOULD BE OBVIOUS TO THE USER THAT IF THE FIT TO WINDOW IMAGES SHOWED THE FULL REGION OF INTEREST AND THE ZOOMED IMAGES DID NOT AS THEY ARE VIEWED AT THE SAME TIME ON THE MONITOR; LINKED IMAGES ARE VIEWED AS A SUPPLEMENT REFERENCE ONLY TO INDIVIDUALLY VIEWED IMAGES WHICH ARE PRIMARILY USED FOR DIAGNOSIS; THERE IS A REFERENCE IMAGE AVAILABLE WHICH WOULD IMMEDIATELY INDICATE TO THE USER THAT SOME OF THE IMAGES ARE ZOOMED. THE ZOOM MAGNIFICATION VALUE IS SHOWN ON THE IMAGES. A ¨LINKED ICON¨ IS ALSO SHOWN ON THE SERIES/IMAGES THAT ARE LINKED TO REMIND THE RADIOLOGIST THEY ARE VIEWING LINKED IMAGES. IN CONCLUSION, THE RISK REPORTED ISSUE HAS BEEN REVIEWED AND IT DOES NOT POSE A RISK THAT WOULD CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT, IF IT WERE TO RECUR, WOULD NOT BE LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | -fail to setup configuration -incorrect configuration -inappropriate fit | -leak -device inoperable -Device displays error message | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED A DELAY IN SURGERY (X 8 DAYS) FOLLOWING A MISSED FRACTURE. PATIENT REQUIRED HOSPITALIZATION (UNKNOWN TIME) POST SURGERY. THE PATIENT PRESENTED IN THE ER ON (B)(6) 2018 WITH ACUTE TRAUMA. AN XR ANKLE RIGHT COMPLETE WAS ORDERED BY EMERGENCY ROOM CLINICIAN. THREE VIEWS WERE OBTAINED ON A LONG CASSETTE, PROVIDING ANATOMY BEYOND ANKLE TO INCLUDE MID SHAFT CALF. THESE VIEWS WERE SUBMITTED FOR INTERPRETATION ON THE RA 1000 AFTER THEY WERE MAGNIFIED BY THE TECHNOLOGIST AND MANUALLY SAVED. THE RADIOLOGIST RECEIVED THE EXAM IN THE MOST RECENT SAVED STATE (MAGNIFIED) AND PROVIDED REPORT BASED ON THIS PRESENTATION. THE RADIOLOGIST HAD THE OPTION TO VIEW IMAGES AS ACQUIRED AT ALL TIMES. THE FINDINGS IDENTIFIED NO RADIOGRAPHIC EVIDENCE OF ACUTE DISPLACED "OSSIOUS" FRACTURE OR DISLOCATION INVOLVING THE RIGHT ANKLE. THE PATIENT SOUGHT REPEATED CARE ON (B)(6) 2018, WITH A HISTORY OF RIGHT LOWER LEG PAIN. AS A RESULT, THE CLINICIAN ORDERED A XR LEG RIGHT LOWER TIBIA/FIBULA EXAM. FOUR VIEWS WERE OBTAINED. THESE VIEWS WERE SUBMITTED FOR INTERPRETATION ON THE RA 1000, REVEALING "A TRANSVERSE FRACTURE THROUGH THE TIBIAL SHAFT." THE INITIAL IMAGES FROM (B)(6) 2018 WERE RE-ASSESSED ON (B)(6) "2017" AND ADDENDED AS FOLLOWS: "AT THE TIME OF INITIAL INTERPRETATION, THE IMAGES WERE SUBMITTED IN A MAGNIFIED STATE THAT OBSCURED VISUALIZATION OF THE MID SHAFT TIBIAL FRACTURE. UPON REPROCESSING IMAGES, THERE IS A NONDISPLACED MIDSHAFT TIBIAL FRACTURE APPARENT ON THESE VIEWS." THE PATIENT REQUIRED SURGERY ON (B)(6) 2018 FOR ROD PLACEMENT TO REPAIR A RIGHT NONDISPLACED TIBIAL SHAFT FRACTURE. POST SURGERY FINDINGS: THERE HAS BEEN OPEN REDUCTION AND INTERNAL FIXATION OF THE TIBIAL DIAPHYSIS FRACTURE WITH NO EVIDENCE OF COMPLICATION STATUS POST SURGERY. Manufacturer Narrative: UDI NOT REQUIRED. WEIGHT COULD NOT BE OBTAINED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | -patient experienced adverse event -rehabilitation delayed -delayed monitoring | -failure of device to self-test -device alarm system issue -no display or display failure | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED A PATIENT CARDIAC SCAN WAS PROCESSED ON A XELERIS 3.1 WORKSTATION. THE PATIENT WAS DIAGNOSED AS "NORMAL" FOLLOWING AN INTERPRETATION THE PHYSICIAN HAD MADE. THE CARDIAC EXAM WAS PERFORMED AS A PRE-OPERATIVE TEST FOR A GALLBLADDER SURGERY THAT WAS PRESCRIBED FOR THE PATIENT. A WEEK AFTER THE CARDIAC EXAM READING, A GALLBLADDER SURGERY WAS PERFORMED. ON THE SAME DAY OF THE SURGERY, THE PATIENT SUFFERED A HEART ATTACK. AS A RESULT, A CATHETERIZATION WAS PERFORMED THAT DAY, AND A BYPASS SURGERY WAS PERFORMED ON THE PATIENT THE FOLLOWING DAY. THE EVENT IS CURRENTLY UNDER INVESTIGATION. Manufacturer Narrative: PATIENT WEIGHT UNKNOWN AT THIS TIME. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE INVESTIGATION REVEALED THE DEVICE AND INTERFACE PERFORMED AS DESIGNED AND INTENDED WITH A CORRECT RESPONSE. THERE WAS NO MALFUNCTION OF THE DEVICE AND THE DATA WAS ACCURATE. THE CAMERA AND SOFTWARE PERFORMED TECHNICALLY AS INTENDED WITH NO MALFUNCTION OR FAILURE. THE ROOT CAUSE WAS IDENTIFIED AS UNINTENDED MISINTERPRETATION OF SIGNIFICANT CLINICAL INFORMATION BY USER. | -fail to process medical imaging -wrong scan condition -error message | -moisture or humidity problem -Contamination During Use -Device Damaged Prior to Use | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED CANCER WAS NOT DETECTED IN A TIMELY MANNER DUE TO NOT VIEWING THE ADDITIONAL SERIES THAT WAS OBTAINED. AN MRI STUDY OF THE LEFT KNEE WAS PERFORMED ON (B)(6) 2019. IT WAS ORDERED STAT. THE FULL STUDY, CONTAINING 9 SERIES WAS SENT INTO THE SITES CPACS SYSTEM WITHIN 30 MINUTES OF ACQUISITION. USING UNIVERSAL VIEWER, THE RADIOLOGIST OPENED THE STUDY ON (B)(6) WHERE A LESION IN THE DISTAL FEMUR SHAFT WAS IDENTIFIED. DEDICATED RADIOLOGY OF FEMUR WAS RECOMMENDED. IN (B)(6) OF 2019, THE RADIOLOGIST WHO READ THE REPORT IDENTIFIED NEED FOR MORE URGENT FOLLOW UP IMAGING AFTER A RE-REVIEW OF MRI KNEE STUDY. IT WAS IDENTIFIED THAT THE 9TH SERIES HAD NOT BEEN VIEWED ON (B)(6). FURTHER TESTING WAS PERFORMED ON (B)(6), WHERE A PET SCAN WAS DONE TO CLASSIFY THE LESION AS CANCEROUS. Manufacturer Narrative: (B)(4). LEGAL MANUFACTURER: HCIT (B)(4) - (B)(4). INITIAL REPORTER: (B)(6). GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | -fail to visualize cancer location -loss of detection -unable to reset viewing panel | -Loss of Power -Decreased Sensitivity -inadequate storage | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT A PATIENT MAY HAVE HAD AN UNNECESSARY SURGERY. THE CUSTOMER ALLEGES THE EVENT OCCURRED DUE TO THE SURGEON NOT FINDING THE RADIOLOGY REPORT IN PACS. THE CUSTOMER STATED THE SURGEON SAW THE IMAGES IN [?]VERIFIED' STATUS IN PACS AND CHOSE TO REVIEW HIMSELF. THIS RESULTED IN THE SURGEON DETERMINING SURGERY WAS APPROPRIATE. AT THE TIME OF THIS EVENT THE RADIOLOGIST'S REPORT WAS COMPLETE AND AVAILABLE IN RISI AND THE CUSTOMER'S HIS. NO FURTHER INFORMATION AVAILABLE AT THIS TIME. Manufacturer Narrative: PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS. DATE OF EVENT IS UNKNOWN. THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. DATE OF MANUFACTURE IS NOT KNOWN. Manufacturer Narrative: THE INITIAL REPORTED EVENT DETAILS WERE DETERMINED TO BE INCORRECT. THE CUSTOMER HAD INITIALLY REPORTED THAT AN EXISTING REPORT WAS NOT TRANSFERRED FROM THE RIS TO PACS AND WAS THEREFORE NOT VISIBLE IN PACS. THE PHYSICIAN ONLY CHECKED THE PACS AND ASSUMED THAT THERE WAS NO REPORT. THE CUSTOMER STATED THE SURGEON SAW THE IMAGES IN VERIFIED STATUS IN PACS AND CHOSE TO REVIEW HIMSELF, RESULTING IN THE DETERMINATION THAT SURGERY WAS APPROPRIATE. THE INITIAL INFORMATION INDICATED THAT THE DOCTOR PERFORMED SURGERY AND IT ENDED IN A CRITICAL SITUATION FOR THE PATIENT. INVESTIGATION DETERMINED THAT AN ANGIOGRAM WAS PERFORMED WITH THE RESULTS REPORT RELEASED ON THE SAME DAY AND TRANSFERRED FROM RISI TO KIS SAP SYSTEM. THE CONTENT OF THE REPORT WAS NOT VISIBLE IN PACS AND THEREFORE ALSO NOT VISIBLE IN CENTRICITY WEB. THE WARD PHYSICIAN CALLED THE RADIOLOGY DEPARTMENT FOR THE REPORT AND WAS VERBALLY PROVIDED REPORT DETAILS. THERE WAS DIVERGENCE IN THE DESCRIPTION OF VASCULAR CALCINATION BETWEEN THE VERBAL AND WRITTEN REPORTS. ON (B)(6) 2015 SURGERY WAS PERFORMED ON THE PATIENT WITHOUT INCIDENT. THE CUSTOMER REPORTED THE SURGEON WOULD HAVE PERFORMED SURGERY IN THE SAME MANNER PER THE WRITTEN REPORT RESULT; IT WAS ONLY THE SUTURES THAT WERE DONE DIFFERENTLY AS ORIGINALLY PLANNED. THE ANGIOGRAM WAS PERFORMED ON THE PATIENT 10 DAYS BEFORE THE SURGERY WAS PERFORMED. THE REPORT WAS AVAILABLE AT THE TIME OF SURGERY FOR FURTHER REVIEW IF NEEDED. NO HARM TO PATIENT RESULTED. THE INVESTIGATION DETERMINED THAT THE PACS SOFTWARE PERFORMED AS DESIGNED, BUT THAT THERE WAS A CONFIGURATION SETTING FOR THE PACS/RIS INTERFACE THAT RESULTED IN THE REPORT NOT BEING PRESENT IN THE PACS DATABASE. THE USER COULD OBTAIN THE REPORTS FROM THE RIS SYSTEM AS WELL AS THE HOSPITAL INFORMATION SYSTEM (HIS). THE CONFIGURATION HAS BEEN CORRECTED AND ALL HISTORICAL REPORTS HAVE BEEN RESENT TO THE PACS. THE CORRECT USER NAMES WERE ALSO PROPERLY CONFIGURED IN THE PACS. IT IS SPECIFIED IN THE CENTRICITY RIS-I 4.2 PLUS HIS-RIS /RIS-PACS INTERFACE CONFIGURATION SERVICE MANUAL THAT THE PARAMETER FOR THE APPROVER USER CODE MUST BE SET ON THE PACS SIDE SO AS TO ALLOW TO RECEIVE REPORTS WITH UNKNOWN USER CODES: RIS REQUEST ACCEPT REPORT WITH UNKNOWN APPROVING STAFF ID =(B)(6). | -patient experienced adverse event -incorrect patient information -patient notifier malfunction | -leak -Device Damaged Prior to Use -disconnection | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER USED A SCENARIO WHICH WAS NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | -incorrect patient information -incorrect message display -incorrect display value | -no display or display failure -incorrect display -human-device interface issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT PANORAMIC X-RAY IMAGES ACQUIRED AT THE END OF THE WORD DAY WERE DUPLICATED AND SAVED IN THE WRONG PATIENT PROFILE. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER CONTACTED GENDEX TECHNICAL SUPPORT ON (B)(6) 2015. GENDEX TECHNICAL SUPPORT UPDATED SOFTWARE DRIVER GXPICTURE FROM VERSION 3.4.1 TO 3.5.3. THE UPDATE TO THE SOFTWARE DRIVER RESOLVED THE ISSUE OF DUPLICATING IMAGES AND THE CUSTOMER REPORTED THAT EVERYTHING IS NOW WORKING CORRECTLY. NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | -duplicated requirement -unable to record event -unable to scroll image | -no display or display failure -inadequate instructions for healthcare professional -break | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT PANORAMIC X-RAY IMAGES FROM ONE PATIENT HAD APPEARED IN THE WRONG PATIENT PROFILE. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER CONTACTED GENDEX TECHNICAL SUPPORT ON (B)(6) 2016. TECHNICAL SUPPORT WAS ABLE TO RESOLVE THE PROBLEM BY MOVING THE IMAGE TO THE CORRECT PATIENT PROFILE. THE IMAGE WAS TAKEN OVER TWO YEARS AGO AND NO FURTHER TROUBLESHOOTING OR ROOT-CAUSE ANALYSIS COULD BE PERFORMED DUE TO THE LACK OF AVAILABLE INFORMATION. NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. DATA NOT AVAILABLE TO PERFORM ANALYSIS. | -incorrect profile for patient care -image not in dicom format -inappropriate image rendered | -Device Clogged -disconnection -Imprecision | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE COMPUTER FAILED TO BOOT UP. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE KEYBOARD WAS REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -fail to setup configuration -fail to generate indication | -failure to calibrate -Failure of device to Self-Test -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED (LIST IS JUMPING) WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: INVESTIGATION REVEALED THAT THE CUSTOMER USED THE SOFTWARE IN A MULTI-WORKSTATION ENVIRONMENT WHICH IS A SCENARIO NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. TECHNICAL SUPPORT REBUILT THE SQL DATABASE AND RETESTED WITH SUCCESSFUL RETREIVAL OF CORRECT PATIENT FILE. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. Manufacturer Narrative: INITIAL INVESTIGATION CONFIRMED REPORT OF "INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE (B)(6) (PATIENT) ADMINISTRATION SCREEN". A RESOLUTION WAS PROVIDED TO THE CUSTOMER. HOWEVER, FURTHER INVESTIGATION IS NEEDED TO DETERMINE ROOT CAUSE OF THE REPORTED ISSUE. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON INVESTIGATION COMPLETION. | -incorrect patient information -incorrect message display -incorrect display | -no display or display failure -incorrect display -human-device interface issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: INVESTIGATION REVEALED THAT THE CUSTOMER USED THE SOFTWARE IN A MULTI-WORKSTATION ENVIRONMENT WHICH IS A SCENARIO NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | -incorrect patient information -incorrect message display -incorrect display | -no display or display failure -incorrect display -human-device interface issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER USED A SCENARIO WHICH WAS NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION. | -incorrect patient information -incorrect message display -incorrect display | -no display or display failure -incorrect display -human-device interface issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. Manufacturer Narrative: THE CUSTOMER USED A SCENARIO WHICH WAS NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS APPEARS TO BE AN ISOLATED INCIDENT AND WILL BE TRACKED FOR TRENDING. THIS CONCLUDES OUR INVESTIGATION. | -incorrect patient information -incorrect message display -incorrect display value | -no display or display failure -incorrect display -human-device interface issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE MOUSE NAVIGATES THROUGH SYSTEM WAS FROZEN. IF THE MOUSE AND THE KEYBOARD ARE NOT WORKING, THE SYSTEM WILL NOT BOOT. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN OVER THE PHONE INVESTIGATION. THE REPORTED ISSUE COULD NOT BE DUPLICATED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -device frozen -frozen screen -unresponsive mouse click | -leak -break -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE NAV SYSTEM WOULD NOT BOOT UP. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ON SITE INVESTIGATION. THE COMPUTER AND THE SOFTWARE WERE INSTALLED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -system failure -unable to compress | -Device Alarm System Issue -failure to deliver -no flow | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE STAT PRIORITY ICON WAS NOT DISPLAYED IN THE STUDY LIST WHEN THE STUDY WAS OPENED. NO PATIENT HARM WAS REPORTED DUE TO THIS ISSUE. Manufacturer Narrative: MCKESSON COMPLETED AN INVESTIGATION AND CONCLUDED THAT THE ROOT CAUSE OF THE PROBLEM WAS A SOFTWARE DEFECT. MCKESSON FOUND THAT THE STAT PRIORITY ALERT ICON MAY NOT BE DISPLAYED IN THE LIST, FOR ANY USER, WHEN ALL STAT STUDIES ARE OPENED AT THE SAME TIME AND THERE ARE NO "HIGH" PRIORITY STUDIES IN THE UNREPORTED STUDY LIST. UNDER THESE CONDITIONS, THE STAT PRIORITY BUTTON WILL NOT DISPLAY THE CORRECT NUMBER OF STAT STUDIES. MCKESSON WILL WORK WITH AFFECTED CUSTOMERS TO APPLY A SOFTWARE UPDATE TO PREVENT THE RECURRENCE OF THE PROBLEM. | -incorrect message display -fail to display message -fail to indicate priority | -insufficient information -No Known Device Problem -infusion or flow problem | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT AND EXHIBITED A NON-RECOVERABLE LOSS OF FUNCTIONALITY. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE UNIVERSAL PCB ASSEMBLY WAS EVALUATED AND REPLACED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -fail to activate mode -fail to resist normal activities -fail to suspend therapy | -failure to calibrate -Device Operates Differently than Expected -Failure of device to Self-Test | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT AND EXHIBITED A NON-RECOVERABLE LOSS OF FUNCTIONALITY. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN OVER THE PHONE INVESTIGATION. SYSTEM FILE CORRUPTION WAS IDENTIFIED AS THE CAUSE OF THE EVENT. THE REPORTED PROBLEM WAS RESOLVED BY THE CUSTOMER WITH ASSISTANCE FROM THE ENGINEER. NO ADDITIONAL SERVICE INFORMATION WAS PROVIDED. | -fail to activate mode -fail to resist normal activities -fail to suspend therapy | -failure to calibrate -Device Operates Differently than Expected -Failure of device to Self-Test | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT AND EXHIBITED A NON-RECOVERABLE LOSS OF SYSTEM FUNCTIONALITY. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE ULTRA 10 CPU WAS EVALUATED AND REPLACED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -fail to boot up -fail to initiate ventilation -fail to resume | -failure to calibrate -disconnection -inaccurate delivery | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT TO A USABLE STATE AND EXHIBITED A NON-RECOVERABLE LOSS OF FUNCTIONALITY. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE UPS (UNINTERRUPTIBLE POWER SUPPLY) WAS EVALUATED AND IDENTIFIED AS REQUIRING REPLACEMENT. NO CONCLUSION CAN BE DRAWN AS FURTHER SYSTEM ANALYSIS OR REPAIR INFORMATION IS UNAVAILABLE AT THIS TIME AND NO ADDITIONAL SERVICE INFORMATION WAS PROVIDED. SEE SCANNED PAGE. | -fail to boot up -fail to initiate ventilation -fail to activate mode | -failure to calibrate -disconnection -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT TO A USABLE STATE AND EXHIBITED A NON-RECOVERABLE LOSS OF FUNCTIONALITY. NO PT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ONSITE INVESTIGATION. THE USER WAS COACHED THROUGH STEP BY STEP GUIDANCE ON HOW TO ADD THE CALLERS FILES FOR THE SENSOR CABLES TO THE SYSTEM. AFTER PROVIDING THE NEEDED ASSISTANCE THE SYSTEM BOOTED UP SUCCESSFULLY. | -fail to boot up -fail to deliver -fail to initiate ventilation | -failure to calibrate -disconnection -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT UP. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: THE INVESTIGATION INTO THE REPORTED EVENT DETERMINED THAT THERE WAS NO DEVICE MALFUNCTION. THIS IS NOT A REPORTABLE EVENT. COMPLAINT INVESTIGATION: THE FSE CONFIRMED THE PROBLEM REPORTED BY THE CUSTOMER OF SYSTEM WILL NOT BOOT. THE FSE DETERMINED THE CAUSE OF THE PROBLEM TO BE FAULTY HARD DRIVE WITH BAD SECTORS. THE FSE VERIFIED SYSTEM FUNCTIONALITY. THE PARTITIONED SERIAL ATA HARD-DRIVE THAT CONTAINS OPERATING SYSTEMS, AND ALSO CONTAINS THE SYSTEM SOFTWARE. THE HARD DRIVE HOLDS SOFTWARE THAT IS A SET OF INSTRUCTIONS THAT DIRECTS THE SYSTEM PROCESSOR TO PERFORM SPECIFIC OPERATIONS. A BAD SECTOR IS A SECTOR ON A COMPUTER'S DISK DRIVE OR FLASH MEMORY THAT IS EITHER INACCESSIBLE OR UNWRITEABLE DUE TO PERMANENT DAMAGE, SUCH AS PHYSICAL DAMAGE TO THE DISK SURFACE OR FAILED FLASH MEMORY TRANSISTORS. REPLACED OTHER COMPONENTS AS PART OF THE TROUBLESHOOTING AND REPAIR. BASED ON AGE OF THE SYSTEM, MANUFACTURED BEFORE 2006, THIS TYPE OF FAILURE WOULD BE EXPECTED AND REFLECTIVE OF THE NORMAL WEAR AND TEAR THAT IS ANTICIPATED AS THE SYSTEM IS USED. THIS COMPLAINT DOES NOT REPRESENT A MALFUNCTION. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE HARD DISK DRIVE AND MOTHERBOARD WERE EVALUATED AND REPLACED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -fail to boot up -fail to setup configuration -system failure | -failure to calibrate -Failure of device to Self-Test -Device Alarm System Issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN OVER THE PHONE INVESTIGATION. THE SYSTEM DATABASE WAS REBUILT WITH ASSISTANCE FROM THE USER. THE REPORTED PROBLEM WAS RESOLVED BY THE CUSTOMER. NO ADDITIONAL SERVICE INFORMATION WAS PROVIDED. | -fail to boot up -system failure -fail to insert | -failure to calibrate -Device Operates Differently than Expected -Failure of device to Self-Test | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM FAILED TO BOOT. THERE WAS NO PATIENT INJURY OR DEATH REPORTED. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE CONNECTORS WERE RESEATED AND THE SOFTWARE WAS INSTALLED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -system failure -fail to unlatch | -reset problem -Device displays error message -inappropriate or unexpected reset | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM HAD A BLUE SCREEN AND IT SHUTDOWN BY ITSELF, RESULTING IN A LOSS OF NAVIGATION FUNCTIONALITY. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ON SITE INVESTIGATION. THE POWER SUPPLY AND THE MONITOR CABLE WERE REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -blank screen -incorrect screen display -device shut down | -no display or display failure -reset problem -no pressure | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM INTERMITTENTLY FAILED TO BOOT TO A USABLE STATE AND EXHIBITED INTERMITTENT FUNCTIONALITY. NO PT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN OVER THE PHONE INVESTIGATION. THE KEYBOARD WAS IDENTIFIED AS A POSSIBLE CAUSE OF THE ISSUE. NO CONCLUSION CAN BE DRAWN AS FURTHER SYSTEM ANALYSIS OR REPAIR INFO IS UNAVAILABLE AT THIS TIME AND NO ADDITIONAL SERVICE INFO WAS PROVIDED. | -fail to boot up -fail to insert -device works intermittently | -failure to calibrate -disconnection -Disconnection | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM INTERMITTENTLY LOCKED UP. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ON SITE INVESTIGATION. THE REPORTED ISSUE COULD NOT BE DUPLICATED. THE SYSTEM WAS REBOOTED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -unable to lock screen -stop intermittently -fail to lock | -reset problem -Device Operates Differently than Expected -failure to deliver | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM LOCKED UP. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE PATIENT DATABASE SOFTWARE WAS EVALUATED AND REBUILT. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -unable to lock screen -lock to incorrect pressure -fail to lock | -Device Alarm System Issue -Inflation Issue -Device Operates Differently than Expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM LOCKED UP. NO PATIENT SERIOUS INJURY OR DEATH WAS REPORTED RELATED TO THIS EVENT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE SYSTEM BIOS (SOFTWARE) WAS EVALUATED AND RESET. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND RETURNED TO SERVICE. | -unable to lock screen -lock to incorrect pressure -fail to lock | -Device Alarm System Issue -Inflation Issue -Device Operates Differently than Expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT THE SYSTEM WOULD NOT RESPOND TO KEYBOARD OR MOUSE INPUTS AND THE SYSTEM WOULD NOT BOOT UP. THERE WAS NO PT INJURY OR DEATH REPORTED. Manufacturer Narrative: A GE SERVICE REP PERFORMED AN ONSITE INVESTIGATION. THE KEYBOARD WAS REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -fail to setup configuration -fail to support accessory | -Device Alarm System Issue -no flow -failure to deliver | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT WHEN USING Q-STATION, DATA FROM A PREVIOUS PATIENT APPEARS ON THE CURRENT PATIENT¨S DATA REPORT. THERE WAS NO REPORT OF INJURY TO PATIENT OR NEGATIVE IMPACT ON CLINICAL JUDGMENT. Manufacturer Narrative: THIS ISSUE IS CURRENTLY UNDER INVESTIGATION. FINAL RESULTS OF THE EVALUATION WILL BE SUBMITTED IN A FOLLOW UP REPORT. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS RESULTING FROM THIS ISSUE. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: IT WAS DETERMINED THAT THE ISSUE WAS A KNOWN SOFTWARE DEFECT. MANDATORY FIELD CHANGE ORDER, (B)(4), WAS DEVELOPED AND RELEASED TO CORRECT THIS ISSUE. THE FIX IS INCLUDED IN Q-STATION VERSION 3.3.2 AND HIGHER. THE AFFECTED CUSTOMERS WILL BE PROVIDED WITH THE NEWEST VERSION OF Q-STATION (3.3.2) ONCE THE NEW ORDERS HAVE BEEN CREATED AND SHIPPED. | -patient data breached -incorrect patient information -incorrect device orientation | -Connection Issue -data issue -application security issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT WHEN USING Q-STATION, DATA FROM A PREVIOUS PATIENT APPEARS ON THE CURRENT PATIENT¨S DATA REPORT. THERE WAS NO REPORT OF INJURY TO PATIENT OR NEGATIVE IMPACT ON CLINICAL JUDGMENT. Manufacturer Narrative: THIS ISSUE IS CURRENTLY UNDER INVESTIGATION. FINAL RESULTS OF THE EVALUATION WILL BE SUBMITTED IN A FOLLOW UP REPORT. THERE HAVE BEEN NO REPORTS OF ADVERSE EVENTS RESULTING FROM THIS ISSUE. ISSUE IDENTIFIED AS SOFTWARE RELATED; NO SYSTEM RETURN FOR EVALUATION IS EXPECTED. Manufacturer Narrative: IT WAS DETERMINED THAT THE ISSUE WAS A KNOWN SOFTWARE DEFECT. MANDATORY FIELD CHANGE ORDER, (B)(4), WAS DEVELOPED AND RELEASED TO CORRECT THIS ISSUE. THE FIX IS INCLUDED IN Q-STATION VERSION 3.3.2 AND HIGHER. THE AFFECTED CUSTOMERS WILL BE PROVIDED WITH THE NEWEST VERSION OF Q-STATION (3.3.2) ONCE THE NEW ORDERS HAVE BEEN CREATED AND SHIPPED. | -patient data breached -incorrect patient information -incorrect device orientation | -Connection Issue -data issue -application security issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THAT WHEN VIEWING STUDIES IN CHEV, THEY DID NOT RETAIN THE PRESENTATION STATE FROM THE SAME STUDY IN PACS. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. Manufacturer Narrative: THE CUSTOMER REPORTED THAT WHEN VIEWING STUDIES IN CHANGE HEALTHCARE ENTERPRISE VIEWER (CHEV), IT DID NOT RETAIN THE PRESENTATION STATE FOR THE SAME STUDY SAVED IN PACS. THERE WAS NO PATIENT HARM REPORTED AS A RESULT OF THIS ISSUE. THE INVESTIGATION CONCLUDED THAT WHILE THIS ISSUE WAS THE RESULT OF A MALFUNCTION, IT WAS NOT LIKELY TO RESULT IN A REPORTABLE ADVERSE EVENT OR SERIOUS INJURY. FURTHER INVESTIGATIONS DETERMINED THIS IS CAUSED BY A SOFTWARE DEFECT THAT WHEN IMAGING STUDIES ARE VIEWED IN CHEV VERSION 2.1, THE STUDIES DO NOT RETAIN THE IMAGE STYLES DEFINED BY PRESENTATION STATES CREATED AND SAVED IN CHANGE HEALTHCARE RADIOLOGY SOLUTIONS (CHRS). IMAGE STYLES INCLUDE THE FOLLOWING ATTRIBUTES OF IMAGE DISPLAY: ORIENTATION, WINDOW/LEVEL, AND ROI (ZOOM/PAN). CHANGE HEALTHCARE WILL BE NOTIFYING AFFECTED CUSTOMERS OF THIS ISSUE AND PROVIDING A SOFTWARE PATCH TO CORRECT THE ISSUE. | -loss of state indication -fail to retain position -fail to retain force | -No Flow -Device Handling Issue -insufficient information | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THE INSTATRAK 3500 COMPUTER TOWER WOULD NOT TURN ON AND FAILED TO BOOT UP. THERE IS NO REPORT OF DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE POWER SUPPLY WAS REPLACED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -fail to turn on device -fail to activate | -Device Alarm System Issue -Failure to Power-Up -device operates differently than expected | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THE MOUSE WAS NOT WORKING AND NO KEYBOARD WAS ATTACHED. WHEN THE MOUSE AND THE KEYBOARD FAILED, THE SYSTEM WILL NOT BOOT. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ONSITE INVESTIGATION. THE FAILURE COULD NOT BE DUPLICATED. AS A PRECAUTION THE MOUSE WAS REPLACED. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -fail to lock -fail to activate alarm | -device operates differently than expected -failure to deliver -Device Alarm System Issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THE PATIENT WAS DISCHARGED FROM INPATIENT HOSPITAL STAY FOLLOWING RADIOLOGIST REVIEW OF HISTORIC ABDOMINAL CT EXAM FROM (B)(6) 2008 THAT CLEARED PATIENT OF ANY PATHOLOGY. AN ABDOMINAL CT EXAM WAS ALSO PERFORMED ON PATIENT ON (B)(6) 2017. THE RADIOLOGIST REVIEWED THE (B)(6) 2008 EXAM AS THE CURRENT EXAM AND PROVIDED REPORT OF FINDINGS OF THIS IMAGE REVIEW ON THE (B)(6) 2017 REPORT. THE PATIENT THEN CAME BACK INTO ER (EMERGENCY ROOM) 2 DAYS LATER, ON (B)(6) 2017 WHERE SHE WAS DIAGNOSED WITH A LARGE ABSCESS WHICH REQUIRED IMMEDIATE SURGICAL INTERVENTION AND 2 DAYS HOSPITALIZATION. THE CUSTOMER REPORTS THAT THE PATIENT IS FULLY RECOVERED. THE CUSTOMER REPORTS THAT THE EXAMS DISPLAYED ON MONITORS IN A GE DEFAULT DDP (DEFAULT DISPLAY PROTOCOL) WHICH WAS NOT WHAT USER EXPECTED. Manufacturer Narrative: CORRECTED DATA - REPORT SOURCE IDENTIFIED AS HEALTHCARE PROFESSIONAL. CORRECTED DATA - DATE OF DEVICE MANUFACTURE CHANGED TO DEC. 2, 2014. ADDITIONAL MANUFACTURER NARRATIVE - EVENT PROBLEM EVALUATION CODES ADDED. INVESTIGATION COMPLETED. ROOT CAUSE WAS DETERMINED TO BE A USE ERROR. NO PRODUCT DEFECT IDENTIFIED. THE SITE INSTALLED A NEW THIRD PARTY HIS (HOSPITAL INFORMATION SYSTEM) AND WITH THIS SYSTEM ADDED MANY NEW EXAM PROCEDURES DESCRIPTIONS IN WHICH GE WAS NOT INVOLVED. THESE NEW EXAM PROCEDURES WERE NOT LINKED TO THE EXISTING DDP'S (DEFAULT DISPLAY PROTOCOLS) BEING USED BY THE SITE. WHEN A STUDY WAS VIEWED WITH ANY OF THESE NEW EXAM PROCEDURES, SINCE A LOCAL DDP WAS NOT DEFINED, THE GE DEFAULT DDP WAS UTILIZED TO DISPLAY THE STUDY AND ITS COMPARISONS. THE GE DEFAULT DISPLAYS THE CURRENT AND PRIOR STUDIES ON THE OPPOSITE SIDE OF THE MONITOR AS THE RADIOLOGIST WAS EXPECTING. THE RADIOLOGIST RECEIVED A POP-UP WARNING THAT THE GE DEFAULT DDP WAS BEING DISPLAYED, IN ADDITION THERE WAS A NOTICE ON THE WORKSTATION TOOLBAR THAT THE DEFAULT WAS IN USE, THE CORRECT STUDY DATE WAS DISPLAYED ON EACH OPEN STUDY AND THE COMPARISON STUDY HAD WHITE BORDER FRAMES AT THE CORNERS OF EACH IMAGE SHEET WHICH DESIGNATES THAT THE IMAGE(S) IS FROM A HISTORICAL STUDY AND NOT PART OF THE STUDY CURRENTLY BEING READ. THE SITE HAS ADDED THE NEW EXAM PROCEDURE DESCRIPTIONS TO THE RELEVANT DDP'S, HOWEVER, THEY REFUSED DDP ADVANCED TRAINING OFFERED BY GE. Manufacturer Narrative: PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS. THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | -patient hospitalized -patient data exposured -inappropriate patient type | -Detachment of Device Component -loose or intermittent connection -detachment of device component | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED THE PATIENT WAS HOSPITALIZED FOR A SMALL BOWEL OBSTRUCTION AND PAIN. A CT WAS ORDERED FOR THE PATIENT. THE CT SCAN WAS PERFORMED ON (B)(6) 2016. THIS STUDY WAS VIEWED BY RADIOLOGIST ON (B)(6) 2016 WHEN AN EXAM NOTE WAS PERFORMED BY THE RADIOLOGIST THAT INDICATED: SBO WITH FREE AIR AND FLUID. THE WARD DOCTOR WAS NOTIFIED OF THESE FINDINGS ON (B)(6) 2016. THE WARD DOCTOR DETERMINED THAT THE PATIENT WAS TOO FRAIL FOR SURGERY, SO NO SURGERY WAS PERFORMED. THE PATIENT DIED ON (B)(6) 2016. THE FINAL REPORT (SIGNED (B)(6) 2016) INDICATED THAT THERE WAS A DELAY IN READING BECAUSE THE STUDY WAS INCORRECTLY DATE AND TIME STAMPED Manufacturer Narrative: THE INVESTIGATION REVEALED THAT THE MERGING OF TWO PATIENT STUDIES WAS PERFORMED CORRECTLY AND THAT THE STUDY HAD THE APPROPRIATE TIME AND DATE STAMP AFTER THE MERGE. THE INVESTIGATION ALSO REVEALED THAT THE EXAM WAS PRESENT ON THE WORKLIST AND AVAILABLE FOR USE BY THE RADIOLOGIST. THE INVESTIGATION DID NOT IDENTIFY ANY DEFICIENCY WITH THE CPACS PRODUCT. THE CPACS PRODUCT IS WORKING AS DESIGNED. Manufacturer Narrative: PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS. THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. MANUFACTURE DATE IS UNKNOWN. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | -patient hospitalized -patient experienced adverse event -patient infected | -fluid leak -Detachment of Device Component -Contamination During Use | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THAT THE DEMOGRAPHICS MANAGER (VERSION 12.0.1-BUILD-03) CAUSED REPORT DATA MISATTRIBUTION. SEVERAL RESULT PDFS DID NOT CROSS TO THE CUSTOMER'S HIS, MEDITECH. ROOT CAUSE AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV. HOWEVER, THE PATIENT DEMOGRAPHIC DATA MAY CHANGE AFTER THE REPORT IS SIGNED AND THIS NEW INFORMATION SHOULD BE REFLECTED IN THOSE PREVIOUSLY-SIGNED SYSTEM REPORTS. THE DEMOGRAPHICS MANAGER UPDATES PREVIOUSLY SIGNED REPORTS (XML AND PDF FILES) WITH THE REVISED PATIENT DEMOGRAPHICS INFORMATION AND RESENDS THE INFORMATION TO THE PROPER EXTERNAL SYSTEMS. THE DEMOGRAPHICS MANAGER IS ABLE TO REVISE THE REPORTS WITHOUT HAVING PHYSICIANS RE-SIGN THEIR PREVIOUSLY SIGNED REPORTS. THE DEMOGRAPHICS MANAGER SERVICE WILL ALWAYS REGENERATE THE PDFS OF ALL SIGNED REPORTS FOR A PATIENT AND/OR STUDY. THE ARCHIVE VERSIONS OF THE REPORT ARE THE TWO COPIES OF XML FILES LOCATED ON THE REPORT STORAGE SERVER AND THE BACKUP SERVER. A PRODUCT DEFECT IN THE DEMOGRAPHICS MANAGER CODE MAY CAUSE SOME REPORTING OBJECTS TO BE POPULATED WITH ANOTHER REPORT'S DATA CAUSING MISATTRIBUTION. UNDER NORMAL CONFIGURATION/OPERATION, THE DM SERVICE FINISHES THE DEMOGRAPHIC UPDATES WITHIN A 120 SECOND WINDOW. IF HOWEVER A LARGE NUMBER OF JOBS ARE QUEUED UP AND THEY DO NOT FINISH PROCESSING WITHIN THE 120 SECOND WINDOW, MISATTRIBUTION MAY OCCUR. ON (B)(4) 2015, AGFA HAS IMPLEMENTED A CORRECTION VIA AGFA ID - (B)(4). MANDATORY SERVICE BULLETINS HAVE BEEN RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR DEMOGRAPHICS MANAGER (DM) FOR SITES WITH IMPAX CV 7.8.X AND CV12 12.2.SU1 AND SU2 INSTALLED. THERE HAVE BEEN NO REPORTS OF PHYSICAL HARM OR HEALTH DAMAGE RESULTING FROM THIS ISSUE. THIS MEDICAL DEVICE REPORT INCLUDES 348 REPORTS AFFECTED ON (B)(4) 2015, WHEN THE ERROR OCCURRED. (B)(6). Manufacturer Narrative: (B)(4). | -fail to record intraoperative data -inappropriate user -incorrect user manual | -detachment of device component -no flow -Connection Issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THAT THE DEMOGRAPHICS MANAGER (VERSION 12.0.1-BUILD-03) CAUSED REPORT DATA MISATTRIBUTION. SEVERAL RESULT PDFS DID NOT CROSS TO THE CUSTOMER'S HIS, MEDITECH. ROOT CAUSE AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV. HOWEVER, THE PATIENT DEMOGRAPHIC DATA MAY CHANGE AFTER THE REPORT IS SIGNED AND THIS NEW INFORMATION SHOULD BE REFLECTED IN THOSE PREVIOUSLY-SIGNED SYSTEM REPORTS. THE DEMOGRAPHICS MANAGER UPDATES PREVIOUSLY SIGNED REPORTS (XML AND PDF FILES) WITH THE REVISED PATIENT DEMOGRAPHICS INFORMATION AND RESENDS THE INFORMATION TO THE PROPER EXTERNAL SYSTEMS. THE DEMOGRAPHICS MANAGER IS ABLE TO REVISE THE REPORTS WITHOUT HAVING PHYSICIANS RE-SIGN THEIR PREVIOUSLY SIGNED REPORTS. THE DEMOGRAPHICS MANAGER SERVICE WILL ALWAYS REGENERATE THE PDFS OF ALL SIGNED REPORTS FOR A PATIENT AND/OR STUDY. THE ARCHIVE VERSIONS OF THE REPORT ARE THE TWO COPIES OF XML FILES LOCATED ON THE REPORT STORAGE SERVER AND THE BACKUP SERVER. A PRODUCT DEFECT IN THE DEMOGRAPHICS MANAGER CODE MAY CAUSE SOME REPORTING OBJECTS TO BE POPULATED WITH ANOTHER REPORT'S DATA CAUSING MISATTRIBUTION. UNDER NORMAL CONFIGURATION/OPERATION, THE DM SERVICE FINISHES THE DEMOGRAPHIC UPDATES WITHIN A 120 SECOND WINDOW. IF HOWEVER A LARGE NUMBER OF JOBS ARE QUEUED UP AND THEY DO NOT FINISH PROCESSING WITHIN THE 120 SECOND WINDOW, MISATTRIBUTION MAY OCCUR. ON (B)(6) 2015, AGFA HAS IMPLEMENTED A CORRECTION VIA AGFA ID (B)(4). MANDATORY SERVICE BULLETINS HAVE BEEN RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR DEMOGRAPHICS MANAGER (DM) FOR SITES WITH IMPAX CV 7.8.X AND CV12 12.2.SU1 AND SU2 INSTALLED. THERE HAVE BEEN NO REPORTS OF PHYSICAL HARM OR HEALTH DAMAGE RESULTING FROM THIS ISSUE. THIS MEDICAL DEVICE REPORT INCLUDES 1 REPORT AFFECTED ON (B)(6) 2015, WHEN THE ERROR OCCURRED. REPORT ID: (B)(4), DATE ERROR OCCURRED: 08/17/15.Event Description: WE ARE REPORTING THIS SUPPLEMENTAL REPORT ON OCTOBER 11, 2016. AFTER AN FDA INSPECTION, CONDUCTED AT ONE OF OUR AGFA SITES, AND THE INSPECTION CLOSED ON (B)(6) 2016, WE REALIZED THE DATE OF EVENT FOR THIS MDR 1225058-2015-14123 WAS INCORRECT. THIS SUPPLEMENT REPORT #1 IS BEING SUBMITTED TO PROVIDE THE CORRECTION FOR THE DATE OF THE EVENT. Manufacturer Narrative: DATE OF EVENT WAS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. | -fail to record intraoperative data -inappropriate user -incorrect user manual | -detachment of device component -no flow -Connection Issue | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THAT THEY DISCOVERED ON (B)(6) 2015, CARDIO ECHO RESULTS FROM THEIR AGFA SYSTEM PROCESSED UNSUCCESSFULLY INTO THEIR EPIC SYSTEM. REPORTS AT THE TIME OF PHYSICIAN SIGNATURE WERE CORRECT, BUT THE IMPAX CV 12.1 DEMOGRAPHICS MANAGER DID NOT PRODUCE THE CORRECT PDF REPORTS. AGFA CARDIO TECHNICAL SUPPORT CONFIGURED AN AUTO CORRECTION ON AGFA REPORT AND THE SYSTEM PERFORMED BETTER. AT THIS TIME, AGFA'S INITIAL RISK MANAGEMENT PROCESS AND EVALUATION OF SEVERITY/FREQUENCY RATING DID NOT RESULT IN CFR PART 803 REPORTING. AGFA TECHNICAL SUPPORT PROCEEDED AND CHANGED THE STORED PROCEDURE THAT DEMOGRAPHICS MANAGER USES AND RESTARTED DEMOGRAPHICS MANAGER. ALL PDF REPORTS OUTLINED BY THE CUSTOMER WERE REVIEWED AND CORRECTED. ON (B)(6) 2015, AFTER MONITORING THE SYSTEM, FURTHER REVIEW AND REPROCESSING OF REPORTS, AGFA TECHNICAL SUPPORT DETERMINED SOME REPORTS STILL CONTAINED INACCURATE DATA DUE TO HOW DEMOGRAPHICS MANAGER PROCESSES UPDATES. AGFA TECHNICAL SUPPORT EXECUTED A CLEANUP OF PROBLEM ACCESSION NUMBERS SO THE DEMOGRAPHICS MANAGER WOULD PROCESS UPDATES AND START THE OUTBOUND PROCESS TO RECREATE PDF REPORTS. ON (B)(6) 2015, A SECOND COMPLAINT WAS RECEIVED FROM A DIFFERENT CUSTOMER, RESULTING IN AVAILABILITY OF FURTHER INFORMATION USED FOR RE-ASSESSMENT OF RISK IN RELATION TO THE EVENT DESCRIBED IN THIS MDR REPORT. THE RISK RE-ASSESSMENT RESULTED IN A RATING OF VIGILANCE REPORTING, AND AGFA IS USING THE (B)(6) 2015 DATE AS THE DATE OF AWARENESS FOR THE EVENT THAT OCCURRED ON (B)(6) 2015. THE ROOT CAUSE HAS BEEN IDENTIFIED AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. DEMOGRAPHICS MANAGER MUST BE ENABLED FOR THIS SITUATION TO OCCUR. A POLLING TIMER WITHIN (DM) CALLS A FUNCTION. THIS FUNCTION PROCESSES THE DM WORK. THE FUNCTION DOES A SELF-CHECK TO SEE IF IT IS ALREADY BUSY BEFORE IT STARTS PROCESSING THE NEXT SET OF WORK. IF IT IS ALREADY BUSY, IT SKIPS THE WORK BUT SETS ITSELF TO "NOT BUSY" WHICH IS WRONG BECAUSE IT IS IN FACT "BUSY". THIS DEFECT THEN IS THAT THE FUNCTION SETS ITSELF TO "NOT BUSY" WHEN IN FACT IT IS "BUSY". THE CONSEQUENCE OF THIS IS THAT THE NEXT TIME THE POLLING TIMER CALLS THE FUNCTION, THE FUNCTION SELF-CHECK COMES BACK AS "NOT BUSY" (WHICH IS WRONG) AND IF IT IS IN THE MIDDLE OF PROCESSING A REPORT, THAT WORK WILL BE STOPPED, THE NEW WORK WILL BE STARTED, AND THE PROCESSING MAY HAVE SOME OBJECTS POPULATED WITH ANOTHER REPORT'S DATA. AGFA SOLUTION DEVELOPMENT IS UNDERWAY AND A CORRECTION WILL BE IMPLEMENTED FOR THIS ISSUE. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS EVENT.Event Description: ROOT CAUSE AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV. HOWEVER, THE PATIENT DEMOGRAPHIC DATA MAY CHANGE AFTER THE REPORT IS SIGNED AND THIS NEW INFORMATION SHOULD BE REFLECTED IN THOSE PREVIOUSLY-SIGNED SYSTEM REPORTS. THE DEMOGRAPHICS MANAGER UPDATES PREVIOUSLY SIGNED REPORTS (XML AND PDF FILES) WITH THE REVISED PATIENT DEMOGRAPHICS INFORMATION AND RESENDS THE INFORMATION TO THE PROPER EXTERNAL SYSTEMS. THE DEMOGRAPHICS MANAGER IS ABLE TO REVISE THE REPORTS WITHOUT HAVING PHYSICIANS RE-SIGN THEIR PREVIOUSLY SIGNED REPORTS. THE DEMOGRAPHICS MANAGER SERVICE WILL ALWAYS REGENERATE THE PDFS OF ALL SIGNED REPORTS FOR A PATIENT AND/OR STUDY. THE ARCHIVE VERSIONS OF THE REPORT ARE THE TWO COPIES OF XML FILES LOCATED ON THE REPORT STORAGE SERVER AND THE BACKUP SERVER. A PRODUCT DEFECT IN THE DEMOGRAPHICS MANAGER CODE MAY CAUSE SOME REPORTING OBJECTS TO BE POPULATED WITH ANOTHER REPORT'S DATA CAUSING MISATTRIBUTION. UNDER NORMAL CONFIGURATION/OPERATION, THE DM SERVICE FINISHES THE DEMOGRAPHIC UPDATES WITHIN A 120 SECOND WINDOW. IF HOWEVER A LARGE NUMBER OF JOBS ARE QUEUED UP AND THEY DO NOT FINISH PROCESSING WITHIN THE 120 SECOND WINDOW, MISATTRIBUTION MAY OCCUR. ON SEPTEMBER 1, 2015, AGFA IMPLEMENTED A CORRECTION VIA AGFA ID - RRC 15-003 (1225058-09-01-2015-001-C) MANDATORY SERVICE BULLETINS HAVE BEEN RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR DEMOGRAPHICS MANAGER (DM) FOR SITES WITH IMPAX CV 7.8.X AND CV12 12.2.SU1 AND SU2 INSTALLED.Event Description: WE ARE REPORTING THIS SUPPLEMENTAL REPORT ON OCTOBER 11, 2016. AFTER AN FDA INSPECTION, CONDUCTED AT ONE OF OUR AGFA SITES, AND THE INSPECTION CLOSED ON SEPTEMBER 29, 2016, WE REALIZED THE DATE OF EVENT AND DATE RECEIVED BY MANUFACTURER FOR THIS MDR 1225058-2015-14119 WERE INCORRECT. THIS SUPPLEMENT REPORT #2 IS BEING SUBMITTED TO PROVIDE THE CORRECTIONS FOR THE DATE OF THE EVENT AND THE DATE RECEIVED BY MANUFACTURER. Manufacturer Narrative: . Manufacturer Narrative: DATE OF EVENT WAS CORRECTED FROM (B)(6) 2015. DATE RECEIVED BY MANUFACTURER WAS CORRECTED FROM 03/09/2015 TO 04/20/2015. | -incorrect output -fail to generate sound -error message | -break -Detachment of Device Component -detachment of device component | Viewer
| FDA | Event Description: THE CUSTOMER REPORTED TO AGFA ON (B)(6) 2015 THEY WERE EXPERIENCING AN ISSUE WITH INCORRECT DEMOGRAPHICS WITH CV REPORTS CROSSING INTO THEIR EPIC SYSTEM. DEMOGRAPHICS MANAGER HAD BEEN INSTALLED (B)(6) 2015 AND THE SITE DISCOVERED THE NAME ON A CV REPORT HAD BEEN CHANGED ON THE .PDF REPORT. THE HIS_DATA FOR THE PATIENT WAS CORRECT, BUT THE PATIENT NAME AT THE TOP OF THE DOCUMENT WAS ANOTHER CV PATIENT. INVESTIGATION REVEALED YOU CAN GO INTO THE PATIENT IN RESULTS MANAGEMENT AND SELECT SHOW REPORT AND ALL OF THE INFORMATION ON THE REPORT IS CORRECT INCLUDING THE PATIENT NAME. A FULL REPROCESS OF POTENTIALLY AFFECTED REPORTS WAS PERFORMED BY AGFA TECHNICAL SUPPORT. THE ROOT CAUSE HAS BEEN IDENTIFIED AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. DEMOGRAPHICS MANAGER MUST BE ENABLED FOR THIS SITUATION TO OCCUR. A POLLING TIMER WITHIN (DM) CALLS A FUNCTION. THIS FUNCTION PROCESSES THE DM WORK. THE FUNCTION DOES A SELF-CHECK TO SEE IF IT IS ALREADY BUSY BEFORE IT STARTS PROCESSING THE NEXT SET OF WORK. IF IT IS ALREADY BUSY, IT SKIPS THE WORK BUT SETS ITSELF TO "NOT BUSY" WHICH IS WRONG BECAUSE IT IS IN FACT "BUSY". THIS DEFECT THEN IS THAT THE FUNCTION SETS ITSELF TO "NOT BUSY" WHEN IN FACT IT IS "BUSY". THE CONSEQUENCE OF THIS IS THAT THE NEXT TIME THE POLLING TIMER CALLS THE FUNCTION, THE FUNCTION SELF-CHECK COMES BACK AS "NOT BUSY" (WHICH IS WRONG) AND IF IT IS IN THE MIDDLE OF PROCESSING A REPORT, THAT WORK WILL BE STOPPED, THE NEW WORK WILL BE STARTED, AND THE PROCESSING MAY HAVE SOME OBJECTS POPULATED WITH ANOTHER REPORT'S DATA. AGFA SOLUTION DEVELOPMENT IS UNDERWAY AND A CORRECTION WILL BE IMPLEMENTED FOR THIS ISSUE. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS EVENT.Event Description: ROOT CAUSE AS A PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV. HOWEVER, THE PATIENT DEMOGRAPHIC DATA MAY CHANGE AFTER THE REPORT IS SIGNED AND THIS NEW INFORMATION SHOULD BE REFLECTED IN THOSE PREVIOUSLY-SIGNED SYSTEM REPORTS. THE DEMOGRAPHICS MANAGER UPDATES PREVIOUSLY SIGNED REPORTS (XML AND PDF FILES) WITH THE REVISED PATIENT DEMOGRAPHICS INFORMATION AND RESENDS THE INFORMATION TO THE PROPER EXTERNAL SYSTEMS. THE DEMOGRAPHICS MANAGER IS ABLE TO REVISE THE REPORTS WITHOUT HAVING PHYSICIANS RE-SIGN THEIR PREVIOUSLY SIGNED REPORTS. THE DEMOGRAPHICS MANAGER SERVICE WILL ALWAYS REGENERATE THE PDFS OF ALL SIGNED REPORTS FOR A PATIENT AND/OR STUDY. THE ARCHIVE VERSIONS OF THE REPORT ARE THE TWO COPIES OF XML FILES LOCATED ON THE REPORT STORAGE SERVER AND THE BACKUP SERVER. A PRODUCT DEFECT IN THE DEMOGRAPHICS MANAGER CODE MAY CAUSE SOME REPORTING OBJECTS TO BE POPULATED WITH ANOTHER REPORT'S DATA CAUSING MISATTRIBUTION. UNDER NORMAL CONFIGURATION/OPERATION, THE DM SERVICE FINISHES THE DEMOGRAPHIC UPDATES WITHIN A 120 SECOND WINDOW. IF HOWEVER A LARGE NUMBER OF JOBS ARE QUEUED UP AND THEY DO NOT FINISH PROCESSING WITHIN THE 120 SECOND WINDOW, MISATTRIBUTION MAY OCCUR. ON SEPTEMBER 1, 2015, AGFA IMPLEMENTED A CORRECTION VIA AGFA ID - RRC 15-003 (1225058-09-01-2015-001-C) MANDATORY SERVICE BULLETINS HAVE BEEN RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR DEMOGRAPHICS MANAGER (DM) FOR SITES WITH (B)(4) AND (B)(4) INSTALLED. | -incorrect user manual -incorrect process -incorrect message display | -break -Detachment of Device Component -Protective Measure Issue | Viewer
| FDA | Event Description: THE FE REPORTED THE SYSTEM FAILED TO BOOT UP. THERE IS NO REPORT OF INJURY OR DEATH ASSOCIATED WITH THE EVENT DESCRIBED IN THIS COMPLAINT. Manufacturer Narrative: A GE SERVICE REPRESENTATIVE PERFORMED AN ON SITE INVESTIGATION. THE CONNECTORS WERE RESEATED DURING THE SERVICE CALL. THE SYSTEM WAS TESTED AND FOUND TO BE WORKING AS INTENDED AND PUT BACK INTO SERVICE. | -fail to boot up -system failure -fail to activate alarm | -reset problem -inadequate instructions for healthcare professional -failure to calibrate | Viewer
| FDA | Event Description: THE HOSPITAL SITE REPORTED THAT THE CARESTREAM VUE PACS ALLEGEDLY CONTRIBUTED TO A PATIENT DEATH DUE TO THE SYSTEM BEING DOWN AND PATIENT IMAGES NOT ACCESSIBLE. Manufacturer Narrative: CARESTREAM RECEIVED A COMPLAINT ON (B)(6) 2020 AND HAS STARTED THE INVESTIGATION PROCESS. WHEN ADDITIONAL INFORMATION AND FACTS BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED WITHIN 30 DAYS. | -unable to access -loss of patient data -fail to access | -device operates differently than expected -insufficient flow or underinfusion -device stops intermittently | Viewer
| FDA | Event Description: THE ISSUE OCCURRED DURING A SIEMENS-HEALTHCARE INSTALLATION UPGRADE OF THE SITE'S (B)(6) A (MLA) SYSTEM FROM SW VERSION VA20 TO VA30. AN ERROR OCCURRED DURING THE INSTALLATION IN THE CONFIGURATION BETWEEN THE MLA AND SYNGO DYNAMICS (SD) PRODUCTS. IT RESULTED IN A CROSSED PATIENT DATA SITUATION FOR 350 PATIENTS. THE PATIENT DEMOGRAPHIC DATA DISPLAYED FOR THE STUDIES IN THE SYNGO DYNAMICS PRODUCT WAS AFFECTED BUT THERE WAS NO CHANGE IN THE IMAGING MODALITIES STUDY PATIENT DEMOGRAPHICS DATA. THE ERROR WAS IDENTIFIED AND THE CONFIGURATION WAS CORRECTED. THERE IS NO REASONABLY FORESEEABLE CLINICAL IMPACT TO THE CROSSED DATA AS THE SITE IS AWARE TO CROSS-CHECK THE PATIENT DEMOGRAPHIC INFORMATION WITH OTHER AVAILABLE SOURCES WHILE THE MANUAL CORRECTIONS ARE TAKING PLACE. THE IMAGING MODALITY STUDY AND PATIENT DEMOGRAPHIC DATA IS NOT AFFECTED. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. SIEMENS SERVICE AND INSTALLATION PERSONNEL ARE WORKING WITH THE SITE PERSONNEL TO MANUALLY CORRECT DEMOGRAPHIC DATA OF THE 350 PATIENT STUDIES. THE EFFORT IS ONGOING AND THE SITE IS CROSSCHECKING THE DISPLAYED PATIENT DEMOGRAPHIC STUDY INFORMATION WITH OTHER SOURCES DURING THE MANUAL CORRECTIONS. THERE ARE NO REPORTED INSTANCES OF HARM OCCURRING FROM THIS ISSUE. PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH WOULD RESULT IN THE NEED FOR A PATIENT RESCAN. THE SYNGO DYNAMICS (OR THE MLA) PRODUCTS HAVE NOT CAUSED OR CONTRIBUTED TO THE CROSSED DATA. NO REASONABLY FORESEEABLE PATH TO OCCURRENCE OF HARM CAN BE DETERMINED. | -patient experienced adverse event -incorrect device orientation -device malfunction | -Device Handling Issue -Human-Device Interface Issue -fluid leak | Viewer
| FDA | Event Description: THE ISSUE OCCURRED DURING A SIEMENS-HEALTHCARE PROFESSIONAL SERVICES MIGRATION OF TWO HOSPITAL'S ARCHIVE SYSTEMS INTO ONE (THE NEW PARENT ORGANIZATION SYSTEM). UNRECOVERABLE LOSS OF 13 IMAGE STUDIES DURING THE MIGRATION OCCURRED. THE ISSUE APPEARS TO BE RELATED TO A MISUNDERSTANDING OF THE STATE OF REPORTS BETWEEN THE "CHILD" SITE'S SOURCE SYSTEM AND THE "PARENT" SITE'S. PRODUCTION DESTINATION SYSTEM. FOR THIS MIGRATION, THE EXPECTATION WAS THAT THE "CHILD" STUDIES AND REPORTS WOULD BE INSERTED INTO THE "PARENT'S" PRODUCTION SYSTEM. THE WORKFLOW DISCREPANCY HAS BEEN IDENTIFIED BY SIEMENS. THE CLINICAL RELEVANCE OF THE LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES ARE STILL AVAILABLE IN THE SITES HIS. THERE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. NO REASONABLY FORESEEABLE PATH TO OCCURRENCE OF HARM CAN BE DETERMINED. | -patient experienced adverse event -unable to record event -fail to transport patient | -leak -disconnection -insufficient flow or underinfusion | Viewer
| FDA | Event Description: THE ISSUE WAS DISCOVERED DURING A SIMPLE REQUEST FOR SERVICE SUPPORT IN DETERMINING WHY ONE PATIENT'S INFORMATION WAS NOT ABLE TO BE CALLED UP FROM THE ARCHIVE BY THE SYNGO DYNAMICS PRODUCT. THE ISSUE WAS ESCALATED BY THE CUSTOMER IN (B)(6) 2017. IT WAS FOUND THAT 155 DIRECTORIES (5170 OBJECTS) WERE MISSING FROM THE SITE'S LONG TERM ARCHIVE SYSTEM. THERE WAS NO INDICATION OR INFORMATION FROM THE SITE OF ANY POLICY OR EVENT WHICH WOULD HAVE CAUSED THIS LOSS. SUBSEQUENT SERVICE INVESTIGATION FOUND A CLOSED SERVICE TICKET DOCUMENTING A HARDWARE FAILURE OF THE SITE'S LONG TERM ARCHIVE SYSTEM (NAS), WHICH WAS ADDRESSED BY A DATA RECOVERY COMPANY IN JANUARY 2017. THE SITE'S NAS IS NOT A SIEMEN'S PRODUCT OR RESPONSIBILITY. IN APRIL 2017, THE SITE REPORTED TO SIEMENS THAT THE SITE'S NAS WAS RECOVERED AND ARCHIVING AND RETRIEVING WAS POSSIBLE AGAIN. THOUGH THE ISSUE IS NOT RELATED TO THE SYNGO DYNAMICS PRODUCT, BASED UPON ALL AVAILABLE INFORMATION, IT IS ASSUMED THAT THE DATA WAS PREVIOUSLY LOST DUE TO A HARDWARE FAILURE OF THE SITE'S LONG TERM ARCHIVE SYSTEM AND THIS LOSS WAS NOT IDENTIFIED DURING THE SITE'S DATA RECOVERY EFFORT, WHICH DID NOT INVOLVE SIEMENS. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. THE REPORTED EVENT OCCURRED IN (B)(6). Manufacturer Narrative: RESUBMISSION OF INITIAL REPORT DUE TO REPORT CODE ERROR. THE CLINICAL RELEVANCE OF THE POTENTIALLY LOST DATA IS THAT THE USER NO LONGER HAS THE HISTORICAL STUDY IMAGES AVAILABLE IF A COMPARISON WITH NEWER IMAGES IS DESIRED, HOWEVER THE CLINICAL STUDIES AND REPORTS CREATED FROM THE READ OF THE IMAGE STUDIES WILL STILL BE AVAILABLE IN THE SYNGO DYNAMICS SYSTEM UNLESS DELETED FROM THE SYSTEM BY SITE IT POLICY OR USER ACTIONS. THE SYNGO DYNAMICS PRODUCT HAS NOT CAUSED OR CONTRIBUTED TO THE LOSS OF DATA. (B)(6). | -fail to enter patient data -missing information -fail to generate indication | -component missing -fluid leak -failure to read input signal | Viewer
| FDA | Event Description: THE MERGE RADSUITE APPLICATION PROVIDES A MEANS TO DISTRIBUTE, DISPLAY, AND STORE DIAGNOSTIC-QUALITY MEDICAL IMAGES IN ELECTRONIC FORMAT. THE SYSTEM DISPLAYS TRADITIONAL 2D AND RECONSTRUCTED 3D RADIOLOGICAL IMAGES USING WEB-ENABLED VIEWERS OVER BOTH LOCAL AND WIDE AREA NETWORKS. A CUSTOMER REPORTED TO MERGE HEALTHCARE SUPPORT ON (B)(6) 2017 THAT THERE WERE "BAD IMAGES PINK/GREN/YELLOW/RED" FOR AN ULTRASOUND IMAGE THAT WAS BEING VIEWED FROM THE ARCHIVE. THE CUSTOMER SITE HAS NOT REPORTED ANY PATIENT SAFETY CONCERNS OR DELAYS, HOWEVER IF THE IMAGES ARE BLURRY OR ARE NOT CONSIDERED TO BE DIAGNOSTIC QUALITY, THERE IS A POTENTIAL FOR PATIENT HARM, DELAY IN TREATMENT OR DIAGNOSIS. (B)(4). Manufacturer Narrative: INITIAL INVESTIGATION FINDINGS DETERMINED THAT THE SITE WAS UTILIZING NEW STORAGE AND OPTIMAL IMAGE/DATA COMPRESSION RULES WERE NOT BEING APPLIED WHEN THE STUDY WAS TRANSFERRED TO THE ARCHIVE. THE STORAGE RULES HAVE BEEN RESOLVED. MERGE HEALTHCARE IS CONTINUING TO INVESTIGATE THE CUSTOMER'S ALLEGATION TO DETERMINE IF ANY ADDITIONAL CORRECTIONS OR CORRECTIVE ACTIONS ARE REQUIRED. THIS IS NOT HARDWARE. | -fail to store medical imaging -fail to process medical imaging -inadequate image types | -use of device issue -data issue -patient data issue | Viewer
| FDA | Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT SEVERAL 6 MONTH AND 1 YEAR FOLLOWUPS WERE NOT SHOWING UP IN THE "NOTIFY" DIALOG. NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | -fail to process medical imaging -fail to interpret images -fail to upload images | -data back-up problem -filling problem -use of device issue | Viewer
| FDA | Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER SITE REPORTED THAT THEY COULD NOT GET A REPORT OUT FOR AN URGENTLY ILL PATIENT UNTIL THE IMAGING CAPTURE SITE'S IDS WERE ADDED. IT TOOK APPROXIMATELY 9 HOURS TO RESOLVE THE PROBLEM WITH THE REPORT. NO REPORT OF INJURY WAS RECEIVED. Manufacturer Narrative: AS PART OF THE NORMAL PROCESS, WHEN THE CUSTOMER ADDS A NEW SECONDARY CAPTURE SITE ID FOR A NEW SITE THEY NEED TO REQUEST IN ADVANCE THAT A TOPLOGO BE CREATED SO THAT THEY CAN GET REPORTS OUT. IF NOT, AS LONG AS THE SITE IS ABLE TO READ/DICTATE CASES THEY SHOULD BE ABLE TO GET A REPORT ON THE CASE ALTHOUGH IT MAY NOT HAVE THE CORRECT ORIGINATING SITE LISTED UNTIL THE CAPTURE SITE ID IS ENTERED - THIS COULD BE EXPLAINED IN THE BODY OF THE REPORT. THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. PROBLEM NOT CAUSED BY A DEVICE PROBLEM. | -fail to process medical imaging -fail to interpret images -fail to upload images | -data back-up problem -filling problem -use of device issue | Viewer
| FDA | Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT A PATIENT ID WAS ORIGINALLY FOR ONE PATIENT, BUT DUE TO A DIGITIZED MAMMO FROM OUTSIDE FACILITY COMING ACROSS WITH SAME PATIENT ID BUT FOR A DIFFERENT PATIENT, IT CAUSED A DISCREPANCY WITH PRIORS COMING UP FOR BOTH PATIENTS AND PATIENTS BEING LINKED. NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | -fail to process medical imaging -fail to interpret images -fail to upload images | -data back-up problem -filling problem -use of device issue | Viewer
| FDA | Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT USING THE "UPDATE DICOM HEADER" OPTION IN UNIVERSAL MANAGER RESULTS IN ALL IMAGES BEING UNVIEWABLE. PANELS ALL DISPLAY "INVALID IMAGE SIGNATURE". NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | -fail to process medical imaging -fail to interpret images -fail to upload images | -data back-up problem -filling problem -use of device issue | Viewer
| FDA | Event Description: THE MERGE UNITY PACS IS A MEDICAL IMAGE AND INFORMATION MANAGEMENT SYSTEM THAT ALLOWS VIEWING, SELECTION, PROCESSING, PRINTING, TELECOMMUNICATIONS, AND MEDIA INTERCHANGE OF MEDICAL IMAGES FROM A VARIETY OF DIAGNOSTIC IMAGING SYSTEMS. MERGE UNITY PACS INTERFACES TO VARIOUS STORAGE AND PRINTING DEVICES USING DICOM OR SIMILAR INTERFACE STANDARDS. A CUSTOMER OBSERVED THAT IMAGES FROM EXAM A WERE SHOWING UP IN MONTAGES FOR EXAM B. NO REPORTS OF INJURIES WERE RECEIVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED AS THE RESULT OF A RETROSPECTIVE REVIEW. | -fail to process medical imaging -fail to interpret images -fail to upload images | -data back-up problem -filling problem -use of device issue | Viewer
| FDA | Event Description: THE MONITOR IMAGE BECAME NOISY WHEN THE USER FACILITY CHECKED THE IMAGE FROM THE ENDOEYE FLEX DEFLECTABLE VIDEOSCOPE LTF-S190-10 IN PREPARATION FOR THE PROCEDURE AFTER THE PATIENT WAS ANESTHETIZED. THE USER FACILITY TURNED OFF AND ON THE SUBJECT DEVICE, AND DISCONNECTED AND CONNECTED THE VIDEO CONNECTOR TO THE SUBJECT DEVICE. HOWEVER, THE PHENOMENON STILL OCCURRED. THE FACILITY COMPLETED THE PROCEDURE WITH REPLACING THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE AND CONCOMITANT THE LTF-S190-10 WERE RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE SUBJECT DEVICE WORKED PROPERLY. ALTHOUGH THE SUBJECT DEVICE CONNECTED TO THE LTF-S190-10, THE REPORTED PHENOMENON COULD NOT BE REPRODUCED. THE SUBJECT DEVICE WORKED PROPERLY AFTER THE USER FACILITY REPLACED THE LTF-S190-10 TO ANOTHER VIDEOSCPE. SO THERE WAS A POSSIBILITY THAT CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE LTF-S190-10 MIGHT OCCUR TEMPORARILY DUE TO ADHESION OF FOREIGN SUBSTANCES ON THE ELECTRICAL CONTACT OF THE VIDEO CONNECTER IN THE LTF-S190-10. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF OTV-S190 AND LTF-S190-10 ALREADY CAUTIONS FOR VIDEO CONNECTOR HANDLING. THERE WERE NO FURTHER DETAILS PROVIDED AT THIS TIME. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00232. | -noisy valve -loss of image -inappropriate image format | -human-device interface issue -no display or display failure -medical gas supply problem | Viewer
| FDA | Event Description: THE MONITOR IMAGE WAS DISTURBED DURING THE UNKNOWN PROCEDURE. THE FACILITY COMPLETED THE PROCEDURE TO REPLACE THE SUBJECT OTV-S7PRO TO ANOTHER DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. THE FACILITY WAS AWARE THAT THE PHENOMENON WAS CAUSED BY THEIR HANDLING. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS AND THERE WAS CORROSION ON CONTACT OF THE VIDEO CONNECTOR SOCKET. THE CONTACT WAS CORRODED BECAUSE THE FACILITY MIGHT CONNECT A VIDEO CONNECTOR WITH WET CONDITION OF A CAMERA HEAD TO THE SUBJECT DEVICE. SO THE MONITOR IMAGE WAS FINALLY DISTURBED DUE TO CONTACT FAILURE CAUSED FROM THE CORROSION. OLYMPUS ALSO CHECKED THE MANUFACTURING HISTORY OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS CAUTIONS FOR VIDEO CONNECTOR HANDLING. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -loss of image -image display artifact -device alarmed | -fluid leak -Inaccurate Flowrate -leak | Viewer
| FDA | Event Description: THE MONITOR IMAGE WAS LOST WHEN THE USER FACILITY INSERTED THE ENDOEYE FLEX DEFLECTABLE VIDEOSCOPE LTF-S190-10 TO THE PATIENT'S BODY CAVITY. THE USER FACILITY TURNED OFF AND ON THE SUBJECT DEVICE, AND DISCONNECTED AND CONNECTED THE VIDEO CONNECTOR TO THE SUBJECT DEVICE. HOWEVER, THE PHENOMENON STILL OCCURRED. THE FACILITY COMPLETED THE PROCEDURE WITH REPLACING THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE AND CONCOMITANT THE LTF-S190-10 WERE RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE SUBJECT DEVICE WORKED PROPERLY. OLYMPUS COULD REPRODUCE THE PHENOMENON WHEN THE SUBJECT DEVICE WAS CONNECTED TO THE LTF-S190-10. THE SUBJECT DEVICE DETECTED CCD OVER CURRENT WHEN THE LTF-S190-10 WAS CONNECTED TO THE SUBJECT DEVICE. THEREFORE, THE LTF-S190-10 MIGHT HAVE SOME MALFUNCTIONS. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THERE WERE NO FURTHER DETAILS PROVIDED AT THIS TIME. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. PLEASE CROSS REFERENCE THE ASSOCIATED COMPLAINT FILES: MFR REPORT#: 8010047-2016-00231. | -fail to turn on device -device malfunction -loss of video input | -fluid leak -device operates differently than expected -use of device issue | Viewer
| FDA | Event Description: THE PHYSICIAN LAUNCHED A CASE IN A FUSION WORKSTATION (FW) INTEGRATED WITH MERGE MAMMO AND ONLY PRIORS CAME UP, NOT THE STUDY REQUESTED. THE PHYSICIAN CLICKED TO VIEW THE (B)(6) 2013 STUDY AND THE FIRST CASE TO LOAD WAS THE (B)(6) 2011 STUDY. THEREFORE THE CURRENT STUDY DID NOT LAUNCH IN THE VIEWER. THERE WAS NO MESSAGE TO THE USER THAT THE RECENT CASE WAS NOT LOADED HOWEVER ALL IMAGES MAINTAINED THE DATE AND TIME STAMP. THE REPORTER IDENTIFIED THIS AS A POTENTIAL SAFETY ISSUE IN THE EVENT THE PHYSICIAN MISDIAGNOSED THE PATIENT, BEING UNAWARE OF THE MISSING STUDIES. Manufacturer Narrative: THE CORRECTION REPORT (B)(4) THAT WAS REFERENCED IN THE INITIAL MDR WAS CANCELLED BY FDA. Manufacturer Narrative: UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR. THE ISSUE WAS OCCURRING AT THIS INSTITUTION BECAUSE OF THE PARTICULAR WORKFLOW THEY USE TO LOAD STUDIES. THE NORMAL WORKFLOW IS TO DOUBLE-CLICK ON THE WORKLIST AND ALL OF THE STUDIES WILL LOAD. THE PHYSICIAN WAS TRYING TO SELECT PARTICULAR STUDIES TO LOAD. THE CODE WAS FIXED IN (B)(4) VERSION 3.3.4., AND WAS APPLIED TO THE SYSTEM AT THE INSTITUTION ON (B)(4) 2014. | -unable to view display -fail to view -device malfunction | -fitting problem -improper or incorrect procedure or method -medication error | Viewer
| FDA | Event Description: THE REPORTED FACILITY CONTACTED MCKESSON TO REQUEST ASSISTANCE WITH AN INVESTIGATION INTO INTERMITTENT DOWNTIME OF THE HORIZON MEDICAL IMAGING (HMI) SYSTEM FOR A SPECIFIC WORKSTATION AT THE SITE. THE SITE REPORTED THAT A PATIENT WAS KEPT UNDER ANESTHESIA 30 MINUTES LONGER THAN PLANNED PRIOR TO TH START OF A SURGICAL PROCEDURE. THE WORKSTATION FORM THE OPERATING ROOM WAS NOT ABLE TO CONNECT TO THE HMI PACS SYSTEM LEADING TO IMAGES FOR THE PATIENT BEING UNAVAILABLE. SOMEONE FROM THE RADIOLOGY DEPARTMENT RETRIEVED THE IMAGES ON FILM FROM THE RADIOLOGY DEPARTMENT. NO PATIENT BEING UNAVAILABLE. SOMEONE FROM THE RADIOLOGY DEPARTMENT RETRIEVED THE IMAGES ON FILM FROM THE RADIOLOGY DEPARTMENT. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: THE REPORTING FACILITY IS MULTI-FACILITY SITE. A SATELLITE FACILITY OF THE REPORTING FACILITY NEEDED A NEW WORKSTATION CONFIGURED AND MCKESSON SUPPORT HAD INADVERTENTLY PROVIDED INCORRECT INFORMATION DURING THE CONFIGURATION OF THE NEW WORKSTATION CONTAINED THE INCORRECT USER NAME AND PASSWORD FOR THE NEW WORKSTATION AT THE SATELLITE FACILITY, THE LOG-IN INFORMATION WAS FOR THE REPORTING FACILITY. WHEN USERS LOG-IN TO THE NEWLY CONFIGURED WORKSTATION AND THE LOG-ON WAS AUTHENTICATED, IT BROUGHT DOWN THE HMI SYSTEM AT THE REPORTING FACILITY SINCE THE USER NAME AND PASSWORD WERE USED TO LOG-IN TO ANOTHER SITE. BOTH THE FACILITIES HAVE CONFIRMED TO MCKESSON THAT THEY HAVE RESOLVED THAT THEY HAVE RESOLVED THIS ISSUE BY THEMSELVES BY CORRECTING THE CONFIGURATION. | -intermittent communication -intermittent sensing -intermittent continuity | -fitting problem -instruction for use issue -Communication or Transmission Issue | Viewer
| FDA | Event Description: THE REPORTED ISSUE WAS IDENTIFIED INTERNALLY BY SIEMENS R&D DEPARTMENT ON THE SYNGOVIA SYSTEMS WITH VERSION VB50 IN THE LUNGCAD NAVIGATION TOOL. AN INCORRECT NODULE GETS CORRECTED IN THE LUNGCAD NAVIGATION TOOL AFTER ACCEPTING THE CURRENTLY EDITED NODULE FOR THE 2ND TIME. CORRECTION FOCUS DOES NOT SWITCH TO THE NEW NODULE AND AFFECTS PREVIOUS NODULE, ON WHICH CORRECTION PEN WAS INVOKED. AS A RESULT, INCORRECT INFORMATION MAY CONTRIBUTE TO MEDICAL DIAGNOSIS AS A DIFFERENT NODULE QUANTIFICATION IS CHANGED AND IT COULD LEAD TO A WRONG NODULE ASSESSMENT. WRONG NODULE ASSESSMENT MAY CONTRIBUTE TO INADEQUATE CLINICAL PATIENT MANAGEMENT. THE DESCRIBED ISSUE IS DETECTIBLE TO THE USER; HOWEVER, THE USER MAY NOT REALIZE THAT VALUES HAD BEEN CHANGED FOR THE ACCEPTED FINDINGS. Manufacturer Narrative: SIEMENS STOPPED DISTRIBUTION OF THE AFFECTED SOFTWARE AND IS WORKING ON A CORRECTIVE ACTION FOR THE REPORTED ISSUE. INTERNAL ID#: (B)(4). | -incorrect snap size -incorrect occlusion -incorrect display value | -device alarm system issue -failure to align -human-device interface issue | Viewer
| FDA | Event Description: THE REPORTING FACILITY ALLEGED THAT THE ECHO IMAGES FOR PATIENT A WERE INADVERTENTLY SENT TO THE FOLDER OF PATIENT B, WHOSE FOLDER SUBSEQUENTLY CONTAINED THE WRONG ECHO IMAGES. THE ISSUE OCCURRED AT 4:20 PM. AS A RESULT OF THIS ISSUE, PATIENT B WAS TRANSFERRED TO THE CARDIAC CARE UNIT (CCU) AT APPROXIMATELY 6:30 PM, AFTER THE DEPARTMENT RECEIVED THE WRONG ECHO IMAGES. THE ISSUE WAS DISCOVERED WHEN THE REPORTING FACILITY STAFF REVIEWED PATIENT B'S RECORDS THAT EVENING. PATIENT A'S ECHO IMAGES HAVE NOT BEEN AFFECTED, SO THERE ARE NO CONSEQUENCES TO PATIENT A AS A RESULT OF THIS ISSUE AND THERE HAS BEEN NO REPORT OF INJURY OR HARM TO THE PATIENT B. Manufacturer Narrative: THE ROOT CAUSE OF THE ISSUE IS CURRENTLY UNKNOWN AND IS UNDER INVESTIGATION. (B)(4) WILL PROVIDE A FOLLOW UP REPORT TO THE AGENCY ONCE ADDITIONAL INFORMATION IS AVAILABLE. | -incorrect patient information -inadvertent motion -error message | -no display or display failure -inadequate user interface -failure to read input signal | Viewer
| FDA | Event Description: THE REPORTING FACILITY CONTACTED CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM TO DETERMINE WHETHER AN IMAGING STUDY PERFORMED FOR AN EMERGENCY PATIENT HAD BEEN IMPORTED INTO MCKESSON RADIOLOGY PACS FROM THE MODALITY, AS THE REPORTING FACILITY HAD BEEN UNABLE TO LOCATE THE STUDY IN THE SYSTEM. THE REPORTING FACILITY INDICATED THE PATIENT HAD SUBSEQUENTLY EXPIRED. Manufacturer Narrative: CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM SUCCESSFULLY LOCATED THE IMAGING STUDY FOR THE PATIENT IN MCKESSON RADIOLOGY PACS. UPON FURTHER INVESTIGATION, CHANGE HEALTHCARE'S CUSTOMER SUPPORT TEAM DETERMINED THAT THE STUDY HAD BEEN SENT TO MCKESSON RADIOLOGY FROM AN UPSTREAM SYSTEM WITHOUT AN ASSIGNED MEDICAL RECORD NUMBER (MRN). IN THIS SITUATION, MCKESSON RADIOLOGY PACS IS DESIGNED TO ASSIGN A TEMPORARY MRN USING A DESIGNATED PREFIX UPON IMPORT TO CLEARLY IDENTIFY IMAGING STUDIES NEEDING FURTHER DEMOGRAPHIC REVIEW. USERS ARE EXPECTED TO IDENTIFY AND UPDATE THE TEMPORARY MRN FIELD DURING THE REVIEW WORKFLOW FOLLOWING IMPORT OF AN IMAGING STUDY. MCKESSON RADIOLOGY PACS PROVIDES ADDITIONAL FEATURES TO ENABLE USERS TO SEARCH FOR AND LOCATE IMAGING STUDIES IN THE SYSTEM USING DIFFERENT METHODS; FOR EXAMPLE, BY USING THE ACCESSION NUMBER SEARCH OR USING THE "RECENTLY PERFORMED" STUDY LIST FILTER. MCKESSON RADIOLOGY PACS PERFORMED AS DESIGNED. | -fail to perform inspection -foreign material inside -fail to display data | -fail-safe design failure -no fail-safe mechanism -human-device interface issue | Viewer
| FDA | Event Description: THE REPORTING FACILITY HAD REPORTED A ONE HOUR DELAY IN CARE FOR A (B)(6) YEAR OLD STROKE PATIENT DUE TO A FAILURE TO IMPORT ALL STUDY IMAGES INTO THE PACS SYSTEM. THE REPORTING FACILITY PERFORMED ALTERNATIVE MEASURES TO VIEW THE STUDY IMAGES TO TREAT THE STROKE PATIENT. NO HARM TO THE PATIENT WAS REPORTED AS A RESULT OF THE DELAY. Manufacturer Narrative: MCKESSON HAS IDENTIFIED A SOFTWARE DEFECT THAT IS ENCOUNTERED WHEN USERS LOGIN TO THE PACS FROM AN EMR, OR USING LEGACY WEB URL. AS A RESULT OF THE SOFTWARE DEFECT, THE USER MAY EXPERIENCE MISSING IMAGES IN A NEWLY IMPORTED STUDY, AND/OR STUDY IMPORTS THAT REMAIN IN AN "IN-PROGRESS" STATUS. MCKESSON WILL WORK WITH AFFECTED CUSTOMERS TO APPLY A SOFTWARE UPDATE TO PREVENT THE RECURRENCE OF THE PROBLEM. | -fail to transport patient -fail to sort patient list -delayed monitoring | -fitting problem -Human-Device Interface Issue -detachment of device component | Viewer
| FDA | Event Description: THE SITE REPORTED THAT WHEN THEY ASSIGN A STUDY AS "NEEDS-OVER READ" USING THE AUTO-ASSIGN FUNCTIONALITY IN MCKESSON RADIOLOGY PACS, THE STUDY IS NOT AUTOMATICALLY ASSIGNED TO THE EMERGENCY DEPARTMENT PHYSICIAN GROUP. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: MCKESSON HAS INVESTIGATED THIS REPORTED ISSUE AND CONCLUDED THAT THE ROOT CAUSE OF THE PROBLEM IS A SOFTWARE DEFECT. AS PART OF THE SITE SPECIFIC WORKFLOW, FOR PATIENT IN THE EMERGENCY DEPARTMENT, RADIOLOGISTS SAVE A PREFERENCE SETTING TO AUTOMATICALLY ASSIGN STUDIES WITH POSITIVE OR SUSPICIOUS FINDINGS, TO THE EMERGENCY DEPARTMENT PHYSICIAN GROUP FOR FURTHER REVIEW. DUE TO THE SOFTWARE DEFECT THE PREFERENCE SETTING WAS NOT SAVED RESULTING IN NOT AUTOMATICALLY ASSIGNING SELECTED STUDIES AND ALERTING THE EMERGENCY DEPARTMENT PHYSICIANS FOR FURTHER REVIEW. THIS COULD POTENTIALLY CAUSE A DELAY IN CARE FOR PATIENT REQUIRING FURTHER TREATMENT. A SOFTWARE UPDATE HAS BEEN PROVIDED TO THE REPORTING FACILITY TO RESOLVE THE ISSUE. | -fail to label measurement automatically -incorrect readings -fail to initiate self-inspection | -inadequate instructions for healthcare professional -use of device issue -detachment of device component | Viewer
| FDA | Event Description: THE SITE'S SMART CR READER (S/N: (B)(4)) BEGAN TO EXPERIENCE SMOKE BEFORE CAR -102 WAS INSTALLED. THE READER UNIT HAS BEEN TAKEN OUT OF COMMISSION UNTIL THE POWER SUPPLY IS REPLACED. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. Manufacturer Narrative: THIS REPORT IS BEING MADE OUT OF AN ABUNDANCE OF CAUTION. SERVICE ENGINEER VISITED SITE AND EVALUATED DEVICE TO DETERMINE CAUSE OF EVENT. HE REPLACED POWER SUPPLY, CHECKED DEVICE FOR PROPER OPERATION AFTER REPAIR AND THE DEVICE WAS RETURNED TO USE. POWER SUPPLY WAS RETURNED TO FUJIFILM IN (B)(4) FOR ANALYSIS BY MANUFACTURER. A CR READER UNIT IS USED FOR READING X-RAY IMAGES THAT HAVE BEEN EXPOSED ON A CR IMAGE PLATE. THE SMARTCR IS ALSO KNOWN AS AN XG-1 OR A CR-IR 346. THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. (B)(4). THIS ADDRESSES A CORRECTIVE ACTION THAT WILL REPLACE THE POWER SUPPLY IN ALL SMART CR SYSTEMS. NOTIFICATION TO USERS WAS MADE. | -device burned -fail to produce mist -part burned | -fluid leak -loss of power -device operates differently than expected | Viewer
| FDA | Event Description: THE SITE'S SMART CR READER (S/N: (B)(4)) BEGAN TO SMOKE BEFORE CAR -102 WAS INSTALLED. THE READER UNIT HAS BEEN TAKEN OUT OF COMMISSION UNTIL THE POWER SUPPLY IS REPLACED. THERE WAS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. Manufacturer Narrative: THIS REPORT IS BEING MADE OUT OF AN ABUNDANCE OF CAUTION. SERVICE ENGINEER VISITED SITE AND EVALUATED DEVICE TO DETERMINE CAUSE OF EVENT. HE REPLACED POWER SUPPLY, CHECKED DEVICE FOR PROPER OPERATION AFTER REPAIR, AND THE DEVICE WAS RETURNED TO USE. A CR READER UNIT IS USED FOR READING X-RAY IMAGES THAT HAVE BEEN EXPOSED ON A CR IMAGE PLATE. THE SMARTCR IS ALSO KNOWN AS AN XG-1 OR A CR-IR 346. THERE IS NO INDICATION THE SYSTEM WAS BEING USED TO READ IMAGE PLATES AT THE TIME OF THE EVENT. NO PATIENTS WERE REPORTED AS BEING INVOLVED IN THE EVENT AND NO REPORTS OF INJURY WERE ALLEGED OR PROVIDED BY THE REPORTER. FUJIFILM SUBMITTED THE CORRECTION REPORT ON MARCH 22, 2016. THIS ADDRESSES A CORRECTIVE ACTION THAT WILL REPLACE THE POWER SUPPLY IN ALL SMART CR SYSTEMS. NOTIFICATION TO USERS WAS MADE. | -device burned -driver noise -part burned | -fluid leak -electrical issue -loss of power | Viewer
| FDA | Event Description: THE STEREOTACTIC IMAGING ON THE CDS WAS NOT WORKING. REQUIRED A TOTAL OF 3 DISCS WITH IMAGING TO BE RETRIEVED, WHILE THE PATIENT WAS UNDER ANESTHESIA (APPROX. 14 HOURS). THIS IS WHAT THE PATIENT SAFETY REPORT STATED: PA WAS ATTEMPTING TO BRING UP THE STEREOTACTIC IMAGE ON DISC ON THE STRYKER UNIT PRIOR TO DR BEGINNING SURGERY. THE DISC WOULD NOT FUNCTION OR OPEN THE IMAGE. THE PATIENT WAS UNDER ANESTHESIA AT THIS POINT WITHOUT AVAILABLE IMAGE. PA CALLED THE FILE ROOM. THE CD BURNER WAS NOT WORKING. HE WAS CONNECTED TO MEDICAL RECORD DEPARTMENT WHO BURNT A DISC FOR HIM. ONCE BACK IN THE OPERATING ROOM, IT WAS REALIZED THAT THE SECOND DISC HAD THE INCORRECT IMAGE BURNT ON IT. WHILE PA WAS RETRIEVING THE 3RD DISC FROM MEDICAL RECORDS, THE STRYKER REP MADE AVAILABLE A SECOND STRYKER UNIT. PRIOR TO PA RETURNING TO OPERATING ROOM WITH A THIRD DISC, THE FIRST DISC WAS ABLE TO BE UTILIZED TO OPEN THE IMAGE ON THE SECOND STRYKER UNIT. DR WAS ABLE TO BEGIN THE SURGERY. IT IS REPORTED THAT ONCE SURGERY BEGAN, THE PATIENT HAD BEEN UNDER ANESTHESIA WITHOUT ANY SURGICAL INTERVENTION FOR APPROXIMATELY ONE HOUR. IT WAS LATER LEARNED THAT THE DISC DRIVE IN THE FIRST STRYKER UNIT WAS DYSFUNCTIONAL. IT IS POSSIBLE THAT IF THE DISC DRIVE WERE TO HAVE WORKED IN THE FIRST STRYKER UNIT, THAT THE FIRST DISC MAY HAVE OPENED THE IMAGE. OVERALL, THREE CONCERNS: THE NON FUNCTIONING DISC DRIVE IN THE STRYKER UNIT. THE INABILITY TO BEGIN SURGICAL INTERVENTION WHILE THE PATIENT IS UNDER ANESTHESIA. AN INABILITY TO UTILIZE THE PAC SYSTEM TO OPEN IMAGE, REQUIRING THE IMAGE TO BE BURNT ON A DISC (ACKNOWLEDGING THIS IS NOT A NEW PROCESS). IT IS REPORTED THAT IT IS NOT UNCOMMON THAT THE INITIAL DISC BROUGHT TO THE OPERATING ROOM IS NOT FUNCTIONAL FOR VARIOUS REASONS AND A SECOND DISC HAS TO BE RETRIEVED DURING OR CASES. THE STRYKER IMAGING REP WAS PRESENT FOR THE CASE. | -unable to scroll image -fail to retrieve data -loss of image | -device operates differently than expected -imprecision -environmental control or utility issue | Viewer
| FDA | Event Description: THE SUV VALUES THAT ARE BEING CALCULATED IN THE PET/CT FUSION TOOL ARE INCORRECT. THE ANOMALY PRESENTED ITSELF WHEN THE SPHERE ANNOTATION WAS DRAWN OVER THE IMAGE, THE VALUES FOR MINIMUM, AVERAGE, AND MAXIMUM WERE ALL EQUAL OR IN SOME CASES VERY CLOSE. THIS IS AN INDICATION THAT THE VALUES ARE INCORRECT. THE VOI (VOLUME OF INTEREST) ONLY REFLECTS THE SMALL REGION WHERE THE USER FIRST BEGAN DRAGGING THE RIGHT-MOUSE BUTTON TO DRAW THE SPHERE. THIS COULD LEAD TO MISDIAGNOSIS IN PET CASES WHERE THE SUV (MAXIMUM) WOULD DISPLAY A LOWER VALUE THAN IT SHOULD. Manufacturer Narrative: CUSTOMERS HAVE BEEN NOTIFIED OF THE MALFUNCTION AND HAVE BEEN GIVEN CLEAR INSTRUCTIONS ON HOW TO PREVENT THE MALFUNCTION FROM OCCURRING AS WELL AS INSTRUCTIONS FOR REQUESTING A FREE UPGRADE TO A CORRECTED VERSION OF THE SOFTWARE. | -incorrect bolus calculation -dose calculation incorrect -incorrect measurement | -fluid leak -failure to read input signal -device difficult to maintain | Viewer
| FDA | Event Description: THE USER FACILITY INFORMED OLYMPUS THAT AN ABNORMAL MONITOR IMAGE OCCURRED WITH USING THE SUBJECT DEVICE ON (B)(6) 2015. OLYMPUS ASKED THE USER FACILITY ABOUT THE DETAILED SITUATION WHEN THE PHENOMENON OCCURRED, AND OBTAINED FOLLOWING ADDITIONAL INFORMATION ON FEBRUARY 19, 2016. DURING HYSTEROSCOPY, AN ABNORMAL MONITOR IMAGE OCCURRED WHEN THE USER FACILITY TURNED ON THE SUBJECT DEVICE AFTER THE PATIENT WAS ANESTHETIZED. THE USER FACILITY ABORTED THE PROCEDURE BECAUSE OF THIS PHENOMENON. THERE WAS NO REPORT OF PATIENT INJURY. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. THE REPORTED PHENOMENON WAS REPRODUCED WHEN OLYMPUS JIGGLED THE VIDEO CONNECTOR PLUG OF A CAMERA HEAD UP AND DOWN WITH CONNECTING A CAMERA HEAD TO THE SUBJECT DEVICE. THE SUBJECT DEVICE WORKED PROPERLY WITH REPLACING THE VIDEO CONNECTOR SOCKET TO ANOTHER ONE. THERE WAS CORROSION ON THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET. THE USER FACILITY MIGHT CONNECT THE SUBJECT DEVICE WITH A CAMERA HEAD WHOSE ELECTRICAL CONTACT OF THE VIDEO PLUG WAS NOT DRY AFTER THE CAMERA HEAD WAS REPROCESSED. THEREFORE, THE ELECTRICAL CONTACT OF THE VIDEO CONNECTOR SOCKET MIGHT GET CORRODED, RESULTING IN CONTACT FAILURE BETWEEN THE SUBJECT DEVICE AND THE CAMERA HEAD. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS THAT AN OPERATOR SHOULD MAKE SURE THE VIDEO PLUG AND ITS ELECTRICAL CONTACTS OF A CAMERA HEAD ARE COMPLETELY DRY BEFORE CONNECTING THE PLUG TO THIS DEVICE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -abnormal results -inappropriate image rendered -distorted images | -break -application program issue -device markings issue | Viewer
| FDA | Event Description: THE VIDISTAR PACS AND DICOM VIEWER SOFTWARE SYSTEM IS USED FOR THE TRANSFER, RECORDING, STORAGE, DISPLAY, AND POST-PROCESSING OF DIGITAL MEDICAL IMAGES, INCLUDING DIGITAL IMAGING AND COMMUNICATIONS IN MEDICINE (DICOM) DATA, MEASUREMENT QUANTIFICATION AND DATA MANIPULATION. VIDISTAR PACS IS SUPPORTING AN INTERFACE THAT PROVIDES THE IMAGES AND MEASUREMENTS TO A REMOTE EHR SYSTEM. SSHFS IS A SECURE FILE SYSTEM CLIENT USED TO MOUNT AND INTERACT WITH DIRECTORIES AND FILES LOCATED ON A REMOTE SERVER. DURING A STUDY TRANSFER TO THE REMOTE EHR SYSTEM, THE PATIENT IMAGES INCLUDED AN IMAGE FROM ANOTHER PATIENT'S STUDY. THE ADDITIONAL IMAGE WAS LABELED CORRECTLY, BUT WAS MISFILED. THE PROBLEM WAS CAUGHT BY THE TECHNICIAN BEFORE THE STUDY WAS INTERPRETED. Manufacturer Narrative: OUR FIRM INVESTIGATED THE CUSTOMER COMPLAINT OF INTERMIXED IMAGES AND WERE ABLE TO DETERMINE THAT THE PIECE OF SOFTWARE RESPONSIBLE FOR TRANSFERRING THE IMAGES FROM ONE SYSTEM TO ANOTHER (SSHFS VERSION 2.4) WAS INTERMIXING IMAGES UNDER SPECIFIC CIRCUMSTANCES. IMPLEMENTING CHECKSUM LOGIC TO THE APPLICATION'S LOGS REVEALED WHEN IN THE TRANSFER PROCESS IMAGES WERE INTERMIXED. THIS INDICATED THAT THE CULPRIT WAS A BUG IN THE TRANSFER SOFTWARE (SSHFS VERSION 2.4) AND NOT THE MODALITY, VIDISTAR PACS, NOR THE EHR SYSTEM. THE TRANSFER SOFTWARE (SSHFS) IS A THIRD-PARTY OPEN SOURCE SOFTWARE THAT WAS ORIGINALLY CHOSEN AND VALIDATED FOR THE INTERFACE BY VIDISTAR AND THE EHR SYSTEM. FURTHER INVESTIGATION SHOWED INSTABILITY OF THE MOUNT CREATED BY THE TRANSFER SOFTWARE (SSHFS VERSION 2.4) AT THE LEVELS OF TRAFFIC CREATED BY THE INTERFACE THUS OFFERING AN INCREASED LIKELIHOOD OF THE SPECIFIC CIRCUMSTANCES NECESSARY FOR THE BUG TO PRESENT. | -fail to process medical imaging -inadequate image manipulation -inadequate message display | -use of device issue -pacing inadequately -disconnection | Viewer
| FDA | Event Description: THE WORKFLOW TRIGGERING THE EVENT INVOLVES MULTIPLE SYSTEMS: MEASUREMENTS ARE DONE ON THE ULTRASOUND MODALITY AND THEN STUDIES ARE SENT IN TO ENTERPRISE IMAGING (EI) FROM AGFA. FROM EI CLIENT THE STUDIES ARE OPENED IN TOMTEC-ARENA FROM TOMTEC AND THE MEASUREMENTS ARE THEN EXPORTED VIA XML FILE TO THE CUPID SYSTEM FROM EPIC FOR REPORTING PURPOSE. IN SOME CIRCUMSTANCES IT HAPPENS THAT THE MEASUREMENTS FROM THE PREVIOUS PATIENT\STUDY IS SENT ADDITIONALLY TO EPIC CUPID FOR THE CURRENT ONE HOWEVER WITH THE CURRENT PATIENTS\STUDY DETAILS. THE SOFTWARE IS USED IN CLINICAL ROUTINE, THERE WAS NO INJURY OR SERIOUS DETERIORATION OF HEALTH STATE REPORTED WITH THIS EVENT. Manufacturer Narrative: AN INVESTIGATION IS UNDERWAY TO DETERMINE THE ROOT CAUSE OF THE ISSUE. RESULTS OF THE INVESTIGATION WILL BE INCLUDED IN A FOLLOW UP REPORT UPON ITS COMPLETION. | -fail to place measurements -request trigger failed -fail to follow process | -leak -fluid leak -failure to read input signal | Viewer
| FDA | Event Description: THIS IS A COMMON AND WIDESPREAD PROBLEM AFFECTING MANY CASES. WHEN THE PATIENT'S EHR IS OPEN AND BEING SEARCHED FOR RADIOLOGY RESULTS, THERE IS A CAMERA ICON ON WHICH TO CLICK WHICH OPENS THE RADIOGRAPH IMAGES TO BE SEEN. THE PATIENT'S CHEST RADIOGRAPH WAS THE SUBJECT OF THIS QUERY. THE ICON WAS CLICKED AND A NORMAL APPEARING CHEST RADIOGRAPH APPEARED AND OTHER RESULTS WERE THEN ACCESSED. THE PROBLEM IS THAT THE RADIOGRAPH THAT APPEARED AND WAS NORMAL, WAS NOT OF THIS PATIENT, BUT OF SOME OTHER PATIENT THAT WAS NOT EVEN ON MY PATIENT LIST. THE PATIENT IN QUESTION HAD A RADIOGRAPH THAT INDICATED PULMONARY EDEMA AND WAS SOUGHT AFTER THE PHYSICAL LUNG EXAM WAS FOUND TO BE INCONSISTENT WITH A NORMAL CHEST RADIOGRAPH. THE DEFECTIVE PRESENTATION OF ANOTHER PATIENT'S RADIOGRAPH IS COMMON WITH THIS INTERFACE OF DEVICES BUT IF THE IMAGE WAS ALLOWED TO STAY ON THE SCREEN, THE CORRECT IMAGE OF THE RIGHT PATIENT WOULD LOAD. THIS DEFECT CAUSES MISDIAGNOSIS AND WRONG TREATMENTS. | -fail to provide patient diagnosis info -fail to store medical imaging -loss of image | -failure to read input signal -therapeutic or diagnostic output failure -device operates differently than expected | Viewer
| FDA | Event Description: THIS IS A TEST SUBMISSION FOR MIGRATION FROM A TEST ACCOUNT TO A PRODUCTION ACCOUNT. NO ACTUAL EVENT OCCURRED. Manufacturer Narrative: THIS IS A FICTITIOUS SUBMISSION FOR PURPOSES OF ESCALATION TO A PRODUCTION ACCOUNT. | -fail to initiate test -usability test failed -inappropriate user | -use of device issue -incorrect or inadequate result -medication error | Viewer
| FDA | Event Description: THIS ISSUE WAS FOUND DURING A RETROSPECTIVE REVIEW CONDUCTED BY SIEMENS. DURING THE REVIEW A COMPLAINT FOR AN ISSUE IDENTICAL TO THE PREVIOUSLY REPORTED ADVERSE EVENT WAS IDENTIFIED. AFTER CONTENT OF THE COMPLAINT WAS EVALUATED, SIEMENS DECIDED TO SUBMIT AN ADVERSE EVENT REPORT. THE EARLIER EXISTING MDR CASE WAS INTERNALLY TRACKED WITH (B)(4), AND WAS REPORTED AS AN ADVERSE EVENT UNDER REPORT # 2240869-2014-00003625. THE SOLUTION FOR THE REPORTED ISSUE WAS PROVIDED BY SIEMENS IN 2015. ALL POTENTIALLY AFFECTED CUSTOMERS WERE NOTIFIED ABOUT THE ISSUE VIA A CUSTOMER SAFETY ADVISORY NOTICE (UPDATE INSTRUCTION SY099/14/S) REPORTED UNDER C&R REPORT # 2240869-01/15/15-0002-C (Z-1027-2015) IN CONJUNCTION WITH SOFTWARE SOLUTION (DISTRIBUTED VIA AN UPDATE INSTRUCTION SY004/15/S). ADDITIONAL SOFTWARE FIXES WERE RELEASED VIA UPDATE INSTRUCTIONS; ú SY0018/15/S (C&R REPORT 2240869-08/12/15-0025-C; Z-2719-2015) ú SY024/15/S ( C&R REPORT #2240869-02/23/15-0006-C; Z-1354-2015) DURING A SITE CHECK IN (B)(6) ("ATRIUM (B)(6)", (B)(6)), WHERE THE (B)(6) SERVER CRASHED, (B)(4) MIXED UP IMAGES WERE DETECTED BY SIEMENS SERVICE PERSONNEL (B)(6) 2014. THIS ISSUE IS IDENTICAL TO THE EARLIER CASE TRACKED WITH COMPLAINT (B)(4), REPORTED AS ADVERSE EVENT UNDER REPORT # 2240869-2014-00003625. FOR THE SITE IN (B)(6) THE FOLLOWING WAS DETERMINED : "TODAY WE HAD AN UNEXPECTED CRASH AT (B)(6). SO ONCE THE SYSTEM WAS OK I RUNNED THE QUERY TO CHECK FOR IMAGE FILENAME DUPLICATES AND WE ALSO HAVE HERE SOME DUPLICATES, BUT FROM THE PAST (B)(6)." (B)(4) ADDITIONAL IMAGES WITH DUPLICATE REFERENCE WERE IDENTIFIED FROM AN EARLIER CRASH (B)(6) 2014. . IN CASE OF AN OPERATING SYSTEM CRASH (BLUESCREEN) OR AN UNEXPECTED CRASH OF THE (B)(6) DICOM SERVICES, EXACTLY DURING WORK-IN-PROGRESS OF STORING RECEIVED IMAGES, IT MAY HAPPEN (ARBITRARILY, DEPENDING ON THE CACHE BEHAVIOR OF THE OPERATING SYSTEM) THAT SOME IMAGE FILENAMES WHICH USED LAST BEFORE THE CRASH ARE USED AGAIN FOR THE NEXT IMAGES RECEIVED AFTER THE SYSTEM RESUMED WORKING. IF THIS HAPPENS, A SITUATION MAY ARISE WHEN IN THE IMAGE TABLE WITHIN THE (B)(6) DATABASE TWO REFERENCES WILL BE POINTING TO THE SAME IMAGE FILENAME. THE REFERENCE FROM BEFORE THE CRASH NOW ALSO POINTS TO THE IMAGE WRITTEN AFTER THE CRASH, WHICH HAS CONTENT FROM AN EXAMINATION/PATIENT AFTER CRASH. THESE REFERENCES ARE UNIDIRECTIONAL, THEREFORE NO TRACE BACK FROM THE IMAGE TO THE PATIENT IS POSSIBLE. Manufacturer Narrative: (B)(4). ROOT CAUSE : PRIOR TO SENDING AN ACKNOWLEDGMENT TO THE OUTGOING MODALITY THAT THE IMAGE HAS BEEN RECEIVED, THE (B)(6) SYSTEM RELIES ON A SIGNAL FROM THE HARDWARE TO ENSURE THAT A FILE IS AVAILABLE IN THE STS (SHORT TERM STORAGE). A REFERENCE WITH A FILE NAME FOR THE IMAGE WILL BE ADDED TO THE DATABASE. BUT THE HARDWARE SIGNAL "AVAILABLE IN STS" DOES NOT ENSURE IMAGE BEING PHYSICALLY WRITTEN TO A HARD DISC; THIS SIGNAL WILL ALREADY BE SENT AFTER AVAILABILITY ONLY IN THE TEMPORARY CACHE. WHEN A DIFFERENT IMAGE IS RECEIVED AFTER THE SYSTEM IS WORKING AGAIN, THE (B)(6) WILL ALLOCATE THE SAME FILE NAME FOR THE NEW IMAGE, WILL CHECK THAT THE FILE NAME DOES NOT EXIST IN THE FILE SYSTEM AND AFTER THE SIGNAL "AVAILABLE IN STS" IS RECEIVED, IT WILL ADD AN ENTRY TO THE DATABASE FOR THE NEW IMAGE. THEREFORE, THE DATABASE WILL END UP WITH HAVING TWO REFERENCES FOR THE SAME IMAGE. AS LONG AS THE NEW IMAGE IS PHYSICALLY AVAILABLE IN THE STS, LOADING AND ARCHIVING FOR BOTH IMAGES WILL BE INDICATED AS SUCCESSFUL, ALTHOUGH A PATIENT DATA MIX-UP MAY HAVE ALREADY OCCURRED. THIS LEADS TO THE FOLLOWING RESULTS: FOR THE FIRST RECEIVED SERIES AN INCORRECT IMAGE IS BEING DISPLAYED (E.G. WRONG EXAMINATION PARAMETER, WRONG BODY PART EXAMINATION, WRONG PATIENT, ETC). THE IMAGE TEXT IN THIS CASE WILL NOT REFLECT CORRECT INFORMATION EITHER. FOR THE LATER RECEIVED STUDY, THE DISPLAYED IMAGE IS CORRECT, IMPACT OF THE ISSUE FOR THE SITE IN (B)(6) : (B)(4) FOUND IMAGES COULD NOT BE ASSIGNED TO THE CORRECT PATIENTS , THEREFORE THEY POSED RISK THAT THEY WILL BE SHOWN INSIDE A SERIES TO WHICH THEY DO NOT BELONG. FOR THESE "FOREIGN-CONTENT-IMAGES" IN THE IMAGE SEQUENCE INSIDE SOME SERIES IT CANNOT BE COMPLETELY PREDICTED, HOW THEY ACTUALLY COULD CONTRIBUTE TO A WRONG DIAGNOSTIC DECISION. IN CASE THEY WOULD BE FROM A DIFFERENT BODY REGION OR FROM A DIFFERENT SCANNER, IT WOULD BE CLEARLY OBVIOUS THAT THEY ARE WRONG. IF THEY ARE FROM SAME BODY REGION, THEN TYPICALLY THESE "WRONGLY INSERTED" IMAGES WOULD HAVE DIFFERENT SPATIAL ALIGNMENT (BECAUSE OF EVEN SLIGHT DIFFERING ADJUSTMENT OF THE SCANNING LOCATION AND PARAMETERS) AND WOULD CAUSE A "JUMPING" WHEN SCROLLING THROUGH THE SERIES. THIS IS NOTICABLE TO THE USER. THEREFORE THE RISK IS SEEN AS RATHER IMPROBABLE THAT THIS ISSUE COULD LEAD TO A WRONG DIAGNOSIS, AN INJURY (OR EVEN DEATH). NO HARM HAS BEEN REPORTED FROM THIS SITE, THE ISSUE WAS NOT DETECTED BY THE CUSTOMER BUT BY PROACTIVE SERVICE ACTION. THIS REPORT WAS SUBMITTED OCTOBER 3, 2017. | -patient experienced adverse event -lead to adverse trend -fail to identify hazardous situation | -Device Handling Issue -label -output problem | Viewer
| FDA | Event Description: THIS ISSUE WAS FOUND DURING A RETROSPECTIVE REVIEW CONDUCTED BY SIEMENS. DURING THE REVIEW A COMPLAINT FOR AN ISSUE IDENTICAL TO THE PREVIOUSLY REPORTED ADVERSE EVENT WAS IDENTIFIED. AFTER CONTENT OF THE COMPLAINT WAS EVALUATED, SIEMENS DECIDED TO SUBMIT AN ADVERSE EVENT REPORT. THE EARLIER EXISTING MDR CASE WAS INTERNALLY TRACKED WITH PM00003625, AND WAS REPORTED AS AN ADVERSE EVENT UNDER REPORT # 2240869-2014-00003625. THE SOLUTION FOR THE REPORTED ISSUE WAS PROVIDED BY SIEMENS IN 2015. ALL POTENTIALLY AFFECTED CUSTOMERS WERE NOTIFIED ABOUT THE ISSUE VIA A CUSTOMER SAFETY ADVISORY NOTICE (UPDATE INSTRUCTION SY099/14/S) REPORTED UNDER C&R REPORT # 2240869-01/15/15-0002-C (Z-1027-2015) IN CONJUNCTION WITH SOFTWARE SOLUTION (DISTRIBUTED VIA AN UPDATE INSTRUCTION SY004/15/S). ADDITIONAL SOFTWARE FIXES WERE RELEASED VIA UPDATE INSTRUCTIONS; ú SY0018/15/S (C&R REPORT 2240869-08/12/15-0025-C; Z-2719-2015). ú SY024/15/S ( C&R REPORT #2240869-02/23/15-0006-C; Z-1354-2015). DURING SITE CHECK WHERE SYNGO.PLAZA SERVER CRASHES WERE OBSERVED, FOUR MIXED UP IMAGES WERE DETECTED BY SIEMENS SERVICE PERSONNEL IN (B)(4) (SITE "HOSPITAL (B)(6)) ON THE (B)(4) 2014. FOR THE SITE IN (B)(4) FOUR COLLIDING IMAGE FILENAMES WERE DETERMINED. SIEMENS VERIFIED ALL FOUR CASES, HOWEVER, THE MOST OBVIOUS ONE WAS THE FIRST CASE - MR IMAGE ON CR THORAX STUDY. FOR THE OTHER 3 CASES IT COULD NOT BE DETERMINED WHICH IMAGE BELONGS TO WHICH PATIENT/STUDY. IN CASE OF AN OPERATING SYSTEM CRASH (BLUESCREEN) OR AN UNEXPECTED CRASH OF THE SYNGO.PLAZA DICOM SERVICES, EXACTLY DURING WORK-IN-PROGRESS OF STORING RECEIVED IMAGES, IT MAY HAPPEN (ARBITRARILY, DEPENDING ON THE CACHE BEHAVIOR OF THE OPERATING SYSTEM) THAT SOME IMAGE FILENAMES, WHICH WERE USED LAST BEFORE THE CRASH, ARE USED AGAIN FOR THE NEXT IMAGES RECEIVED AFTER THE SYSTEM RESUMED WORKING. IF THIS OCCURS, A SITUATION THAT IN THE IMAGE TABLE WITHIN THE SYNGO.PLAZA DATABASE TWO REFERENCES ARE POINTING TO THE SAME IMAGE FILENAME MAY ARISE. THE REFERENCE FROM BEFORE THE CRASH ALSO POINTS TO THE IMAGE WRITTEN AFTER THE CRASH, WHICH HAS CONTENT FROM THE EXAMINATION/PATIENT AFTER CRASH. THESE REFERENCES ARE UNIDIRECTIONAL, THEREFORE NO TRACE BACK FROM THE IMAGE TO THE PATIENT IS POSSIBLE. Manufacturer Narrative: EXEMPTION NUMBER (B)(4). (B)(4) (IMPORTER) IS SUBMITTING THE REPORT ON BEHALF OF SIEMENS (B)(4), (MANUFACTURER). ROOT CAUSE : PRIOR TO SENDING AN ACKNOWLEDGMENT TO THE OUTGOING MODALITY THAT THE IMAGE HAS BEEN RECEIVED, THE SYNGE.PLAZA SYSTEM RELIES ON A SIGNAL FROM THE HARDWARE TO ENSURE THAT A FILE IS AVAILABLE IN THE STS (SHORT TERM STORAGE). A REFERENCE WITH A FILE NAME FOR THE IMAGE WILL BE ADDED TO THE DATABASE. BUT THE HARDWARE SIGNAL "AVAILABLE IN STS" DOES NOT ENSURE IMAGE BEING PHYSICALLY WRITTEN TO A HARD DISC; THIS SIGNAL WILL ALREADY BE SENT AFTER AVAILABILITY ONLY IN THE TEMPORARY CACHE. WHEN A DIFFERENT IMAGE IS RECEIVED AFTER THE SYSTEM IS WORKING AGAIN, THE SYNGO.PLAZA WILL ALLOCATE THE SAME FILE NAME FOR THE NEW IMAGE, WILL CHECK THAT THE FILE NAME DOES NOT EXIST IN THE FILE SYSTEM AND AFTER THE SIGNAL "AVAILABLE IN STS" IS RECEIVED, IT WILL ADD AN ENTRY TO THE DATABASE FOR THE NEW IMAGE. THEREFORE, THE DATABASE WILL END UP WITH HAVING TWO REFERENCES FOR THE SAME IMAGE. AS LONG AS THE NEW IMAGE IS PHYSICALLY AVAILABLE IN THE STS, LOADING AND ARCHIVING FOR BOTH IMAGES WILL BE INDICATED AS SUCCESSFUL, ALTHOUGH A PATIENT DATA MIX-UP MAY HAVE ALREADY OCCURRED. THIS LEADS TO THE FOLLOWING RESULTS: - FOR THE FIRST RECEIVED SERIES AN INCORRECT IMAGE IS BEING DISPLAYED (E.G. WRONG EXAMINATION PARAMETER, WRONG BODY PART EXAMINATION, WRONG PATIENT, ETC). THE IMAGE TEXT IN THIS CASE WILL NOT REFLECT CORRECT INFORMATION EITHER. - FOR THE LATER RECEIVED STUDY, THE DISPLAYED IMAGE IS CORRECT, IMPACT OF THE ISSUE FOR THE SITE IN (B)(4): AFTER ONE WRONG REFERENCE (NO MORE LOST DATA / PATIENT DATA MIX-UP FOR THIS ONE) WAS CLARIFIED, 3 IMAGES WERE LEFT, WHICH COULD NOT BE ASSIGNED TO THE CORRECT PATIENT ANY LONGER. THEREFORE POSED RISK THAT THEY WILL BE SHOWN INSIDE A SERIES TO WHICH THEY DO NOT BELONG. FOR THESE THREE "FOREIGN-CONTENT-IMAGES" IN THE IMAGE SEQUENCE INSIDE SOME SERIES IT CANNOT BE COMPLETELY PREDICTED, HOW THEY ACTUALLY COULD CONTRIBUTE TO A WRONG DIAGNOSTIC DECISION. IN CASE THEY WOULD BE FROM A DIFFERENT BODY REGION OR FROM A DIFFERENT SCANNER, IT WOULD BE CLEARLY OBVIOUS THAT THEY ARE WRONG. IF THEY ARE FROM SAME BODY REGION, THEN TYPICALLY THESE "WRONGLY INSERTED" IMAGES WOULD HAVE DIFFERENT SPATIAL ALIGNMENT (BECAUSE OF EVEN SLIGHT DIFFERING ADJUSTMENT OF THE SCANNING LOCATION AND PARAMETERS) AND WOULD CAUSE A "JUMPING" WHEN SCROLLING THROUGH THE SERIES. THIS IS NOTICEABLE TO THE USER. . THEREFORE THE RISK IS SEEN AS RATHER IMPROBABLE THAT THIS ISSUE COULD LEAD TO WRONG DIAGNOSIS, INJURY (OR EVEN DEATH). NO HARM HAS BEEN REPORTED FROM THIS SITE, THE ISSUE WAS NOT DETECTED BY THE CUSTOMER BUT BY PROACTIVE SERVICE ACTION. | -patient experienced adverse event -lead to adverse trend -fail to identify hazardous situation | -Device Handling Issue -label -output problem | Viewer
| FDA | Event Description: THIS ISSUE WAS FOUND DURING A RETROSPECTIVE REVIEW CONDUCTED BY SIEMENS. DURING THE REVIEW A COMPLAINT FOR AN ISSUE IDENTICAL TO THE PREVIOUSLY REPORTED ADVERSE EVENT WAS IDENTIFIED. AFTER CONTENT OF THE COMPLAINT WAS EVALUATED, SIEMENS DECIDED TO SUBMIT AN ADVERSE EVENT REPORT. THE EARLIER EXISTING MDR CASE WAS INTERNALLY TRACKED WITH (B)(4), AND WAS REPORTED AS AN ADVERSE EVENT UNDER REPORT # (B)(4). THE SOLUTION FOR THE REPORTED ISSUE WAS PROVIDED BY SIEMENS IN 2015. ALL POTENTIALLY AFFECTED CUSTOMERS WERE NOTIFIED ABOUT THE ISSUE VIA A CUSTOMER SAFETY ADVISORY NOTICE (UPDATE INSTRUCTION SY099/14/S) REPORTED UNDER C&R REPORT # 2240869-01/15/15-0002-C (Z-1027-2015) IN CONJUNCTION WITH SOFTWARE SOLUTION (DISTRIBUTED VIA AN UPDATE INSTRUCTION SY004/15/S). ADDITIONAL SOFTWARE FIXES WERE RELEASED VIA UPDATE INSTRUCTIONS; ú SY0018/15/S (C&R REPORT 2240869-08/12/15-0025-C; Z-2719-2015). ú SY024/15/S ( C&R REPORT #2240869-02/23/15-0006-C; Z-1354-2015). DURING SITE CHECKS WHERE SYNGO.PLAZA SERVER CRASHES WERE OBSERVED, 13 MIXED UP IMAGES WERE DETECTED BY SIEMENS SERVICE PERSONNEL IN (B)(6) ON THE (B)(6), 2017. THIS ISSUE TURNED OUT TO BE IDENTICAL TO THE EARLIER CASE TRACKED WITH (B)(4), REPORTED AS ADVERSE EVENT UNDER REPORT # (B)(4). FOR THE SITE IN (B)(4) THE FOLLOWING WAS DETERMINED: "IMAGE FILE NAMES HAS BEEN REUSED FOR DIFFERENT STUDIES/SERIES AND DIFFERENT PATIENTS. IN THIS SITE WE HAVE A SERIES 1.2.840.113704.1.111.2224.1405102009.11, WHICH HAS 13 IMAGES FROM DIFFERENT STUDIES/SERIES." IN CASE OF AN OPERATING SYSTEM CRASH (BLUESCREEN) OR AN UNEXPECTED CRASH OF THE SYNGO.PLAZA DICOM SERVICES, EXACTLY DURING WORK-IN-PROGRESS OF STORING RECEIVED IMAGES, IT MAY HAPPEN (ARBITRARILY, DEPENDING ON THE CACHE BEHAVIOR OF THE OPERATING SYSTEM) THAT SOME IMAGE FILENAMES WHICH USED LAST BEFORE THE CRASH ARE USED AGAIN FOR THE NEXT IMAGES RECEIVED AFTER THE SYSTEM RESUMED WORKING. IF THIS HAPPENS, A SITUATION MAY ARISE WHEN IN THE IMAGE TABLE WITHIN THE SYNGO.PLAZA DATABASE TWO REFERENCES WILL BE POINTING TO THE SAME IMAGE FILENAME. THE REFERENCE FROM BEFORE THE CRASH NOW ALSO POINTS TO THE IMAGE WRITTEN AFTER THE CRASH, WHICH HAS CONTENT FROM THE EXAMINATION/PATIENT AFTER CRASH. THESE REFERENCES ARE UNIDIRECTIONAL, THEREFORE NO TRACE BACK FROM THE IMAGE TO THE PATIENT IS POSSIBLE. Manufacturer Narrative: (B)(4). ROOT CAUSE : PRIOR TO SENDING AN ACKNOWLEDGMENT TO THE OUTGOING MODALITY THAT THE IMAGE HAS BEEN RECEIVED, THE SYNGE.PLAZA SYSTEM RELIES ON A SIGNAL FROM THE HARDWARE TO ENSURE THAT A FILE IS AVAILABLE IN THE STS (SHORT TERM STORAGE). A REFERENCE WITH A FILE NAME FOR THE IMAGE WILL BE ADDED TO THE DATABASE. BUT THE HARDWARE SIGNAL "AVAILABLE IN STS" DOES NOT ENSURE IMAGE BEING PHYSICALLY WRITTEN TO A HARD DISC; THIS SIGNAL WILL ALREADY BE SENT AFTER AVAILABILITY ONLY IN THE TEMPORARY CACHE. WHEN A DIFFERENT IMAGE IS RECEIVED AFTER THE SYSTEM IS WORKING AGAIN, THE SYNGO.PLAZA WILL ALLOCATE THE SAME FILE NAME FOR THE NEW IMAGE, WILL CHECK THAT THE FILE NAME DOES NOT EXIST IN THE FILE SYSTEM AND AFTER THE SIGNAL "AVAILABLE IN STS" IS RECEIVED, IT WILL ADD AN ENTRY TO THE DATABASE FOR THE NEW IMAGE. THEREFORE, THE DATABASE WILL END UP WITH HAVING TWO REFERENCES FOR THE SAME IMAGE. AS LONG AS THE NEW IMAGE IS PHYSICALLY AVAILABLE IN THE STS, LOADING AND ARCHIVING FOR BOTH IMAGES WILL BE INDICATED AS SUCCESSFUL, ALTHOUGH A PATIENT DATA MIX-UP MAY HAVE ALREADY OCCURRED. THIS LEADS TO THE FOLLOWING RESULTS: - FOR THE FIRST RECEIVED SERIES AN INCORRECT IMAGE IS BEING DISPLAYED (E.G. WRONG EXAMINATION PARAMETER, WRONG BODY PART EXAMINATION, WRONG PATIENT, ETC). THE IMAGE TEXT IN THIS CASE WILL NOT REFLECT CORRECT INFORMATION EITHER. - FOR THE LATER RECEIVED STUDY, THE DISPLAYED IMAGE IS CORRECT, IMPACT OF THE ISSUE FOR THE SITE IN (B)(4): 13 IMAGES WERE FOUND WHICH COULD NOT BE ASSIGNED TO THE CORRECT PATIENT , THEREFORE THEY POSED A RISK THAT THEY WILL BE SHOWN INSIDE A SERIES TO WHICH THEY DO NOT BELONG. FOR THESE 13 "FOREIGN-CONTENT-IMAGES" IN THE IMAGE SEQUENCE INSIDE SOME SERIES IT CANNOT BE COMPLETELY PREDICTED, HOW THEY ACTUALLY COULD CONTRIBUTE TO A WRONG DIAGNOSTIC DECISION. IN CASE THEY WOULD BE FROM A DIFFERENT BODY REGION OR FROM A DIFFERENT SCANNER, IT WOULD BE CLEARLY OBVIOUS THAT THEY ARE WRONG. IF THEY ARE FROM SAME BODY REGION, THEN TYPICALLY THESE "WRONGLY INSERTED" IMAGES WOULD HAVE DIFFERENT SPATIAL ALIGNMENT (BECAUSE OF EVEN SLIGHT DIFFERING ADJUSTMENT OF THE SCANNING LOCATION AND PARAMETERS) AND WOULD CAUSE A "JUMPING" WHEN SCROLLING THROUGH THE SERIES. SO USER COULD NOTICE. THEREFORE THE RISK IS SEEN AS RATHER IMPROBABLE THAT THIS ISSUE COULD LEAD TO A WRONG DIAGNOSIS, AN INJURY (OR EVEN DEATH). NO HARM HAS BEEN REPORTED FROM THIS SITE, THE ISSUE WAS NOT DETECTED BY THE CUSTOMER BUT BY PROACTIVE SERVICE ACTION. | -patient experienced adverse event -lead to adverse trend -fail to identify hazardous situation | -Device Handling Issue -label -output problem | Viewer
| FDA | Event Description: THIS MORNING, THE TELEMETRY SYSTEM (SCOTTCARE) FROZE. STAFF WAS IN THE PROCESS OF ADDING PATIENTS TO THE SYSTEM. AFTER THE THIRD WAS ENTERED THE SYSTEM DID NOT RESPOND. SYSTEM REBOOT OCCURRED PER STAFF. BIOMED WAS CALLED AND THEY CAME TO WORK ON THE SYSTEM. AFTER 3 HOURS OF TIME, THE SYSTEM BECAME FUNCTIONAL. STAFF CALLED TO CANCEL PATIENT FOR THE DAY, AS WE DID NOT HAVE AN IDEA TO HOW LONG THE SYSTEM WOULD BE DOWN, WHILE THE FIXES WERE BEING WORKED ON. | -telemetry interruption -device frozen -fail to disable sensor | -device operates differently than expected -electrical issue -device inoperable | Viewer
| FDA | Event Description: THIS SUPPLEMENT REPORT #2 IS BEING SUBMITTED TO PROVIDE CORRECTED INFORMATION. Manufacturer Narrative: AGFA HAS SUBMITTED FDA MDR SUPPLEMENTAL REPORT # 1225058-2017-00002 SUP 1 ON JULY 03, 2017. IN THIS REPORT WE HAVE PROVIDED THE FOLLOWING INFORMATION: "PREVIOUSLY THE CUSTOMER HAS DECLINED THE CORRECTION FOR THE PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. ON AUGUST 22, 2017, THE CUSTOMER HAS ACCEPTED THE CORRECTION AND REQUESTED THAT AGFA TURN OFF THE DEMOGRAPHICS MANAGER COMPONENT. THE CUSTOMER ALSO AGREED TO UPGRADE THEIR SYSTEM TO CV12.2 SU8 AND ON AUGUST 22, 2017 WORK RELATED TO THIS UPGRADE BEGAN IN THE CUSTOMER'S TEST ENVIRONMENT. THE UPGRADE TO THE CUSTOMER'S PRODUCT ENVIRONMENT IS SCHEDULED FOR AUGUST 2017." THIS SUPPLEMENTAL REPORT # 2 IS BEING SUBMITTED TO PROVIDE THE CORRECT INFORMATION. THE DATE OF THE CORRECTION AT THE CUSTOMER SITE HAS BEEN CORRECTED FROM AUGUST 22, 2017 TO JUNE 22, 2017. THE CORRECTED INFORMATION IS BELOW: "PREVIOUSLY THE CUSTOMER HAS DECLINED THE CORRECTION FOR THE PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. ON JUNE, 2017, THE CUSTOMER HAS ACCEPTED THE CORRECTION AND REQUESTED THAT AGFA TURN OFF THE DEMOGRAPHICS MANAGER COMPONENT. THE CUSTOMER ALSO AGREED TO UPGRADE THEIR SYSTEM TO CV12.2 SU8 AND ON JUNE 22, 2017 WORK RELATED TO THIS UPGRADE BEGAN IN THE CUSTOMER'S TEST ENVIRONMENT. THE UPGRADE TO THE CUSTOMER'S PRODUCT ENVIRONMENT IS SCHEDULED FOR AUGUST 2017." Manufacturer Narrative: THIS SUPPLEMENT REPORT #1 IS BEING SUBMITTED TO PROVIDE CORRECTIVE ACTIONS TAKEN. PREVIOUSLY THE CUSTOMER HAS DECLINED THE CORRECTION FOR THE PRODUCT DEFECT WITHIN THE DEMOGRAPHICS MANAGER COMPONENT OF IMPAX CV 12.1. ON (B)(6) 2017, THE CUSTOMER HAS ACCEPTED THE CORRECTION AND REQUESTED THAT AGFA TURN OFF THE DEMOGRAPHICS MANAGER COMPONENT. THE CUSTOMER ALSO AGREED TO UPGRADE THEIR SYSTEM TO CV12.2 SU8 AND ON (B)(6) 2017 WORK RELATED TO THIS UPGRADE BEGAN IN THE CUSTOMER'S TEST ENVIRONMENT. THE UPGRADE TO THE CUSTOMER'S PRODUCT ENVIRONMENT IS SCHEDULED FOR (B)(6) 2017. ANY ADDITONAL SUPPLEMENTAL REPORTING WILL BE REPORTED TO FDA VIA 21 CFR 806 REPORTING AS AN ADDENDUM TO Z-0283-2016. THERE HAS BEEN NO REPORTED HARM TO PATIENT OR USER DURING THIS EVENT. |  |  | Viewer
| FDA | Event Description: THIS WAS A (B)(6) PATIENT WITH A HISTORY OF SEIZURES AND A ASTROCYTOMA. THE PATIENT WAS HAVING LASER INTERSTITIAL THERMAL THERAPY PROCEDURE ON HER ASTROCYTOMA. SHE WAS TAKEN TO THE SIEMENS REALTIME MIR FOR THE PROCEDURE WHICH WAS PERFORMED WITHOUT COMPLICATION. THE COMPLICATION OCCURRED WHEN THE LASER PROBE WAS BEING REMOVED. DURING THE COMPLETION OF THE ABLATION PROCEDURE, THE LASER TIP WAS WITHDRAWN FROM THE CRANIUM BOLT AND IT WAS NOTICED THAT THE LASER CRYSTAL, OR STYLUS, WAS NOT ATTACHED TO THE WAND. IT WAS NECESSARY FOR THE PATIENT TO BE TAKEN TO THE OPERATING ROOM AND THE BURR HOLE ENLARGED TO REMOVE THE STYLUS. THE PATIENT WENT HOME THE FOLLOWING DAY AND POST OP FOLLOW UP WAS VERY POSITIVE. THERE WERE NO COMPLICATIONS WITH THE SECOND OPERATING ROOM VISIT TO RETRIEVE THE STYLUS. | -fail to deliver therapy -fail to enable laser source output -thermal damaged tissue | -fluid leak -Device Handling Issue -failure to recalibrate | Viewer
| FDA | Event Description: UNABLE TO LOCATE FILMS ON MCKESSON SYSTEM. PATIENT WAS DONE IN CATHETER LAB. WORK ORDER PUT INTO IS, BUT THERE WAS AN ERROR MESSAGE ON THE COMPUTER DURING THIS TIME. THE FILMS COULD NOT BE TRANSFERRED OVER UNTIL THE ERROR MESSAGE WAS CLEARED. ONCE THE ERROR MESSAGE WAS CORRECTED, THE FILMS WERE AVAILABLE. ALL FILMS FOR THIS PATIENT ARE PRESENT. THERE IS A PROCESS IN PLACE SO THAT THIS WILL NOT HAPPEN. FILMS ARE TO BE CHECKED DAILY TO BE SURE THEY HAVE TRANSFERRED OVER TO MCKESSON. ANY MISSING FILMS DUE TO ERROR MESSAGES SHOULD BE ADDRESSED IN REAL TIME. CATH LAB ASSISTANT NURSE MANAGER HUDDLED WITH TEAM MEMBERS REGARDING THIS PROCESS AND TO BE SURE THAT THEY ARE FOLLOWING IT DAILY AND WITH ALL CASES. | -error message -patient experienced adverse event -fail to associate patient | -Device displays error message -crack -failure to read input signal | Viewer
| FDA | Event Description: VITAL IMAGES HAS BECOME AWARE OF A CHANGE IN HOW SOME CANON MEDICAL SYSTEMS (FORMERLY TOSHIBA MEDICAL SYSTEMS) SCANNERS HANDLE DICOM INFORMATION. THIS CHANGE MAY AFFECT THE ACCURACY OF PERFUSION RESULTS IN OUR 4D BRAIN PERFUSION SOFTWARE. VITAL IMAGES IS WORKING WITH CANON MEDICAL SYSTEMS TO BETTER UNDERSTAND THIS ISSUE AND DETERMINE ITS POTENTIAL IMPACT. TO DATE, THERE HAVE BEEN NO COMPLAINTS FILED AT VITAL ASSOCIATED WITH THIS ISSUE. Manufacturer Narrative: THIS INVESTIGATION IS ONGOING. | -inadequate image manipulation -loss of image -data corrupted | -Pressure Issue -Human-Device Interface Issue -Inflation Issue | Viewer
| FDA | Event Description: WE ARE REPORTING THIS MDR REPORTABLE EVENT ON DECEMBER 22, 2016. ON (B)(6) 2016, THE CUSTOMER REPORTED TO AGFA THAT WHEN USING MPAX CV REPORTING FOR NON-INVASIVE VASCULAR ULTRASOUND (NIV), WHEN BUILDING A NIV REPORT, IN THE CONCLUSION SECTION OF THE UPPER ARTERIAL DUPLEX, IF YOU SELECT ANY CHECK BOXES UNDER THE LEFT ANEURYSM IT PUTS IN THE REPORT THAT THERE WAS NO ANEURYSM. THIS REPORT IS FOR THE 3ND OF THREE EVENTS FOR THE SAME ISSUE AT THIS SITE. THE FIRST EVENT IS REPORTED IN FDA MDR 1225058-2016-00001. THE SECOND EVENT IS REPORTED IN FDA MDR 1225058-2016-00002. ON (B)(6) 2016, AGFA TECHNICAL SUPPORT CONTACTED THE CUSTOMER AND UPDATED HIM THAT ONLY THREE (3) REPORTS WERE AFFECTED AND THOSE THREE (3) REPORTS WERE ALL LISTED AS TEST REPORTS. THE CUSTOMER WAS PROVIDED THE LIST OF THE AFFECTED REPORTS. AGFA TECHNICAL SUPPORT RE-CONFIGURED THE REPORTS TO PRODUCE CORRECT STATEMENTS IN BOTH THE TEST AND PRODUCTION ENVIRONMENT. A PROBLEM RECORD WAS CREATED ON NOVEMBER 23, 2016 TO ADDRESS THE ISSUE VIA (B)(4). ON DECEMBER 22, 2016 AGFA HAS IMPLEMENTED A CORRECTION VIA AGFA ID - (B)(4) (1225058-12-22-2016-001-C). AGFA SENT URGENT SAFETY NOTICES TO AFFECTED CONSIGNEES FOR THIS ISSUE DESCRIBED IN (B)(4) (1225058-12-22-2016-001-C) ON DECEMBER 22, 2016. THE LETTER INFORMED THE CONSIGNEES THAT AGFA SERVICES WILL CONTACT THE AFFECTED CONSIGNEES TO: EXECUTE A QUERY TO FIND ALL AFFECTED REPORTS AT THE AND TO PROVIDE THE RESULTS. THE QUERY WILL NOT REQUIRE DOWNTIME AND WILL NOT IMPACT SYSTEM PERFORMANCE. CORRECT THE NIV REPORTING CLINICAL CONTENT. THE CORRECTION ACTIVITY WILL NOT REQUIRE DOWNTIME. THE LETTER ALSO INFORMED CUSTOMERS TO CONTACT AGFA IF THEY HAD ANY QUESTIONS RELATED TO THE COMMUNICATION LETTER. A MANDATORY SERVICE BULLETIN WILL BE RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR THIS ISSUE WITH IMPAX CV REPORTING. ANY ADDITIONAL REPORTING FOR THIS CORRECTION ACTIVITY WILL BE REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA AGFA ID - (B)(4) (1225058-12-22-2016-001-C). THERE HAVE BEEN NO REPORTS OF PHYSICAL HARM OR HEALTH DAMAGE RESULTING FROM THIS ISSUE. THIS MEDICAL DEVICE REPORT INCLUDES 1 REPORT AFFECTED ON (B)(6) 2016, WHEN THE ERROR OCCURRED. REPORT ID DATE ERROR OCCURRED (B)(4) (B)(6) 2016. | -device malfunction -unable to connect -inappropriate interface design | -use of device issue -Connection Issue -detachment of device component | Viewer
| FDA | Event Description: WE ARE REPORTING THIS MDR REPORTABLE EVENT ON DECEMBER 22, 2016. ON NOVEMBER 22, 2016, THE CUSTOMER REPORTED TO AGFA THAT WHEN USING MPAX CV REPORTING FOR NON-INVASIVE VASCULAR ULTRASOUND (NIV), WHEN BUILDING A NIV REPORT, IN THE CONCLUSION SECTION OF THE UPPER ARTERIAL DUPLEX, IF YOU SELECT ANY CHECK BOXES UNDER THE LEFT ANEURYSM IT PUTS IN THE REPORT THAT THERE WAS NO ANEURYSM. THIS REPORT IS FOR THE 2ND OF THREE EVENTS FOR THE SAME ISSUE AT THIS SITE. THE FIRST EVENT IS REPORTED IN FDA MDR 1225058-2016-00001. ON NOVEMBER 22, 2016, AGFA TECHNICAL SUPPORT CONTACTED THE CUSTOMER AND UPDATED HIM THAT ONLY THREE (3) REPORTS WERE AFFECTED AND THOSE THREE (3) REPORTS WERE ALL LISTED AS TEST REPORTS. THE CUSTOMER WAS PROVIDED THE LIST OF THE AFFECTED REPORTS. AGFA TECHNICAL SUPPORT RE-CONFIGURED THE REPORTS TO PRODUCE CORRECT STATEMENTS IN BOTH THE TEST AND PRODUCTION ENVIRONMENT. A PROBLEM RECORD WAS CREATED ON NOVEMBER 23, 2016 TO ADDRESS THE ISSUE VIA (B)(4). ON DECEMBER 22, 2016 AGFA HAS IMPLEMENTED A CORRECTION VIA AGFA ID - (B)(4) (1225058-12-22-2016-001-C). AGFA SENT URGENT SAFETY NOTICES TO AFFECTED CONSIGNEES FOR THIS ISSUE DESCRIBED IN (B)(4) (1225058-12-22-2016-001-C) ON DECEMBER 22, 2016. THE LETTER INFORMED THE CONSIGNEES THAT AGFA SERVICES WILL CONTACT THE AFFECTED CONSIGNEES TO: EXECUTE A QUERY TO FIND ALL AFFECTED REPORTS AT THE AND TO PROVIDE THE RESULTS. THE QUERY WILL NOT REQUIRE DOWNTIME AND WILL NOT IMPACT SYSTEM PERFORMANCE. CORRECT THE NIV REPORTING CLINICAL CONTENT. THE CORRECTION ACTIVITY WILL NOT REQUIRE DOWNTIME. THE LETTER ALSO INFORMED CUSTOMERS TO CONTACT AGFA IF THEY HAD ANY QUESTIONS RELATED TO THE COMMUNICATION LETTER. A MANDATORY SERVICE BULLETIN WILL BE RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR THIS ISSUE WITH IMPAX CV REPORTING. ANY ADDITIONAL REPORTING FOR THIS CORRECTION ACTIVITY WILL BE REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA AGFA ID - (B)(4) (1225058-12-22-2016-001-C). THERE HAVE BEEN NO REPORTS OF PHYSICAL HARM OR HEALTH DAMAGE RESULTING FROM THIS ISSUE. THIS MEDICAL DEVICE REPORT INCLUDES 1 REPORT AFFECTED ON (B)(6) 2016, WHEN THE ERROR OCCURRED. REPORT ID DATE ERROR OCCURRED: (B)(4) (B)(6) 2016. | -inadequate device information -fail to generate bolus -incorrect device orientation | -use of device issue -Connection Issue -Protective Measure Issue | Viewer
| FDA | Event Description: WE ARE SUBMITTING THIS MEDICAL DEVICE REPORT TODAY DECEMBER 13, 2016. ON NOVEMBER 22, 2016, THE CUSTOMER REPORTED THAT WHILE USING AN NX WORKSTATION WITH SOFTWARE VERSION NX 3.0.8950 SOFTWARE AND SELECTING THE AFFECTED PATIENT/EXAM FROM CLOSED EXAMS, INITIALLY THE WRONG IMAGE WAS LINKED TO THE EXAM AND APPEARED. AFTER A SHORT TIME THE WRONG IMAGE WAS REPLACED BY THE CORRECT IMAGE, HOWEVER THE WRONG IMAGE WAS USED FOR TRANSMITTING TO PACS. TO CORRECT THIS ISSUE, THE USER BACKED UP ALL LOGS AND REBOOTED THE SYSTEM. THE CORRECT IMAGE WAS THEN TRANSFERRED TO PACS. DURING THE INVESTIGATION BY AGFA NOVEMBER 24, 2016 - DECEMBER 1, 2016, THE ROOT CAUSE WAS IDENTIFIED AS A MEMORY LEAK WHICH WAS INTRODUCED IN THE IMAGE PROCESSING SOFTWARE IN NX 3.0.8950. THE USER RECEIVED TWO WARNING MESSAGES OF POTENTIAL MALFUNCTIONING PRIOR TO THE ISSUE. THE USER DID NOT FOLLOW THE WARNING INSTRUCTIONS AND THIS RESULTED IN THE IMAGE OF A PATIENT BEING ARCHIVED IN ANOTHER PATIENT EXAM RECORD. A RESTART OF THE SYSTEM AT THAT TIME WOULD HAVE PREVENTED THE MALFUNCTION FROM HAPPENING. AS A PREVENTIVE ACTION, AGFA HAS RECOMMENDED THE FOLLOWING TO THE CUSTOMER: RESTART THE SYSTEM ON A MORE REGULAR BASIS. A DAILY RESTART WOULD CERTAINLY BE SUFFICIENT TO PREVENT MEMORY ISSUES. THE SYSTEM IS CONFIGURED TO DISPLAY IMAGES BY DEFAULT IN PRINT MODE ALTHOUGH PRINT AUTOMATION IS NOT ACTIVE. RECONFIGURING THE SYSTEM TO DISPLAY IMAGES IN SOFTCOPY MODE FURTHER REDUCES THE IMPACT OF THE MEMORY LEAK ON MEMORY CONSUMPTION. AGFA SOLUTION DEVELOPMENT IS UNDERWAY AND A CORRECTION WILL BE IMPLEMENTED FOR THIS ISSUE. THE MEMORY LEAK IS RESOLVED IN NX 3.0.9000. A SUPPLEMENTAL REPORT WILL BE PROVIDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS EVENT.Event Description: ON DECEMBER 22, 2016 AGFA HAS IMPLEMENTED A CORRECTION VIA AGFA (B)(4) ((B)(4)) AND AN URGENT SAFETY NOTICE WAS SENT TO AFFECTED CONSIGNEES FOR THIS ISSUE. ON JANUARY 26, 2017, A MANDATORY SERVICE BULLETIN WAS RELEASED TO PROVIDE DOCUMENTATION ON HOW TO DEPLOY THE CORRECTION FOR THIS ISSUE. ANY ADDITIONAL REPORTING FOR THIS CORRECTION ACTIVITY WILL BE REPORTED TO FDA VIA 21 CFR 806 REPORTING VIA AGFA ID - (B)(4) ((B)(4)). THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS EVENT. | -software version incorrect -lack of software version -incompatible software versions | -Connection Issue -Protective Measure Issue -inadequate instructions for healthcare professional | Viewer
| FDA | Event Description: WHEN AN FCR IMAGE WAS FORWARDED FROM SYNAPSE PACS TO ANOTHER MANUFACTURER'S PACS, THE IMAGE FOR PATIENT A WAS ALSO FOUND IN PATIENT B'S STUDY IN THE DESTINATION PACS. (PATIENT B'S ORIGINAL IMAGE WAS NO LONGER IN THE STUDY). PATIENT A'S STUDY WAS CORRECT IN THE DESTINATION PACS. BOTH PATIENT'S STUDIES REMAINED CORRECT IN SYNAPSE PACS. NO INJURY WAS REPORTED; DIAGNOSIS COMPLETED USING SYNAPSE PACS, PRIOR TO SENDING. Manufacturer Narrative: THIS ISSUE WAS REVIEWED FOR MDR REPORTABILITY WHEN IT WAS FIRST RECEIVED. BECAUSE THE IMAGES RETAINED IN SYNAPSE PACS WERE STILL CORRECT AND DIAGNOSIS WAS PERFORMED IN SYNAPSE, FMSU DETERMINED THAT THIS WAS A MALFUNCTION THAT WOULD NOT BE LIKELY TO RESULT IN SERIOUS INJURY IN DEATH. MOVING FCR IMAGES TO ANOTHER SYSTEM IS TYPICALLY DONE FOR STORAGE OR TO SEND IMAGES AND REPORT TO A REFERRING PHYSICIAN. IN THE EVENT IMAGE PIXEL DATA WAS IMPROPERLY WRITTEN TO A PATIENT, THE REFERRING PHYSICIAN WOULD BE ABLE TO LIKELY IDENTIFY THE ISSUE WHEN REVIEWING THE RADIOLOGIST'S REPORT. (B)(4) LATER IDENTIFIED THAT THERE ARE SITES WHERE IMAGES ARE SENT TO OTHER PACS FOR DIAGNOSIS (FOR EXAMPLE FROM REMOTE LOCATIONS WHERE A RADIOLOGIST IS NOT AVAILABLE). IN THIS CASE A RADIOLOGIST MAY NOT HAVE ANY INDICATION THAT THE PATIENT IMAGE IS AN INCORRECT ONE, E.G., IF PATIENT A AND PATIENT B IMAGES WERE OF THE SAME ANATOMY. | -loss of image -image not in dicom format -maintenance does not comply to manufacturers recommendations | -Device Handling Issue -fitting problem -Contamination During Use | Viewer
| FDA | Event Description: WHEN ENTERING THE CORRECT ID NUMBER OF A PATIENT WITH A SLASH, X-RAY IMAGES ARE ASSIGNED TO A DIFFERENT PATIENT. IF THE LABORATORY TECHNICIAN DOES NOT NOTICE THIS IMMEDIATELY, THE X-RAY IMAGES RETURN TO THE PACS AS A STUDY WITH THE WRONG NAME AND BIRTH-DATE NUMBER. THIS ISSUE EXISTS ONLY WHEN A NEW PATIENT WITH A SLASH IS ENTERED IN THE PATIENT ID IN THE DATABASE AND STORED FOR USE LATER (E.G: XXX/XXX/XXX,XXX.XXX). LATER WHEN THE USER DESIRES TO OPEN A NEW STUDY AND SEARCH BY NAME, THE NAME CAN BE SEEN BUT WHEN IT IS RETRIEVED EITHER NOTHING IS RETRIEVED OR A DIFFERENT PATIENT IS RETRIEVED. THE PATIENT LIST USES A "/" TO SPLIT NAMES AND WHEN THE NAME ITSELF CONTAINS A "/", THE SOFTWARE FAILS TO SPLIT THE PATIENT NAMES. Manufacturer Narrative: CARESTREAM PERFORMED AN INVESTIGATION AND IDENTIFIED THE ROOT CAUSE ANALYSIS IN CAPA (B)(4). TO MITIGATE RISK TO PATIENT FROM USE OF THE DEVICE, A SOFTWARE DESIGN CHANGE WILL BE IMPLEMENTED IN IMAGE SUITE SOFTWARE VMR8 AND RELEASED IN MANUFACTURING. NO FURTHER REPORTING IS REQUIRED. | -incorrect patient information -incorrect model number -image orientation incorrect | -improper or incorrect procedure or method -failure of device to self-test -compliance issue | Viewer
| FDA | Event Description: WHEN THE PATIENT UNDERWENT GENERAL ANESTHESIA DURING UNSPECIFIED PROCEDURE, THE USER FACILITY FOUND A VIDEOSCOPE IMAGE ON THE MONITOR BECAME DARK EVEN THOUGH MENU DISPLAY WAS SHOWN ON THE MONITOR PROPERLY. THE USER FACILITY REPLACED THE VIDEOSCOPE TO ANOTHER VIDEOSCOPE, BUT THE PHENOMENON STILL OCCURRED. AFTER THAT, THE USER FACILITY REPLACED THE SUBJECT DEVICE TO ANOTHER DEVICE AND THE PHENOMENON DID NOT OCCUR. THE USER FACILITY COMPLETED THE PROCEDURE USING THE REPLACED DEVICE. THERE WAS NO REPORT OF PATIENT INJURY IN THIS EVENT. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. OLYMPUS CONFIRMED THE VIDEOSCOPE LOCKING MECHANISM OF THE VIDEO CONNECTOR BROKE AND A VIDEOSCOPE WAS CONNECTED TO THE SUBJECT DEVICE UNSTABLY. THE USER FACILITY MIGHT DISCONNECT A VIDEOSCOPE FROM THE SUBJECT DEVICE WITHOUT PRESSING THE LOCKING LEVER. THEREFORE, THE SPRING WHICH COULD KEEP THE VIDEOSCOPE CONNECTION LOCKED BROKE AND THE VIDEOSCOPE COULD NOT CONNECT THE SUBJECT DEVICE FIRMLY. OLYMPUS ALSO CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE INSTRUCTION MANUAL OF THIS DEVICE ALREADY MENTIONS THAT AN OPERATOR SHOULD PUSH THE LOCKING LEVER DOWN TO DISCONNECT THE VIDEO PLUG OF THE VIDEOSCOPE FROM THIS DEVICE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | -dark patches -loss of image -user confused | -incorrect display -No Display or Display Failure -no display or display failure | Viewer
| FDA | Event Description: WHILE EVALUATING A REPORTED ISSUE OF A SINGLE INSTANCE OF DATA LOAD FAILURE, SIEMENS FACTORY AND SERVICE EXPERTS FOUND A DATA LOSS OF 1858 CLINICAL IMAGE STUDIES OVER A TIME PERIOD OF 2012 TO 2016 (ENTIRE DIRECTORIES MISSING BETWEEN JUNE 2015 AND DECEMBER 2016 WITH A FEW OUTLIERS MISSING FROM 2012-2014). THE EXACT CAUSE OF THE DATA LOSS COULD NOT BE DETERMINED. THE HISTORICAL STUDY IMAGES WERE NOT FOUND IN THE PERMANENT ARCHIVE, THEREFORE, THEY ARE CONSIDERED TO BE UNRECOVERABLE AT THIS POINT IN TIME. THE CLINICAL RELEVANCE IS THAT THE HISTORICAL IMAGES ARE NOT AVAILABLE FOR COMPARISON WITH NEWER IMAGES HOWEVER THE CLINICAL STUDY REPORTS GENERATED FROM THE READ OF THE IMAGE STUDIES ARE AVAILABLE. THERE ARE NO INJURIES ATTRIBUTED TO THIS EVENT. Manufacturer Narrative: PER THE POTENTIAL COMPLAINT REPORT, THERE WAS NO INJURY ASSOCIATED WITH THIS ISSUE. NO DATA MIX-UP OR LOSS OCCURRED WHICH WOULD RESULT IN THE NEED FOR A PATIENT RESCAN. WHILE THE SYNGO DYNAMICS PRODUCT CANNOT CAUSE THE ISSUE FOUND, EXAMPLES OF EVENTS WHICH COULD HAVE CAUSED THE LOSS OF MULTIPLE CACHE DIRECTORIES INCLUDE A RAID FAILURE, A DATABASE MIGRATION FROM ONE SERVER TO ANOTHER WITHOUT MOVING THE ACTUAL CACHE CONTENTS, SOME TYPE OF DATABASE MIGRATION THAT POPULATED THE DATABASE WITH DIRECTORIES AND FILES THAT DON'T ACTUALLY EXIST, OR FREEING UP SERVER SPACE IN AN INCORRECT MANNER. SIEMENS LOCAL SERVICE ORGANIZATION JOINTLY WITH THE CUSTOMER IT GROUP INVESTIGATED TO SEE IF ANY OF THE ABOVE ACTIVITIES OCCURRED AT THIS SIT, HOWEVER, THE RESULTS OF NONE DOCUMENTED. THE SYNGO DYNAMICS PRODUCT DID NOT CAUSE OR CONTRIBUTE TO THE LOSS OF DATA. THERE ARE NO REMEDIAL OR CORRECTIVE ACTIONS REQUIRED FROM SIEMENS HEALTHCARE. | -loss of data -loss of image -data corrupted | -data issue -incorrect or inadequate result -application security issue | Viewer
| FDA | Event Description: WHILE PERFORMING PROACTIVE MAINTENANCE OF UNARCHIVED STUDIES, MCKESSON IDENTIFIED 3 UNREPORTED STUDIES FROM THE SHORT TERM STORAGE IN THE HMI PACS THAT FAILED TO ARCHIVE TO LONG TERM STORAGE. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: DURING INTERNAL REVIEW, MCKESSON DETERMINED THAT THIS ISSUE IS SIMILAR TO 8022257-2013-00004, AND IS REPORTABLE UNDER PER 21 CFR 803. MCKESSON HAS IDENTIFIED AN ISSUE WHERE APPLICATION LOGIC LEADS TO DISCREPANCIES BETWEEN THE IMAGE INDEX AND THE CONTENT OF THE IMAGE FOLDER WHICH CAN CAUSE IMAGE LOSS. THE IMAGE FOLDER IS A FILE SYSTEM DIRECTORY WHERE IMAGES ARE STORED; IT ALSO CONTAINS THE IMAGE INDEX. THE IMAGE INDEX INDICATES TO THE APPLICATION A LIST OF IMAGE FILES TO EXPECT FOR USE WITHIN THE APPLICATION. IF THE IMAGE INDEX AND THE CONTENTS OF THE IMAGE FOLDER ARE NOT SYNCHRONIZED, THE FOLLOWING MAY OCCUR: IMAGE MAY BE LOST OR UNAVAILABLE, LEADING TO CLINICAL DECISIONS BEING MADE ON INCOMPLETE STUDIES AND/OR THE STUDY CANNOT BE ARCHIVED. IN THE EVENT OF A HARDWARE FAILURE, UNARCHIVED STUDIES MAY BE LOST LEADING TO PRIOR STUDIES NOT BEING AVAILABLE FOR COMPARISON. THE FOLLOWING FACTORS MAY CONTRIBUTE TO THE OCCURRENCE OF THE DISCREPANCY: NETWORK STABILITY ISSUES (FOR EXAMPLE: POOR PERFORMANCE, INSTABILITY AND COMMUNICATION ERRORS (DROPPED CONNECTIONS, PACKET LOSS)); DISK WRITE FAILURES (CANNOT WRITE TO DISK). THE PRODUCT MAY FAIL TO HANDLE THE OUTCOME OF SUCH FACTORS RESULTING IN THE DISCREPANCIES DESCRIBED ABOVE. | -fail to record data -fail to activate test -fail to analyze signal | -leak -blockage within device or device component -medical gas supply problem | Viewer
| FDA | Event Description: X-RAY IMAGES WERE ADDITIONALLY LINKED TO WRONG PATIENT RECORD. Manufacturer Narrative: THE INVESTIGATION FROM THE R&D DEPARTMENT HAS RESULTED: SIDEXIS 4 IS USED IN COMBINATION WITH SIDICOM WLS AND A CENTRAL PACS SYSTEM. SIDICOM WLS IS CONFIGURATED TO SEND AUTOMATICALLY THE IMAGES TO THE PACS SYSTEM AFTER THE ACQUISITION. MULTIPLE SIDEXIS 4 USERS AT ONE WINDOWS PC'S WITH INDIVIDUAL USER ACCOUNTS LOGGED-IN, IN PARALLEL. AFTER AN EXAMINATION THE USER DOES NO LOGOUT OF HIS WINDOWS SESSION AND A DIFFERENT USER STARTS WITH HIS ADDITIONAL LOG-IN. ADDITIONALLY, SIDEXIS 4 IS NOT CLOSED FROM THE PREVIOUS USER. IF THE ABOVE LISTED CONDITIONS ARE GIVEN THE ISSUE DESCRIBED WILL OCCUR. THE IMMEDIATE ACTION IS TO ADD AN ADDITIONAL INSTRUCTION INTO THE SIDICOM WLS PACKAGES THAT ARE SHIPPED. FURTHER CAPA ACTION IS ONGOING. | -incorrect patient information -inadequate image types -loss of image | -use of device issue -user used incorrect product for intended use -device operates differently than expected | Viewer
| FDA | Fetal patient report was automatically pulling prior measurement data for a prior fetus since the fetal study is based on the mother's MRN, resulting in the wrong fetal measurements getting referenced. | -fail to create related measurement -loss of measurement -inaccurate measurement item | -incorrect measurement -imprecision -break | Viewer
| FDA | Following update from 3.3 to 3.5 of the DynaCAD software it was noted that the Ktrans map was not rendered correctly on the DynaCad Client. The defect causes the pharmacokinetic (PK) color maps to display incorrectly when viewed from remote DynaCAD client computers and could result in visually underestimating calculated Ktrans, Kep, and iAUGC values. The defect also impacts DynaCAD s on-the-fly calculation of Apparent Diffusion Coefficient (ADC) maps. If the ADC values are computed on-the-fly by DynaCAD, the ADC values and colors will also display incorrectly if viewed on a remote DynaCAD client. | -inappropriate color code -missing color code -fail to set value correctly | -device operates differently than expected -use of device issue -difficult to open or close | Viewer
| FDA | For sites using the Issuer of Patient ID (IPID), the system will display the study list and images from different patients with different IPIDs as though they are for the same patient if they all have the first name, last name, and medical record number in common. | -inadequate patient data -incorrect patient information -fail to enter patient data | -Image Resolution Poor -incorrect display -image display error | Viewer
| FDA | FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected. | -incorrect 3d model -incomplete 3d model -adjustment out of specification | -therapeutic or diagnostic output failure -use of device issue -improper or incorrect procedure or method | Viewer
| FDA | Functionality in the report sections "Findings Information" and "Summary of Measured Findings" not functioning properly. The corrected values modified in the report are neither saved, nor printed, nor sent to the information system. The system will still use original values. | -incorrect measurement -inadequate ifu -incorrect readings | -imprecision -Inaccurate Flowrate -leak | Viewer
| FDA | GE Healthcare is conducting a recall due to the potential fall of aged CRT monitors - Prestige 2, Prestilix, Prestige VH, Prestige SI, RFX/SFX, Legacy, Precision 500D, ADVANTX and Innova 2000 systems. | -ageing -loss of function -poor telemetry | -incorrect measurement -user used incorrect product for intended use -use of device issue | Viewer
| FDA | If a reader selects Multi-study review prior to the original study completing its loading process and then immediately returns to image review of the original study, it is possible that not all current images will be present for review, which could result in incorrect treatment/diagnosis. | -inappropriate program selected -lack of review function -unable to review loop | -device cleaning issue -unstable -device handling issue | Viewer
| FDA | If images are sent without an order in the system, they will be matched with the latest order on the current patient/modality matching potentially resulting in the matching of the report to the incorrect accession number (but still associated to the correct patient). | -fail to display message -fail to generate notification -fail to follow plan | -device markings issue -no display or display failure -application program issue | Viewer
| FDA | Image data for a patients image may not be correct | -loss of image -data corrupted -inadequate image manipulation | -data issue -pacing inadequately -application security issue | Viewer
| FDA | Images may potentially become corrupt while using the system. | -system abused -software corrupted -data corrupted | -Shelf Life Exceeded -pacing inadequately -malposition of device | Viewer
| FDA | Inaccuracy of MR thermometry during MRI-guided laser ablation procedures using the Medtronic Visualase Thermal Therapy System may result in unaccounted spread of thermal energy to the surrounding tissue. | -poor thermal expansion -thermal event -incorrect measurement | -incorrect measurement -electromagnetic interference (emi) -optical issue | Viewer
| FDA | Incorrect values for the volume calculation. Software update VB30B via Update Instructions SY018/16/P to resolve software errors. | -incorrect volume calculation -error message -fail to scan model volume | -use of device issue -failure to service -lack of maintenance documentation or guidelines | Viewer
| FDA | Interventional Radiology (IR) images are stored as JPEG2k Lossless in Merge Enterprise Archive (EA) and are not displaying correctly in RadSuite. | -inadequate image types -loss of image -inadequate image manipulation | -therapeutic or diagnostic output failure -patient data issue -solder joint failure | Viewer
| FDA | Issue for customers that use an EMR login or legacy web URL login or legacy web URL login for McKesson Radiology PACS that may result in missing images in a newly imported study, and/or study imports that remain in an "in-progress" status. | -missing data -missing maintenance -fail to transfer data | -use of device issue -human-device interface issue -device alarm system issue | Viewer
| FDA | It was reported by a customer that RadSuite images are not appearing as readable images. This product malfunction has the potential to result in delay in care and a possible re-exposure of the patient to radiation (x-ray). | -loss of image -inadequate image manipulation -unreadable display | -fail-safe design failure -device stops intermittently -break | Viewer
| FDA | It was reported from a complaint that study category labels of "Prior" and "Baseline" and hanging position appeared incorrectly swapped when the hanging protocol for three comparison studies is manually applied from the patient folder. | -labeling misaligned -incorrect labeling -inadequate labeling | -human-device interface issue -No Known Device Problem -failure to sense | Viewer
| FDA | Lateromedial (LM) and Lateromedial Oblique (LMO) Orientation Markers may be displayed incorrectly on the 3D TOMO slider bar. In addition, the slice location of the image and/or the slice direction could be incorrect as well, and result in misdiagnosis | -image orientation incorrect -incorrect display format -incorrect 3d model | -application program issue -device sensing issue -no display or display failure | Viewer
| FDA | Measurements done on the Merge PACS generated MPR s may have incorrect measurements. | -incorrect measurement -fail to measure -inaccurate measurement | -application program issue -imprecision -use of device issue | Viewer
| FDA | Measurements in the software are changing after saving a plan with a measurement result, which may result in potential patient injury or delay in diagnosis or treatment. | -loss of measurement -fail to measure -incorrect measurement | -user used incorrect product for intended use -therapeutic or diagnostic output failure -defective item | Viewer
| FDA | Merge PACS did not show unviewed images when the last view was skipped. There is a potential risk to health of a physician misdiagnosis because not all images are available for viewing. | -loss of image -fail to interpret images -unable to view display | -unstable -device alarm system issue -device unsafe to use in environment | Viewer
| FDA | Multiple studies were missing at a customer location that never made it to the archive before watermarking | -missing maintenance -missing filter -part missing | -component missing -device maintenance issue -device markings issue | Viewer
| FDA | Novarad Corporation announces a voluntary field action for the NovaPACS versions 8.4.7, 8.5.3, and 8.5.6 due to a "pop-up screen" that was prematurely released. | -release delayed -fail to pop-up -inappropriate release | -defective component -device operates differently than expected -Application Program Issue | Viewer
| FDA | Observed an internal repetitive software glitch in ImmersiveView. | -software corrupted -software corruption -software task fault displayed | -material disintegration -device handling problem -Device displays error message | Viewer
| FDA | Ortho Kinematics Inc. sent a Notice of Correction to Released Testing Results, Radiological Read Report for the Vertebral Motion Analyzer (VMA) test because it contained an error. The error occurred due to a software bug that has been corrected. | -incorrect readings -fail to disclose error message -error message | -Communication or Transmission Issue -Device displays error message -Incorrect Measurement | Viewer
| FDA | Ortho Kinematics notified customers that errors were contained in Radiological Read Report and VMA Report for the VMA test. The error resulted from an administrative error of VMA user account settings and has been fixed. | -incorrect readings -error message -inappropriate symbol | -Improper or Incorrect Procedure or Method -no display or display failure -Defective Component | Viewer
| FDA | Ortho Kinematics, Inc. notified customers of erroneous results of VMA tests due to a software bug that has been corrected. | -software defect -fail to monitor feedback -error message | -output issue -device maintenance issue -operating system version or upgrade problem | Viewer
| FDA | Panasonic has issued a Mandatory Product Safety Notice, indicating a possibility of battery ignition (fire) in CF-C2 "Toughbook" laptops. Since Konica Minolta uses this laptop in some of its product configurations, it notified its customers to utilize a BIOS Utility for the laptops to minimize the risk of an incident while Panasonic develops a final solution. | -fail to indicate battery charging status -ignition -inadequate safety information | -market issue -improper or incorrect procedure or method -compliance issue | Viewer
| FDA | Patient information in the header is only found on the first page of the report and not on the subsequent pages. | -missing information -loss of patient data -incorrect message display | -failure to read input signal -patient data issue -no display or display failure | Viewer
| FDA | Potential data loss occurs as a result of product archiving not working properly. | -loss of data -fail to record data -data corrupted | -patient data issue -environmental control or utility issue -market issue | Viewer
| FDA | Potential data loss occurs as a result of the software archiving not working properly. | -loss of data -data corrupted -incorrect data | -patient data issue -Decreased Sensitivity -not audible alarm | Viewer
| FDA | Potential exists for an incorrect patient image being displayed which could result in the delay in diagnosis or treatment. | -image orientation incorrect -incorrect display -incorrect message display | -application program issue -no display or display failure -failure to align | Viewer
| FDA | Potential for images to be flipped while streaming, which could incorrectly display image orientation markers. | -image orientation incorrect -fail to interpret images -inappropriate image format | -poor quality image -application program issue -chemical issue | Viewer
| FDA | Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). | -unable to scroll image -image orientation incorrect -incorrect display format | -use of device issue -dose calculation error due to software problem -not audible alarm | Viewer
| FDA | Potential for RadSuite AV viewer to skip image slices when the user presses page up or page down when the viewer displays multiple viewports and the user has mouse focused on any viewport other than the first viewport. | -skip failed -fail to toggle image visibility -fail to read images | -use of device issue -device alarm system issue -not audible alarm | Viewer
| FDA | Potential that one or more image series may be missing from an exam without a user warning displayed in the Viewer. | -fail to verify image -missing symbol -loss of warning | -no display or display failure -device alarm system issue -device markings issue | Viewer
| FDA | Remote comparison studies displayed may be associated with the wrong patient, without a caution icon present. | -image orientation incorrect -missing symbol -fail to generate indication | -device markings issue -inadequate instructions for healthcare professional -no display or display failure | Viewer
| FDA | Reporting feature times out after inactivity for more than an hour sending the user back to the study list, which causes all reporting data being entered to be lost. | -incorrect time configuration -lack of time selection -lack of time indication | -use of device issue -use of incorrect control settings -device difficult to setup or prepare | Viewer
| FDA | Reports were missing demographic header information when they were faxed to referring physicians. | -missing information -incorrect patient information -missing data | -component missing -device markings issue -inadequate instructions for healthcare professional | Viewer
| FDA | Siemens has identified changes in default carry forward workflow using VA10B_HF03 version of Syngo Dynamics. If the template design has measurements inside a carry forward-configured report section, this will result in carrying forward the entire section of the report including data from previous studies. This may result in the physician making analysis based on the old measurements. | -fail to identify change -lack of change control -fail to measure | -optical issue -material disintegration -compliance issue | Viewer
| FDA | Siemens is releasing a Customer Advisory Notice to explain meaning of a suffix o.p. that shows on some DICOM images on area, distance or perimeter values. This suffix stands for on projection and indicates that a shown measurement value may be different from the real value. | -inadequate warning -incorrect display -lack of orientation | -application program issue -break -use of incorrect control settings | Viewer
| FDA | Siemens is releasing a letter to inform about potential data loss relevant to diagnosis. | -loss of patient data -loss of data -loss of message | -communication or transmission issue -application security issue -failure to sense | Viewer
| FDA | Siemens is releasing a non-medical software application LTA Incomplete Archive Check Tool V1.0. It is intended to eliminate an issue that occurs during syngo.plaza de-archiving from Dicom LTA. It has been defined that the number of de-archived images is less than the count of the archived images for the series. The issues were discovered through Siemens Global Complaint Handling System. | -incomplete software upgrade -fail to schedule check -software version incorrect | -use of device issue -device alarm system issue -deployment issue | Viewer
| FDA | Software changes now available to address several issues | -fail to address problem -fail to implement change -inappropriate change | -compliance issue -Mechanical Issue -fail-safe design failure | Viewer
| FDA | Software defect which may potentially result in one or more images missing in a study | -component missing -missing symbol -fail to identify defect | -defective item -device markings issue -size incorrect for patient | Viewer
| FDA | Software displayed incorrect prior reports in the viewport area, only when more than one prior study (2 or more) was viewed. | -incorrect display format -incorrect message display -incorrect display | -device expiration issue -inappropriate prompts -false alarm | Viewer
| FDA | Software error in previous software versions in which two references for the same image may exist in the database. | -software version incorrect -fail to interpret images -error message | -no pacing -failure to pace or properly pace -computer software issue | Viewer
| FDA | Software error. Incorrect values for the volume calculation from a freehand VOI at the customer site. In volume calculations of prostate as well as in volume calculations of liver were too high. | -incorrect volume calculation -fail to calculate volume -fail to set value correctly | -use of device issue -output below specifications -parameter calculation error due to software problem | Viewer
| FDA | Software issues. Siemens is voluntarily initiating a recall after they became aware of the following system behavior: 1) At times, the view and laterality marker is overlaid by patient demographic information, thus becoming difficult to read. 2) The facility name and address are not shown per default (in the current software versions this information is displayed only when the reader chooses an appropriate image text display mode). | -fail to specify patient population -device malfunction -fail to reattach | -no display or display failure -data back-up problem -human-device interface issue | Viewer
| FDA | Software update for improvements and to resolve several issues | -fail to update software -fail to update indicator -fail to update | -use of device issue -device contamination with biological material -device cleaning issue | Viewer
| FDA | Software update to correct several issues that include (1)Potential data loss, (2) Study mix-up, (3) Incorrect measurements on multi-frame images, (4) Dearchiving issue, and (5) Unauthorized access of data due to inadequate permissions for shared folders. | -incorrect measurement -unauthorized access -incorrect data | -unauthorized access to computer system -human-device interface issue -User used incorrect product for intended use | Viewer
| FDA | Software upgrade to correct format of study dates and issues with Legacy Presentation States (annotations) in order to prevent potential patient misdiagnosis. | -incorrect data format -software version incorrect -incorrect display format | -compliance issue -therapeutic or diagnostic output failure -use of device issue | Viewer
| FDA | Studies coming over via telmed were missing patient's DOB, procedure, and referring physician. | -missing review of pre-operative patient images -patient trigger failed -incorrect patient information | -component missing -device contamination with biological material -compliance issue | Viewer
| FDA | Studies that are viewed in iConnect Access that originate on Merge PACS / Ortho PACS may not display the correct current patient demographics if there has been a PDE (Patient / Study Demographic Edit) performed after the images were ingested in PACS. | -incorrect patient information -fail to enter patient data -fail to display data | -inadequate user interface -human-device interface issue -use of device issue | Viewer
| FDA | Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | -fail to update configuration -fail to update -fail to display data | -disconnection -improper flow or infusion -use of device issue | Viewer
| FDA | Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | -fail to update configuration -fail to update -fail to display data | -disconnection -improper flow or infusion -use of device issue | Viewer
| FDA | Synapse cannot display image files, DICOM SR files, and/or Annotation files. The "Image Not Loaded" message displays instead | -incorrect message display -inadequate message display -fail to display message | -incorrect display -difficult to interrogate -poor quality image | Viewer
| FDA | The bug causes BOLD activation maps to be visualized as overlays without taking the coregistration into account. Any output created from these activation maps as overlays will not be adjusted according to the coregistration. | -fail to adjust image brightness -fail to adjust size -fail to adjust structure opacity | -defective item -device difficult to setup or prepare -device cleaning issue | Viewer
| FDA | The error occurs in the following situations: A group of fibers has been selected using the VOI functionality and put into a fiber group. This is done with images series A visualized in the 3D viewer. Then a new images series B is visualized in the 3D viewer. If image series A and B have the same geometrical resolution (pixel sizes and slice distances), the coregistration matrix taking A to B will not be applied to the fiber group, and therefore the fiber group will in general not be positioned correctly on B in the 3D viewer. Furthermore, if the fiber group is exported as a new image series the fibers could be misplaced, and similarly if the fibers are visualized in the MPR on B, the positioning will not be correct. The misalignment will be equal to the rotations/translations necessary to align A with B. | -incorrect geometry -loss of image -image orientation incorrect | -improper or incorrect procedure or method -failure to sense -fitting problem | Viewer
| FDA | The firm discovered a software issue associated with the Image-Com module/clinical application package (CAP) interface for 3D application in the ARENA software, where measurements from one study might be stored to another study. | -software installation problem -loss of measurement -inappropriate measurement algorithm | -therapeutic or diagnostic output failure -material disintegration -human-device interface issue | Viewer
| FDA | The firm received a report of an atypical dataset being generated. When using the cross-localization feature and images from a modality that generates asymmetrical images, the cross-localization reference line may be inaccurately placed on any of the corresponding images that are open. | -unclear image -inappropriate image format -illegible data | -Human-Device Interface Issue -no display or display failure -filling problem | Viewer
| FDA | The Neuroreader has been distributed with a reference normative database that is different from the version reviewed and cleared by FDA. | -data corrupted -unauthorized access -inappropriate certificate scope | -application program version or upgrade problem -operating system version or upgrade problem -problem with software installation | Viewer
| FDA | The patient name in the Halo title bar and the thumbnails do not match the name on displayed images. | -incorrect patient information -incorrect message display -incorrect display format | -device markings issue -no display or display failure -incorrect display | Viewer
| FDA | The software did not show unviewed images when the last view was skipped. The physician may potentially misdiagnose when not all images are available for viewing. | -loss of image -unable to view display -fail to interpret images | -unstable -no display or display failure -alarm not visible | Viewer
| FDA | The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user. | -fail to update indicator -fail to update configuration -software version incorrect | -alarm not visible -patient data issue -not audible alarm | Viewer
| FDA | The software fails to associate to the correct MG image if there are two images for the same view. | -fail to interpret images -loss of image -fail to identify use error | -device alarm system issue -connection issue -poor quality image | Viewer
| FDA | The software is not adding the correct tomo image. | -image orientation incorrect -fail to support correct use -inappropriate image rendered | -use of device issue -pacing inadequately -reset issue | Viewer
| FDA | The software is not identifying the patient as having atypical hyperplasia, resulting in an incorrect Gail Risk calculation. | -incorrect patient information -incorrect ratio calculation -inadequate risk analysis | -improper or incorrect procedure or method -use of device issue -user used incorrect product for intended use | Viewer
| FDA | The software is unable to accurately determine the calcium score of scans with a slice thickness not equal to 3 mm. | -wrong scan condition -inaccurate measurement item -incorrect sample size | -unstable -failure to recalibrate -poor quality image | Viewer
| FDA | The software was not presenting the PURE filter user interface message on DICOM transfer studies. If a site starts using PURE filter, the user interface message will not display. This may result in incorrect patient follow-up. | -incorrect message display -inadequate message display -fail to display message | -failure to deploy -improper or incorrect procedure or method -reset issue | Viewer
| FDA | The STAT priority icon was not displayed in the Study List when the study was opened. | -fail to mark icon -incorrect screen display -unreadable display | -No Display or Display Failure -Air Leak -incorrect display | Viewer
| FDA | The SUV values that are being calculated in the PET/CT fusion tool are incorrect. | -incorrect value -incorrect measurement -incorrect data | -therapeutic or diagnostic output failure -failure to read input signal -device alarm system issue | Viewer
| FDA | The values provided from the Pixel Value tool do not appear to be correct, which may result in potential patient injury or delay in diagnosis or treatment. | -incorrect tool identification -incorrect value -incorrect display value | -use of device issue -user used incorrect product for intended use -incorrect display | Viewer
| FDA | There is a potential for the wrong patient information may be displayed in the viewer or PowerJacket. Additionally, the incorrect max standard uptake values (SUV) for PET ad CT studies may be returned for Siemens Modality. | -incorrect display value -incorrect display format -incorrect message display | -device markings issue -pacing inadequately -incorrect measurement | Viewer
| FDA | There is a potential that one or more image series (i.e., all images within an image set) may be missing from an exam without a warning displayed in the viewer. | -fail to verify image -loss of warning -missing symbol | -no display or display failure -device alarm system issue -human-device interface issue | Viewer
| FDA | There is a potential that one or more images or image series may be missing from exams without a warning displayed in the viewer. | -component missing -missing symbol -loss of warning | -no display or display failure -device alarm system issue -human-device interface issue | Viewer
| FDA | There is the possibility of viewing studies directly from the Enterprise Archive or VNA with incorrect patient images because the updated series or study is not present in the archive. | -incorrect patient information -inappropriate patient type -inadequate patient data | -parameter calculation error due to software problem -therapeutic or diagnostic output failure -electronic property issue | Viewer
| FDA | This email is to provide notification that, due to a software bug that has been corrected, the Radiological Read Report for the VMA test(s) contained an error. | -software defect -error message -incorrect readings | -Defective Component -break -application program issue | Viewer
| FDA | To inform users about the possible incorrect values for Distance Measurements when using certain modalities in combination with syngo Imaging. | -incorrect distance measurement -incorrect measurement -incorrect data | -imprecision -Communication or Transmission Issue -sticking | Viewer
| FDA | Users can merge a device import file with an image study that already has a confirmed report, which may result in including information not present at the time of physician interpretation. | -lack of confirmation -fail to save surgical plan -unable to switch | -device markings issue -failure of device to self-test -component missing | Viewer
| FDA | When a prior study is being replaced in the workflow step, in certain scenarios (based on the Display Protocol configuration) the prior study will only be replaced in the active Workflow Step and not show in all other workflow steps. The other workflow steps will continue to show the initially loaded study. | -fail to display previous data -fail to analyze new use error -lack of intended users | -user used incorrect product for intended use -incorrect display -no display or display failure | Viewer
| FDA | When an image is flipped horizontally and vertically, the orientation markers appear correctly on the image to which they were applied. However, once new image rendering happens on this image or series (with scrolling; playing cin¨; maximizing or changing viewports/layouts; or synchronizing viewports and resizing the browser), the new image orientation is correctly maintained but the display of orientation markers can become incorrect. | -image orientation incorrect -inappropriate software release -incorrect display format | -device sensing issue -therapeutic or diagnostic output failure -failure to read input signal | Viewer
| FDA | When cut lines are shown on an image, the lines may go from being vertical and then shift to horizontal causing confusion to the physician. This may cause a misdiagnosis. | -confused indication -loss of image -confused interface | -improper or incorrect procedure or method -use of device issue -device markings issue | Viewer
| FDA | When measuring a lesion on an unmagnified mammography image then performing the same measurement on an image magnified by the Hologic imager (different image), the measurements are not the same. | -incorrect measurement -inaccurate measurement -fail to interpret images | -application program issue -incorrect measurement -device alarm system issue | Viewer
| FDA | When RadSuite is used with IPID (Issuer of Patient ID) as a part of the "Patient Identifier," it is possible in some circumstances that the demographics of one patient will be applied to a study or studies for another patient. | -incorrect patient information -inappropriate patient type -incorrect identification | -instruction for use issue -defective item -no display or display failure | Viewer
| FDA | When switching back & forth between multiple UV instances in the Windows taskbar, the patient images displayed on the mammo high resolution monitors may not show images of the patient selected on UV, which could result in an incorrect diagnosis. | -switch back and forth -image orientation incorrect -inadequate message display | -no display or display failure -application program issue -Image Resolution Poor | Viewer
| FDA | When taking measurements from images on the Cardio workstation or from the US cart, numbers are not crossing to the report in the correct unit of measure. | -inadequate measure unit -incorrect measurement -fail to place measurements | -parameter calculation error due to software problem -incorrect measurement -use of device issue | Viewer
| FDA | When the archiving configuration is changed, data received/created after the upgrade may be flagged as "Not to be archived". This is caused by the automatic function for cleaning up temporary data being disabled by the software upgrade. Due to the disabled cleanup function, disc capacity for free space decreases faster than usual. Unless the archiving configuration and the cleanup automatic function are reverted back to the original state (prior to the software upgrade), data that was incorrectly flagged "Not to be archived" must be manually prevented from being deleted by either clinical administrator or service engineer. | -fail to update configuration -fail to extract data -fail to delete data | -failure to deliver -application security issue -patient data issue | Viewer
| FDA | While merging exams in a test instance, merges carry forward in the production system, when the middle tier is configured to the wrong IP address of the Centricity Exam Manager. When a current and historical exam are opened at the same time, the system asynchronously tries to access the operation specifying how each image should be grouped, causing random image-series grouping errors. | -fail to setup configuration -incorrect configuration -incorrect test procedure | -use of device issue -device operates differently than expected -failure of device to self-test | Viewer

## iPlatform
category | swRA UID | SW Design Element, SW Function, or SW Item (Name of the software Sub-system, Interface, Communication, Module, Item, Unit, SOUP component, or source code file) | Potential Software Failure Mode | Software Cause for Failure Mode (Refer to TAB 3, 4, and 5) | Potential Effect of the Software Failure mode on System | HH Map ID | Hazardous Situation / Reference to System Risk Analysis, if available or Hazards and Harms List | Software contribution to a hazardous situation (required by the standard IEC 62304:2015, Clauses 7.1.2, 7.1.3, and 7.1.4) | Probability of Hazardous Situation (P1 - initial) | Harm | Severity | P2 | Initial Risk | Description of Risk Control Measure
Lap-App | LapApp-swRA-00001 | LapApp - Startup | LapApp fails to correctly initialize software components and its internal modules or the comm layer are inoperable | Data corruption, incorrect SW installation | During initialization, LapApp shuts down, and can't command a power up signal to hw app and corresponding robot arm | HHM-0255 | Physician is unable to begin procedure on time | During app boot up, software components are not checked for completed / or passed initialization | Always - 100% | No harm | No Harm - (S0) | Always |  | 1. During boot up/Initiialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the application for successful startup
Lap-App | LapApp-swRA-00002 | LapApp - Configuration Data | LapApp receives configuration data from any of the multiple sources incorrectly or incompletely | Data corruption, incorrct SW installation, interface design errors | Incorrect or incomplete system configuration is used for motion commands, may result in incorrect or undesired motion calculation | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Configuration data is read incorrectly or incompletely from any of the mutiple sources | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Configuration data will be checked with CRC and values will be sanity checked against applicable limits 2. LapApp will not completely intialize until configuration information is fully received and hardware telemetry is fully received
Lap-App | LapApp-swRA-00003 | LapApp - Apply Configuration | LapApp fails to correctly apply configuration | incorrect data definition, interface design errors, errors in handling or interpreting data | Incorrect initialization of internal components, may result in incorrect or undesired motion calculation | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Configuration is used incorrectly to initialize internal components and models | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | During startup, the software will check that all internal components are initialized correctly and are of correct type/behavior
Lap-App | LapApp-swRA-00004 | LapApp - Apply Kinematics Calibration | LapApp fails to apply the kinematics calibration | Data corruption or pasring error for kinematics calibration information | Software fails to apply kinematic calibration and operates with reduced accuracy | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Using incorrect calibration, the app computes wrong motions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Configuration data will be checked with CRC and values will be sanity checked against applicable limits
Lap-App | LapApp-swRA-00005 | LapApp - Apply Kinematics Calibration | LapApp fails to apply the kinematics calibration correctly | Incorrect algorithm, incorrect data definition, interface design errors, errors in handling or interpreting data | Software fails to correctly apply kinematic calibration, can casue undesired motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Using incorrect calibration, the app computes wrong motions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Configuration data will be checked with CRC and values will be sanity checked against applicable limits
Lap-App | LapApp-swRA-00006 | LapApp - Does not recieve base pose | LapApp fails to receive the base pose | networking issues preventing communication | Software fails to correct place robotic manipulator in the world coordinate frame, resulting in unintendned motion during active control | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Using incorrect calibration, the app computes wrong motions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | LapApp requires base pose to be received before initializing
Lap-App | LapApp-swRA-00007 | LapApp - Powerup Robot | LapApp fails to send the power up signal to the hw application for the corresponding robot arm | Incorrect algorithm, control flow errors | Robotic arm inoperable due to sw failure to command it to power up | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | During app startup, software fails to transition to the powerup state or fails to send the powerup signal to the hardware | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | Ensure state change to powerup as well as receipt of powerup acknowledgement from hw in a timely manner
Lap-App | LapApp-swRA-00008 | LapApp - State Transitions | LapApp fails to transition to appropriate states, or transitions to wrong state(s) when commanded | Failure to set the correct initial and later states, control flow errors | Robotic arm not moving in response to commanded motions or moving in an unintended way | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to compute or computes incorrect motions due to being in an incorrect state | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Unhandled state transitions lead to a minor fault
Lap-App | LapApp-swRA-00009 | LapApp - Receive Cannula Latch Detection Signal | LapApp fails to receive the cannula latch detection signal | Data corruption, interface design errors | Software fails to transition to docked state, Cannot load an instrument | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to recognize the docked state and does not respect RCM position hold, and cannula shifts as a result | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | 1. Fault when tool attached but not cannula 2. Appropriate GUI and sound indicators when cannula is latched/unlatched
Lap-App | LapApp-swRA-00010 | LapApp - Receive Instrument RFID Information | LapApp fails to receive instrument RFID information | Data Corruption, interface design errors | Cannot load an instrument | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to recognize the attached instrument and does not transition accordingly | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | 1. Fault if tool not loaded within timeout since attaching 2. Appropriate GUI and sound indicators when instrument is attached and homed
Lap-App | LapApp-swRA-00011 | LapApp - Receive Instrument RFID Information | LapApp receives incorrect instrument RFID information | Data Corruption, incorrect data definition, interface design errors, errors in handling or interpreting data | Wrong instrument is loaded or homed | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to recognize or home the correct instrument | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Sanity checks on received RFID information, including CRC 2. Fault if not loaded within timeout since attaching 3. Appropriate GUI and sound indicators when instrument is loaded/homed
Lap-App | LapApp-swRA-00012 | LapApp - Recieve Hardware Telemetry | LapApp fails to receive hardware telemetry | Data Corruption, interface design errors | Software fails to initialize or update its inernal states | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software failure to initialize its internal components leads to its inability to control the robot | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Timeout faults for failing to receive hardware telemetry
Lap-App | LapApp-swRA-00013 | LapApp - Apply Hardware Telemetry | LapApp receives hardware telemetry incorrectly | Incorrect data definition, interface design errors, errors in handling or interpreting data | Software initializes or updates its inernal states incorrectly | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software failure to initialize its internal components incorrectly leads to it commanding bad motions to the robot | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Timeout faults for failing to receiving hardware telemetry 2. Collision verification on commanded joints 3. Joint limit checks on hardware side
Lap-App | LapApp-swRA-00014 | LapApp - Receive Button Presses | LapApp fails to receive button presses from hardware feedback | Data corruption, interface design errors | Software fails to react to button presses | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software failure to receive button presses prevents it from commanding appropriate state transitions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. HW telemetry being sent out each cycle, so this can be retried on next cycle 2. fault if hardware telemetry not received
Lap-App | LapApp-swRA-00015 | LapApp - Receive Button Presses | LapApp incorrectly interprets button presses | Incorrect data definition, interface design errors, errors in handling or interpreting data | Software reacts to button presses incorrectly | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software failure to interpret button presses correctly results in incorrect state transitions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Software incorporates audio and LED effects based on the mode that the arm is in
Lap-App | LapApp-swRA-00016 | LapApp - Send Motion Requests | LapApp fails to send motion requests (e.g., impedance start) | Incorrect algorithm, control flow errors | Software does not receive approvals in order to transition accordingly | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software failure to send motion requests renders the arm inoperable | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | Timeout checks for sending/receiving motion requests
Lap-App | LapApp-swRA-00017 | LapApp - Send Motion Requests | LapApp sends incorrect motion requests (e.g., impedance instead of admittance start) | Incorrect algorithm, interface design errors, control flow errors | Software receives approval and transitions to incorrect motion states | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software failure to send correct motion requests results in arm unexpected motions in response to user inputs | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. State transition handling 2. Verify robot hardware is in appropriate state 3. Joint command checks on hardware side 4. Collision verification 5. contact force checking
Lap-App | LapApp-swRA-00018 | LapApp - Send State Transitions To Hardware App | LapApp fails to send appropriate state transitions to hardware app | Incorrect algorithm, control flow errors | Arm does not move according to commanded joint motions (ignores motions) | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Failure to transition hardware to appropriate state will prevent hardware from moving | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | Monitor for timely hardware transitions, fault if they do not occur
Lap-App | LapApp-swRA-00019 | LapApp - Send Joint Commands to Hardware App | LapApp fails to send joint commands to hardware app | Incorrect algorithm, control flow errors | Arm does not move according to commanded joint motions (ignores motions) | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Failure to send joint commands will prevent hardware from moving | Always - 100% | Surgery | Intervention Necessary - (S3) | Always |  | 1. Fault if LapApp does not successfully send joint commands 2. Include timestamp in desired joint commands
Lap-App | LapApp-swRA-00020 | LapApp - Compute incorrect joint commands | LapApp computs and sends incorrect joint commands | Data corruption, interface design errors, control flow errors, incorrect algorithm, incorrect data definition | Wrong motions are commanded to the hardware | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Failure to send correct joint commands will cause uncontrolled robot motions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Joint command/limit check on hardware side 2. Collision verification 3. RCM motion check
Lap-App | LapApp-swRA-00021 | LapApp - Convert HID Inputs | LapApp fails to receive or convert HID inputs into appropriate instrument/laparoscope motions | Data Corruption, incorrect algorithm, interface design errors, errors in handling or interpreting data | Incorrect motions are computed and commanded | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Incorrectly computed motions are sent to hardware | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Perform sanity and incremental motion checks on converted HID pose 2. Haptic feedback on HID 3. Collision verification 4. incorporate delay check for HID input
Lap-App | LapApp-swRA-00022 | LapApp - Publish Telemetry | LapApp fails to publish or publishes incorrect telemetry information | Data Corruption, incorrect algorithm, interface design errors | Other sw components react incorrectly; incorrect feedback sent to surgeon | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Other components react to the bad telemetry, including: bad haptics at the HID; incorrect instrument name/type displayed; motion disallowed because some criteria seemingly unmet | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Telemetry is published each control cycle and a fault is thrown if unsuccessful. 2. External applications monitor heartbeat to identify if app is stalled
Lap-App | LapApp-swRA-00023 | LapApp - Admittance Control | LapApp fails to generate correct motion from user command | 2.2 Incorrect Algorithm FDA 618 | LapApp computed and commanded incorrect motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Incorrectly computed motions are sent to hardware | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Monitor admittance stability and fault if the control goes unstable
Lap-App | LapApp-swRA-00024 | LapApp - Inverse Dynamics Computation | LapApp incorrectly computes dynamics force and moments | Incorrect estimation of motion information of the arm joints due to: - Noise in joint motion data 2.2 Incorrect algorithm | Incorrect dynamics forces and moments are used to compensate sensor reading and command to robot arm | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software provides incorrect dynamics forces and moments information for compensation. | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Software shall apply filtering to minimize the impact of noisy data. 2. Software shall do POST and cross-check computed dynamics force and moment with sensor reading. 3. Software shall monitor the change rate and saturation of the algorithm output and fault if thresholds exceeded
Lap-App | LapApp-swRA-00025 | LapApp - Contact Detection | LapApp incorrectly computes contact force and moment | 2.2 Incorrect Algorithm FDA 618; C92142 | Incorrect contact force and moment are calculated | HHM-0109 | Hospital staff collision with equipment or subsystems | Software computes incorrect distal contact force and moment. | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. Software shall do POST and cross-check computed contact force/moment and expected external load. 2. Software shall monitor the change rate and saturation of the algorithm output and fault if thresholds exceeded
Lap-App | LapApp-swRA-00026 | LapApp - Contact Reaction | LapApp fails to slow down motion under high contact | 2.2 Incorrect Algorithm FDA 618; C92142 | Robot arm at high peak speed when excessive contact is detected, leading to longer stopping distance and larger contact when arm stops. Robot arm frequently applies excessive contact. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to slow down robot arm motion and allow | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare |  | 1. Software shall reduce the joint and end-effector speed with threshold determined by the contact magnitude. 2. Fault if we exceed velocity over a saturation threshold
Lap-App | LapApp-swRA-00027 | LapApp - Contact Reaction | LapApp fails to reduce contact force and/or moment with robot arm redundancy | 2.2 Incorrect Algorithm FDA 618; C92142 | Robot arm frequently applies excessive contact. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to reduce contact. | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare |  | Software shall monitor change of the contact force and/or moment as a result of reaction motion, i.e., if reaction motion results no motion or motion towards the direction of contact, contact force and/or moment is expected to not increase.
Lap-App | LapApp-swRA-00028 | LapApp - Maintain RCM during inverse kinematics | LapApp failts to hold the RCM steady when docked and not in Port Bump | 2.2 Incorrect Algorithm FDA 618; C92142 | Cannula shifts and applies large force to patient body; arm reaches singularity | HHM-0051 | Excessive force applied at incision site in Patient | Incorrectly computed motions are sent to hardware | Always - 100% | Chest Injury | Life Threatening - (S5) | Extremely rare |  | 1. Software shall calculate the RCM using forward kinematics with actual joint position to find RCM and stop motion when RCM error exceeds threshold. 2. Fault if there is excessive load on cannula
Lap-App | LapApp-swRA-00029 | LapApp - Collision Avoidance | LapApp fails to compute appropriate motions that avoid colliding with other components | 2.2 Incorrect Algorithm FDA 618; C92142 | Arm motions cause physical collisions | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Motions that cause collisions are sent to hardware | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Collision verification on joint commands 2. RCM tracking error monitoring 3. Fault when there is collision with an external object
Lap-App | LapApp-swRA-00030 | LapApp - Joint Limit Exceeded | LapApp fails to respect joint limits when computing motions | 2.2 Incorrect Algorithm FDA 618; C92142 | Joints are commanded to move outside limits | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Joint commands outside physical limits are sent to hardware | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Joint limit check (inclusive of velocity) after kinematics calculations, both on LapApp and HwApp 2. Collision verification of commands 2. RCM tracking error monitoring
Lap-App | LapApp-swRA-00031 | LappApp - Singularity Avoidance | LapApp fails to mitigate undesired motions near singularity | 2.2 Incorrect Algorithm FDA 618; C92142 | large RCM and/or instrument tracking error, loss of control in Cartesian control, instrument tip vibrations, Getting locked in singularity | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | software fails to compute appropriate motion commands | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1, Damped least-square algorithm in addition to joint limits to prevent approaching singular configurations, 2. Monitoring RCM tracking error 3. Monitor joint tracking error at each control cycle
Lap-App | LapApp-swRA-00032 | LappApp - Impedance | LapApp includes incorrect friction + gravity compensation | Errors due to unexpected hardware components or failures, Incorrect Algorithm | excessive force required to move the robot or the robot moves more than is expected | HHM-0109 | Hospital staff collision with equipment or subsystems | friction + gravity compensation is much smaller than needed. | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | During POST cross-check the result of gravity+friction compensation motor current computation and the reading of motor currents.
Lap-App | LapApp-swRA-00033 | LapApp - Bar optimization | LapApp fails to apply input from BarOptimization or computs and sends incorrect joint commands | Data parsing error; Interface error; 2.2 Incorrect Algorithm FDA 618; | Incorrect motions are computed and commanded; Arm motion does not follow desired trajectory | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | software fails to compute appropriate motion commands | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. LapApp shall track if the bias from bar optimization is successfully applied 2. Monitoring RCM tracking error 3. Montir joint tracking error at each control cycle
Lap-App | LapApp-swRA-00034 | LapApp - Decrement instrument life | LapApp fails to decrement instrument life when requested | Interface design errors, communication errors | Instruments could be used indefinitely | HHM-0481 | Patient is exposed to inadequate (below limit) therapeutic power from energy Instruments | Software fails to decrease life count | Always - 100% | Hemorrhage | Life Threatening - (S5) | Extremely rare |  | 1. LapApp will fault if RobotManipulatorHwApp does not acknowledge life decrement 2. Require life decrement before enabling motion
Lap-App | LapApp-swRA-00035 | LapApp - Process collision rejection message | LapApp fails to process rejection message | Interface design errors, control flow errors | Software fails to inform Model Apps of collision-causing joint commands | HHM-0446 | User confusion when operating the system | Failure to send rejections to Model Apps | Always - 100% | No harm | No Harm - (S0) | Always |  | Checks for difference between model and hardware joints
Lap-App | LapApp-swRA-00036 | LapApp - Fault response | A fault has occured and Robotic motion is not stopped and fault response does not occur | Failure to perform error processing and recovery | Robotic Manpulator motion is not stopped and unexpected motion will occur | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Fault response sw is corrupted and fails to perform fault response | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Faults within the LapApp are handled locally and at each cycle, as necessary. 2. Fault messages are treated as relaible in the RTI network and are guaranteed to arrive. 3. Faults from outside LapApp will also put RobotManipulatorHwApp into the fault state, forcing no new information to be sent to hardware app (i.e., motion stops)
Lap-App | LapApp-swRA-00037 | LapApp - Fault detection | A fault has occurred in the system and the software does not detect the fault. Subsequently, the SW does not initiate a fault response | Data corruption or parsing error when evaluating faults, Failure to perform error processing | Motion is not stopped and unexpected motion will occur | HHM-0051 | Excessive force applied at incision site in Patient | Fault detection sw is corrupted and fails to respond to a fault reported in the system. | Always - 100% | Chest Injury | Life Threatening - (S5) | Extremely rare |  | Fault checks run within each control cycle; if one is not handled correct, it is checked again on the next cycle.
Arm_Distal FW | ArmDistalFw-swRA-00001 | Arm Distal Firmware - Initialization | The application fails to start | 1. Wrong Version loaded 2. Bootloader software defect 3. Image corruption transport error | Force sensors A3 and D-pad (Touch points) lost before the procedure (i.e. system is not ready for starting the procedure) | HHM-0109 | Hospital staff collision with equipment or subsystems | Software error or corruption causing initialization failure | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Distal firmware is not communicating. SW will read Version and verify response.
Arm_Distal FW | ArmDistalFw-swRA-00002 | Arm Distal Firmware - Operation | The application hangs during operation | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | Force sensors A3 lost during the procedure. No contact sensing. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to perform state transition process correctly due to software runtime errors | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Distal firmware is not communicating. SW will read Version and verify response.
Arm_Distal FW | ArmDistalFw-swRA-00003 | Arm Distal Firmware - SPI communication with Ingenia | Software is unable to initialize SPI slave host communication | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | Force sensors A3 lost during the procedure. No contact sensing. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Distal firmware is not communicating.
Arm_Distal FW | ArmDistalFw-swRA-00004 | Arm Distal Firmware - SPI communication with Ingenia | Software is unable to receive data from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Unable to start system due to not being able to retrieve version | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always |  | 1 - Host SW will not allow surgery to begin if the Distal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Distal FW | ArmDistalFw-swRA-00005 | Arm Distal Firmware - SPI communication with Ingenia | Software is unable to transmit data from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Force sensors A3 lost during the procedure. No contact sensing. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Distal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Distal FW | ArmDistalFw-swRA-00006 | Arm Distal Firmware - Force Sensor Communication | Fails to read force sensors | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout 3. Hardware communication/Config issue | Individual touchpoints lost or intermittent | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. Firmware checks status bits of force sensor data and reports Status to Host SW 2. HWApp will check data status and report failure. 3 - SW App will check data validity with modeling
Arm_Distal FW | ArmDistalFw-swRA-00007 | Arm Distal Firmware - Force Sensor Communication | Missing reading force sensors data | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Firmware will have internal timer to detect delays. 2 - SW Host can check packet ID and detect stale and late data.
Arm_Distal FW | ArmDistalFw-swRA-00008 | Arm Distal Firmware - Force Sensor Communication | Wrong force values reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. HWApp will check data status and report failure. 2 - SW App will check data validitiy with modeling
Arm_Distal FW | ArmDistalFw-swRA-00009 | Arm Distal Firmware - Force Sensor Communication | Wrong index of force value reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. HWApp will check data status and report failure. 2 - SW App will check data validitiy with modeling
Arm_Proximal FW | ArmProximalFw-swRA-00001 | Arm Proximal Firmware - Initialization | The application fails to start | 1. Wrong Version loaded 2. Bootloader software defect 3. Image corruption transport error | Force sensors A3 and D-pad (Touch points) lost before the procedure (i.e. system is not ready for starting the procedure) | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption causing initialization failure | Always - 100% | No harm | No Harm - (S0) | Always |  | 1 - Host SW will not allow surgery to begin if the Proximal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Proximal FW | ArmProximalFw-swRA-00002 | Arm Proximal Firmware - Operation | The application hangs during operation | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | Force sensors A3 and D-pad (Touch points) lost during the procedure. No contact sensing. | HHM-0109 | Hospital staff collision with equipment or subsystems | Faulty state machine, software error or runtime processing failure | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Proximal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Proximal FW | ArmProximalFw-swRA-00003 | Arm Proximal Firmware - SPI communication with Ingenia | Software is unable to initialize SPI slave host communicaton | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | Force sensors A3 and D-pad (Touch points) lost during the procedure. No contact sensing | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Proximal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Proximal FW | ArmProximalFw-swRA-00004 | Arm Proximal Firmware - SPI communication with Ingenia | Software is unable to receive data from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Unable to start system due to not being able to retrieve version | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always |  | 1 - Host SW will not allow surgery to begin if the Proximal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Proximal FW | ArmProximalFw-swRA-00005 | Arm Proximal Firmware - SPI communication with Ingenia | Software is unable to transmit data from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Force sensors A3 and D-pad (Touch points) lost during the procedure. No contact sensing. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process telemetry data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Host SW will not allow surgery to begin if the Proximal firmware is not communicating.Host SW will read version and wait for expected response.
Arm_Proximal FW | ArmProximalFw-swRA-00006 | Arm Proximal Firmware - Force Sensor Communication | Fails to read force sensors | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout 3. Hardware communication/Config issue | Individual touchpoints lost and no contact sensing. It may results in collision with external subjects. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. Firmware checks status bits of force sensor data and reports Status to Host SW 2. HWApp will check data status and report failure. 3 - SW App will check data validity with modeling
Arm_Proximal FW | ArmProximalFw-swRA-00007 | Arm Proximal Firmware - Force Sensor Communication | Missing reading force sensors data | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1 - Firmware will have internal timer to detect delays. 2 - SW Host will check packet ID and detect stale and late data.
Arm_Proximal FW | ArmProximalFw-swRA-00008 | Arm Proximal Firmware - Force Sensor Communication | Wrong force values reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. HWApp will check data status and report failure. 2 - SW App will check data validitiy with modeling
Arm_Proximal FW | ArmProximalFw-swRA-00009 | Arm Proximal Firmware - Force Sensor Communication | Wrong index of force value reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Touchpoint controlled motion impacted. The failure doesn't impact RCM or tool motion. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process force sensing data due to runtime error or communication issue | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare |  | 1. HWApp will check data status and report failure. 2 - SW App will check data validitiy with modeling
Arm_Proximal FW | ArmProximalFw-swRA-00010 | Arm Proximal Firmware - D-Pad Buttons | Fails to detect buttons pressed | Software misbehavior occurs during runtime as a result of a coding error | D-Pad touchpoints lost or intermittent. | HHM-0347 | User confusion, annoyance, or inconvenience when operating the system | Software fails to detect or track D-Pad control status due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always |  | 1 - Check complimentary D-Pad signals and make sure they are opposite values. 2 - HWApp checks complimentary D-Pad signals and make sure they are opposite values.
Arm_Proximal FW | ArmProximalFw-swRA-00011 | Arm Proximal Firmware - D-Pad Buttons | Latency in reading the states of buttons | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Delay cannot be >10ms. D-Pad Touchpoint controlled motion impacted. | HHM-0093 | No hazardous situation | Software fails to detect or track D-Pad control status due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always |  | 1 - SW host checks complimentary button state changes to verify button presses
Arm_Proximal FW | ArmProximalFw-swRA-00012 | Arm Proximal Firmware - D-Pad Buttons | Wrong states of buttons reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | D-Pad Touchpoint controlled motion impacted. | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to detect or track D-Pad control status due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Check complimentary D-Pad signals and make sure they are opposite values. 2 - HWApp checks complimentary D-Pad signals and make sure they are opposite values.
Arm_Proximal FW | ArmProximalFw-swRA-00013 | Arm Proximal Firmware - D-Pad Buttons | Wrong index of states of button reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Mismatched touch point controls | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to detect or track D-Pad control status due to runtime error or communication issue | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Check complimentary D-Pad signals and make sure they are opposite values. 2 - HWApp checks complimentary D-Pad signals and make sure they are opposite values.
Arm Load Cell FW | ArmLoadCellFw-swRA-00001 | Load Cell Firmware - Initialization | Application fails to start | 1. Wrong Version loaded 2. Bootloader software defect 3. Image corruption transport error | System is not able to start. | HHM-0255 | Physician is unable to begin procedure on time | Lights wouldn't come on, no load cell data so no feedback data on load to arm, may impact robotic control. | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00002 | Load Cell Firmware - SPI communication with Ingenia | Software is unable to initialize SPI slave host communicatoin | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | System is not able to start. | HHM-0255 | Physician is unable to begin procedure on time | Lights wouldn't come on, no load cell data so no feedback data on load to arm, may impact robotic control. | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin or continue if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00003 | Load Cell Firmware - SPI communication with Ingenia | Software is unable to receive data (loadcell, version info) from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Unintentional movement during admittance docking, port bump, slave clutch | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Might get sensor data but async communication like get_version wouldn't work. Lights left uncontrolled | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin or continue if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00004 | load Cell Firmware - SPI communication with Ingenia | Software is unable to transmit data (Lights, heater) from SPI | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Light pattern transition on arm won't work. Heater may not start. | HHM-0416 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Lights wouldn't come on, no load cell data so no feedback data on load to arm | Always - 100% | Abdominal Injury | Permanent - (S4) | Extremely rare | High | 1 - Host SW will not allow surgery to begin or continue if the Load Cell firmware is not communicating.
Arm Load Cell FW | ArmLoadCellFw-swRA-00005 | Load Cell Firmware - CLI UART communication | Software fails to respond to request | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | No impact (service use only) | HHM-0093 | No hazardous situation | Minimal impact to functionality may delay sensor readings, will affect debug. | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin or continue if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00006 | Load Cell Firmware - CLI UART communication | Software response error | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Hardware version error | No impact | HHM-0093 | No hazardous situation | Minimal impact to functionality may delay sensor readings, will affect debug. | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - UART communication is not possible during procedure. Disable when there is SPI communication.
Arm Load Cell FW | ArmLoadCellFw-swRA-00007 | Load Cell Firmware - Load Cell Communication | Fails to read load | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout 3. Hardware communication/Config issue | Unintentional movement during admittance docking, port bump, slave clutch | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Lights wouldn't come on, no load cell data so no feedback data on load to arm, may impact robotic control. | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Reads sensors for expected values and send status to SW host 2 - Firmware checks for all zeros or all ones.
Arm Load Cell FW | ArmLoadCellFw-swRA-00008 | Load Cell Firmware - Load Cell Communication | Latency in reading load cell | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Delay cannot be >10ms. No impact. | HHM-0093 | No hazardous situation | Impacts robotic control of arm A5 | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Firmware will have internal timer to detect delays. 2 - SW Host will check packet ID and detect stale and late data.
Arm Load Cell FW | ArmLoadCellFw-swRA-00009 | Load Cell Firmware - Load Cell Communication | Wrong force values reported to host PC | Software misbehavior occurs during runtime as a result of a coding error | Unintentional movement during admittance docking, port bump, slave clutch | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Impacts robotic control of arm A5 | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Reads sensors for expected values and send status to SW host 2 - Firmware checks for all zeros or all ones. 3 - SW host will evaluate data trends over time.
Arm Load Cell FW | ArmLoadCellFw-swRA-00010 | Load Cell Firmware - Heater Control | Heater(s) fail to turn on | Software misbehavior occurs during runtime as a result of a coding error | Unable to heat load cell on power up to reduce boot time and ensure correct reading. | HHM-0399 | Excessive force applied to tissue in Patient | Impacts robotic control of arm A5 | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1 - Read the heater current - need HW mitigation 2 - Read temperature, no change in temperature detects whether sensor on.
Arm Load Cell FW | ArmLoadCellFw-swRA-00011 | Load Cell Firmware - Heater Control | Heater(s) fail to turn off | Software misbehavior occurs during runtime as a result of a coding error | Unable to heat load cell on power up to reduce boot time and ensure correct reading. | HHM-0399 | Excessive force applied to tissue in Patient | Impacts robotic control of arm A5 | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1 - Read the heater current - need HW mitigation 2 - Read temperature, no change in temperature detects whether sensor off.
Arm Load Cell FW | ArmLoadCellFw-swRA-00012 | Load Cell Firmware - Heater Control | Firmware temperature read error | Software misbehavior occurs during runtime as a result of a coding error | Incorrectly heat load cell and impacts reading. | HHM-0416 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Impacts robotic control of arm A5 | Always - 100% | Abdominal Injury | Permanent - (S4) | Extremely rare | High | 1 - Read the heater current - need HW mitigation 2 - Read temperature, no change in temperature detects whether sensor off. Compare second heater reading.
Arm Load Cell FW | ArmLoadCellFw-swRA-00013 | Load Cell Firmware - Heater Control | Heater(s) temperature write error | Software misbehavior occurs during runtime as a result of a coding error | Incorrectly heat load cell and impacts reading. | HHM-0416 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Impacts robotic control of arm A5 | Always - 100% | Abdominal Injury | Permanent - (S4) | Extremely rare | High | 1 - Read the heater current - need HW mitigation 2 - Read temperature, no change in temperature detects whether sensor off. Compare second heater reading.
Arm Load Cell FW | ArmLoadCellFw-swRA-00014 | Load Cell Firmware - LED Lighting | LEDs fail to turn on/off | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | Light pattern transition on arm won't work. | HHM-0259 | Physician is unable to complete procedure in a timely manner | Light pattern transition failure causing delayed or no feedback data on load to arm may impact robotic control | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1 - Read return value in last LED - log error
Arm Load Cell FW | ArmLoadCellFw-swRA-00015 | Load Cell Firmware - LED Lighting | LEDs not synchronized | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority 3. Scheduler timing problem, watchdog timeout | LED pattern is correct but out-of-sync. | HHM-0347 | User confusion, annoyance, or inconvenience when operating the system | Light pattern sync failure causing delayed or no feedback data on load to arm may impact robotic control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | N/A
Arm Load Cell FW | ArmLoadCellFw-swRA-00016 | Load Cell - Voltage Sensing 5v, 3.3V | Wrong Voltage readings | Software misbehavior occurs during runtime as a result of a coding error | Force measurment (Intermittent or incorrect) on joints Heater current intermittent or no current measurement Microcontroller may not receive/send data for motor drivers/sensors/buttons/LEDs Unintentional movement during admittance docking, port bump, slave clutch | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Over/Under-voltage as a result of sensing failure may cause incorrect feedback data on load to arm for robotic control | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Firmware compares reading outside of voltage thresholds, reports error to SW
Arm Load Cell FW | ArmLoadCellFw-swRA-00017 | Load Cell - Firmware Update | Firmware Image corrupted | Data lost or corrupted during processing or transmission | System is not able to start. | HHM-0255 | Physician is unable to begin procedure on time | Lights wouldn't come on, no load cell data so no feedback data on load to arm | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00018 | Load Cell - Firmware Update | Wrong firmware image sent | Software fails to properly install due to the incorrect data migration of user setup data. | System is not able to start. | HHM-0255 | Physician is unable to begin procedure on time | Lights wouldn't come on, no load cell data so no feedback data on load to arm | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Load Cell firmware is not communicating. Host SW will read version and wait for expected response.
Arm Load Cell FW | ArmLoadCellFw-swRA-00019 | Load Cell - NVRAM Storage | Fails to read Storage | 1. Data lost or corrupted during processing or transmission 2. Software misbehavior occurs during runtime as a result of a coding error | No impact | HHM-0093 | No hazardous situation | Config params (Version, checksum) are stored on load cell but config param for robotic manipulator stored on wrist board | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Write and Read pattern at POST, check CRC value
Arm Load Cell FW | ArmLoadCellFw-swRA-00020 | Load Cell - NVRAM Storage | Storage data corrupted | Data lost or corrupted during processing or transmission | No impact | HHM-0093 | No hazardous situation | Config params (Version, checksum) are stored on load cell but config param for robotic manipulator stored on wrist board | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Write and Read pattern at POST, check CRC value
Physician console PDU FW | ConsolePDUFw-swRA-00001 | Console PDU Firmware - Power when commanded | Not sending power on/off during when commanded | Software misbehavior occurs during runtime as a result of a coding error | No power state transition for Console | HHM-0255 | Physician is unable to begin procedure on time | Software performs wrong state transition during runtime process | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Reading output values, comparing sent values and retries 2 - Verify the downstream nodes are on. 3 - Host SW will turn off PSU, if PDU fails to turn off power.
Physician console PDU FW | ConsolePDUFw-swRA-00002 | Console PDU - Firmware Update | Firmware Image corrupted | Data lost or corrupted during processing or transmission | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the PDU firmware is not communicating.
Physician console PDU FW | ConsolePDUFw-swRA-00003 | Console PDU - Firmware Update | Wrong firmware image sent | Software fails to properly install due to the incorrect data migration of user setup data. | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the PDU firmware is not communicating.
Physician console PDU FW | ConsolePDUFw-swRA-00004 | Console PDU Firmware - Telemetry | Reading incorrect telemetry value | Software misbehavior occurs during runtime as a result of a coding error | Unintentional trigger of E-stop or loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Compare the downstream nodes data to detect failure, report to host SW.
Physician console PDU FW | ConsolePDUFw-swRA-00005 | Console PDU Firmware - Reset Control | Sending an incorrect reset | 1. EtherCAT master could have sent incorrect power enable message 2. Software Misbehavior 3. Data Corruption | Unintentional Console state transition | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process telemetry data resulting in state transition failure that accidentally power on/off in middle of procedure. This reboots just the Console. | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will detect the reset with the power turning off
Physician console PDU FW | ConsolePDUFw-swRA-00006 | Console PDU Firmware - Ethercat | PDO/PDI messages are lost | Data lost or corrupted during processing or transmission | Loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Console PDU firmware is not communicating.
Physician console PDU FW | ConsolePDUFw-swRA-00007 | Console PDU Firmware - Ethercat | PDO/PDI messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | E-stop is delayed by 10ms | HHM-0093 | No hazardous situation | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Firmware will have internal timer to detect delays. 2 - SW Host will check packet ID and detect stale and late data.
Physician console PDU FW | ConsolePDUFw-swRA-00008 | Console PDU Firmware - Ethercat | PDO/PDI message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | E-stop is lost | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Console PDU firmware is not communicating.
Physician console PDU FW | ConsolePDUFw-swRA-00009 | Console PDU Firmware - Ethercat | Wrong PDI messages are sent to the RTPC | Software failure or memory corruption | Unintentional trigger of E-stop or loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1. Host SW will check version of EtherCAT and firmware Version
Physician console PDU FW | ConsolePDUFw-swRA-00010 | Console PDU Firmware - Ethercat | FCP messages are corrupted | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Firmware checks CRC on header and data packets to detect corruption 2. SW host checks packet number, CRC on header and data packets
Physician console PDU FW | ConsolePDUFw-swRA-00011 | Console PDU Firmware - Ethercat | FCP messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | Delay in start-up by a few seconds | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol promptly due to runtime error causing delay | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Physician console PDU FW | ConsolePDUFw-swRA-00012 | Console PDU Firmware - Ethercat | FCP message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Physician console PDU FW | ConsolePDUFw-swRA-00013 | Console PDU Firmware - Ethercat | Wrong FCP response messages are sent to the RTPC | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will detect wrong messages and check transaction id
Physician console PDU FW | ConsolePDUFw-swRA-00014 | Console PDU Firmware - E-Stop | Firmware does not detect an eStop | Software misbehavior occurs during runtime as a result of a coding error | E-stop functionality lost | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process E-Stop correctly due to runtime error | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1- Read both E-stop signals and compare. Signals should match. 2 - POST e-stop trigger test
Surgical table PDU FW | Surgical TablePDUFw-swRA-00001 | Surgical Table PDU Firmware - Power Sequencing | Not sending power on/off during sequencing | Software misbehavior occurs during runtime as a result of a coding error | No power state transition for table | HHM-0255 | Physician is unable to begin procedure on time | Software performs wrong state transition during runtime process | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Reading output values, comparing sent values and retries 2 - Verify the downstream nodes are on. 3 - Host SW will turn off PSU, if PDU fails to turn off power.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00002 | Surgical Table PDU - Firmware Update | Firmware Image corrupted | Data lost or corrupted during processing or transmission | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Table PDU firmware is not communicating.Host SW will read version and wait for expected response.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00003 | Surgical Table PDU - Firmware Update | Wrong firmware image sent | Software fails to properly install due to the incorrect data migration of user setup data. | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Table PDU firmware is not communicating.Host SW will read version and wait for expected response.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00004 | Surgical Table PDU Firmware - Telemetry | Reading incorrect telemetry value (current) | 1. Software Misbehavior 2. Data Corruption | Unintended manipulator or instruments movement | HHM-0320 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : 6.3 - 15.3mm (M) | Software fails to process telemetry data for robotic manilulator control | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1 -HW App Compare the downstream nodes data to detect failure.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00005 | Surgical Table PDU Firmware - Reset Control | Sending an incorrect reset | 1. EtherCAT master could have sent incorrect power enable message 2. Software Misbehavior 3. Data Corruption | Unintentional Table state transition (e.g. ) Manipulator unable to maintain position on power loss AI: Robbie to check with Gary on power state transitions | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process telemetry data resulting in state transition failure | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will detect the reset of the Table PDU
Surgical table PDU FW | Surgical TablePDUFw-swRA-00006 | Surgical Table PDU Firmware - Ethercat | PDO/PDI messages are lost | Data lost or corrupted during processing or transmission | Loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Table PDU firmware is not communicating.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00007 | Surgical Table PDU Firmware - Ethercat | PDO/PDI messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | E-stop delayed by 10ms | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Firmware will have internal timer to detect delays. 2 - SW Host will check packet ID and detect stale and late data.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00008 | Surgical Table PDU Firmware - Ethercat | PDO/PDI message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | Loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Console PDU firmware is not communicating.
Surgical table PDU FW | Surgical TablePDUFw-swRA-00009 | Surgical Table PDU Firmware - Ethercat | Wrong PDI messages are sent to the RTPC | Software failure or memory corruption | Unintentional trigger of E-stop or loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1. Host SW will check version of EtherCAT and firmware Version
Surgical table PDU FW | Surgical TablePDUFw-swRA-00010 | Surgical Table PDU Firmware - Ethercat | FCP messages are corrupted | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Firmware checks CRC on header and data packets to detect corruption 2. SW host checks packet number, CRC on header and data packets
Surgical table PDU FW | Surgical TablePDUFw-swRA-00011 | Surgical Table PDU Firmware - Ethercat | FCP messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | No impact (Version FCP command) | HHM-0093 | No hazardous situation | Software fails to process firmware command protocol promptly due to runtime error causing delay | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Surgical table PDU FW | Surgical TablePDUFw-swRA-00012 | Surgical Table PDU Firmware - Ethercat | FCP message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Surgical table PDU FW | Surgical TablePDUFw-swRA-00013 | Surgical Table PDU Firmware - Ethercat | Wrong FCP response messages are sent to the RTPC | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will detect wrong messages and check transaction id
Surgical table PDU FW | Surgical TablePDUFw-swRA-00014 | Surgical Table PDU Firmware - E-Stop | Firmware detects a false eStop | Software misbehavior occurs during runtime as a result of a coding error | E-stop functionality is unintentionally triggered | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process E-Stop correctly due to runtime error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1- Read both E-stop signals and compare. Signals should match. 2 - POST e-stop trigger test
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00001 | Tower ESU & LED Control Firmware - Foot Pedal Spoofing | Detect the wrong foot pedal press/release | 1. Software error or failure 2. Data memory corruption | Unintended and incorrect energy generation | HHM-0043 | Patient is exposed to undesired energy application | Software fails to process foot pedal engagement causing incorrect energy generation | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1 - Debounce signals to make sure the foot pedal state is correct. 2 - Complememtary messaging to confirm press/release
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00002 | Tower ESU & LED Control Firmware - Foot Pedal Spoofing | Fire the wrong instrument | 1. Software error or failure 2. Data memory corruption | Energy generation at wrong instrument | HHM-0043 | Patient is exposed to undesired energy application | Software fails to process foot pedal engagement causing incorrect energy generation | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1 - Read energy relay status lines to verify correct instrument is fired.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00003 | Tower ESU & LED Control Firmware - LED Lighting | LEDs fail to turn on/off | 1. Software misbehavior occurs during runtime as a result of a coding error 2. Re-entrance errors, failure to block or unblock interrupts, invalid restart after an interrupt, wrong returning state assumed, wrong interrupt vector, wrong process priority | Incorrect LED pattern - user confusion when turning on or checking battery level due to lack of LED light | HHM-0347 | User confusion, annoyance, or inconvenience when operating the system | Light pattern transition failure causing incorrect feedback data on load to arm may impact robotic control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Code Review/unit testing
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00004 | Tower ESU & LED Control Firmware - Power Sequencing | Does not sequence power Correctly | Software misbehavior occurs during runtime as a result of a coding error | Trip eFuse - Lose power for Tower during procedure. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software performs wrong state transition during runtime process | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1- Check for 24V/48V before sequencing
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00005 | Tower ESU & LED Control Firmware - Power Sequencing | Not sending power on/off during sequencing | Software misbehavior occurs during runtime as a result of a coding error | No power state transition (ON/OFF) for Tower | HHM-0255 | Physician is unable to begin procedure on time | Software performs wrong state transition during runtime process | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Reading output values, comparing sent values and retries 2 - Verify the downstream nodes are on. 3 - Host SW will turn off PSU, if PDU fails to turn off power.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00006 | Tower PDU - Firmware Update | Firmware Image corrupted | Data lost or corrupted during processing or transmission | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the PDU firmware is not communicating.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00007 | Tower PDU - Firmware Update | Wrong firmware image sent | Software fails to properly install due to the incorrect data migration of user setup data. | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software error or corruption resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the PDU firmware is not communicating.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00008 | Tower ESU & LED Control Firmware - Telemetry | Reading incorrect telemetry value (current) | Software misbehavior occurs during runtime as a result of a coding error | Unintentional trigger of E-stop or loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Compare the downstream nodes data to detect failure, report to host SW.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00009 | Tower ESU & LED Control Firmware - Power Enable | Power is not enabled/disabled | 1. EtherCAT master could have sent incorrect power enable message 2. Software Misbehavior 3. Data Corruption | System is unable to boot up | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process telemetry data resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Read eFuse power good value (HW output) with/without telemetry data.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00010 | Tower ESU & LED Control Firmware - Reset Control | Sending an incorrect reset | 1. EtherCAT master could have sent incorrect power enable message 2. Software Misbehavior 3. Data Corruption | Unintentional Tower state transition during procedure | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process telemetry data resulting in state transition failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will detect the reset
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00011 | Tower ESU & LED Control Firmware - Ethercat | PDO/PDI messages are lost | Data lost or corrupted during processing or transmission | Loss of E-stop Functionality. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Tower PDU firmware is not communicating.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00012 | Tower ESU & LED Control Firmware - Ethercat | PDO/PDI messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | Delay of E-Stop by 10ms | HHM-0093 | No hazardous situation | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Firmware will have internal timer to detect delays. 2 - SW Host will check packet ID and detect stale and late data.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00013 | Tower ESU & LED Control Firmware - Ethercat | PDO/PDI message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | Loss of E-stop Functionality. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to runtime error for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will not allow surgery to begin if the Tower PDU firmware is not communicating.
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00014 | Tower ESU & LED Control Firmware - Ethercat | Wrong PDI messages are sent to the RTPC | Software failure or memory corruption | Unintentional trigger of E-stop or loss of E-stop. | HHM-0490 | Unable to address Operating Room emergency in a timely manner | Software fails to process telemetry data due to memory corruption for robotic manilulator control | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1. Host SW will check version of EtherCAT and firmware Version
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00015 | Tower ESU & LED Control Firmware - Ethercat | FCP messages are corrupted | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Firmware checks CRC on header and data packets to detect corruption 2. SW host checks packet number, CRC on header and data packets
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00016 | Tower ESU & LED Control Firmware - Ethercat | FCP (FW Command Protocol) messages are delayed | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | 10ms delay in retrieving version | HHM-0093 | No hazardous situation | Software fails to process firmware command protocol promptly due to runtime error causing delay | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00017 | Tower ESU & LED Control Firmware - Ethercat | FCP message processing stops | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Timers for transmission to detect delays. Firmware will notify SW of delays. 2- SW will monitor responses with timeouts
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00018 | Tower ESU & LED Control Firmware - Ethercat | Wrong FCP response messages are sent to the RTPC | Scheduler timing problem, watchdog time-outs, application is blocked, applicaton waiting for interrupt or a connection fail to establish | System is unable to boot up (Version FCP command) | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware command protocol due to runtime error resulting in initialization or boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will detect wrong messages and check transaction id
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00019 | Tower ESU & LED Control Firmware - E-Stop | Firmware does not detect an eStop | Software misbehavior occurs during runtime as a result of a coding error | E-stop functionality lost | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process E-Stop correctly due to runtime error | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Read both E-stop signals and compare. Signals should match. POST e-stop trigger test
Tower PDU FW | TowerESU&LEDCtrlFw-swRA-00020 | Tower ESU & LED Control Firmware - E-Stop | Firmware does not trigger an eStop | Software misbehavior occurs during runtime as a result of a coding error | E-stop functionality lost | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process E-Stop correctly due to runtime error | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1- Read both E-stop signals and compare. Signals should match. 2 - POST e-stop trigger test
Collision detection App | CollisionDetection-swRA-00001 | CollisionDetection - Startup | CollisionDetection fails to correctly initialize software components and is inoperable | Data Corruption, Incorrect SW installation | During initialization, CollisionDetection shuts down, and can't transition to an operational state | HHM-0255 | Physician is unable to begin procedure on time | During app boot up, software components are not checked for completed / or passed initialization | Always - 100% | No harm | No Harm - (S0) | Always |  | 1. During boot up/Initialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the app for successfull startup
Collision detection App | CollisionDetection-swRA-00002 | CollisionDetection - Configuration Data | CollisionDetection receives configuration data from any of the multiple sources incorrectly or incompletely | Data Corruption, Incorrect SW installation, Interface design errors | Incorrect or incomplete system configuration is used for collision detection, may result in incorrect or incomplete collision detection | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Configuration data is read incorrectly or incompletely from any of the mutiple sources | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Configuration data are CRC checked for completeness and acceptable range 2. The app will not completely intialize until configuration information is fully received
Collision detection App | CollisionDetection-swRA-00003 | CollisionDetection - Apply Configuration | CollisionDetection fails to correctly apply configuration | Incorrect data definition, interface design errors, errors in handling or interpreting data | Incorrect initialization of internal components, may result in incorrect or incomplete collision detection | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Configuration is used incorrectly to initialize internal components and models | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | During startup, the software will check that all internal components are initialized correctly and are of correct type/behavior
Collision detection App | CollisionDetection-swRA-00004 | CollisionDetection - Apply Kinematics Calibration | CollisionDetection fails to apply the kinematics calibration correctly | Incorrect algorithm, incorrect data definition, interface design errors, errors in handling or interpreting data | Software fails to correctly apply kinematic calibration | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Using incorrect calibration, the app computes wrong motions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | The software will check the kinematic calibration for range and completeness and verifies the result of calibration
Collision detection App | CollisionDetection-swRA-00005 | CollisionDetection - Use Hardware Feedback for Collision Detection (when desired joint commands are not received or when in particular modes) | CollisionDetection fails to receive hardware feedback | Data Corruption, interface design errors | Software fails to use hardware feedback to detect collisions | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Failure to detect collisions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | Timeout check and fault if timed out while waiting to receive hardware feedback
Collision detection App | CollisionDetection-swRA-00006 | CollisionDetection - Use Hardware Feedback for Collision Detection (when desired joint commands are not received or when in particular modes) | CollisionDetection uses hardware feedback incorrectly | Incorrect algorithm, incorrect data definition, interface design errors, errors in handling or interpreting data | Software detects incorrect collisions | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Failure to detect correct collisions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Range check joint values before applying to collision models 2. Contact detection and stop
Collision detection App | CollisionDetection-swRA-00007 | CollisionDetection - Use Joint Commands for Collision Detection | CollisionDetection uses joint commands incorrectly to detect collisions | Incorrect algorithm, incorrect data definition, interface design errors, errors in handling or interpreting data | Software detects incorrect missing collisions | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Failure to detect correct collisions | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Range check joint values before applying to collision models 2. Contact detection and stop
Collision detection App | CollisionDetection-swRA-00008 | CollisionDetection - Route Joint Commands To Hardware Apps | CollisionDetection fails to route non-collision-causing joint commands - includes comm layer failures | Interface design errors, control flow errors | Software fails to send joint commands to hardware apps | HHM-0282 | System freezes and becomes unresponsive to all user commands in the middle of surgery procedure when Patient is in contact with Instruments | Failure to send joint commands to hardware apps | Always - 100% | Hemorrhage | Life Threatening - (S5) | Extremely rare |  | 1. Joint command timeout checks in hw apps 2. Add range and timeout checks in model app for joint commands received by hw apps 3. Add fault if failure at RTI level
Collision detection App | CollisionDetection-swRA-00009 | CollisionDetection - Route Joint Commands To Hardware Apps | CollisionDetection routes incorrect/corrupted joint commands to hardware apps | Data corruption, interface design errors, control flow errors | Software sends incorrect/invalid joint commands to hardware apps | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Sending of incorrect/invalid joint commands to hardware apps | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Add range and velocity/acceleration checks on receiving hw apps for joint commands 2. Add range and timestamp checks in model apps for joint commands received by hw apps 3. Contact detection 4. RCM tracking error
Collision detection App | CollisionDetection-swRA-00010 | CollisionDetection - Route Joint Commands To Hardware Apps | CollisionDetection routes joint commands that cause collision to hardware apps | Interface design errors, control flow errors | Software sends joint commands that result in collision to hardware apps | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Sending of collision-causing joint commands to hardware | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Contact force checking on arms 2. Joint tracking error checks 3. RCM tracking error
Collision detection App | CollisionDetection-swRA-00011 | CollisionDetection - Send Rejection To Model Apps | CollisionDetection fails to send rejection messages when joint commands would cause collisions - includes comm layer failures | Interface design errors, control flow errors | Software fails to inform Model Apps of collision-causing joint commands | HHM-0446 | User confusion when operating the system | Failure to send rejections to Model Apps | Always - 100% | No harm | No Harm - (S0) | Always |  | 1. Add checks in Model Apps for difference between model and hardware joints 2. Add fault if failure at RTI level
Collision detection App | CollisionDetection-swRA-00012 | CollisionDetection - Send Rejection To Model Apps | CollisionDetection sends incorrect rejection messages when joint commands would cause collisions | Data corruption, interface design errors, control flow errors | Software incorrectly informs Model Apps of collision-causing joint commands | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Failure to send correct rejections to Model Apps | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Add checks on receiving Model Apps for validity and range of rejection messages 2. Add checks in Model Apps for difference between model and hardware joints
Collision detection App | CollisionDetection-swRA-00013 | CollisionDetection - Publish Telemetry | CollisionDetection fails to publish or publishes incorrect telemetry information | Data Corruption, incorrect algorithm, interface design errors, errors in handling or interpreting data | Other sw components react incorrectly, including bad colllision avoidance behavior; incorrect feedback sent to surgeon | HHM-0446 | User confusion when operating the system | Failure to publish telemetry | Always - 100% | No harm | No Harm - (S0) | Always |  | 1. Telemetry is published each control cycle and a fault is thrown if unsuccessful. 2. External applications monitor heartbeat to identify if app is stalled
Audio node | AudioNode-swRA-00001 | Receive audio Configuration from Configuration Subsystem | Failed to receive Configuration | 1) Failed Communciation 2) Corrupted Data | Unable to play any audio tones | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive audio configuration from Configuration Subsystem due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely rare | High | Inform the user with a notification.
Audio node | AudioNode-swRA-00002 | Received audio trigger | Failed to receive audio trigger | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive audio trigger due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | Inform the user with a notification.
Audio node | AudioNode-swRA-00003 | Received Telemetry from [LapApp app] | Fails to receive telemetry data from [LapApp app] | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive telemetry from LapApp app due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Prevent from playing incorrect audio by using CRC check from incoming data. 2. Lap Telemetry information is continuously being sent.
Audio node | AudioNode-swRA-00004 | Received Telemetry from [BedModel app] | Fails to receive telemetry data from [BedModel app] | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive telemetry from BedModel app due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Prevent from playing incorrect audio by using CRC check from incoming data. 2. Bed Telemetry information is continuously being sent.
Audio node | AudioNode-swRA-00005 | Received Telemetry from [HapticInterfaceDeviceModel app] | Fails to receive telemetry data from [HapticInterfaceDeviceModel app] | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive telemetry from HapticInterfaceDeviceModel app due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Prevent from playing incorrect audio by using CRC check from incoming data. 2. HapticInterfaceDevice telemetry information is continuously being sent.
Audio node | AudioNode-swRA-00006 | Received Telemetry from [ErgonomicsModel app] | Fails to receive telemetry data from [ErgonomicsModel app] | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive telemetry from ErgonomicsModel app due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Prevent from playing incorrect audio by using CRC check from incoming data. 2. Ergonomics telemetry information is continuously being sent.
Audio node | AudioNode-swRA-00007 | Receive Notification | Fails to receive notification | 1) Failed Communciation 2) Corrupted Data | Unable to play audio tone | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive notification due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely rare | High | Message will be sent with QOS reliable setting
Audio node | AudioNode-swRA-00008 | Receive Audio Configuration | Fails to receive Audio Configuration | 1) Failed Communciation 2) Corrupted Data | System plays no sound or incorrect sound | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to receive audio configuration due to communication failure or corrupted data | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely rare | High | Message will be sent with QOS reliable setting
Aux panel App | AuxPanelApp-swRA-00001 | Send command to [BedMobilityApp] to lock the bed | Fails to communicate with [BedMobility app] | Failed Communciation | User is unable to lock the bed | HHM-0255 | Physician is unable to begin procedure on time | Software fails to send command to [BedMobilityApp] to lock the bed due to communication failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | no harm
Aux panel App | AuxPanelApp-swRA-00002 | Send command to [BedMobilityApp] to unlock the bed | Fails to communicate with [BedMobility app] | Failed Communciation | User is unable to unlock the bed | HHM-0255 | Physician is unable to begin procedure on time | Software fails to send command to [BedMobilityApp] to unlock the bed due to communication failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | no harm
Aux panel App | AuxPanelApp-swRA-00003 | Send command to [BedMobilityApp] to Level the bed | Fails to communicate with [BedMobility app] | Failed Communciation | Bed does not return to level | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to send command to [BedMobilityApp] to level the bed due to communication failure | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting
Aux panel App | AuxPanelApp-swRA-00004 | Send Command to [BedModel App] to move table top Flex up/down motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top Flex up/down | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Flex up/down motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting
Aux panel App | AuxPanelApp-swRA-00005 | Send Command to [BedModel App] to move table top Chest up/down motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top Chest up/down | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Chest up/down motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Aux panel App | AuxPanelApp-swRA-00006 | Send Command to [BedModel App] to move table top Legs up/down motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top Legs up/down | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Legs up/down motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Aux panel App | AuxPanelApp-swRA-00007 | Send Command to [BedModel App] to move table top slide front/back motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top slide front/back | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top slide front/back motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Aux panel App | AuxPanelApp-swRA-00008 | Send Command to [BedModel App] to move table top trendelenburg motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top trendelenburg | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Trendelenburg motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Aux panel App | AuxPanelApp-swRA-00009 | Send Command to [BedModel App] to move table top Height motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top Height | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Height motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Aux panel App | AuxPanelApp-swRA-00010 | Send Command to [BedModel App] to move table top Tilt motion | Fails to communicate with [BedModel app] | Failed Communciation | User is unable to move table top Tilt | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to send command to move table top Tilt motion due to communication failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the AuxPanel controls do not work, secondary controls are on the BedPendant Application. BedPendant can be used to level the bed. Message will also be sent with QOS reliable setting.
Overlay App | OverlayApp-swRA-00001 | Overlay app - shutdown | Application fails to stop | Hang | Negative impact on system restart | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process app shutdown thus causing failure of application stoppage | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | Power down system for full restart.
Overlay App | OverlayApp-swRA-00002 | Receive Robotic Manipulator telemetry data | Fails to receive telemetry data | Failed Communication | Controlled arm not being represented on the overlay | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to represent the controlled arm on the overlay due to failed communication of Robotic Manipulator telemetry data | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. User receives a notification if there is no Lap Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00003 | Display Overlay Data - lost | Application stops | 1) Crash 2) Hang | Overlay graphics lost | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to display overlay graphics due to a crash or hang of the application | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. CRC checks for corrupted/invalid data. 2. Mitigation within SVP card will be to add a "heartbeat" message that will show special UI if no new frames received.
Overlay App | OverlayApp-swRA-00004 | Receive Robotic Manipulator telemetry data for collision detection | Fails to receive collision telemetry data | Failed Communication | Collision indications not displayed | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to display collision indications on the overlay due to failed communication of Robotic Manipulator telemetry data | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. User receives a notification if there is no Collision Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00005 | Receive tool identification telemetry data | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of tool information | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software misrepresents tool information on the overlay due to failed communication or corruption of tool identification telemetry data | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. User receives a notification if there is no Lap Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00006 | Receive scope telemetry data | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of endoscope horizon | HHM-0446 | User confusion when operating the system | Software misrepresents endoscope horizon on the overlay due to failed communication or corruption of scope telemetry data | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. User receives a notification if there is no Lap Telemetry received. NOTE: Eye injury is unlikely. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00007 | Receive tool status telemetry data | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of tool mode | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software misrepresents tool mode on the overlay due to failed communication or corruption of tool status telemetry data | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. User receives a notification if there is no Lap Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00008 | Receive tool energy functions telemetry data | Fails to receive telemetry data | Failed communication | Misrepresentation of tool energy functions on Overlay | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software misrepresents tool energy functions on the overlay due to failed communication of tool energy functions telemetry data | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. EnergyActivationCommand proto and TowerHardwareTelemetry. Defaults to false unless we receive otherwise. 2. User receives a notification if there is no Tower Hardware Telemetry received. 3. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00009 | Receive energy activation status of tools | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of energy activation on Overlay | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software misrepresents tool energy activation status on the overlay due to failed communication or corruption of tool energy activation status data | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. User still commands activation from foot pedals and receives auditory signals that they have activated energy. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00010 | Receive instrument grip force telemetry data | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of grip force on Overlay | HHM-0368 | Low grip force in end effector (in Patient tissue) | Software misrepresents instrument grip force on the overlay due to failed communication or corruption of instrument grip force telemetry data | Always - 100% | Hemorrhage | Life Threatening - (S5) | Extremely rare | High | 1. User receives a notification if there is no Lap Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00011 | Receive tool homing and loading telemetry data | Fails to receive telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of tool readiness and loading progress on Overlay | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software misrepresents tool homing and loading on the overlay due to failed communication or corruption of tool homing and loading telemetry data | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. User receives a notification if there is no Lap Telemetry received. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00012 | Receive notification information | Fails to receive notification display request | 1) Failed Communication 2) Corrupted data | No notifications displayed | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to display notifications on the overlay due to failed communication or corruption of notification information | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Failure to receive notification in Overlay does not prevent other applications from also displaying notifications. 2. Redundant information display. 3. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00013 | Display notifications | Message truncated, corrupted or scrambled | Data corruption | Incomplete or incoherent notification displayed to surgeon | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software displays incomplete or incoherent notifications on the overlay due to corruption of notification information | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Notification will not be published or rescinded if not a registered/known type in configuration file. 2. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00014 | Receive notification removal request | Fails to receive notification removal requests | Failed Communciation | Notification persists on Overlay and occupies scope field of view | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to remove notifications on the overlay due to failed communication of notification removal request | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Notification area restricted and avoids main surgeon field of view. 2. Only one system notification visible at any given time, so will not build up space on screen. 3. CRC checks for corrupted/invalid data.
Overlay App | OverlayApp-swRA-00016 | Display accuracy of indicators (offscreen indicators) | Position in the overlay is not accurate to tool tip | 1) lost telemetry 2) Data corruption | Patient harm due to inaccurate tool location | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software displays inaccurate tool location on the overlay due to lost telemetry or data corruption | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | User cannot move or fire energy if the tool is not active.
Overlay App | OverlayApp-swRA-00017 | Receive request to open quick action menu | Fails to receive request to open quick action menu | 1) Failed Communication 2) Corrupted data | Quick action menu inaccessible to user | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software does not allow user access to quick action menu on the overlay due to failed communication or data corruption | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Redundancy in interaction points. (For example, video recording can be started in the Tower, Console Touchpad, or Overlay)
Overlay App | OverlayApp-swRA-00018 | Receive request to close quick action menu | Fails to receive request to close quick action menu | 1) Failed Communication 2) Corrupted data | Quick action menu persists over scope view and surgical field of view is constrained | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to close the quick action menu on the overlay due to failed communication or data corruption | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Quick action menu rendered separately from "drive" information. 2. You cannot drive tools when the in-viewer menu is active.
Overlay App | OverlayApp-swRA-00019 | Display size of UI elements | Drawn too large, placement accuracy, drawn too small to be read | 1) Failed Communication 2) Corrupted data | Can't clearly understand the UI element | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to display the correct size of UI elements on the overlay due to failed communication or data corruption | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Elements are fixed size and rendered onto fixed-size buffer. 2. DRF studies should also make sure UI elements are legible.
Overlay App | OverlayApp-swRA-00020 | Receive telemetry data for active arm | Fails to receive active arm telemetry data | 1) Failed Communication 2) Corrupted data | Misrepresentation of active arm on Overlay | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to receive active arm telemetry data on the overlay due to failed communication or data corruption | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. User receives a notification if there is no Lap Telemetry received. (and ConsoleTelemetry) 2. CRC checks for corruption/invalid data.
Overlay App | OverlayApp-swRA-00021 | UI Elements (color) | Application initialization fails, leading to wrong UI colors. | Failure to initialize application. | Arm identification is incorrect, leading to unexpected arm motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to display the correct color of UI elements on the overlay due to failed application initialization | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | There is redundancy so you can identify an arm by color or number.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00001 | ConsoleErgonomicsApp - Startup | ConsoleErgonomicsApp fails to correctly initialize software components and is inoperable | Data corruption, incorrect SW installation | During initialization, ConsoleErgonomicsApp shuts down, and can't transition to an operational state, and console ergonomics cannot be adjusted | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | During app boot up, software components are not checked for completed / or passed initialization | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. During boot up/Initiialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the application for successful startup.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00002 | ConsoleErgonomicsApp - Configuration Data | ConsoleErgonomicsApp receives configuration data from any of the multiple sources incorrectly or incompletely | Data corruption, incorrect SW installation, Interface design error | Incorrect or incomplete system configuration is used for console ergonomics motion commands, may result in incorrect or undesired ergonomic motion | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Configuration data is read incorrectly or incompletely from any of the mutiple sources | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Configuration data will be checked with CRC and values will be checked to be within acceptable limits. 2. Ergonomics Model App will not completely initialize until configuration data is fully received.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00003 | ConsoleErgonomicsApp - Configuration Data | Console ErgonomicsApp fails to correctly apply configuration | Incorrect data definition, interface design errors, erros in handling or interpreting data | Incorrect initialization of internal components, may result in incorrect or undesired ergonomic motion | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Configuration is used incorrectly to intialize internal components and models | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | During startup, the software will check that all internal components are initalized correctly and are of correct type/behavior.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00004 | ConsoleErgonomicsApp - Powerup Ergonomics Axes | ConsoleErgonomicsApp fails to send the power up signal to the hw application in response to joystick input or user-saved position input from gui | Network issues, incorrect algorithm, control flow errors | Ergonomics axes are inoperable due to sw failure to power it up | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to send the powerup signal to the hardware | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Ensure receipt of power up acknowledgement from hw in a timely manner.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00005 | ConsoleErgonomicsApp - Power down Ergonomics Axes | ConsoleErgonomicsApp fails to send the power down signal to the hw application when no joystick or user-saved position input | Network issues, incorrect algorithm, control flow errors | Ergonomics axes stay powered up when user is not adjusting ergonomics | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to send the power down signal to the hardware | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Ensure receipt of power down acknowledgement from hw in a timely manner.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00006 | ConsoleErgonomicsApp - State Transitions | ConsoleErgonomicsApp fails to transition to appropriate states, or transitions to wrong state(s) when commanded | Failure to set the correct initial and following states, control flow errors | Console ergonomics axes not moving in response to commanded motions or will have undesired motion | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to compute or computes incorrect motions due to being in an incorrect state | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | Unhandled state transitions lead to a minor fault
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00007 | ConsoleErgonomicsApp - Receive hardware telemetry | ConsoleErgonomicsApp recieves hardware telemetry incorrectly | Incorrect data definition, Interface design errors, Errors in handling or interpreting data | Console ergonomics axes not moving in response to commanded motions or will have undesired motion | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process HW telemetry data needed to compute motion command | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Telemetry data received over RTI will be CRC checked. 2. Timeout faults for failing to receive hardware telemetry. 3. Joint limit/motion checks on hardware side.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00008 | ConsoleErgonomicsApp - Receive hardware telemetry | ConsoleErgonomicsApp fails to recieve Console Ergonomics hardware telemetry | Data corruption, interface design errors, network issues | Console ergonomics axes not moving in response to commanded motions or will have undesired motion | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process HW telemetry data needed to compute motion command | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Timeout faults for failing to receive hardware telemetry.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00009 | ConsoleErgonomicsApp - Receive console gui user-saved position | ConsoleErgonomicsApp fails to receive console gui user-saved position | Data corruption, interface design errors, network issues | Software fails to receive user-saved position and cannot move axes to user-saved ergonomic settings | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to process console GUI position for ergo settings due to runtime or data processing error | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. All ergonomics axes can also be moved via joystick inputs
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00010 | ConsoleErgonomicsApp - Receive console gui user-saved position | ConsoleErgonomicsApp incorrectly receives console gui user-saved position | Incorrect data definition, Interface design errors, Errors in handling or interpreting data | Incorrect ergonomic motion | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to process console GUI configuration for Ergo motion due to runtime or data processing error | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Messages received over RTI will be CRC checked. 2. Check received user-saved position within joint limits and doesn't cause self collisions.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00011 | ConsoleErgonomicsApp - Receive HID telemetry | ConsoleErgonomicsApp fails to receive HID telemetry | Data corruption, interface design errors, network issues | Ergonomics axes do not move | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to receive HID telemetry data for processing Ergo motion command due to runtime or data processing error | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Timeout fault is thrown if HID telemetry data is not received.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00012 | ConsoleErgonomicsApp - Recieve HID telemetry | ConsoleErgonomicsApp incorrectly receives HID telemetry | Incorrect data definition, Interface design errors, Errors in handling or interpreting data | Ergonomics axes do not move | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software incorrectly processes HID telemetry data for Ergo motion control due to runtime or data processing error | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | Messages received over RTI will be CRC checked.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00013 | ConsoleErgonomicsApp - Send joint commands to hardware app | ConsoleErgonmicsApp fails to send joint commands to the hardware app | Control flow errors, network issues | Ergonomics axes don't move according to the commanded joint motions (ignores motions) | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to send telemetry data for processing joint control command due to runtime or data processing error | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Fault if ConsoleErgonomicsApp does not successfully send commands to the hardware app 2. Include timestamps in published joint commands
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00014 | ConsoleErgonomicsApp - Collision avoidance | ConsoleErgonomicsApp fails to compute appropriate motions that avoid colliding with other components | Incorrect algorithm | Console axes motion cause physical collisions between console structures | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software sends joint commands that cause collisions to hardware app. | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | Software checks joint commands don't cause collisions.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00015 | ConsoleErgonomicsApp - Joint limit avoidance | ConsoleErgonomicsApp fails to respect joint limits when computing motions | Incorrect algorithm | Joints are commanded to move outside limits. Console axes motion cause physical collisions. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software sends joint commands that are outside physical limits to hardware app. | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | Joint limit check after kinematics calculations.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00016 | ConsoleErgonomicsApp - Joint command computations | ConsoleErgonomicsApp incorrectly computes joint commands from input joystick commands or desired position | Incorrect algorithm, control flow errors | Incorrect motions are computed and commanded to the hardware. Uncontrolled/undesired motion of console ergonomics motors | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to compute joint command correctly from joystick input | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Joint limit/motion checks on hardware side. 2. Software checks joint command doesn't cause collisions.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00017 | ConsoleErgonomicsApp - Joint command computations | ConsoleErgonomicsApp receives more than one input command for each ergonomics axis (e.g. user-saved config received and joystick actuated simultaneously) | Design error | Incorrect motions are computed and commanded to the hardware. Uncontrolled/undesired motion of console ergonomics motors | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software incorrectly computes and processes joint motion commands due to multi-input error | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | SW state machine design only allows control of each axis via one input
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00018 | ConsoleErgonomicsApp - fault response | A fault has occurred and a fault response does not occur | Failure to perform error processing and recovery | Ergonomics motion is not stopped and unexpected motion will occur | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Fault response software is corrupted and fails to perform fault response | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Faults within the ErgonomicsApp are handled locally and at each cycle, as necessary. 2. Fault messages are treated as reliable in the RTI network and are gauranteed to arrive.
Console ergonomi model App | ConsoleErgonomicsModelApp-swRA-00019 | ConsoleErgonomicsApp - fault detection | A fault has occurred and the software does not detect the fault. | Data corruption or parsing error when evaluating faults, failure to perform error processing | Ergonomics motion is not stopped and unexpected motion will occur | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Fault response software is corrupted and fails to respond to a fault reported in the system | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | Fault checks run within each control cycle; if one is not handled correctly, it is checked again on the next cycle.
Patient position HW App | PatientPositioningHWApp-swRA-00001 | Bed Patient Positioning Hw App - Initialize patient positioning HW and sensors | Bed Patient Positioning Hw App fails to initialize patient postioning HW and sensors | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positioning does not move | HHM-0255 | Physician is unable to begin procedure on time | Software fails to complete initialization process of hardware components due to runtime errors resulting in patient positioning inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. HW App Retry 2. Faults will be thrown 3. *Workflow mitigation (ensure system initialized correctly prior to putting patient on bed)*
Patient position HW App | PatientPositioningHWApp-swRA-00002 | Bed Patient Positioning Hw App - Apply Configuration | BedPatientPositionHwApp fails to apply configuration | Data corruption Data parsing error Interface design error | Patient positioning does not move | HHM-0255 | Physician is unable to begin procedure on time | Software fails to apply configuration to HW control due to data corruption or runtime error resulting in patient positioning inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. HW App Retry 2. Faults will be thrown 3. *Workflow mitigation (ensure system initialized correctly prior to putting patient on bed)*
Patient position HW App | PatientPositioningHWApp-swRA-00003 | Bed Patient Positioning Hw App - Apply Configuration | BedPatientPositionHwApp receives incorrect or incorrectly processes configuration data | Data corruption Data parsing error Interface design error | Patient positioning moves in unintended direction | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process configuration data to HW control due to data corruption or runtime error resulting in undesired patient positioning | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Configuration data from Configuration Server is CRC checked for complete data 2. ConfigVerifier checks that configuration values are within range 3. Application verifies UI/UF Paramaters
Patient position HW App | PatientPositioningHWApp-swRA-00004 | Bed Patient Positioning Hw App - Process joint command message from Collision Detection | Bed Patient Positioning Hw App fails to processes joint command message | Data corruption of joint command message from RTI Data parsing error of joint command message Interface design error for joint command message | Patient positioning does not move | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process joint command from collision detection due to data corruption or runtime error resulting in patient positioning inoperable | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Faults will be thrown
Patient position HW App | PatientPositioningHWApp-swRA-00005 | Bed Patient Positioning Hw App - Process joint command message from Collision Detection | Bed Patient Positioning Hw App incorrectly processes joint command message | Data corruption Data parsing error Interface design error | Patient positioning moves unintentionally potentially without user initiation | HHM-0109 | Hospital staff collision with equipment or subsystems | Software fails to process joint command from collision detection correctly due to data corruption or runtime error resulting in undesired patient positioning | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Returns error if joint command message is sent from anyone but collision detection 2. Checks on Commanded Velocity & Acceleration 3. Faults will be thrown (if checks fail)
Patient position HW App | PatientPositioningHWApp-swRA-00006 | Bed Patient Positioning Hw App - Enable hardware | Bed Patient Positioning Hw App fails to enable hardware | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positioning does not enable hardware | HHM-0255 | Physician is unable to begin procedure on time | Software fails to enable hardware component for control due to runtime error or hardware issue resulting in patient positioning inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. HW App Retry 2. Faults will be thrown 3. *Workflow mitigation (ensure system initialized correctly prior to putting patient on bed)*
Patient position HW App | PatientPositioningHWApp-swRA-00007 | Bed Patient Positioning Hw App - Disabling hardware | Bed Patient Positioning Hw App fails to disable hardware | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positioning does not disable hardware leading to overheating | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to disable hardware component for control due to runtime error or hardware issue resulting in undesired patient positioning or component overheating as a result | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Read current on Elmos, then turn off power rail for Elmos 2. Check the status for 'Disabled' mode
Patient position HW App | PatientPositioningHWApp-swRA-00008 | Bed Patient Positioning Hw App - Initialize Ethercat communication | Bed Patient Positioning Hw App fails to initialize communication | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positioning does not move | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process initialization of EtherCAT communication due to runtime error or hardware issue resulting in patient positioning inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. HW App Retry 2. Faults will be thrown 3. *Workflow mitigation (ensure system initialized correctly prior to putting patient on bed)*
Patient position HW App | PatientPositioningHWApp-swRA-00009 | Bed Patient Positioning Hw App - output commands to EtherCAT | Bed Patient Positioning Hw App fails to communicate with EtherCAT communications and processed commands are not sent to the surgical table EtherCAT bus or system | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positioning does not move | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process output commands to EtherCAT due to runtime error or comm bus issue resulting in patient positioning inoperable | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Check 0X20F1 (Elmo Timeout Window) Object Code is set correctly, if Elmo goes into timeout, faults will be thrown
Patient position HW App | PatientPositioningHWApp-swRA-00010 | Bed Patient Positioning Hw App - output commands to EtherCAT | Bed Patient Positioning Hw App incorrectly outputs commands to EtherCAT | Data corruption Data parsing error Interface design error | Patient positioning moves unintentionally | HHM-0109 | Hospital staff collision with equipment or subsystems | Software incorrectly processes output commands to EtherCAT due to runtime error or data corruption resulting in undesired patient positioning | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. Max Position fault (*lookup*) 2. Max Current fault (0X6073) 3. Max Speed Tracking fault (*lookup*) 4. Max Acceleration fault (*lookup*)
Patient position HW App | PatientPositioningHWApp-swRA-00011 | Bed Patient Positioning Hw App - processes received status information via EtherCAT | Bed Patient Positioning Hw App fails to communicate with EtherCAT communications and status is not received from the EtherCAT bus. | Data corruption Software timing Errors due to unexpected hardware components or failures | Patient positoning information is not displayed on the table pendant and available for the rest of the system (collision detection, etc.). Worst case is internal collision. | HHM-0021 | Capital equipment (Tower, console, Bed) or subassemblies (Robotic Arm) collision with other capital equipment or environment | Software fails to process status feedback due to runtime error or comm bus issue resulting in information unavailable or undesired patient positioning leading to potential collision | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | 1. Major Faults will be thrown
Patient position HW App | PatientPositioningHWApp-swRA-00012 | Bed Patient Positioning Hw App - processes received status information via EtherCAT | Bed Patient Positioning Hw App incorrectly processes status information. | Data corruption Data parsing error Interface design error | Patient positoning information is incorrect for Coordinated Motion, resulting in RCM Motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes status feedback due to runtime error or comm bus issue resulting in incorrect information leading to undesired RCM motion | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Check joint positions against the limits 2. Tower displays GUI of intended Patient Positioning and requires visual confirmation? 3. For flexion joints, compare joint positions between Leader and Follower joints.
Patient position HW App | PatientPositioningHWApp-swRA-00013 | Bed Patient Positioning Hw App - Stop motion | Bed Patient Positioning Hw App fails to stop motion | Data corruption Software timing Errors due to unexpected hardware components or failures Interface design error | Patient positoning information fails to stop resulting in collision between parts of system & affects RCM Position | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process motion stop command due to runtime error or hardware issue resulting in undesired patient positioning leading to potential collision | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Velocity check on the joints when in stopping state, verify 0 motion after TBD seconds. 2. Bed model app checks if bed patient positioning HW is in stopping state.
Patient position HW App | PatientPositioningHWApp-swRA-00014 | Bed Patient Positioning Hw App - Fault response | A fault has occured and Robotic motion is not stopped and fault response does not occur | Data corruption Data parsing error Interface design error | Patient positioning continues to move when a portion of the system is in fault. Arm cannot maintain RCM resulting in RCM motion. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process fault response due to runtime error or data corruption resulting in undesired patient positioning leading to unintended RCM movement | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Velocity check on the joints when in stopping state, verify 0 motion after TBD seconds. 2. Bed model app checks if bed patient positioning HW is in stopping state.
Patient position HW App | PatientPositioningHWApp-swRA-00015 | BedPatientPositiningHwApp - Fault detection | A fault has occurred in the system and the software does not detect the fault. Subsequently, the SW does not initiate a fault response | Data corruption Software timing Errors due to unexpected hardware components or failures Failures to perform the correct error processing and recovery | Patient positioning moves unintentionally ex. motor stuck or motor runaway, but we fail to detect it resulting RCM motion. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to detect or report fault condition due to runtime error, data corruption or hardware issue resulting in undesired patient positioning leading to unintended RCM movement | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Verification check on the individual bits of the status and mode of operation
HID model App | HIDModelApp-swRA-00001 | HIDModelApp - Startup | HIDModelApp fails to correctly initialize software components and is inoperable | Data corruption, incorrect SW installation | During initialization, HIDModelApp shuts down, and can't power up HID arms | HHM-0255 | Physician is unable to begin procedure on time | Software fails to initialize software components due to data corruption or incorrect software installation | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. During boot up/Initiialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the application for successful startup
HID model App | HIDModelApp-swRA-00002 | HIDModelApp - Configuration Data | HIDModelApp receives incorrect or incomplete configuration data from any of the multiple sources | Data corruption, incorrect SW installation, interface design error | Incorrect or incomplete system configuration is used for computing motion commands and may result in incorrect or undesired motion of HID and patient side arm | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software receives incorrect or incomplete configuration data due to data corruption, incorrect software installation, or interface design error | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Configuration data will be checked with CRC and values will be checked to be within acceptable limits. 2. HID Model App will not completely initialize until configuration data is fully received.
HID model App | HIDModelApp-swRA-00003 | HIDModelApp - Configuration Data | HIDModelApp fails to correctly apply configuration | Incorrect data definition, interface design errors, errors in handling or interpreting data, data corruption | Incorrect initialization of internal components, may result in incorrect or undesired motion of HID and patient side arm | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to apply configuration due to incorrect data definition, interface design errors, errors in handling or interpreting data, or data corruption | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | During startup, the software will check that all internal components are initalized correctly and are of correct type/behavior.
HID model App | HIDModelApp-swRA-00004 | HIDModelApp - Powerup HID | HIDModelApp fails to send the power up signal to the corresponding HID arm | Incorrect algorithm, control flow errors, network issues | HID arm inoperable due to sw failure to power it up | HHM-0255 | Physician is unable to begin procedure on time | Software fails to send power up signal to the corresponding HID arm due to ncorrect algorithm, control flow errors, or network issues | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Ensure state change to power up as well as receipt of power up acknowledgement from hw in a timely manner.
HID model App | HIDModelApp-swRA-00005 | HIDModelApp - State transitions | HIDModelApp fails to transition to appropriate states, or transitions to wrong state(s) when commanded | Failure to set the correct initial and following states, control flow errors | HID arm joints do not move according to commanded motion, or may result in incorrect or undesired motion of HID and patient side | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to transition to appropriate states, or transitions to wrong state(s) when commanded due to failure to set the correct initial and following states, or control flow errors | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Unhandled state transitions lead to a minor fault.
HID model App | HIDModelApp-swRA-00006 | HIDModelApp - Send hardware state transitions | HIDModelApp fails to send appropriate state transitions to hardware app | Incorrect algorithm, control flow errors, network issues | HID arm joints do not move according to commanded motion, or may result in incorrect or undesired motion of HID. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to send appropriate state transitions to hardware app due to incorrect algorithm, control flow errors, or network issues | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. State transition messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Monitor for timely hardware transitions and throw fault if they do not occur.
HID model App | HIDModelApp-swRA-00007 | HIDModelApp - Recieve Hardware Telemetry | HIDModelApp fails to receive hardware telemetry | Data corruption, interface design errors, network issues | Incorrect motion calculation can result in incorrect or undesired motion of HID. Can result in incorrect or undesired motion of patient side end effector if in DRIVE state (active tele-op). | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to receive hardware telemetry due to data corruption, interface design errors, or network issues | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Telemetry data received over RTI will be CRC checked. 2. Timeout faults for failing to receive hardware telemetry.
HID model App | HIDModelApp-swRA-00008 | HIDModelApp - Recieve Hardware Telemetry | HIDModelApp receives hardware telemetry incorrectly | Incorrect data definition, Interface design errors, Errors in handling or interpreting data | Incorrect motion calculation can result in incorrect or undesired motion of HID. Can result in incorrect or undesired motion of patient side end effector if in DRIVE state (active tele-op). | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software receives incorrect hardware telemetry due to incorrect data definition, interface design errors, or errors in handling or interpreting data | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Telemetry data received over RTI will be CRC checked. 2. Timeout faults for failing to receive hardware telemetry. 3. Joint limit/motion checks on hardware side. 4. If translational or rotational motions of HID exceed a specified value a fault is thrown and instrument motion is stopped (ignores HID motion).
HID model App | HIDModelApp-swRA-00009 | HIDModelApp - Receive patient side pose | HIDModelApp fails to receive patient side end effector pose from HIDApp | Data corruption, interface design errors, network issues | Incorrect end effector pose is used for computing haptic force (if in DRIVE), resulting in incorrect or undesired motion of HID and patient side end effector. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to receive patient side end effector pose from HIDApp due to data corruption, interface design errors, or network issues | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Timeout fault for failing to receive patient side end effector pose.
HID model App | HIDModelApp-swRA-00010 | HIDModelApp - Receive patient side pose | HIDModelApp incorrectly receives patient side end effector pose from HIDApp | Incorrect data definition, interface design errors, errors in handling or interpreting data, data corruption | Incorrect end effector pose is used for computing haptic force (if in DRIVE), resulting in incorrect or undesired motion of HID and patient side end effector. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software receives incorrect patient side end effector pose from HIDApp due to incorrect data definition, interface design errors, errors in handling or interpreting data, or data corruption | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Messages received over RTI will be CRC checked. 2. Self checks on received side end effector pose (incremental motion when in teleop). 3. Computed haptic torques are saturated.
HID model App | HIDModelApp-swRA-00011 | HIDModelApp - Publish telemetry | HIDModelApp fails to publish telemetry | Data corruption, incorrect algorithm, interface design errors, network issues | Can result in incorrect or undesired motion of patient side end effector if in DRIVE state (active tele-op). | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to publish telemetry due to Data corruption, incorrect algorithm, interface design errors, network issues | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Messages are treated as reliable in the RTI network and are guaranteed to arrive. 2. Timeout fault in HapticMasterApp if telemetry not received from model app, and no new motion commands sent to lap app.
HID model App | HIDModelApp-swRA-00012 | HIDModelApp - Send joint commands to hardware app | HIDModelApp fails to send joint commands to hardware app | Data corruption, incorrect algorithm, interface design errors, network issues | Commands are not received by HID hardware and HID arms cannot be moved. | HHM-0285 | System freezes and becomes unresponsive to all user commands in the middle of surgery procedure when Patient is in contact with Instruments | Software HIDModelApp fails to send joint commands to hardware app due to data corruption, incorrect algorithm, interface design errors, or network issues | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Fault if HIDModelApp does not successfully send commands to the hardware app 2. Include timestamps in published joint commands
HID model App | HIDModelApp-swRA-00013 | HIDModelApp - Inverse kinematics computations for gimbal | HIDModelApp incorrectly computes inverse kinematics for gimbal | Data corruption, interface design errors, control flow errors, incorrect algorithm | Incorrect motions are commanded to HID, resulting in incorrect or undesired motion of HID gimbal and misalignment between gimbal and instrument end effector. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly computes inverse kinematics for gimbal due to data corruption, interface design errors, control flow errors, or incorrect algorithm | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Calculated joint commands are checked to be within joint limits.
HID model App | HIDModelApp-swRA-00014 | HIDModelApp - Gravity compensation | HIDModelApp incorrectly computes torques to compensate for gravity | Data corruption, interface design errors, control flow errors, incorrect algorithm | Incorrect gravity compensation for HID, resulting in incorrect or undesired motion of HID and undesired/uncontrolled motion of patient side end effector. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly computes torques to compensate for gravity due to Data corruption, interface design errors, control flow errors, or incorrect algorithm | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Software shall perform POST to verify correct gravity compensation.
HID model App | HIDModelApp-swRA-00015 | HIDModelApp - Forward kinematics computations | HIdModelApp incorrectly computes forward kinematics | Data corruption, interface design errors, control flow errors, incorrect algorithm | Incorrect motions are computed by HIDModelApp and incorrect motion commands are sent to LapApp, resulting in incorrect or undesired motion of patient side end effector. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly computes forward kinematics due to data corruption, interface design errors, control flow errors, or incorrect algorithm | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. If translational motions of HID exceed a specified value a fault is thrown and instrument motion is stopped (ignores HID motion). 2. Calculated motions are checked to be within bounds for incremental motion
HID model App | HIDModelApp-swRA-00016 | HIDModelApp - Inverse kinematics computations for gimbal | HIDModelApp computes gimbal motions that result in singular joint configurations | Incorrect algorithm | Gimbal lock results in loss of orientation motion in one axis until gimbal is driven out of singularity | HHM-0378 | Loss of jaw motion and/or articulation | Software computes gimbal motions that result in singular joint configurations due to incorrect algorithm | Always - 100% | Surgery | Reversible - (S2) | Sometimes | Medium | Redundancy resolution algorithm computes motion of redundant joint (G3) to prevent singular configurations
HID model App | HIDModelApp-swRA-00017 | HIDModelApp - Haptic Force | HIDModelApp incorrectly computes torques for haptics feedback | Data corruption, interface design errors, control flow errors, incorrect algorithm | Incorrect haptic forces/torques are rendered to the user, resulting in undesired motion of HID and patient side. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly computes torques for haptics feedback due to data corruption, interface design errors, control flow errors, or incorrect algorithm | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Computed haptic torques are saturated.
HID model App | HIDModelApp-swRA-00018 | HIDModelApp - fault response | A fault has occured and fault response does not occur | Failure to perform error processing and recovery | HID or patient side robot motion is not stopped and unexpected motion will occur | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to respond to a fault due to not performing error processing and recovery | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Faults within the HIDModelApp are handled locally and at each cycle, as necessary. 2. Fault messages are treated as reliable in the RTI network and are gauranteed to arrive.
HID model App | HIDModelApp-swRA-00019 | HIDModelApp - fault detection | A fault has occurred in the system and the software does not detect the fault. | Data corruption or parsing error when evaluating faults, failure to perform error processing | HID or patient side robot motion is not stopped and unexpected motion will occur | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to detect a fault due to data corruption or parsing error when evaluating faults, or failure to perform error processing | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Fault checks run within each control cycle; if one is not handled correctly, it is checked again on the next cycle.
ErtherCAT API | AsrEtherCAT-swRA-00001 | EtherCAT API - Interface Functions: Initialize and StartCallbackOperation during boot up and initialization of the etherCAT communication. | AsrEtherCAT fails to startup. | No response or corrupted response from the hardware | Before anesthesia, system becomes non-operational after the initialization sequence fails to complete. | HHM-0255 | Physician is unable to begin procedure on time | NO HAZARDOUS SITUATION and NO HARM because the failure mode occurs before anesthesia, system becomes non-operational during initialization. | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. HW App shall retry based on the MAXIMUM_#_HW_COMM_RETRIES. 2. HW App shall throw a Critical fault. 3. *Workflow mitigation (ensure system initialized correctly prior to putting patient on Surgical Table.)*
ErtherCAT API | AsrEtherCAT-swRA-00002 | EtherCAT API - Mailbox Functions: CoeDownload for sending ACYCLIC messages to the firmware drives, and CoeUpload for receiving messages from the firmware drives. | Failure to send correct calibration data, correct motor parameters, and motion control parameters during tele-op | No response or corrupted response from the EtherCAT hardware. | ADM is unable to hold position, or overshoot and undershoot the robotic arms causing the intended motion. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | This single point of failure is caused by the EtherCAT hardware. | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. AsrEtherCAT shall timeout within 1.5 seconds if there is no response to the CAN application protocol over EtherCAT (CoE) call. 2. HW App shall throw a Critical fault.
ErtherCAT API | AsrEtherCAT-swRA-00003 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | An error in the "Start Callback Operation Failure" causing AsrEtherCAT fails to wake up | 1) Too much jitter on the system 2) CPU utilization limit | Inability to catch errors and stop robotic arm on time | HHM-0051 | Excessive force applied at incision site in Patient | This single point of failure is caused by AsrEtherCAT fails to wake up. | Always - 100% | Chest Injury | Life Threatening - (S5) | Extremely rare | High | 1. AsrEtherCAT shall be able to detect communication loss and jitter. 2. A Major fault will be thrown (When communication loss is detected, AsrEtherCAT shall trigger a major fault and then transition the system to a safe state within 3 seconds: rapidly stop motion, then engage the mechanical brake (if available): Error → Fault → Safe Torque Off (STO) → Stop → Brake.)
ErtherCAT API | AsrEtherCAT-swRA-00004 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | Failure to wake up | QNX Timer failed | The system provides no response to user's inputs. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | This single point of failure is caused by a failed QNX Timer. If the AsrEtherCAT fails to wake up, the system becomes non-functional during tele-op. | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. There will be a EtherCAT Cyclic watchdog, if it detects a failure, EtherCAT API shall throw a Major fault.
ErtherCAT API | AsrEtherCAT-swRA-00005 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | Failure to send motion control data out on time to (ADM Firmware, Arm Distal Firmware, or Arm Proximal Firmware) | 1) The application code mutex locks causing priority promotion on slow threads 2) The application algorithms have an edge case bug where they take longer to calculate | Due to the delay of the robotic control messages, system may have vibration during tele-op. | HHM-0342 | Vibration of RCM/Instrument end effector occurs in Patient | This failure mode is a single point of failure. Due to the software fails to send the robotic control messages on time, surgeon may experience vibration and discomfort during tele-op. | Always - 100% | Abdominal Injury | Reversible - (S2) | Unusual | Low | 1. If the motor controllers does not receive data after 3 (?) EtherCAT cycles, the motor controllers shall fault locally and subsequently HwApp shall trigger a Major fault.
ErtherCAT API | AsrEtherCAT-swRA-00006 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | Incoming motion control data is corrupted | Acontis (SOUP) fails to process the CRC | Incorrect feedback values are received, unintended motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Single point of failure caused by Acontis (SOUP), i.e., failing to detect a bad CRC check or no CRC check | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. AsrEtherCAT shall perform data integrity checks on the feedback data in the EtherCAT API code and update the data status. 2. HwApp shall check for the invalid data status, discard the invalid data and throw a Major fault if the invalid data persists for 3 (?) EtherCAT cycles.
ErtherCAT API | AsrEtherCAT-swRA-00007 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | Outgoing motion control data is corrupted. | EtherCAT hardware failure causing data corruption | The microcontrollers will receive incorrect data, unintended motion | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Single-point failure caused by hardware. | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Motor controllers shall perform data integrity checks on the incoming data. 2. If the data integrity checks fail, the motor controllers shall fault locally and subsequently HwApp shall trigger a Major fault.
ErtherCAT API | AsrEtherCAT-swRA-00008 | EtherCAT API - Cyclic Device Function: Send and receive etherCAT packets in real time to and from firmware and trigger the application to process messages | Feedback motor data from firmware is not received by AsrEtherCAT (frame loss). | EtherCAT hardware failure causing frame loss | During tele-op, system becomes non-functional due no status for the current position of the foot pedals and motor controllers. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Single point failure caused by EtherCAT hardware. | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. AsrEtherCAT shall perform data integrity checks on the feedback data in the EtherCAT API code and update the data status. 2. HwApp shall check for the invalid data status, discard the invalid data and throw a Major fault if the invalid data persists for 3 (?) EtherCAT cycles.
Armrest ergonom FW | ArmrestErgo-swRA-00001 | Console Armrest Ergonomics Firmware - Initialization | The application fails to start | No valid image in flash detected | Unable to activate armrest and foot pedal ergonomics. Unable to use LEDs | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to process initialization of armrest and foot pedal ergo functionality at start-up due to invalid software image | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Configuration check within POST, outside of this module
Armrest ergonom FW | ArmrestErgo-swRA-00002 | Console Armrest Ergonomics Firmware - CANOpen communication | CAN bus not initialized | Software error, erroneous logic | Unable to activate armrest and foot pedal ergonomics. Unable to use LEDs | HHM-0354 | Physician/Hospital staff under stress/discomfort operating system | Software fails to process initialization of armrest and foot pedal ergo functionality at start-up due to software error or comm issue | Always - 100% | Joint Disorder | Intervention Necessary - (S3) | Unusual | High | 1. Check within POST, outside of this module
Armrest ergonom FW | ArmrestErgo-swRA-00003 | Console Armrest Ergonomics Firmware - CANOpen communication | CANOpen messages are lost | Software error, erroneous logic | Latched movement of armrest and foot pedal ergonomics. LED pattern transition failure | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process CAN bus command for armrest and foot pedal ergo control due to software error or comm issue | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Software watchdog will trigger fault if new data from armrest is not received within TBC time.
Armrest ergonom FW | ArmrestErgo-swRA-00004 | Console Armrest Ergonomics Firmware - CANOpen communication | CANOpen messages are delayed | Software error, erroneous logic | Latched movement of armrest and foot pedal ergonomics. Delayed LED pattern transition | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process CAN bus command promptly for armrest and foot pedal ergo control due to software error or comm issue | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Software watchdog will trigger fault if new data from armrest is not received within TBC time.
Armrest ergonom FW | ArmrestErgo-swRA-00005 | Console Armrest Ergonomics Firmware - CANOpen communication | CANOpen message processing is stopped | Software error, erroneous logic | Latched movement of armrest and foot pedal ergonomics. LED pattern transition failure | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to complete CAN bus command process for armrest and foot pedal ergo control due to software error or comm issue | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Software watchdog will trigger fault if new data from armrest is not received within TBC time. 2. Rocker switch data will including sample counter to indicate data is not stale
Armrest ergonom FW | ArmrestErgo-swRA-00006 | Console Armrest Ergonomics Firmware - CANOpen communication | Incorrect CANOpen messages sent to host | Software error, erroneous logic | Unintended movement of armrest and foot pedal ergonomics. | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process CAN bus command correctly for armrest and foot pedal ergo control due to software error or comm issue | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Each rocker switch has two complementary outputs and firmware transmits both so that their values can be verified
Armrest ergonom FW | ArmrestErgo-swRA-00007 | Console Armrest Ergonomics Firmware - Rocker switches | Wrong rocker switch value read | Software error, erroneous logic | Unintended movement of armrest and foot pedal ergonomics. | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process armrest and foot pedal ergo control data due to rocker switches in error | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Each rocker switch has two complementary outputs and firmware transmits both so that their values can be verified
Armrest ergonom FW | ArmrestErgo-swRA-00008 | Console Armrest Ergonomics Firmware - Rocker switches | Rocker switch processing stops | Software error, erroneous logic | Latched movement of armrest and foot pedal ergonomics. | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to complete rocker switch processing for armrest and foot pedal ergo control due to software runtime error | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Rocker switch data will include sample counter to indicate data is not stale
Armrest ergonom FW | ArmrestErgo-swRA-00009 | Console Armrest Ergonomics Firmware - USB Charger | USB Charger is not enabled | Software error, erroneous logic | No USB charging available | HHM-0347 | User annoyance when operating the system | Software fails to enable USB charging process in armrest and foot pedal ergo control due to software error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Read value of output register periodoically
Armrest ergonom FW | ArmrestErgo-swRA-00010 | Console Armrest Ergonomics Firmware - USB Charger | USB Charger fault is read incorrectly | Software error, erroneous logic | No impact, faults do not pose hazard to hardware | HHM-0347 | User annoyance when operating the system | Software fails to handle USB charging fault condition correctly in armrest and foot pedal ergo control due to software error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | No Mitigation
Armrest ergonom FW | ArmrestErgo-swRA-00011 | Console Armrest Ergonomics Firmware - HW Revision | HW version read incorrectly | Software error, erroneous logic | Unable to bringup system | HHM-0255 | Physician is unable to begin procedure on time | Software fails to confirm HW component version causing armrest and foot pedal ergo initialization failure at start-up due to software error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Check CRC of read hardware version. 2. Re-read version if CRC fails 3. Software POST Check
Armrest ergonom FW | ArmrestErgo-swRA-00012 | Console Armrest Ergonomics Firmware - LEDs | LED processing loop fails to execute | Software error, erroneous logic | LED pattern transition failure | HHM-0446 | User confusion when operating the system | Software fails to perform LED processing causing LED state transition failure due to runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Report fault if code is inactive for TBC seconds.
Armrest ergonom FW | ArmrestErgo-swRA-00013 | Console Armrest Ergonomics Firmware - LEDs | Incorrect bit pattern sent to LEDs | Software error, erroneous logic | Incorrect LED pattern displayed | HHM-0446 | User confusion when operating the system | Software fails to process LED bit pattern state transition correctly due to runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Code Review 2. Unit test
Armrest ergonom FW | ArmrestErgo-swRA-00014 | Console Armrest Ergonomics Firmware - Firmware update | Firmware image corrupted | Software error, erroneous logic | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware update due to invalid or corrupted firmware causing system or subsystem boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Retry firmware update
Armrest ergonom FW | ArmrestErgo-swRA-00015 | Console Armrest Ergonomics Firmware - Firmware update | Fails to write to storage | Software error, erroneous logic | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware update due to corrupted memory or runtime error causing system boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Rety firmware update
Fan controller FW | ConsoleFanController-swRA-00001 | Console Fan Controller Firmware - Initialization | The application fails to start | No valid image in flash detected | Unable to set G0 fan speed. Unable to detect leakage current | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform initialization at start-up for fan speed control due to invalid software or corrupted data | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Check within POST, outside of this module
Fan controller FW | ConsoleFanController-swRA-00002 | Console Fan Controller Firmware - SPI communication with Ingenia | SPI messages are lost | Software error, erroneous logic, freeze or crash | Unable to set G0 fan speed. Unable to detect leakage current | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to perform SPI messages process for fan speed control due to runtime error or comm issue | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Software can read back current fan speed 2. Packet sequence counter will stop incrementing if stale data. 3. Software watchdog will trigger fault if new data from fan controller is not received.
Fan controller FW | ConsoleFanController-swRA-00003 | Console Fan Controller Firmware - SPI communication with Ingenia | SPI messages are delayed | Software error, erroneous logic, freeze or crash | Expected delay for 1 to 10 seconds in setting fan speed. Delay in notification of leakage current | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to perform SPI messages process promptly for fan control due to runtime error or comm issue | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Software can read back current fan speed 2. Packet sequence counter will stop incrementing if stale data. 3. Software watchdog will trigger fault if new data from fan controller is not received.
Fan controller FW | ConsoleFanController-swRA-00004 | Console Fan Controller Firmware - SPI communication with Ingenia | SPI messsage processing stops | Software error, erroneous logic, freeze or crash | Unable to set G0 fan speed. Unable to detect leakage current | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to complete SPI messages process for fan speed control due to runtime error or comm issue | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Software can read back current fan speed 2. Packet sequence counter will stop incrementing if stale data. 3. Software watchdog will trigger fault if new data from fan controller is not received.
Fan controller FW | ConsoleFanController-swRA-00005 | Console Fan Controller Firmware - SPI communication with Ingenia | Wrong SPI message is sent to host | Software error, erroneous logic, freeze or crash | Incorrect informaton reported about leakage current | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process fan control status correctly due to runtime error or comm issue | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Software will require TBC number of consecutive packets indicating a short before triggering a fault.
Fan controller FW | ConsoleFanController-swRA-00006 | Console Fan Controller Firmware - Fan Controller | Wrong fan speed set | Software error, erroneous logic, freeze or crash | Hot touch point | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to process fan control status correctly due to runtime error or comm issue | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Software can read back current fan speed
Fan controller FW | ConsoleFanController-swRA-00007 | Console Fan Controller Firmware - Fan Controller | Fan control processing stops | Software error, erroneous logic, freeze or crash | Hot touch point | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to complete SPI messages process for fan control due to runtime error | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Software can read back current fan speed
Fan controller FW | ConsoleFanController-swRA-00008 | Console Fan Controller Firmware - Current leak deteciton | Software reads values from leak detection circuitry incorrectly | Software error, erroneous logic, freeze or crash | Incorrect informaton reported about leakage current | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to perform current leak detection for fan control due to runtime error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Firmware will read sensor data TBC times with TBC percent positive rate to confirm that a short exists before reporitng a fault.
Fan controller FW | ConsoleFanController-swRA-00009 | Console Fan Controller Firmware - Current leak deteciton | Current leak deteciton processing stops | Software error, erroneous logic, freeze or crash | Unable to detect leakage current | HHM-0093 | No hazardous situation | Software fails to perform current leak detect process to complete for fan control due to runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Fan controller data will include incrementing counter to show data is not stale, trigger fault after TBC time of counter not incrementing.
Fan controller FW | ConsoleFanController-swRA-00010 | Console Fan Controller Firmware - Firmware update | Firmware image corrupted | Software error, erroneous logic, freeze or crash | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware update due to invalid or corrupted firmware causing system or subsystem boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Retry firmware update
Fan controller FW | ConsoleFanController-swRA-00011 | Console Fan Controller Firmware - Firmware update | Fails to write to storage | Software error, erroneous logic, freeze or crash | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to process firmware update due to corrupted memory or runtime error causing system or subsystem boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Rety firmware update
Foot pedal FW | FootPedalFw-swRA-00001 | Console Foot Pedal Firmware - Initialization | The application fails to start | No valid image in flash detected | Unable to detect foot pedal presses | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to track or detect foot pedal status due to software error for initialization | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Configuration check within POST, outside of this module
Foot pedal FW | FootPedalFw-swRA-00002 | Console Foot Pedal Firmware - Ethercat | EtherCAT is not configured at startup | No valid eeprom data | Pedal press and release functionality lost. Pedal foot presence detection functionality lost. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to detect foot pedal presence to initialize functionality due to EtherCAT configuration failure at start-up | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Check within POST, outside of this module
Foot pedal FW | FootPedalFw-swRA-00003 | Console Foot Pedal Firmware - Ethercat | PDO/PDI messages are lost | Software error, erroneous logic, freeze or crash | Pedal press and release functionality lost. Pedal foot presence detection functionality lost. | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to process EtherCAT telemetry data due to software error causing foot pedal function lost | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Software watchdog will trigger fault if data from foot pedal is not received within TBC time.
Foot pedal FW | FootPedalFw-swRA-00004 | Console Foot Pedal Firmware - Ethercat | PDO/PDI messages are delayed | Software error, erroneous logic, freeze or crash | Pedal press and release may be delayed. Pedal presense detection may be delayed. | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to process EtherCAT telemetry data promptly causing foot pedal function delayed | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Software watchdog will trigger fault if data from foot pedal is not received within TBC time.
Foot pedal FW | FootPedalFw-swRA-00005 | Console Foot Pedal Firmware - Ethercat | PDO/PDI message processing stops | Software error, erroneous logic, freeze or crash | Pedal press and release functionality lost. Pedal foot presence detection functionality lost. | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to complete EtherCAT telemetry data processing causing foot pedal function lost | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Foot pedal data will include incrementing counter to show data is not stale, trigger fault after TBC time of counter not incrementing.
Foot pedal FW | FootPedalFw-swRA-00006 | Console Foot Pedal Firmware - Ethercat | Wrong PDI messages are sent to the RTPC | Software error, erroneous logic, freeze or crash | Pedal Press and Release may trigger unexpectedly. Pedal presence detection may trigger unexpectedly. | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software incorrectly processes EtherCAT telemetry data causing unintended foot pedal function | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Three bits are sent per foot pedal and foot presence sensor and only one of eight combinations can mean a sensor is activated.
Foot pedal FW | FootPedalFw-swRA-00007 | Console Foot Pedal Firmware - Ethercat | FCP messages are lost | Software error, erroneous logic, freeze or crash | LED pattern transition failure | HHM-0460 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software fails to perform firmware command protocol transaction causing LED light pattern error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. If software does not receive a response to an FCP command within TBC time, it will re-send that command. 2. After re-sending command X times, the system throw a fault
Foot pedal FW | FootPedalFw-swRA-00008 | Console Foot Pedal Firmware - Ethercat | FCP messages are delayed | Software error, erroneous logic, freeze or crash | Delayed LED pattern transition. | HHM-0460 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software fails to perform firmware command protocol transaction promptly causing LED light pattern error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. If software does not receive a response to an FCP command within TBC time, it will re-send that command. 2. After re-sending command X times, the system throw a fault
Foot pedal FW | FootPedalFw-swRA-00009 | Console Foot Pedal Firmware - Ethercat | FCP message processing stops | Software error, erroneous logic, freeze or crash | LED pattern transition failure. | HHM-0460 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software fails to complete firmware command protocol transaction causing LED light pattern error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. If software does not receive a response to an FCP command within TBC time, it will re-send that command. 2. After re-sending command TBC times, the system throw a fault
Foot pedal FW | FootPedalFw-swRA-00010 | Console Foot Pedal Firmware - Ethercat | Wrong FCP response messages are sent to RTPC | Software error, erroneous logic, freeze or crash | No effect on the system. | HHM-0093 | No hazardous situation | Software incorrectly performs firmware command protocol transaction causing LED light pattern error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. If software does not receive exepcted response to an FCP command, it will re-send that command.
Foot pedal FW | FootPedalFw-swRA-00011 | Console Foot Pedal Firmware - Foot presence detection | Presence sensor is read incorrectly | Software error, erroneous logic, freeze or crash | System reads foot presence sensor incorrectly | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software incorrectly performs foot pedal sensing data processing causing its presence error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Each presense sensor has two outputs, and both outputs must be in a specific state for the sensor to be activated
Foot pedal FW | FootPedalFw-swRA-00012 | Console Foot Pedal Firmware - Foot presence detection | Presence sensing processing loop fails to execute | Software error, erroneous logic, freeze or crash | System reads foot presence sensor incorrectly | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software incorrectly performs foot pedal sensing data processing causing its presence error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. If sensor data counter does not change within X time, trigger faults 2. Consider reporting fault if code is inactive for TBC seconds
Foot pedal FW | FootPedalFw-swRA-00013 | Console Foot Pedal Firmware - Foot pedal | Foot pedal is read incorrectly | Software error, erroneous logic, freeze or crash | System reads foot pedal state incorrectly | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to track foot pedal state correctly causing its presence or function error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Each foot pedal has two switches, and both switches must be in a specific state for the pedal to be considered pressed
Foot pedal FW | FootPedalFw-swRA-00014 | Console Foot Pedal Firmware - Foot pedal | Foot pedal processing loop fails to execute | Software error, erroneous logic, freeze or crash | System reads foot pedal state incorrectly | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to perform foot pedal loop process causing its presence or function error | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. Foot pedal data will include incrementing counter to show data is not stale, trigger fault after TBC time of counter not incrementing.
Foot pedal FW | FootPedalFw-swRA-00015 | Console Foot Pedal Firmware - LEDs | Incorrect color or pattern is displayed on foot pedal | Software error, erroneous logic, freeze or crash | Incorrect color or pattern is displayed on foot pedal | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to display commanded foot pedal LED light color or pattern | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. MCU will read state of LEDs from LED driver and this data will be made available to RTPC for checking pedal colors are what they were intended to be 2. RTSW to maintain LED State and run comparison against FW with FCP Check.
Foot pedal FW | FootPedalFw-swRA-00016 | Console Foot Pedal Firmware - LEDs | LED processing loop fails to execute | Software error, erroneous logic, freeze or crash | LED pattern transition failure. | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software fails to execute state transition process loop for foot pedal LED light pattern | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare | High | 1. MCU will read state of LEDs from LED driver and this data will be made available to RTPC for checking pedal colors are what they were intended to be 2. RTSW to maintain LED State and run comparison against FW with FCP Check.
Foot pedal FW | FootPedalFw-swRA-00017 | Console Foot Pedal Firmware - Firmware update | Firmware image corrupted | Software error, erroneous logic, freeze or crash | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform firmware update due to data or memory error causing boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Retry firmware update
Foot pedal FW | FootPedalFw-swRA-00018 | Console Foot Pedal Firmware - Firmware update | Fails to write to storage | Software error, erroneous logic, freeze or crash | Old firmware is run. System unable to bootup | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform firmware update due to data or memory error causing boot-up failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Boot loader checks for valid image 2. CRC checking 3. Rety firmware update
HID HW App | HIDHWApp-swRA-00001 | HID HW App - Initialize HW components | HID HW App fails to correctly initialize hardware components and thus is inoperable. | Data corruption, Software timing, Errors due to unexpected hardware components or failures | HID HW app does not start up and the HID arms cannot be used to drive the patient side arms. | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization due to software file or data corruption resulting in application inoperable for arms control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times. Fault is thrown if cannot restart.
HID HW App | HIDHWApp-swRA-00002 | HID HW App - Initialization EtherCAT comms | HID HW App fails to correctly initialize EtherCAT communication. | Data corruption, incorrect SW installation | HID arms do not power up and cannot be used to drive the patient side arms. | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization of EtherCAT comm due to software file or data corruption resulting in power-up failure leading to arms control inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times, if EtherCAT fails to establish communication the Sofware transitions into fault state.
HID HW App | HIDHWApp-swRA-00003 | HID HW App - Configuration data | HID HW App receives incorrect configuration data from any of the multiple sources. | Data corruption, incorrect SW installation, interface design error. | HID HW App uses incorrect or incomplete system configuration to compute motion commands, resulting in incorrect or undesired motion of HID arms and patient side arms (if in active teleop). Unintended Motion of HID and Patient Side Arms | HMM-0322 | Unable to address Operating Room emergency in a timely manner | Software computes and processes motion commands with incorrect or incomplete configuration data due to sw file or data corruption resulting in undesired HID and patient side arms motion | Always - 100% | Organ Failure | Life Threatening - (S5) | Unusual | RAR | 1. Configuration data from Configuration Server is CRC checked for complete data. NVM on hardware will also have a CRC check. 2. Config Verifier checks that values are within expected range.
HID HW App | HIDHWApp-swRA-00004 | HID HW App - output commands to EtherCAT | HID HW App fails to communicate with EtherCAT communications and processed commands are not sent to HID hardware | Data corruption, Software timing, Errors due to unexpected hardware components or failures | HID arms do not move and cannot be used to drive patient side arms. Unintended or No Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process output commands to EtherCAT due to runtime error, data corruption or comm bus issue resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. When EtherCAT loses communication the software transitions into fault state.
HID HW App | HIDHWApp-swRA-00005 | HID HW App - output commands to EtherCAT | HID HW App incorrectly outputs commands to EtherCAT | Data corruption, Data parsing error, Interface design error | Unintended Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes output commands to EtherCAT due to runtime error or data corruption resulting in undesired HID and patient side arms motion | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Limit checks on joint commands (velocity, acceleration, current limits, tracking error).
HID HW App | HIDHWApp-swRA-00006 | HID HW App - receive joint commands | HID Hw App incorrectly receives joint commands | Data corruption, Data parsing | Unintended Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) |  | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Joint commands received from HIDModelApp over RTI are CRC checked. Sanity checks on joint commands (velocity, acceleration limits, joint limits).
HID HW App | HIDHWApp-swRA-00007 | HID HW App - input from EtherCAT | HID HW App fails to communicate with EtherCAT communications and HW status is not received. | Data corruption, Software timing, Errors due to unexpected hardware components or failures | Unintended or No Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process feedback data from HW component (motor, encoder) due to runtime error, data corruption or comm bus issue resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. When EtherCAT loses communication the software transitions into fault state.
HID HW App | HIDHWApp-swRA-00008 | HID HW App - input from EtherCAT | HID HW App incorrectly processes motor, encoder and grasper hall sensor status information received over EtherCAT | Data corruption, Data parsing, Interface design error | Unintended or No Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes feedback data from HW component (motor, encoder) due to runtime error or data corruption resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. HW status CRC checked for complete data? Check timestamps of received data. There shall be checks on motor and encoder statuses over one time step.
HID HW App | HIDHWApp-swRA-00009 | HID HW App - input from EtherCAT | HID HWApp incorrectly processes clutch pedal, camera pedal, viewer head sensor status | Data corruption, Data parsing, Interface design error | Unintended or No Motion of HID and Patient Side Arms | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes feedback data from HW sensors (pedal/finger clutch, camera, viewer) due to runtime error or data corruption resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. HW status CRC checked for complete data. 2. Footpedal mapping comes from config file which is CRC checked for completeness.
HID HW App | HIDHWApp-swRA-00010 | HID HW App - input from EtherCAT | HID HWApp incorrectly processes energy pedal statuses | Data corruption, Data parsing, Interface design error | Energy cannot be activated OR undesired energy activation | HHM-0035 | Energy delivered to wrong area, or Energy application longer than expected in Patient | Software incorrectly processes status feedback from energy pedal due to runtime error or data corruption resulting in undesired or no energy activation | Always - 100% | Burn | Life Threatening - (S5) | Extremely Rare |  | 1) HW status CRC checked for complete data. 2) Footpedal mapping comes from config file which is CRC checked for completeness.
HID HW App | HIDHWApp-swRA-00011 | HID HW App - input from EtherCAT | HID HWApp incorrectly processes swap pedal status | Data corruption, Data parsing, Interface design error | Cannot switch or unintentionally switch HID control to secondary patient side arm | HHM-0287 or HHM-0257 | System freezes and becomes unresponsive to all user commands in the middle of surgery procedure when Patient is in contact with Instruments | Software incorrectly processes status feedback from swap pedal due to runtime error or data corruption resulting in undesired or no switch for HID and side arms | Always - 100% | Chest Injury | Permanent - (S4) | Unusual |  | 1. HW status CRC checked for complete data. 2. Footpedal mapping comes from config file which is CRC checked for completeness.
HID HW App | HIDHWApp-swRA-00012 | HID HW app - motion state | HID HW app fails to process or incorrectly processes incoming messages and does not process transition data or motion into motion mode commands | Network issues, data parsing error, data corruption | HID arm does not transition to proper motion mode. Can result in inability to move HID arms or unintended motion of HID arms. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process motion state transition due to runtime error or network issue resulting in undesired or no motion of HID or side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. Motion command is a single message with a CRC check. 2. HID HW App state change checked by HID model app.
HID HW App | HIDHWApp-swRA-00013 | HID HW app - fault response | A fault has occured and fault response does not occur | Failure to perform error processing and recovery | HID motion is not stopped and unintended motion will occur, can result in unintended motion of patient side (if in active teleop) | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to process fault response for error handling due to runtime error resulting in undesired patient side arms motion | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Faults within the HID Hw App are handled locally at each cycle. Fault messages are guaranteed to arrive via RTI.
HID HW App | HIDHWApp-swRA-00014 | HID HW app - fault detection | A fault has occurred in the system and the software does not detect the fault. Subsequently, the SW does not initiate a fault response | Data corruption, Data parsing error, Failure to perform error processing, SOUP with Ingenia and/or Ethercat | HID is not stopped and unintended motion will occur, can result in unintended motion of patient side (if in active teleop) | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to detect fault and process for error handling due to runtime error or comm issue resulting in undesired patient side arms motion | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Fault checks run within each control cycle.
HID HW App | HIDHWApp-swRA-00015 | HID HW App - Console Viewer Initialization | HID HW App fails to correctly initialize Console Viewer | Data corruption, incorrect SW installation | During initialization, the HID HW app shuts down and/or is not able to detect head engagement. | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization due to software file or data corruption resulting in application inoperable for arms control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times, if unable to restart, a fault shall be thrown.
HID HW App | HIDHWApp-swRA-00016 | HID HW App - Console Viewer EtherCAT comms | HID HW App fails to correctly initialize EtherCAT communication for the Console Viewer | Data corruption, incorrect SW installation | During initialization, the HID HW app shuts down and/or is not able to detect head engagement. | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization of EtherCAT comm due to software file or data corruption resulting in power-up failure leading to arms control inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times, if EtherCAT fails to establish communication, the Sofware shall transition into a fault state.
HID HW App | HIDHWApp-swRA-00017 | HID HW App - Console Viewer processing input from Ethercat | HID HW App - Console Viewer fails to correctly process the PDI message from firmware | Data corruption, incorrect SW installation, Failure to perform error processing, Software error, erroneous logic, freeze or crash | Head sensor engagement and disengagement detection may trigger unexpectedly. Joystick values for the Viewer device may be detected incorrectly. | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes feedback data from HW sensors (pedal/finger clutch, camera, viewer) due to runtime error or data corruption resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual |  | 1. HW status CRC shall be checked for complete data. 2. Timestamps shall be checked for received data. 3. Viewer mapping shall come from the config file and shall be CRC checked for completeness.
HID HW App | HIDHWApp-swRA-00018 | HID HW App - Console Viewer fault detection | A fault has occurred in the system and the software does not detect the fault. Subsequently, the SW does not initiate a fault response | Data corruption, incorrect SW installation, Failure to perform error processing, Software error, erroneous logic, freeze or crash | Dropped frames, video errors, under and over voltage errors incorrectly reported or not reported at all. HID motion is not stopped and unintended motion will occur, can result in unintended motion of patient side (if in active teleop) | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly processes feedback data from HW sensors (pedal/finger clutch, camera, viewer) due to runtime error or data corruption resulting in undesired or no motion of HID and patient side arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Faults within the HID Hw App/Viewer controller are handled locally at each cycle. Fault messages are guaranteed to arrive via RTI.
HID HW App | HIDHWApp-swRA-00019 | HID HW App - Foot Pedal Controller Initialization | HID HW App fails to correctly initialize Foot Pedal Controller | Data corruption, incorrect SW installation | During initialization, the HID HW app shuts down and/or is not able to detect Foot Pedal Press and Presence values for each foot pedal and presense sensors | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization due to software file or data corruption resulting in application inoperable for arms control, tool control or unintended energy apply/release to control tools and arms | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times, if the App cannot initialize a fault shall be thrown.
HID HW App | HIDHWApp-swRA-00020 | HID HW App - Foot Pedal EtherCAT comms | HID HW App fails to correctly initialize EtherCAT communication for Foot Pedals and Foot presence sensors | Data corruption, incorrect SW installation | During initialization, the HID HW app shuts down and/or is not able to detect Foot Pedal Press and Presence values for each foot pedal and presense sensors. | HHM-0255 | Physician is unable to begin procedure on time | Software fails initialization of EtherCAT comm due to software file or data corruption resulting in power-up failure leading to arms control inoperable, tool control inoperable or unintended energy apply/release to control tools and arms | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Restart X times, if EtherCAT fails to establish communication the Sofware shall transition into a fault state.
HID HW App | HIDHWApp-swRA-00021 | HID HW App - Foot Pedal Controller processing input from Ethercat | HID HW App - Foot Pedal Controller fails to correctly process the PDI message from firmware | Data corruption, incorrect SW installation, Failure to perform error processing, Software error, erroneous logic, freeze or crash. | Foot Pedal Press and Presence values for each foot pedal and presense sensors may be read incorrectly. | HHM-0450 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software incorrectly processes feedback data from foot pedal and pedal sensor HW sensors due to runtime error or data corruption resulting in adverse effect or undesired or no motion of patient side arms or unintended energy apply/release to control tools and arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare |  | 1. Sample Count shall be checked for complete data. 2. Timestamps shall be checked for received data. 3. Foot Pedal mapping shall come from the config file and shall be CRC checked for completeness.
HID HW App | HIDHWApp-swRA-00022 | HID HW App - Foot Pedal Controller fault detection | A fault has occurred on Foot Pedals and the software does not detect the fault. Subsequently, the SW does not initiate a fault response | Data corruption, incorrect SW installation, Failure to perform error processing, Software error, erroneous logic, freeze or crash | Dropped frames, foot pedal faults. unintended motion will occur, can result in unintended motion of patient side(if in active teleop) | HHM-0450 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software incorrectly processes feedback data from foot pedal and pedal sensor HW sensors due to runtime error or data corruption resulting in adverse effect or undesired or no motion of patient side arms or unintended energy apply/release to control tools and arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1. Faults within the Foot Pedal Controller are handled locally at each cycle. Fault messages are guaranteed to arrive via RTI.
HID HW App | HIDHWApp-swRA-00023 | HID HW App - Foot Pedal input staleness check | Stale data sent over PDI message for foot pedals and foot presence sensors | Fault/Error on Firmware,Dropped frames, Software error,erroneous logic, Data corruption | Dropped frames, foot pedal faults., can result in unintended motion/no motion of patient side(if in active teleop) | HHM-0450 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software incorrectly processes feedback data from foot pedal and pedal sensor HW sensors due to runtime error or data corruption resulting in adverse effect or undesired or no motion of patient side arms or unintended energy apply/release to control tools and arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1. Staleness checks on PDI data using sample count. 2. Verify the test pattern is sent periodically from Firmware.
HID HW App | HIDHWApp-swRA-00024 | HID HW App - Publish telemetry | HID Hw App- Foot edal Controller fails to publish foot pedal telemetry data | Network issues, RTI Buffer size issues/overflow/corrupt | Foot pedals press status and foot presence status is not set to current values by HID HW App. Unintended motion will occur, can result in unintended motion of patient side(if in active teleop) | HHM-0450 | Inappropriate energy method/characteristic/category delivered to tissue than desired in Patient | Software fails to send the foot pedal press status and foot presense status due to telemetry transmit or comm issue, data corruption or runtime error resulting in adverse effect or undesired or no motion of patient side arms or unintended energy apply/release to control tools and arms | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1. HID Hw App does not publish telemetry data over RTI that HapticMasterApp picks up to set the end effects of each foot pedal and foot presence sensors based on telemetry data. 2. Messages shall be CRC checked. 3. If HID HW App Fails to write to buffer, a fault shall be propogated. 4. If Foot Pedal is unable to be read for 1 Cycle while Energy Generation is applied, a signal shall be sent to system manager to stop Energy Generation.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00001 | SurgicalTableMobilityModelApp - Initialize | Fails to initialize | Data Corruption, Incorrect SW installation | Bed mobility cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform initialization due to data corruption or runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. During boot up/Initialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the app for successful startup
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00002 | SurgicalTableMobilityModelApp - Apply Configuration | Fails to apply configuration | Data Corruption, Incorrect SW installation, Interface design errors | Bed mobility cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to apply configuration due to data corruption or runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Configuration data are CRC checked for completeness and acceptable range 2. The app will not completely initialize until configuration information is fully received
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00003 | SurgicalTableMobilityModelApp - Apply Configuration | Incorrectly applies configuration | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software incorrectly applies configuration due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | During startup, the software will check that all internal components are initialized correctly and are of correct type/behavior. ConfigVerifier checks that the configuration values are within limits.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00004 | SurgicalTableMobilityModelApp - Initialize position from hardware application | Fails to initialize position from hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Bed mobility cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to initialize positioning data for HW components due to data corruption or runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | Transition to a fault state if the model is not correctly initialized from hardware.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00005 | SurgicalTableMobilityModelApp - Initialize position from hardware application | Incorrectly initializes position from hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software incorrectly initializes positioning data for HW components due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Checks that joints are within limits. Transitions to a fault state if they are not. 2. Checks that hardware telemetry joint positions is greater than or equal to the number of joints in the model. Transitions to a fault state if they are not. 3. During motion, motor controller tracking and current thresholds might be triggered if the wheel initial positions are incorrect.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00006 | SurgicalTableMobilityModelApp - Process command message from bed pendant | Fails to process command message | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Bed mobility cannot move. | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to process table motion command due to data corruption or runtime error | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Transition to a fault state if the command is not successfully processed.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00007 | SurgicalTableMobilityModelApp - Process command message from bed pendant | Incorrectly processes command message | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software incorrectly processes table motion command due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. CRC check is performed on bed pendant command messages 2. User can let go of motion enable button to stop undesired motion.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00008 | SurgicalTableMobilityModelApp - Switches mode between Assisted ActiveMobility, Manual Circular ActiveMobility, and Manual Straight ActiveMobility based on command message | Incorrectly switches mode | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Incorrect Algorithm | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software incorrectly processes table motion mode transition due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Checks desired mode from the Bed Pendant against the actual mode. 2. The current mode is displayed to the user on the Bed Pendant UI.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00009 | SurgicalTableMobilityModelApp - Compute wheel commands for components | Fails to compute wheel commands for components | Data corruption, Incorrect Algorithm | Bed mobility cannot move. | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to compute and process wheel motion command for HW components due to data corruption or runtime error | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Transitions to a fault state if the wheel command is not successfully computed such as if commanded position is not within limits.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00010 | SurgicalTableMobilityModelApp - Compute wheel commands for components | Incorrectly computes wheel commands for components | Data corruption, Incorrect Algorithm | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software incorrectly computes and processes wheel motion command for HW components due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Checks if commanded position is within limits. Transitions to a fault state if it is not. 2. Checks actual position/velocity and compare against the desired. If outside of a threshold, transitions to a fault state. 3. Motor controller tracking and current thresholds might be triggered if the wheel commands are incorrect. 4. User can let go of motion enable button to stop undesired motion.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00011 | SurgicalTableMobilityModelApp - Stop mobility by moving into the X-configuration when the hardware is not ready to move | Fails to stop mobility | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Incorrect Algorithm | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software fails to process and apply stop table motion command due to incorrect configuration data or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | Checks the hardware state and enters position hold if the hardware is not ready to move.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00012 | SurgicalTableMobilityModelApp - Stop mobility by moving into the X-configuration when a freeze or stop command is received including releasing the motion enable switch | Fails to stop mobility | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Incorrect Algorithm | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software fails to process and apply stop table motion command due to incorrect configuration data or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | Checks that the measured drive speeds go to zero and measured steering angles are close to the X-configuration. Transitions to a fault state if they are not.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00013 | SurgicalTableMobilityModelApp - Send signal to enable a component's hardware application | Fails to send enable signal to a component's hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bed mobility cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to enable HW components for start-up due to data corruption or runtime error | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | Tries powering up multiple times before a timeout. Throws a fault if timeout is exceeded.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00014 | SurgicalTableMobilityModelApp - Send signal to power down a component's hardware application | Fails to send power down signal to a component's hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bed battery drain due to mobility stayed powered on | HHM-0287 | System freezes and becomes unresponsive to all user commands in the middle of surgery procedure when Patient is in contact with Instruments | Software fails to process power-down signal to HW components due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | Tries powering down multiple times before a timeout. Throws a fault if timeout is exceeded.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00015 | SurgicalTableMobilityModelApp - Send signal to clear faults to a component's hardware application | Fails to send signal to clear faults to a component's hardware | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bed mobility cannot move. | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to monitor (detect and response) and process fault condition from HW components due to data corruption or runtime error | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Tries to clear faults multiple times before a timeout. Throws a fault if timeout is exceeded.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00016 | SurgicalTableMobilityModelApp - Stop motion and realign wheels if the scrub angle is greater than 5 degrees. | Stops mobility and realigns wheels when scrub angle is less than 5 degrees | Data corruption, Incorrect Algorithm | Stops mobility motion unintentionally | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to monitor and perform table motion stop for scrub angle check correctly due to data corruption or runtime error | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | After the bed stops due to a large scrub angle, the steering angle is reset to zero and the user can continue driving.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00018 | SurgicalTableMobilityModelApp - Fault Response | A fault has occurred and fault response does not occur | Failure to perform error processing and recovery | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software fails to perform error handling for fault response due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Faults within the SurgicalTableMobilityModelApp are handled locally and at each cycle, as necessary. 2. Fault messages are treated as reliable in the RTI network and are guaranteed to arrive.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00019 | SurgicalTableMobilityModelApp - Fault Detection | A fault has occurred in the system and the software does not detect the fault. | Data corruption or parsing error when evaluating faults, failure to perform error processing | Unintended bed mobility motion | HHM-0236 | Patient moves unexpectedly or falls from surgical table | Software fails to perform error handling for fault detection due to data corruption or runtime error | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | Fault checks run within each control cycle; if one is not handled correctly, it is checked again on the next cycle.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00020 | SurgicalTableMobilityModelApp - Mobilizes bed | Fails to mobilize bed | Data corruption, Incorrect Algorithm | Bed is not mobile | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to perform bed mobilization process for table motion control | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Checks that the bed is mobilized after commanded within a timeout. Transition to a fault state if that is not true.
Surgical table mobility model App | SurgicalTableMobilityModelApp-swRA-00021 | SurgicalTableMobilityModelApp - Immobilizes bed | Fails to immobilize bed | Data corruption, Incorrect Algorithm | Bed remains mobile | HHM-0229 | Patient moves unexpectedly or falls from surgical table | Software fails to perform bed immobilization process for table motion control | Always - 100% | Abdominal Injury | Intervention Necessary - (S3) | Unusual | High | 1. Checks that the bed is immobilized after commanded within a timeout. Transition to a fault state if that is not true. 2. Wheels will be in a position hold state, so the bed will not move on its own.
Surgical table model App | SurgicalTableModelApp-swRA-00001 | SurgicalTableModelApp - Initialize | Fails to initialize | Data Corruption, Incorrect SW installation, Software Timing Problem | Bar(s) and/or patient positioning cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform initialization process for table positioning control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. During boot up/Initialization, the system software will check that all run time software components are started, complete, and ready for full functionality 2. NodeApp will monitor the app for successful startup
Surgical table model App | SurgicalTableModelApp-swRA-00002 | SurgicalTableModelApp - Apply Configuration | Fails to apply configuration | Data Corruption, Incorrect SW installation, Interface design errors, Software Timing Problem | Bar(s) and/or patient positioning cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to apply configuration data for table positioning control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Configuration data are CRC checked for completeness and acceptable range 2. The app will not completely initialize until configuration information is fully received
Surgical table model App | SurgicalTableModelApp-swRA-00003 | SurgicalTableModelApp - Apply Configuration | Incorrectly applies configuration | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended bar(s) and/or patient positioning motion | HHM-0109 | Hospital staff collision with equipment or subsystems | Software incorrectly applies configuration data for table positioning control | Always - 100% | Bone Disorder | Permanent - (S4) | Extremely Rare | High | 1. During startup, the software will check that all internal components are initialized correctly and are of correct type/behavior. 2. ConfigVerifier checks that the configuration values are within limits.
Surgical table model App | SurgicalTableModelApp-swRA-00004 | SurgicalTableModelApp - Initialize position from hardware application | Fails to initialize position from hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Bar(s) and/or patient positioning cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform initialization commanded from HW App for table positioning control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | Throws a fault if the model is not correctly initialized from hardware.
Surgical table model App | SurgicalTableModelApp-swRA-00005 | SurgicalTableModelApp - Initialize position from hardware application | Incorrectly initializes position from hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Collision detection fails to prevent collision due to incorrect measured position. Undesired RCM motion due to incorrect bar position | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly performs initialization commanded from HW App for table positioning control | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Checks that joints are within limits. Throws a fault if they are not. 2. Checks that hardware telemetry joint positions is greater than or equal to the number of joints in the model. Throws a fault if they are not. 3. Checks expected gravity vector against those reported by the IMU on the bed bar breakout board. Throws a fault if the mismatch exceeds a threshold.
Surgical table model App | SurgicalTableModelApp-swRA-00006 | SurgicalTableModelApp - Process command message | Fails to process command message | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Bar(s) and/or patient positioning cannot move. Causes limited workspace and cannot be used for non-robotic | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process motion command for table positioning control | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | Throws a fault if the command is not successfully processed.
Surgical table model App | SurgicalTableModelApp-swRA-00007 | SurgicalTableModelApp - Process command message | Incorrectly processes command message | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended lateral bar motion causes patient collision. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software incorrectly processes motion command for table positioning control | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1. Checks that target component is a bed component. Throws a fault if not true. 2. Checks that motion mode is supported by the application. Throws a fault if it is not supported. 3. Collision verification prevents collision between subsystems.
Surgical table model App | SurgicalTableModelApp-swRA-00008 | SurgicalTableModelApp - Compute joint commands for components | Fails to compute joint commands for components | Data corruption, Incorrect Algorithm | Bar(s) and/or patient positioning cannot move. Causes limited workspace and cannot be used for non-robotic | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to compute joint positioning commands for table components motion control | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | Resets to the original joint positions when the SW fails to compute new joint commands and enters a stopped (Freeze) state.
Surgical table model App | SurgicalTableModelApp-swRA-00009 | SurgicalTableModelApp - Compute joint commands for components | Incorrectly computes joint commands for components | Data corruption, Incorrect Algorithm | Unintended lateral bar motion causes patient collision. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software incorrectly computes joint positioning commands for table components motion control | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1. Checks that joint command is not further violating joint limits. 2. Uses alpha limiting to limit the velocity of commands. 3. Collision verification prevents collision between subsystems.
Surgical table model App | SurgicalTableModelApp-swRA-00010 | SurgicalTableModelApp - Coordinate motion between bed components | Fails to coordinate motion between bed components | Data corruption, Software timing problem, Incorrect Algorithm | Lack of bar motion results in limited workspace. Will need to undock, move components, and redock. | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software fails to perform table components coordination process for table motion control due to runtime error or data corruption | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | Tries multiple times to put bars in drive_joint mode until timeout. Throws a fault if we exceed the timeout.
Surgical table model App | SurgicalTableModelApp-swRA-00011 | SurgicalTableModelApp - Coordinate motion between bed components | Incorrectly coordinates motion between bed components | Data corruption, Software timing problem, Incorrect Algorithm | Unintended lateral bar motion causes patient collision. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software incorrectly performs table components coordination process for table motion control due to runtime error or data corruption | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1. Ensures that bars are in drive_joint mode 2. Checks that patient position DOF command size is correct. Transitions to fault state if not true. 3. Uses alpha limiting to limit the velocity of commands 4. Dry runs coordinated motion commands. If HW cannot follow it, resets patient positioning command. 5. Uses the most conservative alpha limit for coordinated motion. 6. Collision verification prevents collisions between subsystems.
Surgical table model App | SurgicalTableModelApp-swRA-00012 | SurgicalTableModelApp - Send joint commands to collision detection | Fails to send joint commands to collision detection | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Bar(s) and/or patient positioning cannot move. Causes limited workspace and cannot be used for non-robotic | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to perform joint command sending process for table components collision detection | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | Throws a fault if the joint command is not successfully sent.
Surgical table model App | SurgicalTableModelApp-swRA-00013 | SurgicalTableModelApp - Send joint commands to collision detection | Incorrectly sends joint commands to collision detection | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Unintended lateral bar motion causes patient collision. | HHM-0404 | Patient collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software incorrectly performs joint command sending process for table components collision detection | Always - 100% | Soft Tissue Injury | Permanent - (S4) | Extremely rare | High | 1. Compares positions between model and model from feedback. If they diverge too far, throws a fault. 2. The motor controller has position and speed tracking checks which will protect against discontinuities in the command. If the check fails, the bed throws a fault.
Surgical table model App | SurgicalTableModelApp-swRA-00014 | SurgicalTableModelApp - Send last known collision-free joint command when collision rejection is received | Fails to set joints to collision-free command | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data | Collision occurs | HHM-0021 | Capital equipment (Tower, console, Bed) or subassemblies (Robotic Arm) collision with other capital equipment or environment | Software fails to perform joints setting process correctly for collision-free event resulting in table components collision | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | Collision verification uses a buffer to stop motion prior to actual collision occurring.
Surgical table model App | SurgicalTableModelApp-swRA-00015 | SurgicalTableModelApp -Send freeze (stop) command to hardware when a freeze command is received | Fails to stop motion of a component | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Collision occurs | HHM-0021 | Capital equipment (Tower, console, Bed) or subassemblies (Robotic Arm) collision with other capital equipment or environment | Software fails to perform motion freeze (stop) process correctly resulting in table components collision | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | Checks that the RobotControllers are in the Ready state after a freeze command is sent. Throws a fault if they do not transition.
Surgical table model App | SurgicalTableModelApp-swRA-00016 | SurgicalTableModelApp - Send signal to power up a component's hardware application | Fails to send power up signal to a component's hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bar(s) and/or patient positioning cannot move. | HHM-0255 | Physician is unable to begin procedure on time | Software fails to perform power-up command process to components hardware app for table motion control | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Checks if the RobotControllers have powered up. 2. Retries powering up after 5 seconds if the robot controllers have not powered up. 3. Throws a fault if powering up fails after the retry.
Surgical table model App | SurgicalTableModelApp-swRA-00017 | SurgicalTableModelApp - Send signal to power down a component's hardware application | Fails to send power down signal to a component's hardware application | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bed component's motors stays powered on. Producing heat. | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Software fails to perform power-down command process to components hardware app for table motion control | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1. Checks if the RobotControllers have powered down. 2. Retries powering down after 5 seconds if the robot controllers have not powered down. 3. Throws a fault if powering down fails after the retry.
Surgical table model App | SurgicalTableModelApp-swRA-00018 | SurgicalTableModelApp - Send signal to clear faults to a component's hardware application | Fails to send signal to clear faults to a component's hardware | Data corruption, Incorrect data definition, interface design errors, errors in handling or interpreting data, Software timing problem | Bar(s) and/or patient positioning cannot move. | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to perform error handling process to clear fault event from components hardware app | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Checks that RobotControllers have cleared their faults before moving on from a fault state 2. Retries clearing the fault if the state has not changed after TBD seconds.
Surgical table model App | SurgicalTableModelApp-swRA-00019 | SurgicalTableModelApp - Log telemetry | Fails to log telemetry for bed arm bars | Data corruption Software timing problem | Bed arm bar poses not sent to Lap Apps. Undesired RCM motion due to incorrect bar position | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software fails to perform components telemetry data logging process resulting in bed arm positioning status not reported to control host app | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | Lap App checks if bed hardware telemetry has been received
Surgical table model App | SurgicalTableModelApp-swRA-00020 | SurgicalTableModelApp - Log telemetry | Incorrectly logs telemetry for bed arm bars | Data corruption Software timing problem | Incorrect bed arm bar poses sent to Lap Apps. Undesired RCM motion due to incorrect bar position | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Software incorrectly performs components telemetry data logging process resulting in incorrect bed arm positioning status reported to control host app | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1. Compares positions between model and model from feedback. If they diverge too far, Throws a fault. 2. Lap App checks if the bed and bar have moved too far from their last positions. If they do, throws a fault.
Stow and deploy App | StowDeployApp-swRA-00001 | StowDeployApp-Initialize | StowDeployApp fails to Initialize | Data Corruption | Cannot perform stow deploy | HHM-0255 | Physician is unable to begin procedure on time | Software performs incorrect initialization due to data corruption | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | During boot up, system software will check that all run time software components are started, complete, and ready for full functionality
Stow and deploy App | StowDeployApp-swRA-00002 | StowDeployApp-Apply Configuration | Incorrect configuration is applied | data parsing error | StowDeploy app executes incorrect trajectory, and may cause collision | HHM-0013 | Breach in sterility | Software applies incorrect configuration due to data processing error | Always - 100% | Infection | Permanent - (S4) | Extremely Rare | High | 1. Config data are CRC checked for completeness and acceptable range.
Stow and deploy App | StowDeployApp-swRA-00003 | StowDeployApp-State Transition | StowDeploy state machine fails to transit to CommadStowDeploy state | Response time or delay, data parsing error, data Corruption | Cannot perform stow deploy | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software performs incorrect state transition due to runtime or data processing error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | If the system fails to transit to the CommandStowDeploy state in the predefined time threshold, a fault is triggered.
Stow and deploy App | StowDeployApp-swRA-00004 | StowDeploy App-Target Transition | Trajectory fails to change target pose when commanded | Incorrect Algorithm, Data parsing error | Incomplete execution of stow/deploy | HHM-0259 | Physician is unable to complete procedure in a timely manner | Software performs incorrect trajectory transition due to data or algorithm processing error | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. We are sending the trajectory command at every control tick. If the command message is not delivered on the current tick, we retry on the next tick. If the target transition doesn't happen in a predefined time threshold, a target transition timeout fault will be triggered. 2. If the system has deviated from the entrance criteria of the chosen target pose, motion is rejected. 3. Confirm that BedModelApp and LapApp are in position control mode to receive the motion request
Stow and deploy App | StowDeployApp-swRA-00005 | StowdeployApp-Trajectory Generation | Incorrect trajectory is generated | Incorrect Algorithm, Data Corruption | StowDeploy app executes incorrect stow/deploy trajectory, cause collision | HHM-0013 | Breach in sterility | Software performs incorrect trajectory generation due to data or algorithm processing error | Always - 100% | Infection | Permanent - (S4) | Extremely Rare | High | 1. All the trajectory data are associated with a target pose signature. Unless the desired target pose matches the trajectory data's target pose signature, no command is generated and a fault is triggered. 2. Undesired motion due to incorrect trajectory data can be stopped early by either releasing continuous activation button or pressing bedside touchpoints.
Stow and deploy App | StowDeployApp-swRA-00006 | StowDeployApp-Sending Trajectory | Fails to send trajectory to LapApp and BedModelApp | Response time or delay | Cannot perform stow deploy | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process trajectory transition data due to runtime error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | We are sending the trajectory command at every control tick. If the command fails on the current tick, we retry on the next tick.
Stow and deploy App | StowDeployApp-swRA-00007 | StowDeployApp-Sending Trajectory | Partially sends trajectory to LapApp and BedModelApp, not sending trajectories to all three lap apps and one bed arm bar | Response time or delay | Arms and bars move along an incorrect trajectory, can cause collision | HHM-0013 | Breach in sterility | Software fails to process trajectory transition data due to runtime error | Always - 100% | Infection | Permanent - (S4) | Extremely Rare | High | 1. All robots motion must be synchronized on predefined waypoints before executing the trajectory to move to the next waypoint. If not all robots are synchronized on that waypoint, then all robots will be commanded to go to current waypoint. In other words, the robots already reached the waypoint will stay where they are. 2. To mitigate further damage due to collision, stop motion upon contact detection. 3. Collision Verification, where enabled, prevents collisions between subsystems
Stow and deploy App | StowDeployApp-swRA-00008 | StowDeployApp-Publish Telemetry | Fails to publish StowDeploy telemetry | Response time or delay, data Corruption | Cannot perform stow deploy | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process trajectory telemetry data due to runtime or data processing error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. If StowDeployTelemetry was not published, the software treats all conditions to either false or unknown. Then the StowDeploy workflow model stops StowDeploy due to the met of exit criteria. 2. We are sending the StowDeployTelemetry at every control tick. If the StowDeployTelemetry publishing fails on the current tick, we retry on the next tick.
Stow and deploy App | StowDeployApp-swRA-00009 | StowDeployApp-Publish Telemetry | Published wrong StowDeploy telemetry | Data Corruption | Cannot perform stow deploy, or stow deploy does not stop timely | HHM-0013 | Breach in sterility | Software fails to process trajectory telemetry data due to data processing error | Always - 100% | Infection | Permanent - (S4) | Extremely Rare | High | 1. Undesired motion due to incorrect telemetry data can be stopped early by either releasing continuous activation button or pressing bedside touchpoints.
Stow and deploy App | StowDeployApp-swRA-00010 | StowDeployApp-Receive StowDeploy Command | Fails to receive stow deploy command from tower/pendant | Response time or delay, data corruption | Cannot perform stow/deploy | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Software fails to process bar stow/deploy motion command due to runtime or data processing error | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1.The software requires a periodical stow deploy command. If the command is missed in this control cycle, it will be picked up in the next one. 2. The software clears the current control command after a time limit, and if a new command is not received, a fault is triggered.
Stow and deploy App | StowDeployApp-swRA-00011 | StowDeployApp-Process StowDeploy Command | Failed to processes stow/deploy command | Response time or delay, data parsing error, data Corruption | The motion doesn't stop when commanded to and may cause collisions with bed-sides accessories or support staff.. | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software fails to process bar stow/deploy motion command due to runtime or data processing error | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1.The software requires a periodical StowDeployCmd. If the command is missed in this control cycle, it will be picked up in the next one. 2. The software clears the current control command after a time limit, and if a new command is not received, a fault is triggered 3. To mitigate further damage in case of collision, stop motion upon contact detection.
Stow and deploy App | StowDeployApp-swRA-00012 | StowDeployApp - Fault Detection | A fault has occurred and StowDeployApp doesn't recognize it | Error parsing failure, Data Corruption | Cannot stop stow/deploy timely, collision may happen | HHM-0021 | Capital equipment (Tower, console, Bed) or subassemblies (Robotic Arm) collision with other capital equipment or environment | Software fails to monitor for detection of internal fault condition due to runtime or data processing error | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | Faults within StowDeploy App are handled locally and at each cycle, as necessary. Fault messages are treated as reliable in the RTI network and are guaranteed to arrive.
Stow and deploy App | StowDeployApp-swRA-00013 | StowDeployApp - Fault Response | A fault is detected but StowDeployApp doesn't initiate an appropriate response | Error parsing failure, Response time or delay, Incorrect Algorithm | Cannot stop stow/deploy timely, collision may happen | HHM-0021 | Capital equipment (Tower, console, Bed) or subassemblies (Robotic Arm) collision with other capital equipment or environment | Software fails to monitor for response of internal fault condition due to runtime or algorithm processing error | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | Risk mitigation measure when Faults are detected is to stop all motion. Fault checks run within each control cycle; if one is not handled correct, it is checked again on the next cycle.
Stow and deploy App | StowDeployApp-swRA-00014 | StowDeployApp - Bar Lateral Distance Adjustment | Between Drape to Pre-Dock, assume user moves the Bar inwards too close to the Bed rails/accessories. A potential bar inward roll during Procedure Pose Deployment causes collision. | Incorrect Algorithm | Collison between Bar/Arms and Bed/Bedside Accessories/Patient Arms | HHM-0098 | Hospital staff / Physician collision with moving subassemblies of subsystem (e.g. Arm, Bar, Mantis) | Software offers unrestricted flexibility to the user to adjust lateral distance between the bar and the bed. | Always - 100% | Bone Disorder | Intervention Necessary - (S3) | Extremely Rare | High | 1. Collision can be pre-empted by stopping Continuous Activation ahead of time and increasing the lateral distance between the bar and the bed (or) reversing motion (or) stopping motion early using bedside touchpoints. 2. For dock pose deployment, define a minimum distance threshold between the bar and the table to accommodate inward roll adjustments.
Bed mobility HW App | BedMobilityHWApp-swRA-00001 | Bed Mobility Hw App - Initialize mobility HW and sensors | Bed Mobility Hw App fails to initialize mobility HW and sensors | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobility does not move | HHM-0255 | Column G: =VLOOKUP(F3, HHL!$A$1:$G$1000, 2) | Software fails to complete initialization process of hardware components due to runtime error or hardware issue resulting in bed mobility inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. Retry X times, if App still fails a fault will be thrown
Bed mobility HW App | BedMobilityHWApp-swRA-00002 | Bed Mobility Hw App - Apply Configuration | Bed Mobility Hw App fails to apply configuration | Data corruption Data parsing error Interface design error | Bed mobility does not move | HHM-0255 | #REF! | Software fails to apply configuration to HW control due to data corruption or runtime error resulting in bed mobility inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. App will throw a fault and motors move to disabled status
Bed mobility HW App | BedMobilityHWApp-swRA-00003 | Bed Mobility Hw App - Apply Configuration | Bed Mobility Hw App receives incorrect or incorrectly processes configuration data | Data corruption Data parsing error Interface design error | Bed mobility moves in unintended direction | HHM-0236 | #REF! | Software fails to process configuration data to HW control due to data corruption or runtime error resulting in undesired bed mobility | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Configuration data from Configuration Server is CRC checked for complete data 2. ConfigVerifier checks that configuration values are within range
Bed mobility HW App | BedMobilityHWApp-swRA-00004 | Bed Mobility Hw App - Process motion command from mobility model app | Bed Mobility Hw App fails to processes motion command from mobility model app | Data corruption of joint command message from RTI Data parsing error of joint command message Interface design error for joint command message | Bed mobility does not move | HHM-0259 | #REF! | Software fails to process motion command from mobility model app due to data corruption or runtime error resulting in bed mobility inoperable | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. App will throw a fault and motors move to disabled status
Bed mobility HW App | BedMobilityHWApp-swRA-00005 | Bed Mobility Hw App - Process motion command from mobility model app | Bed Mobility Hw App incorrectly processes motion command from mobility model app | Data corruption Data parsing error Interface design error | Bed mobility moves in unintended direction or unintentionally | HHM-0236 | #REF! | Software fails to process motion command from mobility model app correctly due to data corruption or runtime error resulting in undesired bed mobility | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Model App will check between command and measured, if different by X, a fault is thrown and motors move to disabled status
Bed mobility HW App | BedMobilityHWApp-swRA-00006 | Bed Mobility Hw App - Enable hardware | Bed Mobility Hw App fails to enable hardware | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobility does not enable hardware, so the bed mobility cannot move | HHM-0259 | #REF! | Software fails to enable hardware component for control due to runtime error or hardware issue resulting in bed mobility inoperable | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. App will throw a fault and motors move to disabled status 2. Check to enable HW during POST
Bed mobility HW App | BedMobilityHWApp-swRA-00007 | Bed Mobility Hw App - Disable hardware | Bed Mobility Hw App fails to disable hardware | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobiity does not disable hardware | HHM-0255 | #REF! | Software fails to disable hardware component for control due to runtime error or hardware issue resulting in undesired bed mobility | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. App will throw a fault 2. Drives are configured to move to a fault state if there is a loss of communication from host after X cycles. 3. Cyclic CAN communication will be disabled
Bed mobility HW App | BedMobilityHWApp-swRA-00008 | Bed Mobility Hw App - Initialize CAN communication (See CAN-specific line items) | Bed Mobility Hw App fails to initialize communication | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobility does not enter operational state | HHM-0255 | #REF! | Software fails to process initialization of CAN communication due to runtime error or hardware issue resulting in bed mobility inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. App will throw a fault 2. Drives are configured to move to a fault state if there is a loss of communication from host after X cycles. 3. Cyclic CAN communication will be disabled
Bed mobility HW App | BedMobilityHWApp-swRA-00009 | Bed Mobility Hw App - output commands to CAN | Bed Mobility Hw App fails to communicate with CAN communications and processed commands are not sent to the surgical table CAN bus or system | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobility does not move | HHM-0259 | #REF! | Software fails to process output commands to CAN due to runtime error or comm bus issue resulting in bed mobility inoperable | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. App will throw a fault 2. Drives are configured to move to a fault state if there is a loss of communication from host after X cycles. 3. Cyclic CAN communication will be disabled 4. Model App will have a 5 cycle, 200ms Watchdog on the HW App
Bed mobility HW App | BedMobilityHWApp-swRA-00010 | Bed Mobility Hw App - output commands to CAN | Bed Mobility Hw App incorrectly outputs commands to CAN | Data corruption Data parsing error Interface design error | Bed mobility moves unintentionally or in unintended direction | HHM-0236 | #REF! | Software incorrectly processes output commands to CAN due to runtime error or data corruption resulting in undesired bed mobility | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Check between Input and Output commands to Motor Axis Controller 2. Check between Model App commands and feedback
Bed mobility HW App | BedMobilityHWApp-swRA-00011 | Bed Mobility Hw App - output commands to CAN | Bed Mobility Hw App incorrectly outputs commands to CAN | Data corruption Data parsing error Interface design error | Bed mobility violates joint limit resulting in mechanical damage | HHM-0259 | #REF! | Software incorrectly processes output commands to CAN due to runtime error or data corruption resulting in bed mobility out of limit leading to HW component damage | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. Check between Input and Output commands to Motor Axis Controller 2. Check between Model App commands and feedback
Bed mobility HW App | BedMobilityHWApp-swRA-00012 | Bed Mobility Hw App - processes received status information via CAN | Bed Mobility Hw App fails to communicate with CAN communications and status is not received from the CAN bus. | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed mobility does not move | HHM-0259 | #REF! | Software fails to process status feedback due to runtime error or CAN bus issue resulting in bed mobility inoperable | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. Cyclic communication with CAN shall be stopped. 2. Motor Drives shall be configured such that if sync frames are not received in X cycles, they shall fault.
Bed mobility HW App | BedMobilityHWApp-swRA-00013 | Bed Mobility Hw App - processes received status information via CAN | Bed Mobility Hw App incorrectly processes status information. | Data corruption Data parsing error Interface design error | Bed mobility moves unintentionally (most severe). Can also move in unintended direction or cause large scrub angle leading to wear. | HHM-0236 | #REF! | Software incorrectly processes status feedback from CAN due to runtime error or data corruption resulting in undesired bed mobility | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Check between Input and Output commands to Motor Axis Controller 2. Check between Model App commands and feedback
Bed mobility HW App | BedMobilityHWApp-swRA-00014 | Bed Mobility Hw App - Immobilize bed | Bed Mobility Hw App fails to immobilize the bed | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed remains mobilized | HHM-0236 | #REF! | Software fails to process command for immobilize due to runtime error, data corruption or hardware issue resulting in undesired bed mobility | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. A Fault shall be thrown
Bed mobility HW App | BedMobilityHWApp-swRA-00015 | Bed Mobility Hw App - Mobilize bed | Bed Mobility Hw App fails to mobilize the bed | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed remains immobilized | HHM-0259 | #REF! | Software fails to process command for mobilize due to runtime error, data corruption or hardware issue resulting in bed mobility inoperable | Always - 100% | Surgery | Reversible - (S2) | Extremely Rare | Low | 1. A Fault shall be thrown
Bed mobility HW App | BedMobilityHWApp-swRA-00016 | Bed Mobility Hw App - Stop motion | Bed Mobility Hw App fails to stop motion | Data corruption Software timing Errors due to unexpected hardware components or failures Interface design error | Bed mobility moves unintentionally | HHM-0236 | #REF! | Software fails to process motion stop command due to runtime error or hardware issue resulting in undesired bed mobility leading to potential collision | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Velocity shall be monitored after stop command has been sent, if velocity does not reach 0 after X time, a fault shall be thrown (shutdown command)
Bed mobility HW App | BedMobilityHWApp-swRA-00017 | Bed Mobility Hw App - Fault response | A fault has occured and mobility motion is not stopped and fault response does not occur | Data corruption Data parsing error Interface design error | Bed mobility moves unintentionally | HHM-0236 | #REF! | Software fails to process fault response due to runtime error or data corruption resulting in undesired bed mobility leading to potential collision | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Velocity shall be monitored after a fault has been received, if velocity does not reach 0 after X time, a fault shall be thrown (shutdown command)
Bed mobility HW App | BedMobilityHWApp-swRA-00018 | Bed Mobility Hw App - Get battery status | Bed Mobility Hw App fails to correctly read battery status | Data corruption Software timing Errors due to unexpected hardware components or failures | Battery dies or thinks that it is dead while the bed is still mobile. Patient might fall off bed since it is still mobile. | HHM-0236 | #REF! | Software fails to track or report battery status due to runtime error, data corruption or hardware issue resulting in bed mobility incomplete | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. If HW detects a failure in communicating with battery, a fault shall be thrown, and motors drives shall be commanded to disabled state.
Bed mobility HW App | BedMobilityHWApp-swRA-00019 | Bed Mobility Hw App - Read Z-Burg encoders | Bed Mobility Hw App fails to read Z-burg encoders | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed does not mobilize | HHM-0255 | #REF! | Software fails to read or report sensing status from encoders due to runtime error, data corruption or hardware issue resulting in bed mobility inoperable | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1. If HW detects a failure in communicating with Z-Burg Encoders, a fault shall be thrown, and motors drives shall be commanded to disabled state.
Bed mobility HW App | BedMobilityHWApp-swRA-00020 | Bed Mobility Hw App - Read Z-Burg encoders | Bed Mobility Hw App incorrectly reads Z-burg encoders | Data corruption Data parsing error Interface design error | Bar collides with floor while bed is mobile | HHM-0021 | #REF! | Software incorrectly processes sensing data from encoders due to runtime error or data corruption resulting in undesired bed mobility leading to collision | Always - 100% | Surgery | Intervention Necessary - (S3) | Extremely rare | High | 1. Application shall check position feasibility of Z-Burg, if not feasible a fault shall be thrown
Bed mobility HW App | BedMobilityHWApp-swRA-00021 | Bed Mobility Hw App - Read signals from the Z-Burg emergency release covers | Bed Mobility Hw App incorrectly reads signals from the Z-Burg emergency release covers as open | Data corruption Software timing Errors due to unexpected hardware components or failures | All bed subsystems stop moving | HHM-0256 | #REF! | Software incorrectly processes sensing data as open from cover release signal due to runtime error or hardware issue resulting in bed mobility inoperable unintentionally | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. User shall be notified that emergency release covers are open.
Bed mobility HW App | BedMobilityHWApp-swRA-00022 | Bed Mobility Hw App - Read signals from the Z-Burg emergency release covers | Bed Mobility Hw App incorrectly reads signals from the Z-Burg emergency release covers as closed | Data corruption Software timing Errors due to unexpected hardware components or failures | Bed continuous to move when Z-Burg emergency release cover is open | HHM-0236 | #REF! | Software incorrectly processes sensing data as closed from cover release signal due to runtime error or hardware issue resulting in continued bed mobility for E-Stop | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. User shall be notified that emergency release covers are open. 2. If notification does not show up, user can still press the emergency release button to trigger E-stop.
Bed mobility HW App | BedMobilityHWApp-swRA-00023 | Bed Mobility Hw App - Read signals from the Z-Burg emergency release buttons | Bed Mobility Hw App incorrectly reads signals from Z-Burg emergency release buttons as pressed | Data corruption Software timing Errors due to unexpected hardware components or failures | All bed subsystems stop moving | HHM-0256 | #REF! | Software incorrectly processes sensing data as pressed from emergency button due to runtime error or hardware issue resulting in bed mobility inoperable unintentionally | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1. Throw a fault if there is a mismatch between redundant switch signals. 2. Notify the user if both redundant switch signals go high indicating that the button has been pressed. 3. Check that both redundant switch signals are low during POST.
Bed mobility HW App | BedMobilityHWApp-swRA-00024 | Bed Mobility Hw App - Read signals from the emergency release buttons | Bed Mobility Hw App incorrectly reads signals from Z-Burg emergency release buttons as not pressed | Data corruption Software timing Errors due to unexpected hardware components or failures | No nofication is provided when emergency release is initiated | HHM-0236 | #REF! | Software incorrectly processes sensing data as not pressed from emergency button due to runtime error or hardware issue resulting in continued bed mobility for E-Stop | Always - 100% | Chest Injury | Permanent - (S4) | Unusual | High | 1. Throw a fault if there is a mismatch between redundant switch signals. 2. Notify the user if both redundant switch signals go high indicating that the button has been pressed. 3. Check that both redundant switch signals are low during POST.
Wrist FW | WristFw-swRA-00001 | Wrist Firmware - Initialization | The application fails to start | No valid image in flash detected | No communication to the Wrist board | HHM-0255 | Physician is unable to begin procedure on time | Faulty state machine causing initialization failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the Wrist firmware is not communicating.
Wrist FW | WristFw-swRA-00002 | Wrist Firmware - Cannula presence detection | Cannula is present but not detected or reported to the RTPC | Software error, erroneous logic, freeze or crash | Arm docking failure | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Faulty state machine (CannulaUndocked declared incorrectly) | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will not allow surgery to begin if the Wrist firmware is not communicating.
Wrist FW | WristFw-swRA-00003 | Wrist Firmware - Cannula latch detection | Cannula latch state is not read or reported to the RTPC | Software error, erroneous logic, freeze or crash | Arm docking failure | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | 1. Corrupted latch sensor data or memory 2. EtherCAT comms issues | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will not allow surgery to begin if a cannula is not latched on an arm that is to be used.
Wrist FW | WristFw-swRA-00004 | Wrist Firmware - Cannula type and orientation | Cannula ID is not read or reported to the RTPC | Software error, erroneous logic, freeze or crash | Arm docking failure | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | 1. Corrupted proximity sensor data or memory 2. EtherCAT comms issues | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will not allow surgery to begin if a cannula is not identified on an arm that is to be used. (PMD)
Wrist FW | WristFw-swRA-00005 | Wrist Firmware - Cannula type and orientation | Wrong cannula ID is reported to the RTPC | Software error, erroneous logic, freeze or crash | RCM lateral drift (along the instrument insertion site) | HHM-0047 | Excessive force applied at incision site in Patient | 1. Corrupted proximity sensor data or memory 2. EtherCAT comms issues | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1 - Host SW will get the cannula hall sensor data and run an independent algorithm to verify the cannula type.
Wrist FW | WristFw-swRA-00006 | Wrist Firmware - 48 V Power Control | 48V power to the ADM motors is not enabled | Software error, erroneous logic, freeze or crash | Failure to power on the system | HHM-0255 | Physician is unable to begin procedure on time | Communication or software errors causing power control (enable) process failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - ADM current tracking fault 2 - ADM position tracking fault
Wrist FW | WristFw-swRA-00007 | Wrist Firmware - 48 V Power Control | 48V power to the ADM motors is disabled in the middle of a procedure | Software error, erroneous logic, freeze or crash | Instrument motion stops | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Communication or software errors causing power control (enable) process failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - ADM current tracking fault 2 - ADM position tracking fault
Wrist FW | WristFw-swRA-00008 | Wrist Firmware - Buttons state | Button is not read or reported to the RTPC | Software error, erroneous logic, freeze or crash | Cannot do port bump); Cannot change between RCM1 & RCM2 | HHM-0047 | Excessive force applied at incision site in Patient | Communication issue in button state process resulting in execessive body wall tension for the patient | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Extremely rare | High | 1 - The GUI provides a redundant means to enter the same mode(s) as the button.
Wrist FW | WristFw-swRA-00009 | Wrist Firmware - Buttons state | Button state is reported incorrectly to the RTPC | Software error, erroneous logic, freeze or crash | Unintentional movement of the instrument - RCM drift | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Communication issue in button state process resulting in unintentional instrument movement | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - The Wrist FW sends the state of both button GPIOs along with the logical button state every cycle. The host SW shall ensure the state of the two GPIO both match the logical button state before actiing on the button state.
Wrist FW | WristFw-swRA-00010 | Wrist Firmware - Voltage monitoring | Voltage is not read or read incorrectly | Software error, erroneous logic, freeze or crash | False positive (user annoynace) voltage error check | HHM-0347 | User annoyance when operating the system | Comminucation or software errors causing voltage monitoring failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | For 48V: 1. ADM current tracking fault. 2. ADM position tracking fault. 3. Wrist has ADC sensing of 48V input voltage along with power good signals for the 48V input and 48V output. The hardware will automatically cut the 48V output if it senses that the 48V go out of range. For 24V: 4. Wrist has ADC sensing of 24V input along with a power good signal on 24V input.
Wrist FW | WristFw-swRA-00011 | Wrist Firmware - Voltage monitoring | Voltage is not read or read incorrectly | Software error, erroneous logic, freeze or crash | False negative voltage error check. May lose communication via Wrist board | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Comminucation or software errors causing voltage monitoring failure | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | For 48V: 1. ADM current tracking fault. 2. ADM position tracking fault. 3. Wrist has ADC sensing of 48V input voltage along with power good signals for the 48V input and 48V output. The hardware will automatically cut the 48V output if it senses that the 48V go out of range. For 24V: 4. Wrist has ADC sensing of 24V input along with a power good signal on 24V input.
Wrist FW | WristFw-swRA-00012 | Wrist Firmware - Temperature monitoring | Temperature in not read, or read in correctly | Software error, erroneous logic, freeze or crash | False negative temperature error check | HHM-0298 | Touch temperature exceeds limits between Hospital staff/Physician/Patient and any part of the subsystem/subassemblies | Comminucation or software errors causing temperature monitoring failure | Always - 100% | Burn | Intervention Necessary - (S3) | Extremely rare | High | 1 - Wrist FW reads both the MCU temperature and the PCB temperature through an ADC input. The separate sensors and methods to sample provides some redundancy.
Wrist FW | WristFw-swRA-00013 | Wrist Firmware - Communication to RTPC | Communication is not processed | Software error, erroneous logic, freeze or crash | Unable to interact with the system/touchpoints | HHM-0256 | Physician is unable to complete procedure but has to convert to alternative method (open procedure) | Fails to perform arm docking or change mode process | Always - 100% | Surgery | Intervention Necessary - (S3) | Always | RAR | 1 - Host SW will not allow surgery to begin if the Wrist firmware is not communicating.
Wrist FW | WristFw-swRA-00014 | Wrist Firmware - NVRAM storage | NVRAM storage not written or read | Software error, erroneous logic, freeze or crash | Unable to start procedure | HHM-0255 | Physician is unable to begin procedure on time | Comminucation issue causing start procedure failure | Always - 100% | No harm | No Harm - (S0) | Always | No Harm | 1 - Host SW will not allow surgery to begin if the parameters cannot be read.
Wrist FW | WristFw-swRA-00015 | Wrist Firmware - NVRAM storage | NVRAM storage incorrectly written or read | Software error, erroneous logic, freeze or crash | Wrong parameters results in unintentional motion of the arm | HHM-0322 | Uncontrolled/undesired motion of Instrument end effector occurs in Patient : >15.3 mm (L) | Corrupted NVRAM data resulting in incorrect parameters used in procedure | Always - 100% | Abdominal Injury | Life Threatening - (S5) | Unusual | RAR | 1 - Host SW will include data integrity word in the parameter storage format (e.g. CRC32). The parameter data will not be accepted if the data integrity check fails.
Note on M1: Sam Diraimondo: need drop down for PH3 - sometimes and the rest of PH values (see email from Erik)

## Viewer(sandbox)
Category | Sub_category | No. | Desc
Market | Voice of customer | voc-1 | Pre-op planning is a critical step to ensure safety, efficacy and efficiency of resection procedures.
Market | Voice of customer | voc-2 | Pre-op planning is intended to inform 3 clinical decisions as framed by surgeons: “where to enter”, “how to move” and the “amount of resection”.
Market | Voice of customer | voc-3 | Specifically, when reviewing pre-operative images, surgeons are looking for: 1. Size, location and vascularity of the tumor(s)
Market | Voice of customer | voc-4 | Larger tumors are more concerning for spread. Anything above 6 or 7 cm might mean a full removal and how it sits on the kidney in relationship to other organs.
Market | Voice of customer | voc-5 | Where is the cancer? Has it invaded the outside wall? Where has it invaded? What structures have been invaded?
Market | Voice of customer | voc-6 | My focus is to see the proper anatomy of the mass and its vascular anatomy
Market | Voice of customer | voc-7 | Specifically, when reviewing pre-operative images, surgeons are looking for: 2. Tumor relationship to surrounding critical structures (i.e. vasculature, ureter, chest wall, etc.)
Market | Voice of customer | voc-8 | I need to avoid critical structures like the ureter. Knowing where it is can speed things up for us.
Market | Voice of customer | voc-9 | I’m looking for the relationship of tumor to surrounding structures, particularly the vasculature. The hepatic veins and vena cava are the bigger ones for liver
Market | Voice of customer | voc-10 | Is it close to the vessel? Or is it infiltrated?
Market | Voice of customer | voc-11 | Specifically, when reviewing pre-operative images, surgeons are looking for: 3. Patient-specific anatomical variations or aberrant anatomy.
Market | Voice of customer | voc-12 | There is a great deal of variety within the anatomy. This is what we're looking for in the CT
Market | Voice of customer | voc-13 | Because of the anatomical variability, we are often operating blindly. We don't know how many bronchi we will come into.”
Market | Voice of customer | voc-14 | By integrating 3D models and patient simulation from Visible Patient into J&J ecosystem, we’ll provide the following value to the clinicians:
Market | Voice of customer | voc-15 | we’ll provide the following value to the clinicians: Highly detailed, virtual 3D model, using patients’ pre-op images (CT, MRI)
Market | Voice of customer | voc-16 | we’ll provide the following value to the clinicians: Color-coded, interactive experience that allows you to explore and visualize spatial relationships between tumors, critical structures and surrounding anatomy of your patient
Market | Voice of customer | voc-17 | we’ll provide the following value to the clinicians: Easy-to-use interface that is accessible on your team’s PC/Mac, tablet, and mobile devices for reference in the OR, at home, and on-the-go
Market | Voice of customer | voc-18 | we’ll provide the following value to the clinicians: Allows you to explore anatomy on your schedule, without reliance on interpretation by your radiology team
Market | Voice of customer | voc-19 | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume
Market | Voice of customer | voc-20 | Surgeons use 2D images to “envision a 3D model” in their mind when planning their approach. More experienced surgeons are more likely to be confident in this ability. However, even experienced surgeons acknowledge inefficiency and communication challenges with the surgical team when developing a surgical plan, e.g. reliance on the radiologist to develop a surgical plan.
Market | Voice of customer | voc-21 | In addition to 3D models, 2D visualization, with color, of a 3D model (Anatomical Atlas) is also a valuable tool for clinicians as it is “what I’m used to and know how to interpret, but can see things better”.
Market | Voice of customer | voc-22 | The patient simulation (Clip Applying) tool provides significant value when planning painstaking, tissue-sparing procedures. For many, the patient simulation tool is considered a “breakthrough” because of high-quality image and instant volume calculations that allow for simulation of different surgical plans pre-operatively (e.g. segmentectomies vs a lobectomy).
Market | Voice of customer | voc-23 | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee.
Market | Voice of customer | voc-24 | Surgeons see the flexibility in access across devices and locations (at home, at the office, on the go, in the OR) as a critical attribute of the solution
Market | Voice of customer | voc-25 | All surgeons for the target specialties (thoracic, HBP, colorectal and urology) are procedurally required to review pre-operative patient images before the procedure. However, only a subset of surgeons refers back to pre-operative patient images intra-operatively, mostly due to an unexpected finding or uncertainty. Hence, the primary use case for pre-operative planning is before the procedure begins; intraoperative reference of the preoperative planning is secondary.
Market | program goals |  | The primary clinical goals of this project are to design, develop and deploy a solution that will enable Digital Solution to ● launch surgical decision support, especially medical imaging software application products
Market | program goals |  | enable surgeons to order, manage, and review patient-specific 2D and 3D anatomical models
Market | program goals |  | enable surgeons to view pre-operative images and 2D and 3D anatomical models online
Market | program goals |  | enable surgeons to perform a simulated resection by applying virtual clips to approximated segments in order to simulate post-operative organ volume
Market | program goals |  | enable surgeons to save a viewing and simulation session and continue at a later time
Market | program goals |  | enable surgeons to continue their viewing and simulation session in the OR using the Digital Hub hardware device after the commercial release of the Digital Hub (viewing of data sets on the Digital Hub will be outlined in a future Charter),
Market | program goals |  | enable surgeons to continue viewing their images and simulation in the OR on the web portal,
Market | program goals |  | release an integrated solution to the EU and/or US markets late 2021 or early 2022
Market | program goals |  | release an integrated solution in APAC later in 2022 as outlined in a future Charter,
Market | program goals |  | expand and engage the overall user community post-launch.
Market | Product Description |  | Bandit Phase 2 should be comprised of the following major components: surgeon portal, digital solutions viewer, digital surgery enablement platform.
Market | Product Description |  | Surgeon Portal - a web portal for surgeons to log into, select the appropriate language, order the modeling service, review the status of the modeling service, upload DICOM images, start the Digital Solutions Viewer.
Market | Product Description |  | Digital Solutions Viewer - a DICOM viewer that allows for viewing a DICOM image and/or the 3D anatomical model selected by the surgeon through the surgeon portal. After the surgeon confirms the selected dataset, the image viewer will show the pre-operative image, allow for placement of annotations and measurements, show the anatomical model, allow for interactions with the anatomical model, allow for simulation of resection by placing a virtual clip.
Market | Product Description |  | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables user management, data storage, and provides DICOM studies to the external 3D model creation service as well as to the Digital Solutions Viewer.
Market | Product Description |  | Bandit Phase 2 will deliver an integrated experiences for the customer, which can be categorized in three groups: ● 3D Model Submission and Delivery ● Digital Solution Viewer ● Patient Modeling and Outcome Simulation
Market | Product Description |  | AI-enhanced visualization and personalized insights to help surgeons get a clearer roadmap for surgery
Market | Product Description |  | 3D model submission and delivery: 3D model order submission by clinical user; 3D model order management by order admin; 3D model order billing; PACS integration
Market | Product Description |  | Image viewer: image viewer; meaurement tools
Market | Product Description |  | patient modeling and outcome simulation: integrated 3D models; patient specific simulation of resection
Market | Regulatory assessment |  | Target for the Digital Solutions Viewer is the EU. The Digital Solutions Viewer will be the first CE marked product and first product under the new EU Medical Device Regulation (MDR) for Auris.
Market | Regulatory assessment |  | The Digital Solutions Viewer is considered software as a medical device (SaMD) and will need to conform to the requirements of the Medical Device Regulation (EU) 2017/745. The Viewer will need to provide Technical Documentation (TD) covering the design, development, and state of the art with respect to safety and performance and demonstrate compliance to the General safety and Performance requirements in Annex I.
Market | Regulatory assessment |  | The proposed classification of the Digital Solutions Viewer is a Class IIa Rule 11 (Software intended to provide information which is used to make decisions with diagnosis or therapeutic purposes).
Market | Regulatory assessment |  | Type of conformity assessment route: Quality Management System, Annex IX excluding chapter II.
Market | Regulatory assessment |  | As the Digital Solutions viewer will be the first MDR product for Auris, the Auris QMS must be MDR ready to support certification. The proposed certificate scope will be for image visualization and planning software.
Market | Regulatory assessment |  | An EU Authorised Representative agreement between Auris and J&J Medical GmbH must be in place.
Market | Regulatory assessment |  | The MDR stipulates that the manufacturer must provide sufficient clinical evidence (Article 61) to justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR.
Market | Regulatory assessment |  | The MDR provides a possibility to use clinical data related to a marketed equivalent device (found in the clinical evaluation required for a device) under conformity assessment.
Market | Regulatory assessment |  | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes of the software are intended to be similar for the Auris Viewer.
Market | Regulatory assessment |  | The equivalence requirement to demonstrate sufficient levels of access to comparator data should be met as Visible Patient is collaborating with J&J and the Auris Viewer is intended to interface with the Visible Patient Lab 3D recreated models.
Market | Regulatory assessment |  | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes
Market | Market overview |  | Market Target Segments: ● Cutting Edge Pioneers: Large, urban, teaching hospitals leading the way through technological innovation with a large budget.
Market | Market overview |  | Competitive Strivers: Moderately sized, urban, teaching hospitals that intentionally spend their resources to differentiate from the competition.
Market | Market overview |  | Source of Volume (SOV): Attract potential category users
Market | Market overview |  | Hospital Targets: High procedure volumes in 2+ of the following surgical specialties: Thoracic, Urology, Colorectal, and HPB; typically, urban and / or academic hospitals
Market | Market overview |  | Primary User Targets: Surgeons in Thoracic, Urology, Colorectal, and Hepato-biliary
Market | Market overview |  | Secondary User Targets (secondary expansion within primary user hospitals): Lower volume surgeons in Primary User Target specialties, Surgeons in GYN
Market | Market overview |  | Bandit opportunity immediately establishes an Ethicon position in Advanced Imaging
Market | Market overview |  | Ethicon Franchise: Aligns to Ethicon franchise objectives of developing connected products & solutions and leveraging data-driven insights to improve outcomes.
Market | Market overview |  | Advanced Imaging: Establishes an immediate Ethicon position in Advanced Imaging category, the “North Star” for the Ethicon Digital Solutions portfolio.
Market | Market overview |  | Blunts Competitive Messaging prior to JNJ Robot launch: Bandit blunts ISRG IRIS™ launch into Advanced Imaging, the #1 driver of robot brand choice
Market | Market overview |  | When the unified global Digital Solutions user portal launches in H2 2021, Bandit expands the value of Digital Solutions to Ethicon:
Market | Market overview |  | Builds Loyal Digital Surgery Customer Base: Increases the value of the online Ethicon Digital Solutions video, image & case library, creating network effects & switching costs
Market | Market overview |  | Provides Critical Data Pipeline: Adds pre-op image and 3D model data to feed the Ethicon AI Insights platform
Market | Market overview |  | Pre-operative 3D planning software is in-market today, but most focus on orthopaedics and neurosurgery. Bandit has three main competitors in the US (Fujifilm Synapse, TeraRecon iNtuition, and Intuitive Surgical IRIS™), but only Fujifilm and TeraRecon have a significant presence in EMEA, Bandit’s initial regional foothold market. Only Fujifilm is used today in APAC.
Market | Market overview |  | Fujifilm and TeraRecon require radiologist creation of 3D models and involvement with interpretation and dictation, whereas Bandit eliminates this step, giving surgeons a streamlined pre-op planning sequence that puts the control in their hands on their schedule.
Market | Market overview |  | Intuitive IRIS™ is currently cleared for use with CT only, whereas the others, including Bandit, are applicable for MRI as well.
Market | revenue and cost |  | revenue and cost: Bandit is projected to create ~€200MM in cumulative end user sales through 2025 with peak sales of €75+MM with launches in EMEA, APAC and NA.
Market | Project assumptions |  | An initial project team will be deployed to lead activities needed to complete the deliverables of Phase 1: Planning and Definition Phase per document 700-000049-00, Design Control. The Design and Development Plan (DDP) will ultimately define the organizational and team structure of the project and will be completed prior to exiting the Planning and Definition Phase.
Market | Project assumptions |  | The major project milestones and estimated timing are listed below. This schedule assumes a fully staffed development team and approved program budget by the end of Phase 1. An updated schedule and list of major milestones along with program risks will be provided in the Design and Development Plan (DDP) prior to exiting the Planning and Definition Phase.
Market | Project assumptions |  | The following assumptions were made as basis of this Project Authorization Charter: ● CE Mark is prioritized (EMEA launch) and 510K Approval is performed afterwards
Market | Project assumptions |  | Changes to UI of Digital Solutions Viewer are limited to meet timeline
Market | Project assumptions |  | Contract technical writer (w/SaMD exp) is on-board by mid-January 2021 for DHF support
Market | Project assumptions |  | CIR, DIR (and lower) are written to minimize V&V and released no later than February 2021
Market | Project assumptions |  | No issues are found in the 1st formative; and a 2nd formative is not required
Market | Project assumptions |  | Expedited Review (MDR) has minimal questions and no major revisions needed
Market | Project assumptions |  | FDA acceptance of substantial equivalence with minimal questions
Market | Project assumptions |  | Code the complete the Digital Solutions Viewer is solid with few bugs/issues by mid-March 2021
Market | regulatory |  | Web Portal & DSEP are not part of MDR submission
Market | regulatory |  | QMS audit can completed by regulatory body by end of May (earlier than planned)
Market | regulatory |  | QMS documents can be reviewed and released in Agile within 2 weeks (documents complete to released)
Market | regulatory |  | No issues found during phase reviews or phase exit reviews
Market | regulatory |  | Notified body is available for an expedited review and we meet the date provided to them
Clinical_Workflow | Clinical Hazard Analysis | CHA-01 | The intended use of the device is an aid in pre-surgical planning for procedures requiring spatial understanding of complex patient anatomy.
Clinical_Workflow | Clinical Hazard Analysis | CHA-02 | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan.
Clinical_Workflow | Clinical Hazard Analysis | CHA-03 | Patients undergoing procedures in the general, bariatric, HPN, gynecology, thoracic, and urology specialties, who require medical images to be taken to plan for their surgical procedure.
Clinical_Workflow | Clinical Hazard Analysis | CHA-04 | No, the medical device does not sustain or support life.
Clinical_Workflow | Clinical Hazard Analysis | CHA-05 | If the medical device fails while being used to reference medical images or 3D models during a procedure, the user may need to access the patient images on another PACS viewer to proceed.
Clinical_Workflow | Clinical Hazard Analysis | CHA-06 | Measurements are taken by the medical device. The medical device can take linear measurements between two points displayed on the DICOM or 3D model
Clinical_Workflow | Clinical Hazard Analysis | CHA-07 | Measurements are taken by the medical device. The medical device can calculate a designated area shown on a 2D DICOM slice
Clinical_Workflow | Clinical Hazard Analysis | CHA-08 | Measurements are taken by the medical device. The medical device can through the clip apply feature, calculate volume of a designated portion of an organ to be resected
Clinical_Workflow | Clinical Hazard Analysis | CHA-09 | Measurements taken on a 2D device should be accurate for images acquired with high resolution.
Clinical_Workflow | Clinical Hazard Analysis | CHA-10 | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices.
Clinical_Workflow | Clinical Hazard Analysis | CHA-11 | The medical device may be used with a radiology monitor
Clinical_Workflow | Clinical Hazard Analysis | CHA-12 | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software.
Clinical_Workflow | Clinical Hazard Analysis | CHA-13 | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk
Clinical_Workflow | Clinical Hazard Analysis | CHA-14 | When using a tablet to interface with the device during a procedure, there is a risk that the procedure would coincide with battery depletion.
Clinical_Workflow | Clinical Hazard Analysis | CHA-15 | successful application of the medical device is critically dependent on human factors such as the user interface. The user interface is fundamental to interacting with and displaying measurements of the patient images and anatomic models.
Clinical_Workflow | Clinical Hazard Analysis | CHA-16 | use errors can arise from not understanding or properly navigating the user interface. Symbols for the user interface have been based on those common in other PACS viewer products, but may be arranged differently or applied slightly differently for this medical device.
Clinical_Workflow | Clinical Hazard Analysis | CHA-17 | In general, the medical device is used in the office environment for planning prior to a case. When using the medical device in the operating room, the user may be more distracted and could make use errors such as loading the incorrect patient’s data, or ignoring the orientation indicators.
Clinical_Workflow | Clinical Hazard Analysis | CHA-18 | The medical device displays information. The display relies on text, colors, sizing, etc to relay spatial information and distinguish tissue types. These can be adjusted for accessibility but must take colorblind users into account
Clinical_Workflow | Clinical Hazard Analysis | CHA-19 | the medical device must be used with a display large enough and with high enough resolution so that the user can clearly see details of the patient images and anatomic models.
Clinical_Workflow | Clinical Hazard Analysis | CHA-20 | The device must render the patient models taking all these factors into account.
Clinical_Workflow | Clinical Hazard Analysis | CHA-21 | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar.
Clinical_Workflow | Clinical_procedure | CPR_01 | User shall launch program by openning Aurora viewer during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_02 | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_03 | User shall find patient(s) of interest by searching for patient by name, MRN, surgeon during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_04 | User shall find patient(s) of interest through sorting lists by specialty, patient, surgeon, delivery date, or status to find patient of interest during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_05 | User shall find patient(s) of interest through filters list using "FILTER BY" tool in order to find patient during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_06 | User shall select patient and launch image viewer by touching/clicking on row and row will highlight when selected during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_07 | User shall select patient and launch image viewer by clicking "NEXT" to load patient images and 3D model during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_08 | User shall verify images are correct by double checking patient informaiton during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_09 | User shall view 2D image stacks by monitoring loading of the 2D images on the "Viewing" panel during the load and view desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_10 | After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks.
Clinical_Workflow | Clinical_procedure | CPR_11 | User shall select which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas during setting screen layout phase
Clinical_Workflow | Clinical_procedure | CPR_12 | User shall select whether they would like the segmented anatomic regions to be displayed by infill or contour when turning on and off segmentation layers
Clinical_Workflow | Clinical_procedure | CPR_13 | User shall adjust view of desired image stacks by operating pan, rotate, zoom, and scroll to select stack they would like to manipulate during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_14 | User shall use touchscreen, mouse, or side menu to select "Pan, Rotate, Zoom or Scroll"
Clinical_Workflow | Clinical_procedure | CPR_15 | User shall undo a single adjustment with the "Undo" button during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_16 | Usershall reset the view altogether with the "Re-Center" button during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_17 | User shall navigate to the Windowing menu on the sidebar to asjust the contrast or brightness of the image during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_18 | User shall select from presets to auto-adjust window width and level based on target anatomy the image was acquired to show during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_19 | User shall choose to select "Width/Level" button to make customize adjustments to the brightness and contrast during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_20 | User shall drag slider(s) to desire width and level percentages when adjust windowing during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_21 | User shall reset the view using the "Reset" button when adjust windowing during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_22 | User shall select the annotation menu by navigate to tools from the sidebar during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_23 | User shall select "distance" tool to measure the straight-line distance between two points on the image during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_24 | User shall click on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_25 | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_26 | User shall delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu by using distance tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_27 | User shall add and delete arrows in a similar fashion after clicking the "Arrow" menu option. Arrows added to the screen will also appear in the Annotation sidebar menu list by arrow tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_28 | User shall add a note by selecting the "note" annotation option and clicking anywhere on the image during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_29 | when a textbox appears, user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_30 | User shall return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_31 | User shall remove any added note using the trash icon in the Annotation menu or on the note itself during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_32 | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_33 | User shall reset all annotations using the "Reset" button during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_34 | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_35 | User shall confirm they ar viewing the correct 3D model by checking the information shown onscreen. They togge on "Patient Info" if not already on when openning 3D Model View during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_36 | User shall select "3D Model" menu to see which anatomic structures are available when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_37 | User shall select/deselect the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_38 | User shall scan the volumes of the different models to check against expectations when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_39 | User shall select "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_40 | Alternatively, user shall select from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_41 | User shall enter fullscreen at times for a better view of the model when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_42 | User shall select an individual structure when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_43 | User shall adjust the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_44 | User shall select "Done" button to complete the adjustment when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_45 | User shall select "Clip Apply Model" menu and allows module to load when navigating to Clip Apply feature during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_46 | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_47 | User shall right click on the model the apply the clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_48 | User shall observe the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_49 | If the user is unhappy with the clip placement they chose, they shall either remove a single clip, or remove all clips when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_50 | At any time, the user shall select any of the clips in the menu in order to review the volume simulation details for that clip when adding and removing Clips when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_51 | When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_52 | If they choose to save the plan, user names the plans and confirms. (Need to check) when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_53 | User shall navigate to "My Views" menu when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_54 | User shall select "SAVE NEW VIEW" and pop-up appears when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_55 | User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_56 | User shall press "CONFIRM" a second time and the view is saved when saving current view for later access during saving current view phase
DIR | Requirement | DIR_1 | Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer.
DIR | Image Viewer | DIR_2 | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard
DIR | Image Viewer | DIR_3 | The system shall have the capability to render 2D views of the preoperative patient images.
DIR | Image Viewer | DIR_4 | The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study.
DIR | Image Viewer | DIR_5 | The system shall have the capability to render a view of 3D model segmentations
DIR | Image Viewer | DIR_6 | The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view
DIR | Image Viewer | DIR_7 | The system shall provide a method to zoom in and out on the views.
DIR | Image Viewer | DIR_8 | The system shall provide a method to pan a viewing panel.
DIR | Image Viewer | DIR_9 | The system shall provide a method to rotate a 2D viewing panel.
DIR | Image Viewer | DIR_10 | The system shall provide a method to rotate a 3D viewing panel.
DIR | Image Viewer | DIR_11 | The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel.
DIR | Image Viewer | DIR_12 | The system shall provide a method to rotate the cut plane within an MPR view.
DIR | Image Viewer | DIR_13 | The system shall provide a method to undo and redo user interactions.
DIR | Image Viewer | DIR_14 | The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size.
DIR | Image Viewer | DIR_15 | The system shall provide a method for hiding a viewing panel.
DIR | Image Viewer | DIR_16 | The system shall provide a method to adjust the Window Center and Window Width adjustments using presets.
DIR | Image Viewer | DIR_17 | The system shall provide a method for manually adjusting window width and center values for the intensity range
DIR | Image Viewer | DIR_18 | The system shall provide a method for a reset all viewing panels simultaneously
DIR | Image Viewer | DIR_19 | The system shall provide a method to select the layout for the viewing panels
DIR | Image Viewer | DIR_20 | The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume.
DIR | Image Viewer | DIR_21 | The system shall be able to display metadata contained in the DICOM data.
DIR | Image Viewer | DIR_22 | The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects.
DIR | Image Viewer | DIR_23 | The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model.
DIR | Image Viewer | DIR_24 | The system shall be able to display metadata contained in the 3D model container
DIR | Image Viewer | DIR_25 | The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels
DIR | Image Viewer | DIR_26 | The system shall use SI units and their derivatives
DIR | Image Viewer | DIR_27 | The system shall provide a method for performing one or more measurements within one of the viewing panels
DIR | Image Viewer | DIR_28 | The system shall label and list measurements automatically
DIR | Image Viewer | DIR_29 | The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object.
DIR | Image Viewer | DIR_30 | The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel.
DIR | Image Viewer | DIR_31 | The system shall allow for simulation of resection by applying a clip to approximated segments.
DIR | Image Viewer | DIR_32 | The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip.
DIR | Image Viewer | DIR_33 | The system shall provide the ability to support individual language packs.
DIR | Image Viewer | DIR_34 | The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language.
DIR | Image Viewer | DIR_35 | The system shall, when appropriate and available, use standard internationally recognized symbols.
DIR | Image Viewer | DIR_36 | The system shall, when appropriate and available, use standard internationally recognized icons.
DIR | Image Viewer | DIR_37 | The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System.
DIR | Image Viewer | DIR_38 | The System shall record all functions performed by the System pre, intra, and post operatively.
DIR | Image Viewer | DIR_39 | The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel.
DIR | Image Viewer | DIR_40 | The System shall record system functions and operational events generated by the System.
CIR | Image Viewer | Digital_TXT_11 | The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning.
CIR | Image Viewer | Digital_TXT_12 | The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images.
CIR | Image Viewer |  | The scope of this document does not cover how 3D model are ordered, created and delivered, as this is accomplished by other software components.
CIR | Image Viewer | Digital_TXT_13 | All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated.
CIR | Image Viewer | Digital_TXT_14 | 3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator.
CIR | Image Viewer |  | 3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon.
CIR | Image Viewer |  | Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components.
CIR | Image Viewer |  | DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data.
CIR | Image Viewer |  | PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images.
CIR | Image Viewer |  | Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format.
CIR | Image Viewer |  | Visible Patient (VP): External company, responsible for managing and delivering 3D models.
CIR | Image Viewer |  | Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively
CIR | Image Viewer |  | 3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models
CIR | Image Viewer |  | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital.
CIR | Image Viewer |  | Marketing and Sales Representative: Internal J&J users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies.
CIR | Image Viewer | Digital_CIR_DS-1 | View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings.
CIR | Image Viewer | Digital_CIR_DS-131 | Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings.
CIR | Image Viewer | Digital_CIR_DS-91 | 2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings
CIR | Image Viewer | Digital_CIR_DS-5 | View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings.
CIR | Image Viewer | Digital_CIR_DS-132 | Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings
CIR | Image Viewer | Digital_CIR_DS-92 | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach.
CIR | Image Viewer | Digital_CIR_DS-93 | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure.
CIR | Image Viewer | Digital_CIR_DS-12 | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens.
CIR | Image Viewer | Digital_CIR_DS-14 | Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device.
CIR | Image Viewer | Digital_CIR_DS-104 | Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice.
CIR | Image Viewer | Digital_CIR_DS-24 | Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently.
CIR | Image Viewer | Digital_CIR_DS-109 | Indications for Use for 3D Model Viewing and Manipulation: As Ethicon marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party.
CIR | Image Viewer | Digital_CIR_DS-110 | Countries Supported for 3D Model Viewing and Manipulation: As Ethicon marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model.
CIR | Image Viewer | Digital_CIR_DS-113 | Capture and Storage of System Data: As Ethicon marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development.
CIR | Image Viewer | Digital_CIR_DS-114 | Branding: As Ethicon marketing representative, I need the application to follow JnJ defined branding, so that I can maximize consistency in the look and feel of JnJ digital products that may be used by the same user base and reduce cognitive load on the users when using JnJ digital products.
CIR | Image Viewer | Digital_CIR_DS-115 | Demo Capabilities: As Ethicon marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively.
CIR | Image Viewer | Digital_CIR_DS-116 | Trial Capabilities: As Ethicon marketing and sales representative I need to be provide a trial functionality to the application so that I can enable clinical users evaluate the value and workflow of the application for their specific cases
Clinical_HA | Image Viewer |  | Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division.
Clinical_HA | Image Viewer |  | Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure.
Clinical_HA | Image Viewer |  | Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure.
Clinical_HA | Image Viewer |  | Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body.
Clinical_HA | Image Viewer |  | Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure.
Clinical_HA | Image Viewer |  | Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure.
Clinical_HA | Image Viewer |  | Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure.
User manual | Image Viewer |  | The Digital Solutions Viewer (Viewer) is used by radiologists and surgeons during surgical planning to create, view, and manipulate 3D models created from pre-operative CT or MRI. Staff in the operating room (OR) may also use the Viewer to refer to any saved surgical plans while assisting the surgeon.
The Viewer can be used for the following: • annotate 2D source images • manipulate the 3D model • apply clips to the 3D model to simulate tissue resection • save and re-open resection plans
Surgeons and radiologists review and annotate the 2D source images during pre-operative planning. Annotations are lines or arrows which provide measurements and can be saved. Saved plans are listed in the My Views window.
The patient information is located in the lower left corner of each image stack view. 1. Select the Viewing down arrow. 2. Select the Patient Info toggle to show or to hide the patient information.
In the Viewing window, select the Anatomical Atlas toggle to turn colorization on or off. • Select Infill to color the entire organ. • Select Contour to color just the outline of the organ.
Scroll: within the image view, hover your mouse and scroll the mouse wheel.
Zoom in or Zoom out: Within the image view, left click, hold and drag the mouse wheel upwards.
Pan: Within the image view, right click on the view, hold and drag the image
Re-center: In the Viewing window, click Recenter
Adjust the 2D Image Windowing Levels 1. Select the Windowing down arrow. 2. Select any of the preset views. 3. Select Width/Level and drag the sliders to adjust the percentages. 4. Select Reset, to return to the default view.
Annotate 2D Source Images: 1. In the Viewing window, select the Annotation down arrow. 2. To draw a straight line, select the Distance tool, left click the point of origin and drag the mouse to the destination point, then left click again. 3. To draw a straight line with an arrow, select the Arrow tool, left click the point of origin and drag the mouse to the destination point, then left click again. 4. To save the annotation with a timestamp to open later, select the My Plan toggle and select Save New Plan. 5. To delete an annotation, select the My Plan toggle and click the trashcan icon next to the plan you want to delete. 6. To delete all notations, in the Annotation window, select Reset.
Simulate resection volumes in the 3D Model: Surgeons and radiologists apply clips to the 3D model in order to simulate tumor resectioning during preoperative planning. Clips are placed on the targeted organ's network and estimated volumes for the resected and remaining organ tissue is calculated
Adjust Opacity on the 3D Model: 1. Select the 3D Model or the Clip Apply Model down arrow. 2. Select the structure and use the slider to adjust the opacity.
Manipulate and Apply Clips to a 3D Model: 1. Select the Clip Apply Model down arrow. 2. To manipulate the 3D model 3. Within the 3D model, right click the mouse to apply a clip. 4. To remove a clip, select the trashcan next to the clip you want to delete. 5. To remove all clips, select REMOVE ALL CLIPS. 6. Review the estimated volumes and click the Information icon for additional calculation information.
Open Saved Plans and Manipulate Images: During pre-operative planning, surgeons or radiologist save annotated views that can be reopened later
Open a Saved Plan: 1. Select the My Views down arrow. 2. Select a file to view.
Manipulate the 2D Image Stacks: 1. In the Viewing window, select the Anatomical Atlas toggle to turn colorization on or off. • Select Infill to color the entire organ. • Select Contour to color just the outline of the organ. 2. To manipulate the image stacks:
Adjust the 2D Image Windowing Levels: 1. Select the Windowing down arrow. 2. Select any of the preset views. 3. Select Width/Level and drag the sliders to adjust the percentages. 4. Select Reset, to return to the default view.
Adjust Opacity on the 3D Model: 1. Select the 3D Model or the Clip Apply Model down arrow. 2. Select the structure and use the slider to adjust the opacity.