## TOC
Date | Stem | Run | Stop | Hazard-harm | ProductCode_Lists | HS_collection | Others | Description
2021-03-18 | 4309 | 2663 | 23 | 2014 | 4789 | 68 (pace, brace, stent-graft) | Brace_shop1, IV_Set, reference
2021-07-04 | 4894 | 3346 | 23 | 2014 | 4789 | 73 (+viewer) | (+viewer) | add ventilator+digital solutions viewer | implement in GPT2

## Stem
Requirement Keywords | priority | Failure_Mode | Hazard | Hazardous_Situation | Harm | Severity_Level | Mitigation
aami | 3 | AAMI regulation violated | compliance issue | [5 empty] | a | 2 | 2019-11-24 |  | 585 | 275 | 487 | 2 | for a keyword, if its verb and noun has different stems: keywords shall be listed: verb + noun + pl of noun, stem shall be used from verb (the first one)
abrasion, abrasive | 2 | abrasive wear | device abrasion from instrument or another object; naturally worn | [5 empty] | b | 3 | 2019-12-02 | merge proj2_requirements
abrasion, abrasive | 2 | change in shape | material integrity issue; material deformation | [5 empty] | c | 4 | 2019-12-11 | ML sentence | 878
abrasion, abrasive | 2 | device abrasion from instrument or another object | device abrasion from instrument or another object | [5 empty] | d | 5 | 2020-01-12 | communications | 878
abrasion, abrasive | 2 | insulation short | electrical shorting; spark | [5 empty] | e | 6 | 2020-01-29 | 2nd released | 1056 | 324 | 950
abrasion, abrasive | 2 | loss of electrical isolation | electrical shorting | [5 empty] | f | 7 | 2020-02-02 | correct master lists by rules | 1108
absent, absence, absences | 2 | detection absence failed | high sensing threshold; under-sensing | [5 empty] | g | 8 | 2020-03-17 | add standards | 1389
absorb, absorber, absorbers, absorption | 3 | incomplete absorption | defective item | [5 empty] | H | 9 | 2020-05-17 | FDA Maude partial | 1782
abuse | 2 | data abused | data issue | [5 empty] | I | 10 | 2020-05-25 | Ventilator | 2281
abuse | 2 | system abused | computer system security issue; unauthorized access to computer system | [5 empty] | J | 11 | 2020-07-01 | pump (infusion, dibete, Pneumatic) | 2597
AC | 3 | fail to run on AC | failure to run on AC/DC | [5 empty] | K | 12 | 2020-08-15 | mitigation variation (FM vs keywords) | 3372
accelerate, acceleration, deceleration | 3 | fail to monitor acceleration | electronic property issue; application program issue | [6 empty] | 13 | 2020-08-21 | SRA Assura (HV therapy) | 3633
accelerate, acceleration, deceleration | 3 | falling object | slippage of device or device component; component falling | [6 empty] | 14 | 9/1/52020 | add multiple outputs (hazard,hs)
accelerate, acceleration, deceleration | 3 | inadvertent motion | unintended system motion; unintended movement | [6 empty] | 15 | 2020-10-15 | break into ML1 &ML2 new folder hazard-harm
accelerate, acceleration, deceleration | 3 | loose object translation | migration of device or device component
accelerate, acceleration, deceleration | 3 | sloshing | vibration
accelerometer | 3 | accelerometer input high | failure to read input signal
accelerometer | 3 | inaccurate accelerometer | imprecision
accelerometer | 3 | speed measurement inaccurately | imprecision
access, accessible, accessibility | 2 | fail to access | human-device interface issue
access, accessible, accessibility |  | fail to access across devices and locations | human-device interface issue
access, accessible, accessibility |  | fail to access at home | human-device interface issue
access, accessible, accessibility |  | fail to access at office | human-device interface issue
access, accessible, accessibility |  | fail to access on the go | human-device interface issue
access, accessible, accessibility |  | fail to access in the OR | human-device interface issue
access, accessible, accessibility |  | fail to access through web portal | human-device interface issue
access, accessible, accessibility |  | fail to access functions | computer software issue
access, accessible, accessibility |  | incorrect accession number | application security issue
access, accessible, accessibility | 2 | unable onehand to access | device handling issue
access, accessible, accessibility | 2 | unable to access | human-device interface issue
access, accessible, accessibility | 2 | unable to access battery | human-device interface issue
access, accessible, accessibility | 2 | unable to access power | human-device interface issue
access, accessible, accessibility |  | multiple accesses at once time | application security issue; computer software issue
access, accessible, accessibility | 2 | unauthorized access | unauthorized access to computer system
accessory | 3 | accessory life short | shelf life exceeded
accessory |  | fail to label manufacturer information | device markings issue
accessory |  | fail to label serial number | device markings issue
accessory |  | fail to label batch code | device markings issue
accessory |  | fail to label date of manufacture | device markings issue
accessory |  | fail to label use by date | device markings issue
accessory |  | fail to disclose accessory limitation | instruction for use issue
accessory |  | fail to disclose toxic material | instruction for use issue
accessory |  | accessory broken or damaged | break
accessory |  | inadequate info of accessories | instruction for use issue; device markings issue
accessory | 3 | incompatible accessory | accessory incompatible; device inoperable
accumulate, accumulator | 3
accurate, accuracy, inaccuracy | 1 | accuracy out of specification | calibration issue; imprecision
accurate, accuracy, inaccuracy | 1 | fail to display accuracy | no display or display failure
accurate, accuracy, inaccuracy | 1 | low accuracy rate | inaccurate flowrate
accurate, accuracy, inaccuracy | 1 | low measurement accuracy | incorrect measurement
activate, activation | 2 | activation delayed | difficult or delayed activation
activate, activation | 2 | device remains activated | device remains activated
activate, activation | 2 | difficult to activation | difficult or delayed activation
activate, activation | 2 | energy activation failed | failure to deliver energy
activate, activation | 2 | fail to activate | difficult or delayed activation
activate, activation | 2 | inadvertent energy activation | unintended energization
activate, activation | 2 | Inadvertently activating button accidentally | self-activation or keying
activate, activation | 2 | premature activation | self-activation or keying
activate, activation | 2 | self-activation | self-activation or keying
activate, activation | 2 | tool activation failed | difficult or delayed activation
activity | 3 | inadequate activity | difficult or delayed activation
activity | 3 | incomplete activity | difficult or delayed activation
activity | 3 | lack of activity | difficult or delayed activation
actuate, actuator | 3 | actuation force out of specification | mechanical issue
actuate, actuator | 3 | actuation sensing out of specification | invalid sensing
actuate, actuator |  | inappropriate force generated | mechanical issue
actuate, actuator | 3 | loss of control | electronic property issue
actuate, actuator |  | adaptor leakage | air leak; leak
actuate, actuator |  | adaptor detached | detachment of device component
actuate, actuator |  | adapter mismatched | component incompatible
actuate, actuator |  | adaptor stripped | connection issue
actuate, actuator | 3 | difficult to install adaptor | installation-related problem
actuate, actuator | 3 | fail to output current | electronic property issue
actuate, actuator |  | part broken or damaged | break; defective component
adhere | 2 | fail to adhere | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | adhesive applied to wrong place | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | clean skin insufficiently | device cleaning issue
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | excessive adhesive applied | misassembled
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | excessive adhesive removal | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | insufficient adhesive applied | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | insufficient electrical continuity | intermittent continuity
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | insufficient medical adhesive | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | insufficient medical adhesive removal | misassembled
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | loss of adhesive | failure to adhere or bond
adhesive, adhesion, agglutinate, agglutinated, agglutination, agglutinative | 2 | Medical adhesive blocks part | blocked connection
adjust, adjustment, adjustable, adjustability | 2 | adjustable angle out of specification | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | adjustable volume out of specification | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | adjustment depth out of specification | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | adjustment difficult | device difficult to setup or prepare
adjust, adjustment, adjustable, adjustability | 2 | adjustment distance out of specification | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | adjustment failed | device difficult to setup or prepare
adjust, adjustment, adjustable, adjustability | 2 | adjustment force high | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | adjustment out of specification | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | fail to adjust manually | device operates differently than expected
adjust, adjustment, adjustable, adjustability | 2 | fail to adjust settings | device operates differently than expected
adjust, adjustment, adjustable, adjustability |  | fail to save adjustments | device difficult to setup or prepare
adjust, adjustment, adjustable, adjustability | 2 | inaccurate adjustment | imprecision
adjust, adjustment, adjustable, adjustability | 2 | incorrect adjustment engaged | use of incorrect control settings
adjust, adjustment, adjustable, adjustability | 2 | unable single hand to adjust | device difficult to setup or prepare
adjust, adjustment, adjustable, adjustability | 2 | single member adjustment failed | device difficult to setup or prepare
adjust, adjustment, adjustable, adjustability | 2 | unable auto adjustment | device difficult to setup or prepare
advance | 3 | difficult to advance | difficult to advance
advance |  | excessive force transferred | device handling issue
AF | 2 | AF induction | wireless communication issue
air | 3 | air leak noticed | air leak
air |  | air in line | infusion or flow issue; connection issue; improper flow or infusion; obstruction within device
air |  | Air Embolism | obstruction within device
air |  | unspecified air in line alarm | device alarm system issue; infusion or flow issue
air |  | Proximal Occlusion alarm | device alarm system issue; infusion or flow issue
air |  | Distal Air in Line Alarm | device alarm system issue; infusion or flow issue
air |  | allow air passage | air leak
air |  | air entrainment | infusion or flow issue
air |  | air bubbles | infusion or flow issue
air |  | sucking air in line | infusion or flow issue
air |  | air infiltration | infusion or flow issue
air |  | cold air | temperature issue
airflow | 2 | loss of airflow | no flow
airflow | 2 | partial loss of airflow | insufficient flow or underinfusion
alarm | 2 | alarm delayed | delayed alarm
alarm |  | alarm interrupted | device alarm system issue
alarm | 2 | alarm not visible | alarm not visible
alarm |  | alarm notified | device displays error message; device inoperable
alarm | 2 | alarm stuck off | device alarm system issue
alarm | 2 | alarm stuck on | defective alarm
alarm | 2 | annoying alarm | improper alarm; false alarm
alarm |  | device alarmed | device alarm system issue
alarm | 2 | fail to activate alarm | device alarm system issue; device operates differently than expected; improper alarm
alarm |  | fail to trigger alarm by software | not audible alarm
alarm |  | fail to trigger alarm by hardware | not audible alarm
alarm |  | fail to generate visual alarm | device alarm system issue
alarm |  | inappropriate alarm settings | device alarm system issue
alarm |  | fail to adjust alarm settings | device alarm system issue; use of incorrect control settings
alarm | 2 | fail to clear alarm | defective alarm
alarm | 2 | fail to disable alarm | defective alarm; device alarm system issue
alarm | 2 | fail to enter safe state | device unsafe to use in environment
alarm | 2 | fail to raise high priority alarm | defective alarm; device alarm system issue
alarm | 2 | fail to raise low priority alarm | defective alarm
alarm | 2 | fail to raise medium priority alarm | defective alarm
alarm | 2 | fail to set or reset alarm automatically | reset issue; device alarm system issue
alarm | 2 | fail to set or reset alarm manually | device alarm system issue; reset issue
alarm | 2 | fail to silence alarm | defective alarm
alarm |  | fail to suppress low priority alarms | device alarm system issue
alarm | 2 | fail to view alarm settings | device alarm system issue
alarm |  | fail to disclose alarm test | instruction for use issue
alarm | 2 | false alarm | false alarm
alarm | 2 | improper alarm | improper alarm
alarm |  | incorrect message display | device alarm system issue
alarm |  | incorrect alarm configuration | device alarm system issue
alarm | 2 | inadequate alarm states | device alarm system issue
alarm |  | inadequate alarm log file | patient data issue
alarm | 2 | inappropriate audible tone | defective alarm
alarm |  | intermittent alarm | defective alarm
alarm |  | loss of alarm | device alarm system issue
alarm | 2 | low alarm volume | defective alarm
alarm | 2 | software does not alarms as needs | device alarm system issue
alarm | 2 | software provides false error alarms | false alarm
alarm | 2 | unnoticeable alarm | low audible alarm
alarm | 2 | user is not aware of alarm condition | alarm not visible
alert | 3 | fail to activate alert | device alarm system issue; failure to deploy
alert | 3 | fail to clear alert | prompts will not clear
alert |  | alert notified | device displays error message
alert | 3 | false alert | false alarm
alert | 3 | frequent annoyed clinical alerts | improper alarm
algorithm | 1 | algorithm execution delayed | programming issue
algorithm | 1 | algorithm execution problem | programming issue
algorithm | 1 | confused about the data | data issue
algorithm | 1 | fail to execution algorithm | programming issue
algorithm | 1 | false detections and inappropriate mode | improper or incorrect procedure or method
algorithm | 1 | inconsistent algorithm | programming issue
algorithm | 1 | incorrect algorithm | programming issue
algorithm |  | inappropriate algorithm setting | use of incorrect control settings
algorithm | 2 | incorrect priority level | programming issue
algorithm | 1 | ineffective algorithm | programming issue
algorithm | 1 | intermittent algorithm execution | programming issue
algorithm | 1 | non-optimal therapy | improper device output
align, alignment | 2 | fail to align | failure to align
align, alignment | 2 | fail to guide alignment | failure to align
align, alignment | 2 | misalignment issue | failure to align
align, alignment | 2 | misalignment out of range | failure to align
allergic, allergy | 1 | allergic composition | biocompatibility issue
allergic, allergy | 1 | allergic reaction | biocompatibility issue
allergic, allergy | 1 | inadvertent contact with allergic parts | biocompatibility issue
allergic, allergy | 1 | irritation | patient-device incompatibility
allergic, allergy | 1 | skin allergic reaction | biocompatibility issue
alternator | 3
altitude | 3 | altitude accommodation slow | environmental control or utility issue
altitude | 3 | altitude variation | environmental control or utility issue
altitude |  | lack of barometric compensation | imprecision
altitude | 3 | loss of function due to altitude change | device inoperable
amateur, athlete, adult | 3 | inappropriate user | user used incorrect product for intended use
amplifier, preamplifier | 3
amplitude | 3 | high amplitude | artifact
amplitude | 3 | inaccurate amplitude measurement | incorrect measurement
analyze, analyzer | 3 | fail to analyze change | failure to analyze rhythm
angle, angular | 3 | angle out of specification | defective component
animal |  | inappropriate stunning method | improper or incorrect procedure or method
annunciator | 3
anode | 3 | anode too large | battery issue
anode | 3 | electrical short | battery issue; electrical shorting; spark
anode | 3 | inappropriate anode number | battery issue
anode | 3 | inappropriate anode size | battery issue
anomaly | 1 | fail to evaluate anomaly | use of device issue; improper device output
anomaly | 1 | fail to record anomaly | data back-up problem
antenna | 3 | antenna bridging to another conductor | electrical shorting
antenna | 3 | antenna broken or damaged | break
antenna | 3 | antenna protruding out | material protrusion
antenna | 3 | antenna touching itself | electrical shorting
antenna | 3 | poor telemetry | radio signal problem
antenna | 3 | weak antenna anchor weld | loose or intermittent connection
anticoagulant | 2 | premature use of anticoagulant | improper or incorrect procedure or method
applicator | 3 | applicator failure | device inoperable
apply, application | 3 | misapplication | device handling issue
approve | 3 | approved inappropriately | compliance issue
approve | 3 | fail to approve | compliance issue
approve | 3 | fail to approve change | compliance issue
arc | 2 | arcing at paddles | arcing at paddles
arc | 2 | arcing between barrel pin receptacles | arcing
arc | 2 | arcing occur | arcing
arc | 2 | arcing of electrodes | arcing at electrodes
arc | 2 | circuit broken or damaged | circuit failure; break
arc | 2 | high current | electronic property issue
architecture | 3 | fail to support interface | compliance issue
architecture | 3 | fail to support SOUP | compliance issue
architecture | 3 | inadequate architecture design | component missing
architecture |  | inadequate system design | component missing
architecture | 3 | inappropriate architecture | compliance issue
architecture | 3 | lack of architecture | component missing
area |  | fail to calculate area | incorrect software programming calculations
area | 3 | area out of specification | defective component
arm, c-arm |  | fail to run | device inoperable
arm, c-arm |  | c-arm dropped | component falling
arm, c-arm | 3 | unintended arm motion | unintended arm motion
arrester | 3
arrhythmia | 2 | undetected arrhythmia | failure to sense
artery | 2 | thoracic artery punctured | use of device issue; Material Puncture
assembly, assemble, assembled | 3 | fail to assemble | disassembly
assembly, assemble, assembled |  | assembly incomplete | device misassembled during manufacturing or shipping
assembly, assemble, assembled | 3 | device misassembled | device misassembled during manufacturing or shipping
assembly, assemble, assembled | 3 | fail to engage | deployment issue
assembly, assemble, assembled | 3 | misassembly by user | Misassembled by Users
assembly, assemble, assembled | 3 | misassembly during installation | installation-related problem
assembly, assemble, assembled | 3 | misassembly during maintenance | device maintenance issue; misassembled
assembly, assemble, assembled | 3 | misassembly during repair | misassembled
assembly, assemble, assembled | 3 | scratch on the implant | Scratched Material
astm | 3 | ASTM violated | compliance issue
atomizer, nebulizer, nebulization |  | clogged | device clogged; failure to deliver; no flow
atomizer, nebulizer, nebulization |  | device blocked | device clogged; failure to deliver; no flow
atomizer, nebulizer, nebulization |  | device malfunction | device operates differently than expected; Product Quality Issue
atomizer, nebulizer, nebulization |  | incompatible connection | Detachment of Device or Device Component
atomizer, nebulizer, nebulization |  | improper drug nebulization | device operates differently than expected
atomizer, nebulizer, nebulization |  | fail to spray anesthetic | failure to deliver; failure to fire; no flow
atomizer, nebulizer, nebulization |  | fail to interface with other device | device-device incompatibility
atomizer, nebulizer, nebulization |  | fail to produce sufficient mist | insufficient flow or under infusion
atomizer, nebulizer, nebulization |  | fail to warn use of gas-powered nebulizer | instruction for use issue
atomizer, nebulizer, nebulization |  | fail to produce mist | device clogged; failure to deliver; failure to fire; no flow
atrial | 2 | atrial evoked response small | Failure to Interrogate; Device Sensing Issue
atrial | 2 | atrial overdrive | Over-sensing
atrial | 2 | atrial undersensing | under-sensing
atrial | 2 | inappropriate atrial sensing | invalid sensing
atrial | 2 | loss of capture | failure to sense
atrial | 2 | loss of sensing | failure to sense
atrial | 2 | premature atrial contraction | Over-sensing
attach, attachable, attachment | 3 | attached loosely | detachment of device component
attach, attachable, attachment | 3 | attachment failed | detachment of device component
attach, attachable, attachment | 3 | difficult to be attached | detachment of device component
attach, attachable, attachment | 3 | fail to attach within time frame | detachment of device component
attach, attachable, attachment | 3 | loose part | detachment of device component
attach, attachable, attachment | 3 | unable onehand to attach | detachment of device component
attenuator | 3
audio, auditory, audible, audibility | 3 | audio jack mismatched to plug | fitting problem; component incompatible
audio, auditory, audible, audibility | 3 | audio latency over specification | delayed alarm
audio, auditory, audible, audibility |  | inadequate audibility design | human-device interface issue
audio, auditory, audible, audibility | 3 | loss of audible alarm | not audible alarm
audio, auditory, audible, audibility | 3 | loss of audio | not audible alarm
audio, auditory, audible, audibility | 3 | low audible alarm | low audible alarm
audio, auditory, audible, audibility | 3 | low audio volume | use of device issue
augmentate, augmentation | 3 | loss of augmentation | electrical issue
authenticate, authentication | 1 | authentication failed | unauthorized access to computer system
authenticate, authentication | 1 | authentication key corrupted | unauthorized access to computer system
authorize, authorization, unauthorize, unauthorization | 1 | unauthorized access | unauthorized access to computer system
authorize, authorization, unauthorize, unauthorization | 1 | unauthorized communication with device | unauthorized access to computer system
authorize, authorization, unauthorize, unauthorization |  | inappropriate authorization lists | market issue
authorize, authorization, unauthorize, unauthorization | 1 | unauthorized deletion | unauthorized access to computer system
autotrak | 2 | fail to autotrak | device inoperable
backdrive, backdrivable, back-drive, back-driveable | 2 | backdrive out of range | defective item
backflow | 2 | excess backflow | insufficient flow or underinfusion
backlash | 3 | backlash over limit | Bleed Back
backlight | 3 | difficult to setup | device difficult to setup or prepare; optical issue
backlight |  | fail to adjust brightness | optical issue
backlight | 3 | Insufficient brightness | optical issue
backpannel, backplane, baffle | 3
back-up, backup, redundant, redundancy, back up, dual | 3 | fail to back-up | Failure to Back-Up
back-up, backup, redundant, redundancy, back up, dual | 3 | fail to convert to back-up | failure to convert to back-up
back-up, backup, redundant, redundancy, back up, dual | 3 | switch to backup mode incorrectly | improper or incorrect procedure or method; fail-safe mechanism issue
bag, softpack, backpack, resuscitation bag |  | bag broken or damaged | break
bag, softpack, backpack, resuscitation bag |  | bag deflated | deflation issue; device damaged prior to use
bag, softpack, backpack, resuscitation bag |  | bag inflated | Inflation Issue
bag, softpack, backpack, resuscitation bag | 3 | bag improperly sized | size incorrect for patient
bag, softpack, backpack, resuscitation bag | 3 | inability to hear alarm | not audible alarm
bag, softpack, backpack, resuscitation bag | 3 | inadequate padding in strap | component missing
bag, softpack, backpack, resuscitation bag | 3 | stitching broken or damaged | break
bag, softpack, backpack, resuscitation bag | 3 | strap broken or damaged | break
bag, softpack, backpack, resuscitation bag | 3 | unable to carry in bag | human-device interface issue
bag, softpack, backpack, resuscitation bag |  | unable to fill | filling problem
bag, softpack, backpack, resuscitation bag |  | unable to inflate | Inflation Issue
bag, softpack, backpack, resuscitation bag |  | unable to deflate | deflation issue
bag, softpack, backpack, resuscitation bag |  | unable to ventilate with bag | deflation issue
bag, softpack, backpack, resuscitation bag |  | unable to compress | deflation issue
bag, softpack, backpack, resuscitation bag |  | high resistance to open | difficult to open or close
bag, softpack, backpack, resuscitation bag | 3 | zipper broken or damaged | break
bag, softpack, backpack, resuscitation bag | 3 | zipper stuck | sticking; difficult to open or close
balance | 2 | loss of balance | unstable
balloon | 2 | balloon asymmetrical | deflation issue
balloon | 2 | balloon burst | Burst
balloon |  | balloon detached | detachment of device component; material separation; fluid leak
balloon |  | balloon deflated | deflation issue; pressure issue
balloon | 2 | balloon leakage | deflation issue
balloon | 2 | balloon mushroomed | deflation issue
|  | balloon over inflation | Inflation Issue
balloon | 2 | balloon pinhole | deflation issue
balloon | 2 | balloon ruptured | Burst
balloon |  | fail to integrate balloon to system | fitting problem; component incompatible
balloon |  | fail to inflate balloon evenly | Inflation Issue
balloon |  | inadequate balloon apposition | positioning issue
balloon |  | inappropriate deflation times | deflation issue
balloon |  | inappropriate balloon type | deflation issue; Inflation Issue
balloon |  | inappropriate rated burst pressure | pressure issue; Burst
balloon |  | inappropriate rated burst volume | Burst
balloon |  | low rated fatigue | deflation issue
balloon |  | unable to inflate | Inflation Issue
bandwidth |  | wireless bandwidth limitation | electronic property issue; optical issue
bandwidth | 3 | bandwidth out of range | electronic property issue; optical issue
barometer | 3 | measurement high | incorrect measurement
barometer | 3 | measurement low | incorrect measurement
basal | 3 | delivery delayed | delivery system failure
basal | 3 | incorrect basal calculation | dose calculation error due to software problem
basal | 3 | uneven basal rate | inaccurate delivery
baseline | 3 | Inaccurate baseline | inaccurate delivery
battery |  | battery broken or damaged | battery issue; break
battery | 2 | battery corrosion | Corrosion
battery | 2 | battery depletion | Premature Discharge of Battery
battery | 1 | battery exploded | Misfire
battery |  | battery fault displayed | device displays error message
battery | 2 | battery increment high | battery issue
battery | 2 | battery leakage | battery issue; chemical issue
battery | 2 | battery open | circuit failure
battery | 2 | battery switch failed | defective component; battery issue
battery |  | battery switched inappropriately | battery issue
battery | 2 | chemical release | chemical spillage
battery | 2 | confused battery status | battery issue
battery | 2 | current drain | battery issue
battery |  | drained battery | battery issue
battery | 2 | expired battery used | device expiration issue; battery issue
battery |  | fail to charge | battery issue
battery | 2 | fail to communicate with driver | battery issue
battery |  | fail to full charge battery | battery issue
battery | 2 | fail to indicate battery charging status | no display or display failure
battery |  | fail to label standard symbol | Packaging Issue
battery | 2 | fail to measure battery voltage | battery issue; measurement system incompatibility
battery | 2 | fail to monitor battery temperature | battery issue; temperature issue
battery |  | fail to pass battery test | battery issue
battery | 2 | fail to provide low battery indicator | no display or display failure
battery | 2 | fail to run on battery | battery issue
battery | 2 | fail to supply power | battery issue
battery |  | false battery alarm | defective alarm
battery | 2 | fire | Fire; spark
battery | 2 | high impedance | High Battery Impedance
battery |  | improper disposal of battery | environmental control or utility issue; battery issue
battery | 2 | inaccurate battery data | battery issue; imprecision
battery |  | incorrect battery level displayed | device displays error message
battery | 2 | incorrect polarity connection | battery issue; misconnection
battery | 2 | internal battery short | battery issue
battery | 2 | lack of CB certified | battery issue
battery | 2 | lack of over charge module | charging issue
battery | 2 | lack of over current module | battery issue; no fail-safe mechanism
battery | 2 | lack of over heat module | battery issue
battery | 2 | lack of over voltage module | battery issue
battery | 2 | lack of protection circuit | battery issue
battery | 2 | lack of sea transportation certified | battery issue; Manufacturing or Shipping Issue Associated with Device
battery | 2 | lack of short circuit module | electrical shorting; spark
battery | 2 | loss of battery measurement | incorrect measurement
battery | 2 | low battery | Low Battery
battery | 2 | low carbon black content | battery issue
battery | 2 | low impedance | Low Battery Impedance
battery | 2 | low power output | battery issue
battery | 2 | non-rechargeable battery | battery issue
battery | 2 | overheat | overheating of device or device component
battery | 2 | power out of range | battery issue
battery | 2 | reduced battery life cycle | battery issue
battery | 2 | reduced battery capacity | Low Battery
beat | 2 | aberrant beat | Device Sensing Issue
bed | 3 | bed ripped | Material Rupture
bed | 3 | bed torn | torn material
belt | 3 | accessory missing | component missing
belt | 3 | accessory unattached | detachment of device component
belt | 3 | belt mismatched | component incompatible
belt | 3 | belt size out of specification | size incorrect for patient; patient-device incompatibility
belt | 3 | clinic infection | bacterial contamination of device
belt | 3 | component detached | detachment of device component
belt | 3 | component torn | torn material
belt | 3 | difficult to apply device | use of device issue
belt | 3 | lack of labeling | device markings issue
belt | 3 | low strength | mechanical issue
belt | 3 | material degradation by inappropriate cleaning | device cleaning issue
belt | 3 | material frayed or unraveling | Material Frayed, Unraveled Material
belt | 3 | missing brand | device markings issue
belt | 3 | out of dimension | defective component
belt | 3 | reduced life cycle | shelf life exceeded
belt | 3 | seam torn or broken | torn material
belt | 3 | split coating | Split
belt | 3 | soft good worn | Naturally Worn
belt | 3 | stitching ripped or unraveled | Unraveled Material
belt | 3 | strap shorten | device misassembled during manufacturing or shipping
belt | 3 | strap torn or broken | torn material
belt | 3 | unmatched belt set | patient-device incompatibility
belt | 3 | usage misaligned | improper or incorrect procedure or method
belt | 3 | weight out of specification | patient-device incompatibility; material integrity issue
bend | 3 | over bending | Bent
BG, blood glucose | 2 | BG variation | improper device output
BG, blood glucose | 2 | fail to check blood glucose periodically | failure of device to self-test
BG, blood glucose | 2 | incorrect BG measurement data | incorrect measurement; data issue
BG, blood glucose | 2 | low blood glucose level | improper device output
bias | 3 | no bias | improper device output
bias | 3 | wrong bias | improper device output
bifurcation, bifurcated |  | inappropriate prosthesis geometry | size incorrect for patient; patient-device incompatibility
binder | 3 | high binder content | improper device output
bio | 1 | biohazard contamination | device contamination with biological material
bio | 1 | biological substance | biocompatibility issue
bio | 1 | material biocompatibility compromised | biocompatibility issue
bio | 1 | slow and gradual medicine degradation | degraded
bioburden |  | viable microorganism | bacterial contamination of device
biocompatibility, biocompatible, bio-compatible | 1 | material biocompatibility compromised | biocompatibility issue
biocompatibility, biocompatible, bio-compatible |  | foreign matter | biocompatibility issue
biocompatibility, biocompatible, bio-compatible |  | cross contamination | bacterial contamination of device
biocompatibility, biocompatible, bio-compatible | 1 | material bio-incompatibility | biocompatibility issue
biological | 1 | biohazard contamination | device contamination with biological material
biological | 1 | biological substance | biocompatibility issue
biological | 1 | slow and gradual medicine degradation | degraded
biostability |  | loss of chemical integrity in living tissue | chemical issue
biostability |  | loss of physical integrity in living tissue | material integrity issue
bleed | 1 | anaphylactic reaction | Improper Chemical Reaction
bleed | 1 | excessive bleeding | use of device issue
bleed | 1 | tissue bleed | mechanical issue
block, blockage | 2 | complete blockage | Blockage within Device or Device Component
block, blockage | 2 | fail to block non-essential communication | Blockage within Device or Device Component
block, blockage |  | fail to block | Inflation Issue
block, blockage | 2 | no delivery | Blockage within Device or Device Component
block, blockage | 2 | partial blockage | Blockage within Device or Device Component
block, blockage | 2 | port blocked | blocked connection
block, blockage | 2 | under delivery | Blockage within Device or Device Component
blood | 2 | blood gas measurement erroneous | incorrect measurement
blood |  | damaged blood product | improper flow or infusion
blood |  | Hemolyzed blood delivered | improper flow or infusion
blood | 2 | blood in tubing | Bleed Back
blood |  | blood backflow | Bleed Back
blood |  | blood clotted | chemical issue; human-device interface issue
blood |  | blood leakage | fluid leak
blood |  | blood loss | fluid leak
blood | 2 | blood pooling | Bleed Back
blower | 3 | air delivery compromised | inaccurate delivery
blower | 3 | blower blocked | Blockage within Device or Device Component; obstruction within device
blower |  | blower failed | defective component; no flow
blower | 3 | blower overheat | overheating of device or device component
blower | 3 | blower runaway | device inoperable
blower | 3 | blower stopped | device operates differently than expected; no flow
blower | 3 | less air input | failure to read input signal
blower |  | noisy blower | audible noise
blower | 3 | particulate generated | Particulates
bolt | 3
bolus | 2 | bolus amount out of range | inaccurate delivery
bolus | 2 | bolus oversized | inaccurate delivery
bolus | 2 | delivery delayed | delivery system failure
bolus | 2 | fail to generate bolus | failure to deliver
bolus | 2 | fail to generate immediate bolus | failure to deliver
bolus | 2 | inadequate bolus size | inaccurate delivery
bolus | 2 | inappropriate new bolus generated | failure to deliver
bolus | 2 | incorrect bolus calculation | dose calculation error due to software problem
bolus | 2 | initiate bolus undesired | inaccurate delivery
bolus | 2 | lack of bolus function | failure to deliver
bolus | 2 | slow delivery | delivery system failure
bolus | 2 | unable to cancel bolus | use of device issue
bond | 2 | bond broken or damaged | failure to adhere or bond; break
bond | 2 | bond lift | failure to adhere or bond
bond | 2 | bond mislocated | dislodged or dislocated; failure to adhere or bond
bond | 2 | bond over sized | device misassembled during manufacturing or shipping
bond | 2 | bond pad broken or damaged | Loss of or Failure to Bond; break
bond |  | inadequate bond tensile strength | mechanical issue; failure to adhere or bond
bond |  | inadequate bond torque strength | mechanical issue; failure to adhere or bond
bond | 2 | bond peel | failure to adhere or bond
bond | 2 | fail to bond | failure to adhere or bond
bond | 2 | foreign material inside | failure to adhere or bond; Packaging Issue
bond | 2 | loss of bond | Loss of or Failure to Bond
boundary | 3 | out of limits | defective component
brace | 3 | bio-contamination | device contamination with biological material; delivered as unsterile product
brace | 3 | brace control stop failed | device operates differently than expected
brace | 3 | brace misused | use of device issue
brace | 3 | brace uncomfortable | human-device interface issue
brace | 3 | difficult to apply brace | human-device interface issue; device difficult to setup or prepare
brace | 3 | fail to be washed | device cleaning issue
brace | 3 | fail to deliver adequate immobilization | failure to deliver
brace | 3 | fail to respect reimbursement | Market Issue
brace | 3 | fail to support light sport | use of device issue
brace | 3 | inappropriate distribute weight | use of device issue; mechanical issue
brace | 3 | incompatible accessory | accessory incompatible
brace | 3 | incorrect brace size | size incorrect for patient; patient-device incompatibility
brace | 3 | incorrect to apply brace | device handling issue
brace | 3 | insufficient stability | unstable
brace | 3 | less flexible | material rigid or stiff
brace | 3 | material bio-incompatibility | biocompatibility issue
brace | 3 | missing part | component missing
brace | 3 | missing splint | component missing
brace | 3 | part out of specification | defective component
brace | 3 | radiolucent material compromised | material integrity issue
brace | 3 | reducing inversion or eversion movement failed | unintended movement
brace | 3 | restricting much plantar or dorsi flexion | difficult to remove
brace | 3 | support failed | use of device issue
brake | 1 | loss of brake | unintended movement
brake | 1 | mechanical brakes disengaged | unintended movement; deployment issue
brake | 1 | weak mechanical brake strength | mechanical issue
brand | 3 | incorrect brand | compliance issue
brand |  | branding misaligned | compliance issue
brand | 3 | missing brand | device markings issue
breach | 1 | loss of electrical isolation | electrical shorting
breadth | 3 | breadth out of range | defective component
break | 2 | fail to break | device inoperable
breath | 2 | fail to initiate breath | device inoperable
breath |  | fail to trigger breath | device inoperable
breath |  | fail to disable breath | excess flow or over-infusion
breath | 2 | inadequate breath | use of device issue
breath |  | reduced breath | use of device issue
breath | 2 | incorrect breath type | use of device issue
breath | 2 | initiate engagement failed | deployment issue
breath |  | no breath | infusion or flow issue
breath | 2 | pressure control failed | pressure issue
bright, brightness | 3 | brightness out of specification | optical issue
bright, brightness |  | fail to adjust brightness | application program issue
bright, brightness | 3 | Insufficient brightness | optical issue
brush | 3 | brush wear out | Naturally Worn
bubble | 2 | electrical short | electrical shorting; spark
bubble | 2 | fail to identify bubble | failure to sense
bubble | 2 | loss of electrical isolation | electrical shorting
bubble | 2 | moisture inside device | Moisture or Humidity Problem
bubble | 2 | potential bridging between wires | electrical shorting
buckle | 3 | buckle broken or damaged | break
bug |  | major bug existed | computer software issue
bug | 1 | lack of bug fix procedure | computer software issue
burn |  | black soot | device emits odor
burn |  | incorrect image burnt | data issue
burn |  | device burned | smoking
burn |  | burning odor | device emits odor; smoking
bus | 3
bushing | 3
button, key, keypad | 3 | button detachment | detachment of device component
button, key, keypad | 3 | button hard to press | failure to deploy
button, key, keypad | 3 | button stuck off | sticking
button, key, keypad | 3 | button stuck on | sticking
button, key, keypad |  | button works intermittently | defective component
button, key, keypad | 3 | confused button | inadequate user interface
button, key, keypad |  | delayed button | defective component
button, key, keypad | 3 | fail to indicate button status | no display or display failure
button, key, keypad |  | fail to adjust image contrast | application program issue
button, key, keypad |  | fail to adjust image brightness | application program issue
button, key, keypad | 3 | Inadvertently activating button accidentally | deployment issue; use of device issue
button, key, keypad | 3 | indistinct button | human-device interface issue
button, key, keypad | 3 | key stuck off | sticking
button, key, keypad | 3 | key stuck on | sticking
button, key, keypad | 3 | keypad worn out | Naturally Worn
button, key, keypad | 3 | keypad/button broken or damaged | break
button, key, keypad | 3 | lack of button | inadequate user interface
button, key, keypad | 3 | lack of debounce button design | inadequate user interface
button, key, keypad | 3 | non-momentary press type button applied | use of device issue
button, key, keypad | 3 | self-buttoning | deployment issue
button, key, keypad | 3 | self-keying | deployment issue
button, key, keypad | 3 | sticky button | defective component
button, key, keypad |  | non unique key | human-device interface issue; protective measure issue; misconnection
button, key, keypad | 3 | uncomfortable button | human-device interface issue
cable | 3 | cable abrasion | device abrasion from instrument or another object
cable | 3 | cable broken or damaged | break
cable | 3 | cable disconnection | disconnection
cable | 3 | cable length out of range | defective component
cable | 3 | cable oversized | defective component
cable | 3 | cable short | defective component
cable | 3 | cable stored inappropriately | device handling issue
cable | 2 | compromised insulation | electrical shorting
cable |  | invalid cable | component incompatible
cable |  | weak mechanical strength | mechanical issue
cable | 3 | lack of locking mechanism | protective measure issue
cable | 3 | Trip hazard | tip-over
calcium |  | calcified | sediment, precipitate or deposit in device or device ingredient; calcified
calculate, calculation | 2 | dose calculation incorrect | dose calculation error due to software problem; data issue
calculate, calculation | 2 | fail to generate periodic calculation | incorrect software programming calculations; data issue
calculate, calculation | 2 | incorrect software programming calculation | incorrect software programming calculations; data issue
calculate, calculation |  | inconsistent calculation | incorrect software programming calculations; patient data issue
calculate, calculation |  | inaccurate calculation | incorrect software programming calculations; patient data issue
calculate, calculation |  | incorrect calculation | incorrect software programming calculations; data issue
calculate, calculation | 2 | misuse the drug calculation feature | dose calculation error due to software problem; use of device issue
calculate, calculation | 2 | reprogram incorrect | incorrect software programming calculations
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | calibration failed | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | difficult to calibrate | calibration issue
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | fail to calibrate circuit compensation | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | fail to calibrate manually | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to activate calibration at multiple points | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to activate automatic calibration | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | fail to recalibrate | failure to recalibrate
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to display calibration | image display error
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to disclose calibration | instruction for use issue
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to disclose sensor calibration | instruction for use issue
calibrate, calibration, calibrator, recalibrate, recalibration |  | fail to trigger calibration | Failure to Calibrate
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | miscalibration | Calibration Error
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | out of calibration | imprecision
calibrate, calibration, calibrator, recalibrate, recalibration | 1 | short validation period | calibration issue
camera | 3 | fail to turn on | device inoperable
cancel, cancellation | 3 | fail to cancel | use of device issue
cannula, cannulate, cannulation | 3 | inner out of specification | defective component
cannula, cannulate, cannulation |  | cannula short | defective component
cannula, cannulate, cannulation | 3 | outer out of specification | defective component
capacitor | 3 | capacitor broken or damaged | break
capacitor | 3 | capacitor deformation | defective component
capacitor | 3 | charge time extended | charging issue
capacitor | 3 | excessive leakage current | electrical issue
capacitor | 3 | fail to seal | failure to adhere or bond
capacitor | 3 | inappropriate capacitor charging | charging issue
capacitor | 3 | inappropriate capacitor maintenance | device maintenance issue
capacitor | 3 | out of dimension | defective component
capacitor | 3 | partial energy delivered | output energy incorrect
capacitor | 3 | voltage out of range | electrical issue
capsulotomy | 2 | deflation | deflation issue
capture, autocapture | 2 | autocapture is not enabled | Failure to Capture
capture, autocapture | 2 | capture misinterpreted | capturing issue
capture, autocapture | 2 | capture threshold decreased | unstable Capture Threshold
capture, autocapture | 2 | fail to identify capture threshold | Loss of Threshold
capture, autocapture |  | fail to capture data | data issue
capture, autocapture | 2 | inadequate working margin | capturing issue
capture, autocapture | 2 | intermittent capture | Intermittent Capture
capture, autocapture | 2 | loss of capture | Failure to Capture
capture, autocapture | 2 | signal desynchrony | capturing issue
capture, autocapture | 2 | unrecognized capture | capturing issue
capture, autocapture | 2 | unstable capture threshold | unstable Capture Threshold
carrier | 3
cascade | 3
case | 3 | water ingress | fluid leak
cassette |  | foreign material inside | particulates
cassette |  | inappropriate fit | fitting problem
cassette |  | invalid filter | component incompatible
cassette |  | difficult to insert | Difficult to Insert
cassette |  | difficult to remove | difficult to remove
cassette |  | incompatible cassette | accessory incompatible
cassette | 3 | cassette leakage | fluid leak
cast | 3 | casting not completely removed | device misassembled during manufacturing or shipping
cast | 3 | opaque casting | device misassembled during manufacturing or shipping
cast | 3 | rough casting | device misassembled during manufacturing or shipping
caster | 3
catheter | 2 | catheter cracked | Crack
catheter | 2 | catheter degradation | degraded
catheter | 2 | catheter detached | detachment of device component; fluid leak; disconnection
catheter | 2 | catheter incompatible with liquid | Incompatibility Problem
catheter |  | catheter kinked | Kinked
catheter | 2 | catheter leakage | leak
catheter | 2 | catheter obstructed | blocked connection
catheter | 2 | catheter pull out | Retraction Problem
catheter |  | IV Catheter out of vein | Retraction Problem
catheter | 2 | catheter tip encapsulation | device misassembled during manufacturing or shipping
catheter |  | hole in catheter | Material Puncture
catheter | 2 | lack of radiopaque stripe | device markings issue
catheter | 2 | loss of sterility | Delivered As Unsterile Product; Failure to Sterilize
cathode | 3 | cathode cracked | Crack
cathode | 3 | cathode sharp edge | Material Puncture
cathode | 3 | cathode tab misalignment | failure to align
cathode | 3 | electrical short | electrical shorting; spark
cathode | 3 | inappropriate cathode powder mix | device misassembled during manufacturing or shipping
cathode | 3 | inappropriate cathode size | defective component
caution | 2 | inadequate caution | no display or display failure
caution | 2 | missing caution | no display or display failure
cell | 3
CFR | 2 | CFR regulation violated | compliance issue
chamber | 3 | induced cross chamber rectification | electronic property issue
chamber |  | drip chamber discolored | Material Discolored
chamber |  | drip chamber cloudy | Material Discolored; Material Opacification
chamber |  | fail to warn usage in hyperbaric chamber | instruction for use issue
chamber | 3 | VT/VF induction | electronic property issue
change | 2 | fail to analyze change | compliance issue
change | 2 | fail to change | compliance issue
change |  | fail to save change | loss of data; data issue
change | 2 | fail to inform change | compliance issue
change | 2 | inappropriate change | compliance issue
change | 2 | introduce side effect | use of device issue
change | 2 | lack of additional test | Output Issue
change | 2 | lack of change analysis | compliance issue
change | 2 | lack of change control | compliance issue
change | 2 | lack of change impact analysis | compliance issue
change | 2 | lack of effectiveness check | failure of device to self-test
change | 2 | lack of repeat test | use of device issue
change | 2 | lack of risk management activity | Fail-Safe Design Failure
change | 2 | lack of verification | protective measure issue
change | 2 | unintended change | use of device issue
channel | 3 | crosstalk in channel | radio signal problem
channel | 3 | interval variation in channel | radio signal problem
channel | 3 | noise in channel | radio signal problem
charge, charger | 2 | aborted charge | aborted charge
charge, charger |  | charge fault displayed | device displays error message
charge, charger | 2 | charge leakage | charging issue
charge, charger | 2 | fail to charge | Failure to Charge
charge, charger | 2 | fail to check charge level periodically | failure of device to self-test
charge, charger |  | fail to full charge battery | Failure to Charge
charge, charger | 2 | fail to indicate battery charging status | no display or display failure
charge, charger | 2 | long charge time | delayed charge time
charge, charger | 2 | over charge | charging issue
charge, charger | 2 | residual charge | charging issue
charge, charger | 2 | under charge | charging issue
chassis | 3 | weak mechanical interface strength | mechanical issue
check | 3 | fail to pass check | failure of device to self-test
check | 3 | fail to interval check | failure of device to self-test
check | 3 | fail to periodic check | failure of device to self-test
check | 3 | fail to schedule check | failure of device to self-test
check | 3 | fail to system check | failure of device to self-test
chemical | 2 | chemical reaction | Improper Chemical Reaction
chemical |  | Caustic Substances | Material Erosion
chemical |  | BPA leached | biocompatibility issue
chemical |  | chemical toxin | biocompatibility issue
chemical | 2 | chemical spillage | chemical spillage
chemical | 2 | incorrect chemical composition | chemical issue
circuit |  | circuit blocked | occlusion within device
circuit | 3 | circuit broken or damaged | circuit failure
circuit |  | circuit leakage | leak; air leak
circuit | 3 | fail to open circuit | circuit failure
circuit |  | circuit disconnected | disconnection
circuit | 3 | open circuit | circuit failure
circuit | 3 | short circuit | electrical shorting; spark
claim | 3 | false claim | compliance issue
clamp, roller clamp |  | fail to occlude fluid flow | fluid leak
clamp, roller clamp |  | fail to retain position | positioning issue
clamp, roller clamp |  | restrict flow undesired | improper flow or infusion
clamp, roller clamp |  | tubing broken or damaged | break
clamp, roller clamp | 3 | detach without tool | detachment of device component
class | 3
classification, classify | 3 | inappropriate classification | compliance issue
clean, cleanable, cleanser, cleaning | 1 | inappropriate cleaning cycle | device cleaning issue | cleaning cycle failed
clean, cleanable, cleanser, cleaning | 1 | clinic infection | bacterial contamination of device
clean, cleanable, cleanser, cleaning | 1 | device damaged by inappropriate cleaning | device handling issue
clean, cleanable, cleanser, cleaning | 1 | fail to clean adequately | device cleaning issue
clean, cleanable, cleanser, cleaning |  | fail to remove microbiological contamination | bacterial contamination of device
clean, cleanable, cleanser, cleaning | 1 | foreign material inside | metal shedding debris; contamination during use; Packaging Issue
clean, cleanable, cleanser, cleaning | 1 | inappropriate cleanser | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | material degradation by inappropriate cleaning | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | product damaged by inappropriate washing | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | surface adverse impacted by cleaning agent | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | unable for cleaning | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | unclean device is delivered | device cleaning issue
clean, cleanable, cleanser, cleaning | 1 | unqualified cleaning agent | device cleaning issue
clean, cleanable, cleanser, cleaning |  | water ingress | device maintenance issue
clear | 2 | fail to clear data | data issue
clearance | 2 | inadequate clearance | protective measure issue; improper or incorrect procedure or method
clearance | 2 | insufficient safety clearance | device unsafe to use in environment
clip, scissors |  | clip force out of specification | device handling issue
clip, scissors | 3 | clip disengaged | deployment issue
clock | 2 | clock advanced | Date-related software issue
clock | 2 | clock out of tolerance | Date-related software issue
clock | 2 | early delivery | delivery system failure
clock | 2 | incorrect time and date | imprecision
clock | 2 | late delivery | delivery system failure
clog |  | clogged | device clogged
close | 2 | difficult to close | difficult to open or close; leak
clot | 2 | air introduced into vasculature | use of device issue
clot |  | Vein clotted off | use of device issue
clot | 2 | blood vessel blocked | human-device interface issue
coaptation | 2 | incomplete coaptation | human-device interface issue
coat, coating |  | inable to maintain adequate resistance | physical resistance
coat, coating |  | inable to maintain integrity over time | material integrity issue
coat, coating |  | inadequate coating durability | material integrity issue
coat, coating |  | inadequate coating lubricity | material integrity issue
coat, coating |  | inappropriate coating parameters (porosity, density, distrubution) | use of incorrect control settings
coat, coating |  | inappropriate interaction between coating and prosthesis | Improper Chemical Reaction
coat, coating | 2 | increased abrasion at coating edge | device abrasion from instrument or another object
coat, coating | 2 | insufficient coating | material integrity issue
coat, coating | 2 | over coating | material integrity issue
coat, coating |  | unacceptable particulate level | Particulates
coat, coating | 2 | unacceptable sensing | invalid sensing
code | 2 | failure to override code | programming issue
coefficient | 2 | coefficient error | parameter calculation error due to software problem
coil | 3 | coil absent | material integrity issue
coil | 3 | coil broken or damaged | break
coil | 3 | uncoiled | Uncoiled
cold |  | cold burn | temperature issue
collide, collision | 1 | device collided | Device or Device Component Damaged by another Device
collide, collision | 1 | device collision | Unintended Collision
collide, collision | 1 | unintended collision | Unintended Collision
collimator | 3 | collimator design problem | incorrect measurement; component or accessory incompatibility
color | 3 | color deviation | use of device issue; defective item
color | 3 | confused color scheme | inadequate user interface
color | 3 | display color out of specification | incorrect display
color | 3 | inappropriate color scheme | use of device issue; defective item; incorrect display
color | 3 | low color accuracy | poor quality image
comfort, comfortable, comfortably | 3 | loss of comfort | human-device interface issue
comfort, comfortable, comfortably | 3 | uncomfortable | human-device interface issue
command | 2 | fail to abort command | Communication or Transmission Issue
command | 2 | invalid command | Communication or Transmission Issue
command | 2 | loss of command | Communication or Transmission Issue
commission | 2 | commission time long | imprecision
communicate, communication | 2 | Command timeout | Communication or Transmission Issue
communicate, communication |  | communication error | Communication or Transmission Issue
communicate, communication | 2 | Communication is prevented | Communication or Transmission Issue
communicate, communication | 2 | Component open or short | difficult to open or close
communicate, communication | 2 | data corrupted | Communication or Transmission Issue
communicate, communication | 2 | fail to bidirectional communication | Communication or Transmission Issue
communicate, communication |  | fail to pass communication test | Communication or Transmission Issue
communicate, communication | 2 | fail to unidirectional communication | Communication or Transmission Issue
communicate, communication | 2 | intermittent communication | Communication or Transmission Issue
communicate, communication | 2 | Loss of communication | Communication or Transmission Issue
communicate, communication | 2 | Loss partial power | loss of power
communicate, communication | 2 | received data incorrectly | data issue
communicate, communication | 2 | received data late | data issue
comparator | 3
compatible, compatibility, incompatibility, incompatible | 3 | chemical incompatibility | chemical issue; Incompatibility Problem
compatible, compatibility, incompatibility, incompatible | 3 | device-device incompatibility | Device-Device Incompatibility
compatible, compatibility, incompatibility, incompatible | 3 | electrochemical incompatibility | chemical issue; Incompatibility Problem
compatible, compatibility, incompatibility, incompatible | 3 | environmental incompatibility | Incompatibility Problem
compatible, compatibility, incompatibility, incompatible | 3 | improper component selection | component or accessory incompatibility
compatible, compatibility, incompatibility, incompatible | 3 | improper fit between connection | fitting problem; component or accessory incompatibility
compatible, compatibility, incompatibility, incompatible | 3 | improper power configuration | power conditioning issue
compatible, compatibility, incompatibility, incompatible | 3 | material incompatibility | patient-device incompatibility; Device-Device Incompatibility
compatible, compatibility, incompatibility, incompatible | 3 | part mismatched | component or accessory incompatibility
compatible, compatibility, incompatibility, incompatible | 3 | power source incompatibility | Power Source Issue
compatible, compatibility, incompatibility, incompatible | 3 | unintended compatibility | Incompatibility Problem
compatible, compatibility, incompatibility, incompatible | 3 | unique compatibility violated | Incompatibility Problem
compensate, compensation | 3 | fail to compensate | improper device output
compensate, compensation | 3 | improper compensation | electronic property issue; radio signal problem
compensate, compensation |  | inappropriate leak compensation setting | use of incorrect control settings
compensate, compensation | 3 | lack of compensation | electronic property issue; radio signal problem
compliant, compliance, comply | 2 | fail to follow standard | nonstandard device or device component
compliant, compliance, comply |  | compliance out of specification | material rigid or stiff
compliant, compliance, comply |  | incorrect display range | incorrect display
compliant, compliance, comply |  | low compliance accuracy | incorrect measurement; imprecision
compliant, compliance, comply | 2 | part mismatched | component incompatible
component | 3 | component broken or damaged | break
component |  | component burned | Burn of Device or Device Component; burning odor
component |  | component cracked | crack
component |  | dry component | Moisture or Humidity Problem
component | 3 | component detached | detachment of device component; Material Disintegration
component | 3 | component dropped | component falling
component | 3 | component incompatible | fitting problem; connection issue; component incompatible
component | 3 | component issue | defective component
component | 3 | component misassembled | device misassembled during manufacturing or shipping
component | 3 | component missing | component missing
component |  | component not centered | failure to align
component | 3 | component off pad | detachment of device component
component | 3 | component open | defective component
component | 3 | component overheating | overheating of device or device component
component |  | component position incorrect | failure to align
component | 3 | component short | defective component
component | 3 | component temperature high | overheating of device or device component
component | 3 | component unattached | detachment of device component
component | 3 | component unavailable during implant | component missing
component | 3 | component worn | defective component
component |  | component discolored | Material Discolored
component |  | moisture in component | Moisture or Humidity Problem
component |  | fail to label manufacturer information | device markings issue
component |  | fail to withstand pressure | mechanical issue
component | 3 | components assembled incorrect | device misassembled during manufacturing or shipping
component | 3 | incorrect component shipped | incorrect device or component shipped
component |  | loose component | loose or intermittent connection
component |  | Removable component | detachment of device component
component |  | defective component | defective component
component | 3 | unreliable component | defective component
compress, compression, decompress, decompression |  | fail to compress | mechanical issue; data issue; application program issue
compress, compression, decompress, decompression |  | loss of precision | data issue
compress, compression, decompress, decompression | 3 | fail to decompress | mechanical issue; data issue; application program issue
compressor |  | fail to run | device inoperable
compressor | 3 | clogged | device clogged; device maintenance issue
computer |  | fail to boot up | reset issue
computer | 2 | unexpected shutdown | failure to shut off
concentrate, concentrator | 3
conduct, conductor, conductivity | 3 | conductor broken or damaged | break
conduct, conductor, conductivity | 3 | conductor fatigue fracture | fracture; mechanical issue
conduct, conductor, conductivity | 3 | fail to conduct | failure to conduct
conduct, conductor, conductivity | 3 | inappropriate color light | defective item
conduct, conductor, conductivity | 3 | incorrect incorporated | defective item
conduct, conductor, conductivity | 3 | leakage pathway between conductors | electrical issue; electrical shorting
conduct, conductor, conductivity | 3 | misconduct service activities | Maintenance Does not comply to Manufacturers Recommendations
conduct, conductor, conductivity | 3 | solder-coated conductor | Solder Joint Failure
conduct, conductor, conductivity | 3 | strained conductor | device misassembled during manufacturing or shipping
conduct, conductor, conductivity | 3 | twisting conductor | device misassembled during manufacturing or shipping
conduct, conductor, conductivity | 3 | unsecured conductor | migration of device or device component
configure, configuration, configurable | 3 | configuration changed | use of incorrect control settings; data issue
configure, configuration, configurable | 3 | configuration confused | inadequate user interface; data issue
configure, configuration, configurable | 3 | configuration data corrupted | use of incorrect control settings; data issue
configure, configuration, configurable | 3 | configuration transform slow | use of incorrect control settings; data issue
configure, configuration, configurable | 3 | fail to save configuration | use of device issue; data issue
configure, configuration, configurable | 3 | fail to setup configuration | device difficult to setup or prepare; data issue
configure, configuration, configurable | 3 | incorrect configuration | use of incorrect control settings; data issue
configure, configuration, configurable | 3 | loss of configuration indication | no display or display failure; data issue
configure, configuration, configurable | 3 | software firmware not compatible | application program version or upgrade problem
configure, configuration, configurable | 3 | software version incorrect | Operating System Version or Upgrade Problem
configure, configuration, configurable | 3 | update configuration failed | device maintenance issue; device handling issue; data issue
confirm, confirmation | 3 | fail to confirm | unauthorized access to computer system; computer system security issue
confirm, confirmation |  | incorrect confirmation | unauthorized access to computer system; computer system security issue
confirm, confirmation |  | incorrect confirmation status | data issue
confirm, confirmation | 3 | lack of confirmation | unauthorized access to computer system; computer system security issue
conjoin | 3 | fail to conjoin | use of device issue
connect, connection, connectivity | 2 | blocked connection | blocked connection
connect, connection, connectivity | 2 | connection broken or damaged | disconnection; break; device damaged prior to use
connect, connection, connectivity |  | connection cracked | Crack; fluid leak
connect, connection, connectivity |  | connection detached | disconnection
connect, connection, connectivity |  | connection distorted | Material Distortion
connect, connection, connectivity |  | Connection Issue | connection issue
connect, connection, connectivity | 2 | connection error | misconnection
connect, connection, connectivity |  | connection leakage | leak; air leak; fluid leak
connect, connection, connectivity |  | connection splitted | split
connect, connection, connectivity |  | fail to connect | connection issue
connect, connection, connectivity |  | fail to transmit signal | failure to transmit record
connect, connection, connectivity |  | unable to connect | connection issue
connect, connection, connectivity | 2 | inadequate connection | connection issue
connect, connection, connectivity | 2 | incompatible connection | fitting problem; component or accessory incompatibility
connect, connection, connectivity | 2 | incomplete connection | connection issue
connect, connection, connectivity | 2 | intermittent connection | loose or intermittent connection
connect, connection, connectivity | 2 | lack of both mechanically secure and soldered | Solder Joint Failure
connect, connection, connectivity | 2 | loose connection | loose or intermittent connection; fluid leak
connect, connection, connectivity | 2 | loss of connection | loose or intermittent connection; connection issue
connect, connection, connectivity | 2 | misconnection | misconnection
connect, connection, connectivity | 2 | open connection | failure to disconnect
connect, connection, connectivity | 2 | wrong connection | misconnection
connector |  | connector broken or damaged | break
connector |  | hole in connector | hole in material
connector |  | connector septum broken or damaged | break
connector |  | connector blocked | Blockage within Device or Device Component
connector |  | connector cracked | Crack; fluid leak
connector |  | connector inside scratched | fluid leak
connector |  | connector inside blocked | Blockage within Device or Device Component; fluid leak; Failure to Flush
connector |  | connector disconnected | disconnection; fluid leak
connector |  | connector detached | detachment of device component; connection issue
connector | 3 | loose connector | loose or intermittent connection; decoupling
connector |  | connector leakage | leak; air leak; fluid leak
connector |  | connector housing leakage | fluid leak
connector |  | connector adaptor leakage | fluid leak
connector |  | connector port leakage | fluid leak
connector |  | connector edge leakage | leak; air leak; fluid leak
connector |  | connector falling apart | fluid leak; Material Separation
connector | 3 | connector mismatched | fitting problem; component incompatible; decoupling
connector |  | connector rolled | material integrity issue
connector |  | connector misshaped | fitting problem; component incompatible
connector | 3 | difficult to connect | connection issue
connector |  | fail to detach | device handling issue
connector |  | fail to label text | device markings issue
connector |  | fail to label standard symbol | device markings issue
connector | 3 | ISO regulation violated | nonstandard device or device component
connector | 3 | single connector mechanism failed | connection issue
connector |  | non-standardized connector | nonstandard device or device component
connector | 3 | unsecured connector | migration of device or device component
connector |  | undedicated connector | misconnection
connector | 3 | wrong connector | misconnection
contact | 2 | damaged contact seal | failure to adhere or bond
contact | 2 | insufficient electrical contact | connection issue; intermittent continuity
contact | 2 | lack of electrical contact | connection issue; intermittent continuity
contact | 2 | skin allergic reaction | biocompatibility issue
contact | 2 | undesired contact | misconnection
container |  | container component mismatched | component incompatible
container | 3 | burst container | Burst
contaminate, contamination | 1 | biohazard contamination | device contamination with biological material
contaminate, contamination | 1 | contaminated | device contamination with biological material; device contaminated during manufacture or shipping
contaminate, contamination |  | foreign matter | particulates
contaminate, contamination | 1 | contamination during use | contamination during use
contaminate, contamination | 1 | contamination of device ingredient or reagent | device contamination with biological material
contaminate, contamination | 1 | device contamination with body fluid | contamination during use
contaminate, contamination |  | device contamination with expired gas | contamination during use
contaminate, contamination | 1 | device contamination with chemical or other material | contamination during use
contaminate, contamination | 1 | device contamination within biological material | device contamination with biological material
contaminate, contamination |  | fail to disclose contamination | instruction for use issue
contaminate, contamination | 1 | electrical performance deviated | improper device output
contaminate, contamination | 1 | microorganisms introduced | bacterial contamination of device
content | 3 | content out of specification | improper device output
content | 3 | inadequate contents | device markings issue
content | 3 | incorrect product information | device markings issue
continuity, continuousness | 1 | inconsistent contact | connection issue
continuity, continuousness | 1 | insufficient fatigue life of cables | shelf life exceeded
continuity, continuousness | 1 | intermittent continuity | Intermittent Continuity
continuity, continuousness | 1 | loss of continuity | failure to disconnect
continuity, continuousness | 1 | loss of electrical continuity | intermittent continuity
contrast |  | fail to adjust contrast | application program issue; use of device issue
contrast | 3 | contrast ratio out of specification | optical issue
control, controller | 2 | fail to activate control | device inoperable; device operates differently than expected
control, controller | 2 | fail to disable control | device remains activated; device operates differently than expected
control, controller | 2 | fail to re-enable control | failure to deploy; device operates differently than expected
control, controller |  | fail to control flow | infusion or flow issue
control, controller |  | control out of range | device operates differently than expected
control, controller |  | low control accuracy | imprecision; device operates differently than expected
control, controller | 2 | loss of control | device inoperable; device operates differently than expected
convert, converter | 3 | converter malfunction | device operates differently than expected
convert, converter |  | inappropriate converter designed | component or accessory incompatibility
convert, converter | 3 | data converted from incorrect channel | Communication or Transmission Issue; data issue
convert, converter | 3 | data converted from incorrect count | incorrect software programming calculations
convert, converter | 3 | fail to convert | deployment issue
cool, coolant, cooling | 3 | excessive cooling | Excessive Cooling
cool, coolant, cooling | 3 | fail to cool | Insufficient Cooling
cool, coolant, cooling | 3 | insufficient cooling | Insufficient Cooling
cool, coolant, cooling |  | insufficient insulation | overheating of device or device component
cool, coolant, cooling |  | inadequate coolant supply | insufficient cooling; temperature issue; environmental control or utility issue
cool, coolant, cooling | 3 | contraindicated coolant | temperature issue; environmental control or utility issue
cord |  | arcing occur | arcing
cord |  | cord broken or damaged | break
cord |  | cord torn | torn material
cord | 3 | downsized cord size | defective component
cord |  | exposed wire | break
cord | 3 | oversized cord size | defective component
core | 3
correct, correction | 3 | overcorrection | Overcorrection
correct, correction | 3 | undercorrection | calibration issue; Use of Device
corrode, corrosion | 1 | biological hazards such as toxicity or allergenicity | biocompatibility issue
corrode, corrosion |  | foreign matter | particulates; corrosion
corrode, corrosion | 1 | degradation of parts | Corrosion
corrode, corrosion | 1 | electrical performance deviated | Corrosion
corrode, corrosion |  | surrounding material corrosion | Corrosion
corrode, corrosion | 1 | moisture induced corrosion | Corrosion
corrupt, corruption | 1 | data corruption | computer software issue
corrupt, corruption | 1 | data overwritten | data issue; loss of data
corrupt, corruption | 1 | device memory corruption | computer software issue
corrupt, corruption | 2 | non recoverable corruption | computer software issue
corrupt, corruption | 2 | recoverable corruption | computer software issue
corrupt, corruption | 1 | software corruption | computer software issue
cost | 3 | high cost | Market Issue
cost | 3 | high manufactured cost | Market Issue
cost | 3 | high procedure cost | Market Issue
count, counter | 3 | fail to count | incorrect software programming calculations; failure to service
count, counter | 3 | incorrect count | incorrect software programming calculations; inadequate service
count, counter | 3 | fail to count operating hours | failure to service; inadequate service
counterbalance | 3 | counterbalance break | break
couple, coupler | 3 | fail to couple | Decoupling
couple, coupler | 3 | improper fit to common receptacle | fitting problem; component or accessory incompatibility
couple, coupler | 3 | inappropriate coupling | Decoupling
CPU | 2 | CPU failure | defective component
CPU | 2 | fail to enter safe state | Fail-Safe Mechanism Issue
CPU | 2 | fail to reset CPU | failure to reset
crc, code redundancy check | 2 | CRC error | protective measure issue
crc, code redundancy check | 2 | CRC mismatch | Incompatibility Problem
crc, code redundancy check | 2 | fail to generate periodic test | incorrect or inadequate test results
crc, code redundancy check | 2 | fail to recalculate CRC | incorrect software programming calculations
crc, code redundancy check | 2 | lack of CRC | protective measure issue
creepage | 2 | arcing | arcing
creepage | 2 | short creepage distance | circuit failure
crimp | 3 | inadequate crimp strength | mechanical issue
crimp | 3 | inadequate electrical continuity | intermittent continuity
crimp | 3 | loss of electrical continuity | intermittent continuity
crimp | 3 | part crimped | kinked
criteria | 3 | out of criteria | compliance issue
cuff |  | cuff broken or damaged | air leak; break
cuff |  | fail to inflate part evenly | Inflation Issue
cuff |  | device malfunction | device operates differently than expected
cuff | 3 | cuff degeneration | degraded
cuff |  | cuff leakage | leak; air leak; deflation issue
cuff |  | cuff ruptured | material rupture
cuff |  | unable to deflate | deflation issue
cuff |  | unable to inflate | Inflation Issue
current, DC, direct current, AC, ampere | 2 | current leakage | electrical issue
current, DC, direct current, AC, ampere | 2 | current out of range | electrical issue; therapeutic or diagnostic output failure; unexpected therapeutic results
current, DC, direct current, AC, ampere |  | output current out of range | electrical issue; therapeutic or diagnostic output failure; unexpected therapeutic results
current, DC, direct current, AC, ampere | 2 | gradient induced eddy current | electronic property issue
current, DC, direct current, AC, ampere | 2 | high current | electrical issue; therapeutic or diagnostic output failure; unexpected therapeutic results
current, DC, direct current, AC, ampere | 2 | high current drain | electrical issue
current, DC, direct current, AC, ampere |  | inappropriate current supply | Power Source Issue
current, DC, direct current, AC, ampere |  | fail to disclose input current | instruction for use issue
current, DC, direct current, AC, ampere | 2 | induced current | electronic property issue
current, DC, direct current, AC, ampere | 2 | induced RF current | Radiofrequency Interference (RFI)
current, DC, direct current, AC, ampere | 2 | over current detection | protective measure issue
cursor | 3 | fail to move cursor | computer software issue
cursor | 3 | invisible cursor | incorrect display
cusp | 3
cut | 3 | fail to cut | Failure to Cut
cutter | 3
cybersecurity, security | 1 | existing vulnerabilities | unauthorized access to computer system; application security issue
cybersecurity, security | 1 | hacked access | unauthorized access to computer system; application security issue
cybersecurity, security | 1 | malicious device attack | unauthorized access to computer system; application security issue
cybersecurity, security |  | system security compromised | unauthorized access to computer system; application security issue
cybersecurity, security | 1 | unintentional shutdown | False Device Output; device inoperable
cycle, cycling | 3 | component broken or damaged | break
cycle, cycling | 3 | cycle length variability | defective component
cycle, cycling | 3 | cycle sensitivity low | decreased sensitivity
cycle, cycling | 3 | device malfunction | failure to cycle
cycle, cycling | 3 | duty cycle low | device operates differently than expected
cycle, cycling |  | inappropriate auto-cycling | device operates differently than expected
cycle, cycling |  | incomplete cycle | failure to cycle
cycle, cycling | 3 | fail to cycle | failure to cycle
cylinder | 3
damage | 2 | component broken or damaged | material integrity issue; break
damage | 2 | device broken or damaged | device damaged prior to use; break
damage |  | device broken or damaged | material integrity issue; break
damage | 2 | device damaged by another device | Device or Device Component Damaged by another Device
damage | 2 | device damaged prior to use | device damaged prior to use
damage | 2 | seal damaged | failure to adhere or bond
damage | 2 | unintended soft tissue damage | mechanical issue
damper | 3 | loss of function | device inoperable
data | 2 | data back-up failed | Failure to Back-Up
data |  | fail to transform data | data issue; loss of data
data |  | fail to post process data | data issue
data |  | fail to reconstruct data | data issue
data | 2 | data corrupted | data issue; patient data issue
data | 2 | data insecurity | application security issue
data |  | data mix-up | data issue
data |  | data migrated | data issue
data | 2 | delayed data | data issue
data |  | duplicated data | data issue; patient data issue
data |  | conflicted data | data issue
data | 2 | expired data | device expiration issue; data issue
data | 2 | illegible data | data issue
data | 2 | inadequate data | data issue
data |  | inadequate data storage | data issue
data | 2 | incomplete data | data issue
data | 2 | incorrect data | data issue
data |  | incorrect data imported | data issue; patient data issue
data |  | Inaccurate data | data issue; imprecision
data |  | inconsistent data | patient data issue
data |  | data unavailable | data issue
data | 2 | missing data | loss of data
data |  | unrecoverable data loss | loss of data
data | 2 | low data exchange speed | data issue
data | 2 | unauthorized deletion | unauthorized access to computer system
data | 2 | unsecured data | application security issue
current, DC, direct current, AC, ampere | 3 | fail to run on DC | failure to run on AC/DC
deactivate, deactivation | 2 | fail to deactivate control | device remains activated
debris | 3 | difficult to remove | difficult to remove
debris |  | loose debris | metal shedding debris
decelerate, deceleration | 2 | speed deceleration failed | imprecision
decoder | 3
decommission | 2 | decommission time long | imprecision
decompose, decomposition | 3 | decomposed incorrectly | device maintenance issue
decontaminate, decontamination | 1 | decontamination failed | device contamination with biological material
decouple | 3
defect | 1 | defect introduced | defective component
defect | 1 | fail to identify defect | failure to sense
defect | 1 | lead to unacceptable risk | device unsafe to use in environment
defibrillate, defibrillator, defibrillation | 2 | device broken or damaged | device damaged prior to use; break
defibrillate, defibrillator, defibrillation | 2 | loss of function | no device output
deflate, deflation | 2 | inappropriate deflation | deflation issue
deform, deformation | 1 | dimensional deformation | defective component
deform, deformation | 1 | inadequate hardness | material integrity issue
deform, deformation | 1 | permanent deformation | defective component
degrade, degradation, biodegradation | 2 | degraded weld strength | mechanical issue; degraded
degrade, degradation, biodegradation |  | particulates | Particulates
degrade, degradation, biodegradation | 2 | insufficient efficacy and potency | degraded
degrade, degradation, biodegradation |  | degraded function | degraded
degrade, degradation, biodegradation | 2 | less mechanical strength | mechanical issue
degrade, degradation, biodegradation | 2 | permanent degradation | degraded
degree, deg | 3 | degree out of range | defective component
dehumidifier |  | condensation issue | moisture or humidity problem
dehumidifier | 3 | loss of function | moisture or humidity problem
dehydrator | 3 | loss of function | device inoperable
deionizer | 3 | loss of function | device inoperable
delaminate, delamination | 3 | fluid leak | fluid leak
delaminate, delamination | 3 | loss of electrical continuity | intermittent continuity
delaminate, delamination | 3 | loss of electrical isolation | electrical shorting
delaminate, delamination | 3 | reduced weld strength | delamination
delete, deletion |  | data deleted | loss of data
delete, deletion |  | incorrect deletion | application program issue; use of device issue
delete, deletion | 1 | unauthorized deletion | unauthorized access to computer system; loss of data
deliver, delivery | 2 | air delivery compromised | infusion or flow issue; decrease in pressure; no flow
deliver, delivery | 2 | deliver to wrong place | improper flow or infusion
deliver, delivery | 2 | delivery delayed | infusion or flow issue; improper device output
deliver, delivery | 2 | delivery interrupted | infusion or flow issue; improper device output
deliver, delivery |  | deliver inappropriately | excess flow or over-infusion; improper device output
deliver, delivery | 2 | delivery rate out of range | inaccurate flowrate
deliver, delivery | 2 | fail to cancel delivery | infusion or flow issue
deliver, delivery | 2 | fail to deliver | failure to deliver; no flow; improper device output; therapeutic or diagnostic output failure
deliver, delivery | 2 | fail to follow delivery schedule | infusion or flow issue
deliver, delivery | 2 | inaccurate delivery | improper flow or infusion; improper device output
deliver, delivery | 2 | low dose delivery | insufficient flow or underinfusion
deliver, delivery | 2 | no delivery | no flow; decrease in pressure
deliver, delivery | 2 | over delivery | excess flow or over-infusion
deliver, delivery | 2 | over dose delivery | excess flow or over-infusion
deliver, delivery | 2 | reverse delivery | improper flow or infusion
deliver, delivery | 2 | slow delivery | improper flow or infusion
deliver, delivery | 2 | under delivery | insufficient flow or underinfusion; decrease in pressure
deliverable | 3 | fail to plan deliverable | inaccurate delivery
deliverable | 3 | inadequate deliverable | inaccurate delivery
deliverable | 3 | lack of deliverable | failure to deliver
demagnetizer | 3
demodulator | 3
density | 3 | density out of specification | material integrity issue
deposit | 3 | deposit in device or device ingredient | sediment, precipitate or deposit in device or device ingredient
depressurize, depressurization | 2 | air bubbles | hole in material
depressurize, depressurization | 2 | leakage | leak
depressurize, depressurization | 2 | vacuum failed | suction issue
depth | 3 | less deep | defective component
depth |  | insertion depth less deep | defective component
depth | 3 | out of dimension | defective component
depth | 3 | too deep | defective component
derivative |  | lack of geographical origin | image display error
derivative |  | lack of traceability | data issue
design | 3 | design lack of details | Human Factors Issue
design | 3 | design lack of interface | inadequate user interface; loss of mechanical integrity
design | 3 | design lack of intuitively implemented | inadequate user interface
design | 3 | design out of date | Human Factors Issue
design | 3 | fail to decompose into subsystem | Human Factors Issue; loss of mechanical integrity
design | 3 | fail to design | Human Factors Issue
design | 3 | inadequate mechanical design | loss of mechanical integrity
design |  | inadequate electrical design | electrical issue
design |  | inadequate visibility design | human-device interface issue
design |  | inadequate audibility design | human-device interface issue
design |  | inadequate tactility design | human-device interface issue
design | 3 | lack of environment data | loss of data
design |  | lack of atraumatic design | protective measure issue
design | 3 | lack of test specification | failure of device to self-test
detach, detachable, detachment | 3 | detachment failed | detachment of device component
detach, detachable, detachment | 3 | fail to detach | failure to disconnect
detach, detachable, detachment | 3 | fail to detach within time frame | failure to disconnect
detach, detachable, detachment | 3 | unable onehand to detach | detachment of device component
detect, detection | 3 | absent undetected | failure to sense
detect, detection | 3 | inappropriate detection criteria | Device Sensing Issue
detect, detection | 3 | incorrect detection | invalid sensing
detect, detection | 3 | installation undetected | failure to sense
detect, detection | 3 | loss of detection | failure to sense
detect, detection | 3 | open lead detection failed | failure to sense
detect, detection | 3 | present undetected | failure to sense
detector | 3
deteriorate, deterioration |  | deteriorated function | improper device output
deteriorate, deterioration | 2 | safety deteriorated | device unsafe to use in environment
device |  | device alarmed | device alarm system issue
device |  | device blocked | Blockage within Device or Device Component
device | 3 | device broken or damaged | device damaged prior to use; break; leak
device |  | device burned | Burn of Device or Device Component; burning odor
device |  | device contamination with chemical or other material | Foreign Material Present in Device
device |  | device crushed | collapse
device |  | device detached | detachment of device or device component
device |  | device disconnected | disconnection
device | 2 | device dysfunction | device operates differently than expected
device |  | device expired | device expiration issue
device | 3 | device explant | device handling issue
device | 3 | device falling damage | component falling; tip-over | device falling
device |  | device frozen | temperature issue
device |  | device failsafe failed | protective measure issue
device | 3 | device handling inappropriately | device handling issue
device | 3 | device heating | overheating of device or device component
device | 3 | device ingredient or reagent inappropriately | sediment, precipitate or deposit in device or device ingredient
device | 3 | device inoperable; device stopped; device out of work | application program issue; device inoperable; computer software issue
device |  | device internal rattling | audible noise
device | 3 | device internal short | electrical shorting; spark
device |  | device leakage | air leak; gas leak; leak; fluid leak
device | 2 | device malfunction | device operates differently than expected; imprecision; Product Quality Issue; unexpected therapeutic results
device | 3 | device malposition | Malposition of Device
device | 1 | device microbial contamination | device contamination with biological material
device | 2 | device migrated | migration of device or device component
device | 3 | device misoriented | failure to align
device |  | device paused | application program issue; device inoperable; computer software issue
device | 3 | device reset incorrectly | reset issue
device |  | long response time | device operates differently than expected
device | 3 | device rotated | migration of device or device component
device |  | device shut down | device operates differently than expected; failure to power-up
device | 3 | device slipped | slippage of device or device component
device |  | device stops intermittently | device stops intermittently
device | 3 | device tipped over | tip-over
device | 3 | device unsafe to use in environment | device unsafe to use in environment
device |  | device works intermittently | device stops intermittently
device | 3 | early device replacement | material integrity issue
device | 3 | fail to shut-down device | failure to shut off
device |  | fail to turn on device | failure to power-up
device | 3 | incorrect device shipped | incorrect device or component shipped
device | 3 | occlusion within device | occlusion within device
diagnose, misdiagnose, misdiagnosing, diagnosis, misdiagnosis, diagnostic | 2 | diagnose incorrectly | therapeutic or diagnostic output failure
diagnose, misdiagnose, misdiagnosing, diagnosis, misdiagnosis, diagnostic |  | inconsistent diagnosis | therapeutic or diagnostic output failure; application program issue
diagnose, misdiagnose, misdiagnosing, diagnosis, misdiagnosis, diagnostic | 2 | incorrect diagnostic data | data issue
diagnose, misdiagnose, misdiagnosing, diagnosis, misdiagnosis, diagnostic | 2 | misinterpreted diagnostic data | data issue
dialysate | 3
dialyzer | 3 | loss of function | device inoperable
diameter, OD | 3 | diameter out of specification | defective component; loose or intermittent connection
diaphragm | 3 | difficult to replace | difficult to remove; device maintenance issue
dielectric | 3 | inadequate dielectric standoff | electrical shorting
dielectric |  | dielectric breakdown | electrical shorting
dielectric | 3 | reduced dielectric strength | electrical shorting
dilator |  | insufficient rigidity | material rigid or stiff
dilator |  | size out of range | size incorrect for patient
dilator |  | tip rollback | tip-over
dimension, longitudinal, diagonal | 2 | out of dimension | defective component; use of device issue
dimension, longitudinal, diagonal | 2 | oversized dimension | defective component; use of device issue
diode | 3
disable | 2 | disable failed | device remains activated
disassemble, disassembly | 3 | fail to disassemble device | disassembly
disc | 3 | incomplete to open disc | difficult to open or close
discard | 2 | fail to discard data | data issue
discharge | 2 | carrier fails to retaining set during activation | device difficult to maintain
discharge | 2 | fail to discharge | failure to discharge
discharge | 2 | insufficient discharge time | delayed charge time
disconnect, disconnection | 2 | accidental disconnection | disconnection; use of device issue; device handling issue
disconnect, disconnection | 2 | disconnected communication | disconnection; Communication or Transmission Issue
disconnect, disconnection | 2 | fail to detach | failure to disconnect
disconnect, disconnection | 2 | part disconnected | disconnection; detachment of device component; material separation; use of device issue
disconnect, disconnection | 2 | wrong disconnection | misconnection
disengage, disengagement | 1 | fail to disengage | connection issue
disengage, disengagement | 1 | disengagement slow | connection issue
disinfect, disinfection, disinfectant | 2 | disinfection inefficacy | failure to disinfect
disinfect, disinfection, disinfectant | 2 | improper disinfection | failure to disinfect
disinfect, disinfection, disinfectant |  | inable to reduce viable microorganism | bacterial contamination of device
disinfect, disinfection, disinfectant | 2 | inadequate disinfection | failure to disinfect
disinfect, disinfection, disinfectant | 2 | surface adverse impacted by cleaning agent | device cleaning issue
disinfect, disinfection, disinfectant | 2 | unqualified disinfecting agent | failure to disinfect
disk | 3
dislocate, dislocation | 3 | component dislocated | dislodged or dislocated
dislodge, dislodgement | 1 | dislodge at the distal end | dislodged or dislocated
dislodge, dislodgement | 1 | dislodged lead from device | dislodged or dislocated
dislodge, dislodgement | 1 | lead dislodgement | dislodged or dislocated
dispense, dispenser | 3 | dispense at incorrect location | positioning issue; improper flow or infusion
dispense, dispenser | 3 | dispense incorrect volume | inaccurate dispensing
dispense, dispenser | 3 | inaccurate dispensing | inaccurate dispensing
displace, displacement | 2 | displacement distance over limit | dislodged or dislocated; positioning issue
displace, displacement | 2 | displacement out of specification | dislodged or dislocated; positioning issue
display | 3 | brightness out of specification | incorrect display
display | 3 | contrast ratio out of specification | incorrect display
display |  | colored stripes | erratic display
display | 3 | display color out of specification | incorrect display
display |  | fail to consider colorblind user | incorrect display
display | 3 | display difficult to read | display misread
display |  | dark patches | erratic display
display |  | display value out of range | incorrect display
display | 3 | display distance out of specification | incorrect display
display | 3 | erratic of intermittent display | erratic display
display |  | error message | device displays error message
display | 3 | fail to display default setting | no display or display failure
display | 3 | fail to display message | no display or display failure
display |  | fail to display volume | no display or display failure
display |  | fail to display in the same coordinate system | incorrect display
display | 3 | fail to update display | device displays error message
display | 3 | inadequate measure unit | incorrect measurement
display | 3 | inadequate message display | incorrect display
display | 3 | inappropriate display size | incorrect display
display | 3 | inconsistent display information | incorrect display
display | 3 | incorrect display value | incorrect display
display | 3 | incorrect display format | incorrect display
display |  | incorrect display range | incorrect display
display | 3 | incorrect message display | incorrect display; no known device problem
display | 3 | long response time | imprecision
display | 3 | loss of connection | no display or display failure
display | 3 | low display speed | erratic display
display | 3 | low resolution | Image Resolution Poor
display | 3 | no display | no display or display failure
display | 3 | unable to view display | no display or display failure
display | 3 | view angle out of specification | incorrect display
disposable, disposal |  | Inappropriate disposal of device | environmental control or utility issue; device handling issue
disposable, disposal |  | Inappropriate disposal of contaminated sets | environmental control or utility issue; device handling issue
disposable, disposal | 3 | single use violated | use of device issue; device disinfection or sterilization issue
dissection | 3 | Major hemorrhage of hepatic vessel | use of device issue
distance |  | CL distance out of range | environmental control or utility issue
distance | 3 | distance out of range | positioning issue; incorrect measurement
distort, distortion | 3 | distorted images | image display error
distract, distraction | 2 | fail to follow safe practices | improper or incorrect procedure or method
distract, distraction | 2 | improper or incorrect use of device | use of device issue; improper or incorrect procedure or method
divider | 3
dock | 3 | lack of docking location | positioning issue
dock | 3 | unable onehand to dock | positioning issue
document, documentation | 3 | approved inappropriately | device markings issue
document, documentation | 3 | changed inappropriately | device markings issue
document, documentation | 3 | inadequate information | no information
document, documentation | 3 | insufficient documentation | device markings issue
document, documentation | 3 | issued inappropriately | device markings issue
document, documentation | 3 | lack of document | device markings issue
document, documentation | 3 | lack of document control procedure | device markings issue
document, documentation | 3 | reviewed inappropriately | device markings issue; data issue
dome | 3 | dome crushed | collapse
dose, dosage | 1 | dose calculation error | dose calculation error due to software problem
dose, dosage | 1 | fail to clear dose parameters | dose calculation error due to software problem; use of incorrect control settings
dose, dosage | 1 | fail to generate alert | device alarm system issue; failure to deploy
dose, dosage | 1 | fail to program dosage | dose calculation error due to software problem
dose, dosage | 1 | incorrect dose | dose calculation error due to software problem
dose, dosage |  | incorrect dose displayed | dose calculation error due to software problem
dose, dosage |  | incorrect dose plan | dose calculation error due to software problem
dose, dosage | 1 | over dose | dose calculation error due to software problem
dose, dosage | 1 | under dose | dose calculation error due to software problem
dosimeter, dosimetry |  | fail to response to radiation | imprecision
doubler | 3
download | 3 | download failed | problem with software installation
download |  | download stuck | problem with software installation
download | 3 | unable to download | problem with software installation
drape | 3
driveline | 3 | driveline abrasion | device abrasion from instrument or another object
driveline | 3 | fail to withstand kinking | Kinked
driveline | 3 | loss of airflow | no flow
driveline | 3 | material flexibility compromised | material rigid or stiff
driveline | 3 | wrong connection | misconnection
driver, driver system | 2 | adverse impacted by orientation | improper device output
driver, driver system | 2 | driver heavy | defective component
driver, driver system | 2 | driver less durable | material integrity issue
driver, driver system | 2 | driver noise | audible noise
driver, driver system | 2 | driver poor usability | inadequate user interface
driver, driver system | 2 | driver-device incompatibility | component or accessory incompatibility
driver, driver system | 2 | dust ingress | leak
driver, driver system | 2 | fail to drive artificial heart | failure to deploy
driver, driver system | 2 | fail to drive full eject | unintended ejection; expulsion
driver, driver system | 2 | fail to drive partial fill | filling problem
driver, driver system | 2 | fail to produce flow waveform | electrical issue; failure to spike
driver, driver system | 2 | fail to produce pressure | no pressure
driver, driver system | 2 | inadequate cardiac output | improper device output
driver, driver system | 2 | incapable to monitor battery temperature | battery issue; temperature issue
driver, driver system | 2 | incapable to monitor temperature | temperature issue
driver, driver system | 2 | incapable to support clinically stable patient | use of device issue
driver, driver system | 2 | incapable to support transient condition | use of device issue
driver, driver system | 2 | long start-up time | imprecision
driver, driver system | 2 | loss of power | loss of power
driver, driver system | 2 | low measurement accuracy | incorrect measurement
driver, driver system | 2 | noise over specification | audible noise
driver, driver system | 2 | switch to AC power failed | electrical power problem
driver, driver system | 2 | switch to automobile power failed | electrical power problem
drop, bump | 3 | device broken or damaged | device damaged prior to use; break
drop, bump | 3 | device malfunction | device operates differently than expected
drop, bump | 3 | latent failure | device damaged prior to use
drop, bump | 3 | loss of stability | unstable
drop, bump | 3 | unaware the dropped device is not functioning properly | component falling
drop, bump | 3 | unintended soft tissue damage | mechanical issue
drug | 1 | adminster drug by unauthorized person | medication error; unauthorized access to computer system
drug | 1 | drug availability below specification | medication error
drug | 1 | drug concentration not in safe range | medication error; output problem
drug |  | Drug with altered activity | medication error
drug | 1 | fail to configure drug information | medication error
drug |  | loss of medication | fluid leak
drug | 1 | inappropriate safe limit | medication error
drug | 1 | incorrect concentration | medication error; output problem
drug | 1 | incorrect drug name | medication error
drug | 1 | incorrect drugs | medication error
drug | 1 | mixed drugs | medication error
drug |  | hazardous drug splashed | fluid leak; device handling issue; contamination during use
drug |  | hazardous drug leakage | fluid leak; contamination during use
drug |  | drug leakage | fluid leak
drug |  | drug aerolization | fluid leak
drug |  | drug injected | fluid leak
drug | 1 | usage out of drug safe limit | medication error
drug | 1 | user misuses drug library | use of device issue; medication error
dry | 3 | fail to dry | environmental control or utility issue
duct | 3
duplexer | 3
durability, duration, end of life, life time |  | inappropriate maximum duration | use of incorrect control settings
durability, duration, end of life, life time |  | inadequte reliability | shelf life exceeded; improper device output
durability, duration, end of life, life time |  | excessive use of device | shelf life exceeded; improper device output
durability, duration, end of life, life time | 3 | reduced service life | shelf life exceeded; improper device output
dust | 3 | dust ingress | leak
ECG, EKG | 3 | failure to perform | deployment issue
ECG, EKG | 3 | unable to obtain ECG | failure to read input signal; artifact
edge | 3 | loss of electrical isolation | electrical shorting
edge | 3 | out of boundary | defective item
edge | 3 | user bleed due to sharp edge | human-device interface issue
egress | 3 | egress failed | device clogged
eject, ejection | 2 | fail to eject | unintended ejection; expulsion
eject, ejection | 2 | unintended ejection | unintended ejection; expulsion
elastic, elasticity | 3 | elastic tearing | torn material
elastic, elasticity | 3 | Loss of elasticity | material rigid or stiff
elastomer | 3 | elastomer material delaminates | delamination
electrical | 2 | electrical short | electrical shorting; spark
electrical |  | loss of electrical connection | intermittent continuity
electrical | 2 | overheating | overheating of device or device component
electrical | 2 | power outage | power conditioning issue
electrical | 2 | shock | Inappropriate Shock
electrode | 2 | apply to broken skin | use of device issue
electrode | 2 | electrode degeneration | degraded
electrode | 2 | electrode fractured | fracture; mechanical issue
electrode | 2 | electrode heating | overheating of device or device component
electrode | 2 | electrode migrated | migration of device or device component
electrode | 2 | electrode misplaced | device handling issue
electrode | 2 | electrode out of condition | defective item
electrode | 2 | electrode washed with belt | device cleaning issue
electrode | 2 | expired electrode | device expiration issue
electrode | 2 | group electrodes wrong | component incompatible
electrode | 2 | improper electrode size | defective component
electrode | 2 | inappropriate electrode application | use of device issue
electrode | 2 | inappropriate electrode placement | use of device issue
electrode | 2 | incompatible electrode | component incompatible
electrode | 2 | incorrect electrode placement | use of device issue
electrode | 2 | incorrect lead wire polarity | device misassembled during manufacturing or shipping
electrode | 2 | less sticky | detachment of device component
electrode | 2 | loss of conduction | failure to conduct
electrode | 2 | stored inappropriately | device handling issue
electrolyte |  | Electrochemical Effect | improper chemical reaction
electrolyte | 3 | electrolyte leakage | electrical issue; chemical issue; improper chemical reaction
emergency | 2 | emergency intervention required | use of device issue
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | chemical composition out of specification | chemical issue
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | data corrupted | Electromagnetic Interference (EMI)
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | device broken or damaged | device damaged prior to use; break
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | device malfunction | device operates differently than expected
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | device movement | unintended movement
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | device overheating | overheating of device or device component
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | device vibration | vibration
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | disorder pacing | failure to pace or properly pace
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | force or torque on the device | Electromagnetic Interference (EMI)
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic |  | fail to advice EMI | instruction for use issue
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | hardware broken or damaged | break
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | lead dislodgement | dislodged or dislocated
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | sensor rate maximized | incorrect measurement
EMI, electromagnetic, electromagnet, electrical-magnetic, electrical magnetic, electro-magnetic | 2 | suspect EMI | Electromagnetic Interference (EMI)
emit, emitter | 3 | device emits odor | Device Emits Odor; smoking
EN, European Union |  | EN regulation violated | compliance issue
encapsulate, encapsulation | 3 | encapsulation mis-registered | reset issue
encapsulate, encapsulation | 3 | encapsulation separation | Material Separation
encapsulate, encapsulation | 3 | foreign material inside | Item Contaminated during Manufacturing or Shipping; Packaging Issue
encapsulate, encapsulation | 3 | hole in encapsulation | hole in material
encapsulate, encapsulation | 3 | insufficient encapsulation | device misassembled during manufacturing or shipping
encapsulate, encapsulation | 3 | too much encapsulation | device misassembled during manufacturing or shipping
enclosure | 3 | loose attachment | Detachment of Device or Device Component
enclosure | 3 | enclosure too thin | defective component
enclosure | 3 | integrity compromised | Material Disintegration
enclosure | 3 | lead to unsafe condition | device unsafe to use in environment
enclosure | 3 | missing part | component missing
enclosure | 3 | sharp edges or corners | Material Puncture
encoder
endoleak |  | endoleak between modular graft components | fluid leak
endoleak |  | endoleak from patient branch arteries | fluid leak
endoleak |  | endoleak from sealing zone | failure to adhere or bond
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | endoscope inter-pupil distance out of range | optical issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | fail to comply with standards or regulation | nonstandard device or device component
endoscope, videoscope, laparoscopic, hysteroscopy |  | abnormal image | image display error; incorrect display
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | failed to be disinfected | failure to disinfect
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | failed to be sterilized | Failure to Sterilize; Delivered As Unsterile Product
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | focus distance out of range | Misfocusing
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | horizontal view degree out of range | optical issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | incorrect configuration | use of incorrect control settings
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | input control failed | failure to read input signal
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | internal temperature out of range | temperature issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | loss of visual | human-device interface issue; no display or display failure
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | low resolution image | Image Resolution Poor
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | low rigidity body | material integrity issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | video channel failed | Communication or Transmission Issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | waterproof failed | leak
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | working distance out of range | optical issue
energize | 3 | fail to energize | failure to deliver energy
energize | 3 | stuck in energized position | unintended energization
energy | 3 | energy activation failed | failure to deliver energy
energy | 3 | energy deactivation failed | unintended energization
energy |  | undesired energy path | energy output to patient tissue incorrect
energy |  | undesired energy site | energy output to patient tissue incorrect
energy | 3 | fail to deliver energy | failure to deliver energy
energy | 3 | incorrect energy delivery state | failure to deliver energy
energy |  | stored energy | energy output to patient tissue incorrect; output energy incorrect
energy | 3 | loss of energy | failure to deliver energy
engage, engagement | 1 | engagement slow | connection issue
engage, engagement | 1 | fail to engage | disconnection
engage, engagement |  | incorrect engagement | misconnection
engage, engagement | 1 | fail to retent engagement | connection issue
engage, engagement | 1 | stuck in engagement position | sticking; blocked connection
engage, engagement | 1 | undesired engagement | misconnection
enter | 3 | fail to enter data | data issue
epoxy | 3 | epoxy bleed-out | Material Erosion
epoxy | 3 | epoxy bridge | electrical shorting
epoxy | 3 | epoxy delamination | delamination
epoxy | 3 | epoxy mislocated | dislodged or dislocated
epoxy | 3 | epoxy oxidized | environmental control or utility issue
epoxy | 3 | epoxy smear | Foreign Material Present in Device
epoxy | 3 | excess conduct epoxy | device misassembled during manufacturing or shipping
epoxy | 3 | loss of electrical isolation | electrical shorting
epoxy | 3 | mixed epoxy | device misassembled during manufacturing or shipping
epoxy | 3 | over epoxy to block parts | Blockage within Device or Device Component
equalizer | 3
ergonomic | 3 | adjustment difficult | device difficult to setup or prepare
ergonomic | 3 | users confused | human-device interface issue
escalate | 2 | fail to escalate fault condition | improper or incorrect procedure or method
escape | 3
establish | 3 | lack of establishment process | improper or incorrect procedure or method
ethernet | 3 | ethernet port mismatched | fitting problem; component incompatible
ethernet |  | fail to label symbol | device markings issue
ethernet | 3 | ethernet speed slow | improper device output
ethernet | 3 | loss of communication | Communication or Transmission Issue
evaluate, evaluation | 3 | fail to evaluate | use of device issue
evaluate, evaluation | 3 | fail to evaluate change | use of device issue
evaluate, evaluation | 3 | fail to evaluate risk | Fail-Safe Design Failure
evaluate, evaluation | 3 | inadequate evaluation | use of device issue
evaporator | 3
event | 3 | capture event improperly | capturing issue
event | 3 | event missed | deployment issue
event | 3 | lack of suspend mode event | improper or incorrect procedure or method
event |  | lack of events | Communication or Transmission Issue
event | 3 | unable to record event | data back-up problem
exceed | 3 | decceed | defective item
exceed | 3 | out of specification | defective component
exercise | 3 | exercise over activity limit | use of device issue
exercise | 3 | incorrect prescribed exercise | improper or incorrect procedure or method
exercise | 3 | untimely prescribed exercise | imprecision
exhale, exhalate, exhalated, exhalation | 2 | excessive exhalation resistance | physical resistance
exhale, exhalate, exhalated, exhalation | 3 | fail to control pressure in the circuit | pressure issue
exhale, exhalate, exhalated, exhalation | 3 | fail to cycle to exhalation | failure to cycle
exhale, exhalate, exhalated, exhalation | 3 | fail to cycle to inhalation | failure to cycle
exhale, exhalate, exhalated, exhalation | 3 | fail to enter exhalation phase | device operates differently than expected
exhale, exhalate, exhalated, exhalation | 3 | fail to remain exhalation phase | device operates differently than expected
exhale, exhalate, exhalated, exhalation | 3 | initiate engagement failed | deployment issue
exhale, exhalate, exhalated, exhalation | 3 | loss of pressure | no pressure
expander | 3
expire, expiration |  | expired device used | device expiration issue
expire, expiration | 1 | expired parts used | device expiration issue
expire, expiration | 1 | expired solution used | device expiration issue
explantate, explant, explants, explantation | 2 | premature explanation | device handling issue
expose, exposure | 2 | device malfunction | device operates differently than expected
expose, exposure | 2 | exposure to fluid | environmental particulates; biological environmental factor
expose, exposure | 2 | exposure to water | environmental particulates; biological environmental factor
expose, exposure |  | exposure to gas | environmental particulates; biological environmental factor
expose, exposure |  | exposure to environmental material | environmental particulates; biological environmental factor
expose, exposure | 2 | extension out of specification | mechanical issue
expose, exposure | 2 | insufficient extension during implant | mechanical issue
extender | 3
extract | 3 | fail to extract data | data issue
extrude, extrusion | 2 | impending extrusion | Extrusion
eye |  | eye punctured | Material Puncture
eyelet | 3
facility | 3
fan | 3 | loss of fan speed signal | failure to read input signal
fan | 3 | loud fan noise | audible noise
fan | 3 | low fan speed | improper device output; fire
fan |  | fan broken or damaged | break
fan | 3 | reduced air flow | insufficient flow or underinfusion
fasten, fastener | 3 | fasten broken | break
fasten, fastener | 3 | loose attachment | detachment of device component
fasten, fastener |  | fastener loose due to vibration | detachment of device component
fasten, fastener |  | screw without spring washer | detachment of device component
fasten, fastener |  | screw without thread-lock | detachment of device component
fasten, fastener | 3 | low fasten force | mechanical issue
fatigue | 2 | component breaks | break
fatigue | 2 | component fractured | fracture; mechanical issue
fatigue | 2 | device is not reliable | defective item
fatigue | 2 | loss of electrical continuity | intermittent continuity
fatigue | 2 | reduced life cycle | shelf life exceeded
feedback | 3 | fail to assess feedback | use of device issue
feedback | 3 | fail to document feedback | use of device issue
feedback | 3 | fail to evaluate feedback | use of device issue
feedback | 3 | fail to monitor feedback | device operates differently than expected
feedback | 3 | incorrected feedback | use of device issue
feedback | 3 | loss of feedback | loss of data
ferrule | 3
fiber | 3
fiberguide | 3 | fiberguide incompatibility | component or accessory incompatibility
fiberoptic | 2 | fiberoptic break | break
fiberoptic | 3 | fiberoptic break | break
fiberoptic | 2 | fiberoptic fragmentation | Material Fragmentation
fiberoptic | 2 | fiberoptic separation | Material Separation
fill | 3 | inability to fill medicine | filling problem
fill | 3 | over fill | Overfill
fill | 3 | short fill | Short Fill
fill |  | fill stopped | Short Fill
fill |  | unstable fill | filling problem
fill |  | unable to fill | filling problem
film | 3 | missing film | loss of data
filter, filtration | 2 | filter broken or damaged | fluid leak; leak; particulates; contamination during use; break
filter, filtration |  | filter clogged | device clogged
filter, filtration | 2 | filter leakage | leak; fluid leak
filter, filtration |  | filter line leakage | leak; fluid leak
filter, filtration |  | filter inside cracked | crack
filter, filtration |  | filter cracked | crack
filter, filtration |  | filter proximal side leakage | leak; fluid leak
filter, filtration |  | filter distal end leakage | leak; fluid leak
filter, filtration |  | filter disconnected | disconnection
filter, filtration |  | filter air vent leakage | leak; fluid leak
filter, filtration |  | filter disintegrated | particulates; Material Disintegration
filter, filtration | 2 | filter open | Filtration Issue
filter, filtration | 2 | filter plugged | Filtration Issue
filter, filtration |  | filter mismatched | fitting problem
filter, filtration | 2 | filter ruptured | Filtration Issue
filter, filtration | 2 | filter torn | torn material
filter, filtration |  | hole in filter | fluid leak; leak
filter, filtration |  | filter cut | fluid leak; leak
filter, filtration |  | difficult to replace | difficult to remove; device maintenance issue
filter, filtration |  | allow air passage | air leak
filter, filtration | 2 | inadequate filter | Inadequate Filtration Process
filter, filtration |  | inappropriate pore filtration rate | Filtration Issue
filter, filtration |  | bubble point out of specification | Filtration Issue
filter, filtration | 2 | lack of filtration process | Inadequate Filtration Process
filter, filtration | 2 | inadequate ultra filtration | Ultra-filtration
filter, filtration | 2 | lost of fire protection | Filtration Issue
filter, filtration | 2 | low filter effect | Filtration Issue
filter, filtration | 2 | missing filter | Filtration Issue
filter, filtration |  | single use violated | use of device issue
filter, filtration | 2 | noise passed | Inadequate Filtration Process
filter, filtration |  | wet spot in filter | fluid leak; leak
filter, filtration | 2 | pore size out of range | Filtration Issue
fire, flame, explosive, pop | 1 | circuit overheating | overheating of device or device component
fire, flame, explosive, pop | 2 | continuous firing | Continuous Firing
fire, flame, explosive, pop | 2 | fail to fire | Failure to Fire
fire, flame, explosive, pop |  | catch fire | fire
fire, flame, explosive, pop | 2 | misfire | Misfire
firmware, FW | 3 | inappropriate firmware upgrade | application program version or upgrade problem
firmware, FW | 3 | incorrect firmware version | application program version or upgrade problem
fit | 1 | inappropriate fit | fitting problem
fit | 1 | loss of contact | connection issue
fit | 1 | not grip part firmly leading to inconvenience | loose or intermittent connection
fix | 3 | attached loosely | detachment of device component
fix | 3 | difficult to be attached | detachment of device component
fix | 3 | difficult to be reattached | detachment of device component
fix | 3 | difficult to remove | difficult to remove
fixate, fixation |  | loss of fixation | detachment of device component
fixate, fixation | 3 | fail to be fixated | detachment of device component
flammability | 1 | inappropriate flammability level | Failure to Fire
flammability | 1 | UL94 regulation violated | compliance issue
flange | 3
flash | 2 | flash broken or damaged | break
flash | 2 | loss of data | loss of data
flash | 2 | loss of software application | application program issue
flasher | 3
flex, flexible, flexibility, pliable | 2 | flex abrasion | device abrasion from instrument or another object
flex, flexible, flexibility, pliable | 2 | flex broken or damaged | break
flex, flexible, flexibility, pliable | 2 | flex fatigue fracture | fracture; mechanical issue
flex, flexible, flexibility, pliable | 2 | flex torn | torn material
flex, flexible, flexibility, pliable | 2 | less flexible | material rigid or stiff; improper device output
flow |  | backflow | insufficient flow or underinfusion
flow | 3 | excess flow | excess flow or over-infusion
flow |  | Concurrent Flow | improper flow or infusion
flow |  | fail to flow check | failure of device to self-test; Output Issue
flow | 3 | fail to limit flow | insufficient flow or underinfusion
flow |  | fail to mark flow direction | device markings issue
flow |  | fail to auto-set flow | improper flow or infusion
flow |  | fail to disclose input flow | instruction for use issue
flow |  | fail to combine flow rates proportioned | improper flow or infusion
flow | 3 | free flow | free or unrestricted flow
flow |  | fast flow | improper flow or infusion
flow |  | flow out of range | use of incorrect control settings; improper flow or infusion
flow |  | fluctuated flow | improper flow or infusion
flow |  | inaccurate flow rate | inaccurate flowrate
flow |  | incorrect flow calculation | improper flow or infusion
flow |  | slow flow | improper flow or infusion
flow | 3 | improper flow | improper flow or infusion
flow | 3 | inadequate flow | insufficient flow or underinfusion
flow | 3 | insufficient flow | insufficient flow or underinfusion
flow | 3 | lack of extra flow | insufficient flow or underinfusion
flow |  | low flow accuracy | imprecision; improper flow or infusion
flow | 3 | no flow | no flow; insufficient flow or underinfusion
flow | 3 | obstructed flow | no flow; blocked connection
flow | 3 | over flow | excess flow or over-infusion
flow | 3 | unrestricted flow | free or unrestricted flow
fluid | 1 | circuit short | electrical shorting; spark
fluid | 1 | component broken or damaged | break
fluid |  | no fluid flow | no flow
fluid |  | sterile fluid compromised | contamination during use
fluid |  | fluid path occluded | occlusion within device
fluid | 1 | fluid increased | fluid leak
fluid | 1 | fluid ingress | fluid leak; electrical shorting
fluid |  | fluid unclamped | Free or Unrestricted Flow
fluid | 1 | fluid jet | fluid leak
fluid |  | open fluid path | fluid leak
fluid | 1 | fluid leak | fluid leak
flush, flushing | 3 | difficult to flush | Difficult to Flush
flush, flushing |  | long flushing times | Difficult to Flush
flux | 3 | flux residue | Particulates
focus | 3 | misfocusing | Misfocusing
foil | 3 | anode foil too thick | defective component
foil | 3 | anode foil too thin | defective component
foil | 3 | foil cracked | Crack
foil | 3 | inappropriate foil material | material integrity issue
fold | 3 | difficult to fold | Difficult to Fold or unfold
fold | 3 | fail to fold | Failure to Fold
footprint | 3 | footprint oversized | defective component
footrest | 3 | footrest break | break
footrest | 3 | footrest detachment | detachment of device component
force | 2 | abnormal force | mechanical issue
force | 2 | force out of range | mechanical issue
force | 2 | induced translational force | electronic property issue
force | 2 | low force capacity | mechanical issue
force | 2 | over force | mechanical issue
force | 2 | permanent deformation | mechanical issue
force | 2 | unintended motion | unintended system motion
FPGA | 3 | FPGA failure | defective component
fragment | 3 | device fragments remain in patient | Device or Device Fragments Location Unknown
frame | 3
freeze | 2 | unable to freeze | Insufficient Cooling
frequency, Freq | 3 | frequency out of range | electronic property issue; use of incorrect control settings | frequency out of specification
frequency, Freq |  | High Frequency Leakage | electronic property issue; electrical shorting
frequency, Freq | 3 | high frequency electrical noise | failure to read input signal
frequency, Freq | 3 | inaccurate frequency | electronic property issue
frequency, Freq | 3 | no oscillation | electronic property issue
friction | 2 | excessive friction | physical resistance
friction | 2 | loss of extension or retraction during implant | Retraction Problem
friction | 2 | part worn out | Naturally Worn
function |  | device functions compromised | device operates differently than expected
function |  | loss of function | device inoperable
fungus | 3 | fungus in device environment | Fungus in Device Environment
furnace | 3
fuse, fuse-holder | 2 | fail to use fuse as isolating method | improper or incorrect procedure or method
fuse, fuse-holder |  | fail to label rating | device markings issue
fuse, fuse-holder | 2 | fuse blown | Fail-Safe Mechanism Issue
fuse, fuse-holder | 2 | lack of fuse | no fail-safe mechanism
fuse, fuse-holder | 2 | overcurrent protection fails | protective measure issue
fuse, fuse-holder | 2 | short circuit protection fails | electrical shorting; spark
gain | 3 | gain out of specification | electronic property issue
gain | 3 | inadequate gain | electronic property issue
gap | 3 | inappropriate gap | misassembled
gas | 3 | gas leakage | Gas leak
gas | 3 | gas output problem | improper gas output
gas | 3 | gas scavenging | Gas leak
gas |  | flammable gas | fire
gas |  | fail to identify gas type | failure to sense
gas | 3 | poor gas exchange | Gas leak
gasket | 3
gauge, gage |  | gage broken or damaged | Failure to Calibrate; Incorrect or Inadequate Result
gauge, gage |  | gage malfunction | Incorrect or Inadequate Result; Product Quality Issue
gauge, gage | 2 | gage inoperable | Failure to Calibrate; Incorrect or Inadequate Result
gaze | 3 | gaze volume out of specification | use of device issue
gaze | 3 | inappropriate tracking volume | use of device issue
gear | 3
gearbox | 3 | excessive vibration | vibration
gearhead | 3 | bearing failed | defective component
gearhead | 3 | dowell pin failed | defective component
gearhead | 3 | fail to drive piston | failure to deploy
gearhead | 3 | gear failed | defective component
gearhead | 3 | grease | misassembled
gearhead | 3 | heavy duty retaining ring dislocated | dislodged or dislocated
gearhead | 3 | heavy duty retaining ring failed | disassembly
gearhead | 3 | Input Pinion failed | defective component
gearhead | 3 | output plate failed | defective component
gearhead | 3 | planet bearings failed | defective component
gearhead | 3 | planet carrier failed | defective component
gearhead | 3 | planet washer failed | defective component
gearhead | 3 | ring gear failed | defective component
gearhead | 3 | screw dislocated | dislodged or dislocated
gearhead | 3 | screw shear failed | mechanical issue
gearhead | 3 | screw sheared | mechanical issue
gearhead | 3 | shaft shear failed | mechanical issue
gearhead | 3 | spacer and wave washer dislocated | dislodged or dislocated
gearhead | 3 | spacer and wave washer failed | defective component
gearhead | 3 | stator bonding failed | failure to adhere or bond
gearhead | 3 | stator extension failed | mechanical issue
gearhead | 3 | track bearing failed | defective component
gearhead | 3 | track bearing post dislocated | dislodged or dislocated
gearhead | 3 | track bearing retainer ring failed | disassembly
gel | 3 | clean skin insufficiently | device cleaning issue
gel | 3 | gel contaminated | device contamination with biological material
gel | 3 | gel leak | Gel leak
gel | 3 | inadequate gel | material integrity issue
gel | 3 | loss of adhesive | failure to adhere or bond
gel | 3 | unsticking gel | failure to adhere or bond
generate, regenerate, generator | 3 | fail to generate power | failure to power-up
geometry | 3 | fail to optimize geometry | defective component
geometry | 3 | incorrect geometry | defective component
germicide | 3
glass | 3 | shattered glass | break
glue | 3 | loose attachment | detachment of device component
glue |  | glue flaked | Flaked
gowning |  | aseptic processing environment compromised | environmental control or utility issue
gowning |  | fail to reduce contamination | bacterial contamination of device
grade | 2 | incorrect grade of medicine | degraded; medication error
graft |  | seam-line broken | failure to adhere or bond
graph, graphic, graphical, chart | 3 | confused graph | image display error; data issue
graph, graphic, graphical, chart | 3 | illegible graph | image display error; data issue
graph, graphic, graphical, chart |  | fail to reset graph | image display error; data issue
graph, graphic, graphical, chart |  | fail to indicate site of fault graphically | no display or display failure
graph, graphic, graphical, chart | 3 | incorrect graph | image display error; data issue
graph, graphic, graphical, chart |  | incorrect chart | image display error; data issue
graph, graphic, graphical, chart |  | inadequate graphic display | image display error; incorrect display
graph, graphic, graphical, chart | 3 | incorrect graph format | image display error; data issue
grasp | 3 | grasp difficult | difficult to open or close
grasp | 3 | grasp force low | mechanical issue
grasp | 3 | undesired close grasp | difficult to open or close
grasp | 3 | undesired open grasp | difficult to open or close
gravity | 3 | falling object | component falling
grille | 3
grommet | 3
ground | 3 | component broken or damaged | Grounding Malfunction; break
ground | 3 | insufficient grounding | Grounding Malfunction
ground | 3 | loss of grounding | Grounding Malfunction
guard | 3 | inappropriate guard | Fail-Safe Mechanism Issue
guidance | 3 | inadequate guide | device markings issue
guidance | 3 | incorrect guide | device markings issue
guidance | 3 | missing guide | device markings issue
guide, surgiguide, guideline |  | inappropriate guide | failure to osseointegrate
guide, surgiguide, guideline |  | unstable surgiguide | unstable
guide, surgiguide, guideline |  | part broken or damaged | break
guide, surgiguide, guideline |  | unable to fit in patient mouth | component incompatible
guide, surgiguide, guideline | 3 | lack of guideline | device markings issue
guidewire |  | coating flaked | Flaked
guidewire |  | coil stretched | Stretched
guidewire |  | frature surface | Scratched Material; fracture
guidewire |  | inappropriate coating | defective component
guidewire |  | lack of safety wire | protective measure issue
guidewire |  | loose part | loose or intermittent connection
guidewire |  | low peak tensile force | mechanical issue
guidewire | 3 | radio detectability failed | failure to sense
guidewire |  | size out of range | size incorrect for patient
handle |  | fail to transfer force | device handling issue
handle |  | missing handle | device handling issue
handle |  | handle damaged | break
handle |  | fail to transfer torque | device handling issue
handpiece | 3 | handpiece break | break
handpiece | 3 | handpiece detachment | detachment of device component
handset | 3
haptic | 2 | haptic bent | Bent
haptic | 2 | haptic broken | break
haptic | 2 | haptic damaged in delivery system | Device or Device Component Damaged by another Device
haptic | 2 | haptic detached | detachment of device component
haptic | 2 | haptic dislocated | dislodged or dislocated
haptic | 2 | haptic fractured | fracture; mechanical issue
haptic | 2 | haptic kinked | Kinked
haptic | 2 | haptic missing | component missing
haptic | 2 | haptic stuck in delivery system | Blockage within Device or Device Component
haptic | 2 | haptic stuck to optic | sticking
hardware |  | hardware failsafe failed | protective measure issue
hardware |  | low hardware configuration | use of incorrect control settings
hardware | 3 | malfunction | device operates differently than expected; Product Quality Issue
hazard | 2 | data may not be stored | data issue
hazard | 1 | fail to analyze new hazard | use of device issue
hazard | 1 | fail to identify hazard | failure to sense
hazard | 1 | fail to identify non-user interface related hazard | Human Factors Issue
hazard | 2 | inadequate hazard information | no information
head | 3 | unintended head motion | Unintended Head Motion
header |  | incomplete header info (patient name, patient ID, study date, study time) | data issue
header | 3 | missing header | data issue
headphone | 3
headset | 3
heal | 3 | inappropriate healing process | improper or incorrect procedure or method
heal | 3 | missing removable rigid part | component missing
heart | 1 | complete heart block | Blockage within Device or Device Component
heart |  | heart condition deterioration | no known device problem
heat | 2 | chemical release | chemical spillage
heat | 2 | excessive heating | overheating of device or device component
heat | 2 | fail to heat | Insufficient Heating
heat | 2 | catch fire | Fire
heat | 2 | insufficient heating | Insufficient Heating
heat | 2 | MRI heating | overheating of device or device component
heat | 2 | overheating | overheating of device or device component
heat | 2 | superheat | overheating of device or device component
heat | 2 | thermal injury | temperature issue
heat | 2 | tissue damaged | temperature issue
heater |  | fail to interface with other device | device-device incompatibility
heater |  | low power output | power conditioning issue
heater | 3 | inappropriate operation | device handling issue
height | 3 | out of dimension | defective component
helix | 3 | excessive friction of helix shaft within header | physical resistance
helix | 3 | helix length too long | defective component
helix | 3 | helix length too short | defective component
helix | 3 | helix not fully retracted during implant | Retraction Problem
helix | 3 | helix separated from shaft | Failure to Separate
helix | 3 | inadequate helix weld strength | mechanical issue
helix | 3 | insufficient helix extension after lead is implanted | mechanical issue
helix | 3 | insufficient helix retraction | Retraction Problem
helix | 3 | loss of helix extension | mechanical issue
helix | 3 | loss of helix retraction | Retraction Problem
hemodynamic | 2 | hemodynamic collapse | improper flow or infusion
hemodynamic | 2 | hemodynamics change | improper flow or infusion
hemodynamic | 2 | hemodynamics compromised | improper flow or infusion
hemostasis | 2 | inadequate hemostasis | Free or Unrestricted Flow
hermeticity | 2 | hermetic weld does not meet leakage rate requirement | leak
hermeticity | 2 | loss of hermeticity | leak
hermeticity | 2 | misaligned parts | failure to align; leak
hinge | 3 | unsecured hinge | migration of device or device component
hinge | 3 | unsecured hinge | migration of device or device component
hold, holder | 3 | fail to hold status | device difficult to maintain
hole |  | out of dimension | defective component
hole |  | inappropriate hole location | hole in material
home | 3 | position repeatability out of specification | positioning issue
hose, flexible hose |  | bio-contamination | device contamination with biological material; delivered as unsterile product
hose, flexible hose |  | improper length | defective component
hose, flexible hose | 3 | hose mismatched | component incompatible
house, housing | 3 | housing cracked | Crack
house, housing | 3 | housing splitted | Split
house, housing |  | inappropriate hosing design | human-device interface issue
house, housing | 3 | insufficient space | inadequate storage
hub | 3 | inappropriate fit | fitting problem
hub |  | blood leakage | fluid leak
hub |  | low strength | mechanical issue
humid, humidity, humidifier, humidification, Passive humidification, active humidification | 3 | device malfunction | device operates differently than expected
humid, humidity, humidifier, humidification, Passive humidification, active humidification |  | device overheating | overheating of device or device component
humid, humidity, humidifier, humidification, Passive humidification, active humidification |  | inadequate humidification | Moisture or Humidity Problem
humid, humidity, humidifier, humidification, Passive humidification, active humidification |  | fail to interface with other device | device-device incompatibility
humid, humidity, humidifier, humidification, Passive humidification, active humidification |  | run out of water | improper or incorrect procedure or method
humid, humidity, humidifier, humidification, Passive humidification, active humidification | 3 | high humidity | Moisture or Humidity Problem
humid, humidity, humidifier, humidification, Passive humidification, active humidification |  | patient infected | use of device issue; device handling issue
humid, humidity, humidifier, humidification, Passive humidification, active humidification | 3 | humid out of range | Moisture or Humidity Problem
hydraulic, hydration | 3 | backflow | improper flow or infusion
hydraulic, hydration | 3 | cross flow | improper flow or infusion
hydraulic, hydration | 3 | fluid jet | fluid leak
hydraulic, hydration | 3 | hydaulic ram | Blockage within Device or Device Component
hydraulic, hydration | 3 | implosion | collapse
hydraulic, hydration | 3 | loss of pressure | no pressure
hydraulic, hydration |  | hydration drip | fluid leak
hydraulic, hydration | 3 | loss of vacuum | suction issue
hydraulic, hydration | 3 | miscalibrated relief device | Failure to Calibrate
hydraulic, hydration | 3 | mislocated relief device | dislodged or dislocated
hydraulic, hydration | 3 | pipe or vessel rupture | Material Rupture; use of device issue
hydrogel | 3 | clean skin insufficiently | device cleaning issue
hydrogel | 3 | inadequate hydrogel | material integrity issue
hydrogel | 3 | loss of adhesive | failure to adhere or bond
IC, integrated circuit, chip | 3
ICD | 2 | competitive ventricular pacing | failure to pace or properly pace
ICD | 2 | device rotated | migration of device or device component
ICD | 2 | induced lead voltage | electronic property issue
ICD | 2 | intermittent ventricular pause | failure to pace or properly pace
ICD | 2 | Twiddler’s Syndrome | failure to pace or properly pace
ICD | 2 | VT/VF induction | electronic property issue
ICU, intensive care unit | 3
identify, identification | 3 | fail to identify change | failure to sense; compliance issue
identify, identification | 3 | fail to identify gap | compliance issue
identify, identification | 3 | identification failed | use of device issue
identify, identification | 3 | inadequate identification | use of device issue
identify, identification | 3 | incorrect identification | use of device issue
IEC | 2 | fail to follow standard | compliance issue
IEC | 2 | IEC regulation violated | compliance issue
igniter | 3
illuminate, illumination | 3 | illumination frequency out of specification | optical issue
illuminate, illumination | 3 | illumination out of specification | optical issue
illuminate, illumination | 3 | illumination power out of specification | optical issue
illuminate, illumination | 3 | illumination transmit failed | optical issue
illuminate, illumination | 3 | inappropriate color light | defective item
illuminate, illumination | 3 | loss of illumination | optical issue
image | 3 | distorted MRI images | image display error
image | 3 | illumination out of specification | optical issue
image | 3 | image brightness out of specification | optical issue
image |  | dark image | image display error
image | 3 | image display artifact | image display error
image | 3 | image display error | image display error
image |  | image disordered | image display error
image |  | image frozen | image display error
image |  | image load slow | improper device output; patient data issue
image |  | image orientation marked incorrectly | image orientation incorrect
image | 3 | image orientation incorrect | image orientation Incorrect
image |  | image misinterpreted | use of device issue
image |  | unable to reversal image | image reversal
image |  | unable to burn image on CD | data back-up problem
image | 3 | image reversal | image reversal
image |  | image saved incorrectly | patient data issue
image |  | image flipped | image reversal
image |  | twiddler’s Syndrome | image display error
image |  | image left-right flipped | image reversal
image |  | image up-down flipped | image reversal
image |  | image with 0 kb size | data issue
image |  | image anterior-posterior flipped | image reversal
image |  | image skipped | data issue
image |  | image duplicated | data issue
image |  | image intermixed | data issue
image |  | incomplete image | image display error
image |  | image loss | loss of data
image |  | image deleted | loss of data
image |  | image turning color | image display error
image |  | image link to wrong patient | data issue; patient data issue
image |  | unable to pan image | application program issue
image |  | unable to rotate image | application program issue
image |  | unable to zoom image | application program issue
image |  | unable to scroll image | application program issue
image |  | unable to read image | application program issue; patient data issue
image |  | unable to delete image | application program issue; patient data issue
image |  | inappropriate stitching image | image display error
image |  | inappropriate image rendered | poor quality image
image |  | incorrect image annotation | image display error
image |  | incorrect image laterality position | image orientation Incorrect
image |  | incorrect image anterior position | image orientation Incorrect
image |  | incorrect image posterior position | image orientation Incorrect
image |  | incorrect pixel depth calculated | image display error
image |  | incorrect image name | patient data issue
image |  | incorrect image status | patient data issue
image |  | incorrect image amount | data issue
image |  | lost of annotation | image display error
image |  | incorrect image cut line | image orientation incorrect
image |  | display to different patient image | incorrect display; image display error; patient data issue
image |  | save image to different patient | patient data issue
image |  | display to different image | incorrect display; image display error
image |  | incorrect display value | image display error; data issue
image |  | incorrect patient image displayed | image display error; patient data issue
image |  | incorrect patient data | patient data issue
image |  | invalid image signature | patient data issue
image |  | inaccuracy distance measurement | incorrect measurement
image |  | fail to flag incorrect image | image display error
image |  | fail to adjust image contrast | application program issue; poor quality image
image |  | fail to adjust image | application program issue
image |  | fail to adjust image brightness | application program issue; poor quality image
image |  | fail to display calibration | application program issue; image display error
image |  | fail to display image | no display or display failure
image |  | fail to display unviewed image | image display error
image |  | fail to display label | image display error
image |  | fail to merge multiple images | application program issue
image |  | inconsistent value | patient data issue
image |  | inconsistent measurement value | patient data issue
image |  | fail to resize image | reset issue
image |  | fail to reconstruct image | patient data issue
image |  | fail to register image | patient data issue
image |  | fail to access image | failure to service; data issue; patient data issue
image |  | image load error | improper device output; patient data issue
image |  | fail to load image | application program issue; image display error; patient data issue
image |  | fail to transfer image | communication or transmission issue
image |  | fail to retrieve image | communication or transmission issue; loss of data; data issue
image |  | fail to store image | loss of data
image |  | fail to retain image styles | data issue | orientation, window/level, zoom/pan
image |  | incorrect image measurement | incorrect measurement
image |  | inaccuracy image measurement | incorrect measurement
image | 3 | low color accuracy | poor quality image
image |  | low image accuracy | poor quality image
image | 3 | low resolution image | Image Resolution Poor
image |  | missing image series | image display error; data issue; loss of data
image |  | missing historical study images | loss of data
image |  | missing clinical information | application program issue; image display error
image | 3 | no image | image display error
image |  | invalid image format | data issue
image |  | non-dicom format image | data issue
image | 3 | poor quality image | poor quality image; therapeutic or diagnostic output failure
image |  | washed-out grey image | poor quality image
image |  | image flipped | image reversal
image |  | image corrupted | patient data issue
image |  | wrong image selected | use of device issue
image |  | black image | image display error
image |  | undetected image | data issue
image |  | incorrect image view | image display error
image |  | image missed review | data issue
image |  | image formatted | data issue; incorrect display
image | 3 | unclear fluro image | Image Resolution Poor
imager | 3
immobilize, immobilization, immobility | 1 | immobilization compromised | unstable
immobilize, immobilization, immobility | 1 | weak immobilization | unstable
immunize, immunity, immunization | 2 | loss of function | device inoperable
immunize, immunity, immunization | 2 | permanent degradation | failure to disinfect; degraded
immunize, immunity, immunization | 2 | unwanted movement | unintended movement
impedance | 2 | high impedance | High Impedance
impedance | 2 | impedance fluctuated | Impedance Issue
impedance | 2 | impedance value drifted | Impedance Issue
impedance | 2 | incorrect impedance calculation | Impedance Issue
impedance | 2 | incorrected impedance value | Impedance Issue
impedance | 2 | low impedance | Low Impedance
implant |  | ageing | shelf life exceeded
implant |  | contamination | contamination during use
implant |  | difficult to remove | difficult to remove
implant |  | difficult to replace | difficult to remove
implant |  | electro-magnetic interference (EMI) | electromagnetic interference (EMI)
implant |  | impaired combination | material disintegration
implant |  | implant dislodged | dislodged or dislocated
implant | 2 | implant loosening not related to bone-ingrowth | osseointegration issue
implant | 2 | implant overload | mechanical issue
implant | 2 | implant rejection | biocompatibility issue
implant | 2 | implant removal | difficult to remove
implant | 2 | implant repositioning | positioning issue
implant | 2 | implant reprogramming | programming issue
implant |  | implant mismatched | component incompatible
implant | 2 | improper disposal of explanted material | device handling issue
implant | 2 | improper lead implanted | user used incorrect product for intended use
implant |  | improper implant position | malposition of sevice; positioning issue
implant |  | inadequate markings | device markings issue
implant |  | incorrect implant size | size incorrect for patient
implant |  | incorrect implant type | component incompatible
implant |  | leakage | fluid leak
implant |  | material degraded | degraded
implant |  | patient-device interaction | human-device interface issue
implant |  | radio detectability failed | failure to sense
implant |  | reduced service life | shelf life exceeded; improper device output
implant |  | size out of specification | size incorrect for patient
implant |  | sterility compromising | delivered as unsterile product; failure to sterilize
implement | 3 | fail to implement | use of device issue
implement | 3 | fail to implement change | use of device issue
implement | 3 | lack of implementation process | use of device issue
inactivation |  | fail to activate | difficult or delayed activation
incorporate, incorporation | 3 | fail to incorporate | use of device issue; detachment of device component; material disintegration
incorporate, incorporation |  | fail to interface with other device | component or accessory incompatibility
incorporate, incorporation |  | fail to interface with other component | component or accessory incompatibility
incorporate, incorporation |  | component detached | detachment of device component; material disintegration
incorporate, incorporation |  | incompatible accessory | accessory incompatible; material disintegration; Communication or Transmission Issue
incorporate, incorporation | 3 | invalidated incorporating | protective measure issue; material disintegration
incorporate, incorporation | 3 | lack of characteristics | Instruction for Use Issue
incorporate, incorporation | 3 | lack of configuration | Instruction for Use Issue
incorporate, incorporation | 3 | lack of hazardous situation | device unsafe to use in environment
incorporate, incorporation | 3 | lack of information flow | no information
incorporate, incorporation | 3 | lack of purpose | Instruction for Use Issue
incorporate, incorporation | 3 | lack of security specification | computer system security issue
incorporate, incorporation | 3 | lead to new risk | no fail-safe mechanism
increase, increasion | 3 | unable to increase | failure to deploy
indexer, index | 3 | index to incorrect patient | patient data issue
indicate, indication, indicator | 2 | confused indication | inadequate user interface
indicate, indication, indicator | 2 | fail to update indicator | incorrect display
indicate, indication, indicator | 2 | inappropriate indication | incorrect display
indicate, indication, indicator |  | inadeqaute indication | no display or display failure
indicate, indication, indicator | 2 | indicator stuck off | no display or display failure; Incorrect or Inadequate Result
indicate, indication, indicator | 2 | indicator stuck on | no display or display failure; Incorrect or Inadequate Result
indicate, indication, indicator |  | indistinct indication | incorrect display
indicate, indication, indicator | 2 | invisible indication | no display or display failure
indicate, indication, indicator | 2 | loss of alarm type indication | device alarm system issue
indicate, indication, indicator | 2 | lack of indication | no display or display failure | loss of indication
indicate, indication, indicator | 2 | loss of presence indication | no display or display failure
indicate, indication, indicator | 2 | loss of priority indication | no display or display failure
indicate, indication, indicator | 2 | loss of state indication | no display or display failure
indicate, indication, indicator | 2 | premature elective replacement indicator | no display or display failure
indicate, indication, indicator | 2 | premature end of life indicator | no display or display failure
indicate, indication, indicator | 2 | premature indicator activation | Premature Indicator Activation
induce, induction | 3 | inappropriate induction | electronic property issue
infect, infection, infectant | 2 | adverse impact on exposed surface | bacterial contamination of device
infect, infection, infectant |  | cross infection | bacterial contamination of device
infect, infection, infectant |  | Systemic Sepsis | bacterial contamination of device
infect, infection, infectant | 2 | infection around implant | bacterial contamination of device
infect, infection, infectant | 2 | microorganisms introduced | bacterial contamination of device
inflate, inflation | 2 | unable to inflate | Inflation Issue
information | 3 | illegible information | device markings issue
information | 3 | inaccurate information | imprecision
information | 3 | inadequate information | no information; data issue
information | 3 | insufficient information | no information
information | 3 | missing information | no information; inadequate user interface
information | 3 | unclear information | no information
infuse, infusion, infusate, infuser | 2 | excessive dose | incorrect or inadequate test results
infuse, infusion, infusate, infuser | 2 | excessive rate | Increased Pump Speed
infuse, infusion, infusate, infuser | 2 | exposure to biohazardous materials | environmental particulates; biological environmental factor
infuse, infusion, infusate, infuser | 2 | fail to configure infusion parameters | use of incorrect control settings
infuse, infusion, infusate, infuser | 2 | fail to infuse | Failure to Infuse
infuse, infusion, infusate, infuser | 2 | fail to initiate infusion | Failure to Infuse
infuse, infusion, infusate, infuser | 2 | fail to pause lower precedence infusion | infusion or flow issue
infuse, infusion, infusate, infuser | 2 | fail to resume lower precedence infusion | reset issue
infuse, infusion, infusate, infuser | 2 | fail to stop delivery | Free or Unrestricted Flow
infuse, infusion, infusate, infuser | 2 | improper infusion | improper flow or infusion
infuse, infusion, infusate, infuser | 2 | incorrect infusion data | data issue
infuse, infusion, infusate, infuser | 2 | infusate degradation | improper flow or infusion; degraded
infuse, infusion, infusate, infuser | 2 | infusion of particulates | Particulates
infuse, infusion, infusate, infuser | 2 | insufficient dose | incorrect or inadequate test results
infuse, infusion, infusate, infuser | 2 | insufficient rate | Decreased Pump Speed
infuse, infusion, infusate, infuser | 2 | intermittent infusion | Intermittent Infusion
infuse, infusion, infusate, infuser | 2 | over infusion | excess flow or over-infusion
infuse, infusion, infusate, infuser | 2 | single drug infusion violated | medication error; device disinfection or sterilization issue
infuse, infusion, infusate, infuser | 2 | temporary interruption of infusion | Intermittent Infusion
infuse, infusion, infusate, infuser | 2 | unable to automatic infusion | Failure to Infuse
infuse, infusion, infusate, infuser | 2 | unable to manual infusion | Failure to Infuse
infuse, infusion, infusate, infuser | 2 | under infusion | insufficient flow or underinfusion
infuse, infusion, infusate, infuser | 2 | unintended delay of infusion start | infusion or flow issue
infuse, infusion, infusate, infuser | 2 | unintended termination of infusion | insufficient flow or underinfusion
infuse, infusion, infusate, infuser |  | infuser leakage | fluid leak
infuse, infusion, infusate, infuser | 2 | wrong drug | medication error
infuse, infusion, infusate, infuser | 2 | wrong route | misassembled
ingress | 2 | dust ingress | leak; electrical shorting
ingress | 2 | inadvertent energy activation | unintended energization
ingress |  | incorrect ingress classification | device markings issue
ingress | 2 | short circuit | electrical shorting; spark
ingress | 2 | water proof failed | leak; electrical shorting
inhalate, inhalation, inhale, inhaled | ventilator | fail to cycle to exhalation | failure to cycle
inhalate, inhalation, inhale, inhaled |  | fail to cycle to inhalation | failure to cycle
inhalate, inhalation, inhale, inhaled |  | fail to remain inhalation phase | device operates differently than expected
inhalate, inhalation, inhale, inhaled |  | water inhaled in lung | no known device problem
Inlet, outlet | 3 | blocked | blocked connection
Inlet, outlet |  | diameter out of specification | defective component; loose or intermittent connection
Inlet, outlet | 3 | leakage | leak
input | 3 | invalid input | failure to read input signal
input | 3 | no input | failure to read input signal
input |  | lack of mouse input | human-device interface issue
input |  | lack of keyboard input | human-device interface issue
input |  | lack of touchscreen input | human-device interface issue
input | 3 | undesired multiple inputs | Failure to Select Signal
insert, insertion | 3 | air trapped | Entrapment of Device or Device Component
insert, insertion | 3 | inability to fully insert | Difficult to Insert
insert, insertion | 3 | incomplete fill | Short Fill
insert, insertion | 3 | insert force out of specification | Difficult to Insert
insert, insertion | 3 | insert restriction | Difficult to Insert
insert, insertion |  | Difficult to Insert | Difficult to Insert
insert, insertion | 3 | insertion difficulties | Difficult to Insert
inspect, inspection | 3 | fail to perform inspection | protective measure issue
inspect, inspection | 3 | inappropriate inspection | protective measure issue
install, installation | 3 | fail to install | installation-related problem
install, installation |  | fail to contain installation instruction | instruction for use issue
install, installation | 3 | install force out of specification | installation-related problem
install, installation | 3 | installation time long | installation-related problem
install, installation | 3 | unable onehand to install | installation-related problem
install, installation | 3 | single member installation failed | installation-related problem
instruct, instruction | 3 | confused instruction | Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
instruct, instruction |  | overly complicated instruction | Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
instruct, instruction |  | violation of instruction | improper or incorrect procedure or method; use of incorrect control settings
instruct, instruction |  | abbreviation of instruction | improper or incorrect procedure or method; use of incorrect control settings
instruct, instruction | 3 | inadequate instruction | Instruction for Use Issue; Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
instruct, instruction | 3 | inadequate instructions for healthcare professional | Inadequate instructions for healthcare professional
instruct, instruction | 3 | inadequate instructions for non-healthcare professional | Inadequate instructions for non-healthcare professional
instruct, instruction | 3 | lack of application specification | improper or incorrect procedure or method
insufflate, insufflation | 3 | insufflation failed | expulsion
insulate, insulation | 2 | compromised insulation | electrical shorting; overheating of device or device component
insulate, insulation | 2 | energy delivers to wrong place | output energy incorrect
insulate, insulation | 2 | fail to meet minimum insulation requirement | electrical shorting
insulate, insulation | 2 | hole in insulation | electrical shorting
insulate, insulation | 2 | insulating material compromised | electrical shorting; overheating of device or device component
insulate, insulation | 2 | insulation abrased | electrical shorting
insulate, insulation | 2 | insulation broken or damaged | electrical shorting
insulate, insulation | 2 | insulation degradation | electrical shorting
insulate, insulation | 2 | insulation detached | detachment of device component; electrical shorting
insulate, insulation | 2 | insulation discolored | electrical shorting
insulate, insulation | 2 | insulation failure | electrical shorting
insulate, insulation | 2 | lack of heat-resistant material | material integrity issue
insulate, insulation | 2 | loss of electrical isolation | electrical shorting
insulate, insulation | 2 | no insulation coverage | electrical shorting
insulate, insulation |  | lack of dual insulation | electrical shorting
insulate, insulation | 2 | none insulation | electrical shorting
insulate, insulation | 2 | short circuit | electrical shorting; spark
insulin | 1 | air bubbles | Entrapment of Device or Device Component
insulin | 1 | incorrect grade insulin | degraded; medication error
insulin | 1 | incorrect insulin amount setting | use of incorrect control settings
insulin | 1 | incorrect insulin calculation | dose calculation error due to software problem
insulin | 1 | insulin amount out of range | dose calculation error due to software problem; use of device issue
insulin | 1 | insulin degradation | degraded; medication error
insulin | 1 | insulin delivery too high | inaccurate delivery
insulin | 1 | insulin delivery too low | inaccurate delivery
insulin | 1 | lack of fast acting | deployment issue
insulin | 1 | lack of regular insulin | medication error
insulin | 1 | non-sterile insulin | Failure to Sterilize
integrate, integration, integrator, integrity, combine, combination | 3 | fail to integrate | Material Disintegration; component incompatible; component or accessory incompatibility
integrate, integration, integrator, integrity, combine, combination | 1 | Loss of elasticity | material rigid or stiff
integrate, integration, integrator, integrity, combine, combination | 1 | loss of set integrity | disassembly; detachment of device component
integrate, integration, integrator, integrity, combine, combination | 1 | mechanical broken or damaged | mechanical issue; break
integrate, integration, integrator, integrity, combine, combination | 1 | thermal damage | overheating of device or device component
intensity | 3 | inappropriate intensity step | use of device issue
interact, interaction | 3 | device malfunction | device operates differently than expected
interact, interaction |  | fail to send JSON messages | Communication or Transmission Issue
interact, interaction |  | fail to determine interaction order | Communication or Transmission Issue
interact, interaction |  | fail to determine event time | Communication or Transmission Issue
interact, interaction |  | lack of events | Communication or Transmission Issue
interact, interaction |  | lack of interaction requirement | data issue
interact, interaction | 3 | loss of communication | Communication or Transmission Issue
interact, interaction | 3 | loss of detection | failure to sense
intercom | 3
interface | 2 | confused interface | inadequate user interface
interface | 2 | fail to avoid incorrect connection terminal | misconnection
interface | 2 | fail to support accessory | accessory incompatible
interface |  | fail to display view | no display or display failure
interface |  | fail to detect low resolution | Image Resolution Poor
interface | 2 | inappropriate interface design | inadequate user interface
interface |  | inadequate user interface | inadequate user interface
interface | 2 | loss of connection | connection issue
interface | 2 | part mismatched | component incompatible
interface | 2 | program exit | Application Interface becomes nonfunctional or program exits abnormally
interface | 2 | unstable interface | unstable; inadequate user interface
interface | 2 | weak interface strength | Loss of or Failure to Bond
interfer, interference | 2 | device malfunction | device operates differently than expected
interfer, interference | 2 | interference with circuit | Electromagnetic Compatibility Issue
interfer, interference | 2 | interference with monitoring device | Electromagnetic Compatibility Issue
interfer, interference | 1 | interference with pacemaker | Electromagnetic Compatibility Issue
interfer, interference |  | interference with other device | mechanical issue
interfer, interference | 2 | mechanical interference | mechanical issue
interlock | 3 | fail to interlock | device unsafe to use in environment
internet | 3 | loss of communication | Communication or Transmission Issue
interposer | 3 | interposer solder reflow | improper or incorrect procedure or method
interpret, interpretation, misinterpret, misinterpreting | 3 | incorrect interpretation | Incorrect Interpretation of Rhythm; use of device issue
interrogate, interrogation | 3 | difficult to interrogate | Difficult to Interrogate
interrogate, interrogation | 3 | fail to interrogate | Failure to Interrogate
interrogate, interrogation | 3 | unable to interrogate | Failure to Interrogate
interrupter | 3
interval | 3 | inappropriate interval settings | use of incorrect control settings
interval | 3 | incorrect interval calculation | parameter calculation error due to software problem
introducer |  | corrosion | Corrosion
introducer |  | out of dimension | defective component
introducer |  | radio detectability failed | failure to sense
inverter | 3
IOL, intraocular len | 3
IPX, IPX0, IPX1, IPX2, IPX3, IPX4, IPX5, IPX6, IPX7, IPX8, IPX9k, IP22, IP24 | 2 | dust ingress | leak
IPX, IPX0, IPX1, IPX2, IPX3, IPX4, IPX5, IPX6, IPX7, IPX8, IPX9k, IP22, IP24 | 2 | water proof failed | leak
IPX, IPX0, IPX1, IPX2, IPX3, IPX4, IPX5, IPX6, IPX7, IPX8, IPX9k, IP22, IP24 | 2 | water resistant failed | leak
irrigate | 3 | inability to irrigate | Inability to Irrigate
ISO, international standard organization | 3 | ISO regulation violated | compliance issue
isolate, isolation | 1 | bond delamination | delamination; Loss of or Failure to Bond
isolate, isolation | 1 | electrical leakage | electrical issue
isolate, isolation | 1 | electrical short | electrical shorting; spark
isolate, isolation |  | signal interference | radiofrequency interference (RFI); ambient noise issue
isolate, isolation | 1 | improper isolation | electrical shorting
isolate, isolation | 1 | insulation broken or damaged | electrical shorting
isolate, isolation | 1 | loss of electrical isolation | electrical shorting
isolator | 3 | fail to prevent contaminant ingress | contamination during use
isolator |  | fail to reproduce interior bio-decontamination | contamination during use
issue | 3 | issued inappropriately | improper device output
ista, international safe transit association | 3 | ista testing failed | incorrect or inadequate test results
jack | 3
jaw | 3
joint |  | fail to withstand pressure | mechanical issue
joint | 2 | joint broken or damaged | break
joystick | 3
jumper | 3
junction |  | junction broken or damaged | Detachment of Device or Device Component; break
junction | 3
keyboard |  | unrespond to keyboard | human-device interface issue; application program issue
keyboard | 3 | lack of keyboard | inadequate user interface
keyer | 3
kinetic | 3
kit | 3 | loose kit | slippage of device or device component
knife | 3
knob | 3 | knob broken or damaged | break
knob |  | knob missing | component missing
knob |  | knob stuck | sticking; difficult to open or close
knob |  | loose knob | defective component
knob |  | knob spins freely | defective component
knot | 3
KVO, keep vein open | 3 | inappropriate KVO flow rate | inaccurate flowrate
label, labeling | 2 | illegible label | device markings issue
label, labeling |  | improper ink reaction | device contamination with biological material
label, labeling | 2 | inadequate labeling | device markings issue
label, labeling |  | incomplete labeling | device markings issue
label, labeling |  | inadequate warning | device markings issue
label, labeling |  | incorrect label ink color | device contamination with biological material
label, labeling | 2 | incorrect labeling | device markings issue; patient data issue
label, labeling |  | incorrect case labeling | device markings issue
label, labeling |  | inappropriate labeling orientation | device markings issue
label, labeling |  | label discolored | device markings issue
label, labeling |  | labeling broken or damaged | break; device markings issue
label, labeling |  | label smeared | device markings issue
label, labeling |  | labeled backward | device markings issue
label, labeling |  | labeling expired | device markings issue
label, labeling | 2 | labeling missing | device markings issue
label, labeling |  | labeling misaligned | device markings issue
label, labeling | 2 | lack of local language | Instruction for Use Issue; inadequate user interface
label, labeling |  | lack of CE mark | device markings issue
label, labeling |  | lack of NRTL mark | device markings issue
label, labeling |  | lack of limitations | device markings issue
label, labeling |  | lack of side effect labeling | device markings issue
label, labeling | 2 | mislabeled | device markings issue; size incorrect for patient
label, labeling |  | wrong print | device markings issue
label, labeling | 2 | off label use | use of device issue
laminate, lamination | 3
lamp | 3
language | 3 | inadequate localization configuration | inadequate user interface; human-device interface issue
language |  | inadequate language pack | inadequate user interface; human-device interface issue
language | 3 | inappropriate default language | inadequate user interface; human-device interface issue
language | 3 | lack of local language | inadequate user interface; human-device interface issue
language | 3 | language change inconveniently | inadequate user interface; human-device interface issue
endoscope, videoscope, laparoscopic, hysteroscopy | 3 | sterilization inefficacy | Failure to Sterilize
laser | 2 | laser output above specification | output above specifications
laser | 2 | laser pulse timing error | imprecision
laser | 2 | laser tip separation | Material Separation
laser | 2 | unable laser output | no device output
latch | 3 | bent latch | Bent
latch | 3 | fail to latch | difficult to open or close
latch | 3 | fail to remain latch state | device operates differently than expected
latch | 3 | inefficient latch tool design | Human Factors Issue
latch | 3 | jammed latch | Mechanical Jam
latch | 3 | latch broken or damaged | break
latch | 3 | latch force out of specification | mechanical issue
latch | 3 | latch out of control | deployment issue
latch | 3 | sticky latch | difficult to open or close
latch | 3 | worn latch | Naturally Worn
latency |  | slow internet connection | communication or transmission issue
latency | 1 | latency over specification | communication or transmission issue
latent | 1 | latent weld failure | device damaged prior to use
launch | 3 | inadequate localization plan | Instruction for Use Issue
launch | 3 | launch delayed | Market Issue
LCD, LCD panel | 3 | image color accuracy low | poor quality image
LCD, LCD panel |  | LCD broken or damaged | no display or display failure
LCD, LCD panel | 3 | incorrect display | incorrect display
LCD, LCD panel | 3 | Insufficient brightness | optical issue
LCD, LCD panel |  | low resolution | Image Resolution Poor
LCD, LCD panel | 3 | no display | no display or display failure
lead | 2 | compromised maneuverability | Difficult to Interrogate
lead | 2 | difficult to advance lead through introducer | difficult to advance
lead | 2 | difficult to install lead into device | installation-related problem
lead | 2 | electrocution from lead | Inappropriate Shock
lead | 2 | electrode sensing impedance out of range | Impedance Issue
lead | 2 | excessive lead electrodes heating | overheating of device or device component
lead | 2 | excessive lead insert force | Difficult to Insert
lead | 2 | excessive lead remove force | difficult to remove
lead | 2 | excessive particulate release | Particulates
lead | 2 | inadequate sterilization due to lead design | device disinfection or sterilization issue
lead | 2 | inappropriate lead labeling | device markings issue
lead | 2 | inappropriate lead size | defective component
lead | 2 | incorrect lead placement | use of device issue
lead | 2 | incorrect lead replacement | device maintenance issue
lead | 2 | insulation broken or damaged | electrical shorting
lead | 2 | insulation delaminated | electrical shorting; delamination
lead | 2 | insulation material heated | material integrity issue; electrical shorting
lead | 2 | lead abrasion | device abrasion from instrument or another object
lead | 2 | lead breakage | break
lead | 2 | lead broken or damaged | break
lead | 2 | lead burned | Burn of Device or Device Component; burning odor
lead | 2 | lead dislodgement | dislodged or dislocated
lead | 2 | lead fractured | fracture; mechanical issue
lead | 2 | lead malfunction | device operates differently than expected
lead | 2 | lead maturation period long | device difficult to setup or prepare
lead |  | device migrated | migration of device or device component
lead | 2 | lead oversensing | Over-sensing
lead | 2 | lead short | electrical shorting; spark
lead | 2 | lead sterilization not being effective | Failure to Sterilize; Delivered As Unsterile Product
lead | 2 | loose lead | Detachment of Device or Device Component
lead | 2 | loss of electrical connection | intermittent continuity
lead | 2 | loss of electrical isolation | electrical shorting
lead | 2 | loss of lead retention | Detachment of Device or Device Component
lead | 2 | missing lead | component missing
lead | 2 | MR-induced lead force | Electromagnetic Interference (EMI)
lead | 2 | MR-induced lead heating | Electromagnetic Interference (EMI)
lead | 2 | MR-induced lead torque | Electromagnetic Interference (EMI)
lead | 2 | MR-induced lead vibration | Electromagnetic Interference (EMI); vibration
lead | 2 | MR-induced lead voltage | Electromagnetic Interference (EMI)
lead | 2 | poor fluoroscopic visibility | human-device interface issue
lead | 2 | restriction during lead insertion | Difficult to Insert
lead | 2 | shock from lead | Inappropriate Shock
lead | 2 | stiff lead | material rigid or stiff
lead | 2 | stiff lead tip | material rigid or stiff
lead | 2 | unacceptable electrode polarization | electrical issue
leaflet | 2 | leaflet disruption | Material Puncture
leaflet | 2 | leaflet disruption due to suture abrasion | device abrasion from instrument or another object
leaflet | 2 | leaflet fractured | fracture; mechanical issue
leak, leakage |  | air leak | leak
leak, leakage |  | loud air leak | audible noise
leak, leakage | 1 | battery leakage | battery issue
leak, leakage | 1 | current leakage | electrical issue
leak, leakage |  | fail to pass leak test | failure of device to self-test; leak
leak, leakage | 1 | fluid jet | fluid leak
leak, leakage | 1 | fluid leak | fluid leak
leak, leakage |  | inadequate leak compensation | imprecision
leak, leakage |  | incorrect leak calculation | parameter calculation error due to software problem
leak, leakage |  | incorrect leak range | leak
leak, leakage |  | loss of medication | fluid leak
leak, leakage | 1 | leakage out of range | leak
leak, leakage |  | low leakage parameter accuracy | incorrect measurement; imprecision
leak, leakage | 1 | leaky capacitor | leak
leak, leakage | 1 | medication leakage | medication error
LED, light emitting diode | 3 | inappropriate LED location | positioning issue
LED, light emitting diode | 3 | LED stuck off | no display or display failure
LED, light emitting diode | 3 | LED stuck on | no display or display failure
LED, light emitting diode |  | fail to label symbol | device markings issue
LED, light emitting diode | 3 | loss of indication | no display or display failure
LED, light emitting diode | 3 | loss of state indication | no display or display failure
length | 3 | insufficient extended length | defective component
length | 3 | length out of range | defective component
length | 3 | too long | defective component
length | 3 | too short | defective component
lens, IOL | 2 | color variation | Optical Discoloration
lens, IOL | 2 | crease on lens | Foreign Material Present in Device
lens, IOL | 2 | difficulty loading into cartridge | device difficult to setup or prepare; failure to deploy
lens, IOL | 2 | lens aspherical | defective component
lens, IOL | 2 | lens bifocal | Misfocusing
lens, IOL | 2 | lens clipped | defective component
lens, IOL | 2 | lens cloudy | Material Opacification
lens, IOL | 2 | lens cracked | Crack
lens, IOL | 2 | lens damaged by cartridge | Device or Device Component Damaged by another Device
lens, IOL | 2 | lens damaged in delivery system | Device or Device Component Damaged by another Device
lens, IOL | 2 | lens defective | defective component
lens, IOL | 2 | lens deposits | sediment, precipitate or deposit in device or device ingredient
lens, IOL | 2 | lens discoloration | Optical Discoloration
lens, IOL | 2 | lens dislocated | dislodged or dislocated
lens, IOL | 2 | lens malposition | Malposition of Device
lens, IOL | 2 | lens mark | Optical Obstruction
lens, IOL | 2 | lens migrated | migration of device or device component
lens, IOL | 2 | lens misshaped | defective component
lens, IOL | 2 | lens multifocal | Misfocusing
lens, IOL | 2 | lens nicked | Cut in Material
lens, IOL | 2 | lens not round | defective component
lens, IOL | 2 | lens opacification | Material Opacification
lens, IOL | 2 | lens replacement | material integrity issue
lens, IOL | 2 | lens repositioning | Malposition of Device
lens, IOL | 2 | lens scratch | Scratched Material
lens, IOL | 2 | lens split | Split
lens, IOL | 2 | lens stuck in cartridge | sticking
lens, IOL | 2 | lens stuck in delivery system | Blockage within Device or Device Component
lens, IOL | 2 | lens stuck to itself | sticking
lens, IOL | 2 | lens torn | torn material
lens, IOL | 2 | lens warped | Stretched
lense | 3
life, lifetime, life time | 3 | lack of life time | shelf life exceeded
life, lifetime, life time | 3 | material life short | shelf life exceeded
life, lifetime, life time | 3 | product life short | shelf life exceeded
life, lifetime, life time | 3 | reduced service life | shelf life exceeded; improper device output
light, sunlight | 3 | inadequate lighting | Inadequate Lighting
light, sunlight | 3 | light interference | Light Interference
limit | 3 | out of limits | defective component
limiter | 3
liner | 3 | air flow blocked | Blockage within Device or Device Component
liner | 3 | difficult to remove | difficult to remove
liner | 3 | material biocompatibility compromised | biocompatibility issue
liner | 3 | missing moisture wicking material | component missing
liner | 3 | part washing damaged | device cleaning issue
liner | 3 | unable for cleaning | device cleaning issue
liquid |  | liquid
liquid | 1 | liquid ingress | fluid leak
lithium | 3 | lithium cluster buildup | sediment, precipitate or deposit in device or device ingredient
lithium | 3 | lithium cluster deposited | sediment, precipitate or deposit in device or device ingredient
load | 2 | device malfunction | device operates differently than expected
load | 2 | early load | Premature Deployment
load | 2 | improper loading | Premature Deployment
load | 2 | load out of specification | Premature Deployment
load | 2 | loss of function | device inoperable
load | 2 | low load | Premature Deployment
load | 2 | low load capacity | Device-Device Incompatibility
load | 2 | over load | device damaged prior to use
load | 2 | permanent deformation | mechanical issue
locate, location, localization, local |  | lack of local language | human-device interface issue
locate, location, localization, local | 3 | mislocation | positioning issue
lock | 2 | back out | Fail-Safe Mechanism Issue
lock | 2 | damaged locking mechanism | protective measure issue
lock | 2 | incorrect pin number | device markings issue
lock |  | fail to lock | protective measure issue; improper or incorrect procedure or method; use of incorrect control settings
lock | 2 | locking mechanism failed | Fail-Safe Mechanism Issue; use of incorrect control settings
lock |  | system locked up | application security issue
lock | 2 | pin dissociation | device misassembled during manufacturing or shipping
login | 3 | fail to log in | application security issue
login | 3 | fail to log out | application security issue
log | 3 | inadequate log events | data issue
login | 3 | insecure login | unauthorized access to computer system
log |  | missing data | data issue
log | 3 | insufficient log space | data issue
log |  | log error | device displays error message; application program issue
login | 3 | log in or off slow | human-device interface issue
log | 3 | unable to record event | data back-up problem
logo | 3 | incorrect logo | device markings issue
logo | 3 | mislocated logo | device markings issue
logo | 3 | missing logo | device markings issue
longevity | 1 | reduced device life cycle | shelf life exceeded
lose, loss | 2 | loss of communication | Communication or Transmission Issue
lose, loss | 2 | loss of data | loss of data
lose, loss | 2 | loss of parts | component missing
lubricate, lubrication |  | lack of lubrication | component missing
lubricate, lubrication |  | excessive lubrication | defective component
lubricate, lubrication | 3 | insufficient lubrication | component missing
luer |  | luer cracked | fluid leak; crack
luer |  | closed male luer cracked | fluid leak; crack
luer |  | luer tip detached | disconnection
luer |  | luer lock tip broken or damaged | break
luer |  | inadequate flow | insufficient flow or underinfusion
lumbago | 3 | pain treatment failed | therapeutic or diagnostic output failure
lung |  | lung nodule | therapeutic or diagnostic output failure
lung | 2 | lung punctured | use of device issue; Material Puncture
magazine | 3
magnet, magnetic, magnetic field | 2 | chemical composition out of specification | chemical issue
magnet, magnetic, magnetic field | 2 | data corrupted | Electromagnetic Interference (EMI)
magnet, magnetic, magnetic field | 2 | device malfunction | device operates differently than expected
magnet, magnetic, magnetic field | 2 | device movement | unintended movement
magnet, magnetic, magnetic field | 2 | device overheating | overheating of device or device component
magnet, magnetic, magnetic field | 2 | disorder pacing | failure to pace or properly pace
magnet, magnetic, magnetic field | 2 | hardware broken or damaged | break
magnet, magnetic, magnetic field | 2 | lead dislodgement | dislodged or dislocated
magnet, magnetic, magnetic field | 2 | magnet failure | defective component
magnet, magnetic, magnetic field |  | excess magnetic field generated | electromagnetic interference (EMI)
magnet, magnetic, magnetic field | 2 | magnet mode discrepancy | improper or incorrect procedure or method
magnet, magnetic, magnetic field | 2 | magnetic interference | Electromagnetic Interference (EMI); patient-device incompatibility
magnet, magnetic, magnetic field | 2 | sensor rate maximized | incorrect measurement
magnet, magnetic, magnetic field | 2 | unintended magnet quench | Unintended Magnet Quench
magnetizer | 3
magnetoresistor | 3
magnifier | 3
maintain, maintenance | 3 | difficult to maintain | device operates differently than expected
maintain, maintenance |  | lack of maintenance process | device maintenance issue
maintain, maintenance | 3 | fail to maintain failure records | device maintenance issue; data back-up problem
maintain, maintenance | 3 | fail to maintain record | device maintenance issue; data back-up problem
maintain, maintenance | 3 | fail to maintain status | device operates differently than expected; data back-up problem
maintain, maintenance | 3 | incorrect preventative maintenance | Maintenance Does not comply to Manufacturers Recommendations
maintain, maintenance | 3 | lack of maintenance process | device maintenance issue
maintain, maintenance | 3 | loss of function | device inoperable
maintain, maintenance | 3 | loss of performance | device maintenance issue; no device output
maintain, maintenance | 3 | maintenance does not comply to manufacturers recommendations | Maintenance Does not comply to Manufacturers Recommendations
maintain, maintenance | 3 | missing maintenance | device maintenance issue
maintain, maintenance | 3 | no maintenance | device maintenance issue
manipulation | 3 | excessive manipulation | use of device issue
manual | 3 | inadequate user manual | Instruction for Use Issue
manual | 3 | incorrect user manual | Instruction for Use Issue
manual | 3 | missing contraindication | Instruction for Use Issue
manual | 3 | missing installation information | no information
manual | 3 | missing operation information | no information
manual | 3 | missing side effect notes | Instruction for Use Issue
manual | 3 | missing troubleshooting information | no information
manufacturer, trademark | 3 | incorrect manufacturer address | device markings issue
manufacturer, trademark | 3 | incorrect manufacturer name | device markings issue
manufacturer, trademark |  | fail to label manufacturer information | device markings issue
manufacturer, trademark | 3 | obsolete manufacturer address | device markings issue
mark, marking | 2 | device marked incorrectly | device markings issue
mark, marking | 2 | illegible marking | device markings issue
mark, marking | 2 | inappropriate marking symbol | device markings issue
mark, marking | 2 | incomplete marking | device markings issue
mark, marking | 2 | no marking | device markings issue
mark, marking | 2 | unclear marking | device markings issue
market | 3 | cross-border selling | Market Issue
market | 3 | grey-trade | Market Issue
market | 3 | market share low | Market Issue
mask |  | lack of exhalation valve | component missing
mask |  | inadequate steam | device disinfection or sterilization issue
mask |  | mask blocked | Blockage within Device or Device Component
mask |  | mask deflated | deflation issue; defective component; Product Quality Issue
mask |  | mask detached | detachment of device or device component; device damaged prior to use
mask |  | mask inflated insufficiently | human-device interface issue
mask |  | mask leakage | air leak
mask |  | mask malformed | defective component; Inflation Issue
mask |  | mask torn | defective component
mask |  | patient experienced adverse event | no known device problem
mask |  | tight mask | human-device interface issue; defective component
mass |  | fail to label mass | device markings issue
mass | 3 | mass out of specification | material integrity issue; use of device issue
material | 3 | breathable material compromised | material integrity issue
material | 3 | cut or hole or dent in material | hole in material; Cut in Material; Dent in Material
material | 3 | foreign material inside | Item Contaminated during Manufacturing or Shipping; Packaging Issue
material |  | flammable material | Fire
material |  | incorrect chemical composition | material erosion
material | 3 | inappropriate color scheme | defective item
material |  | inappropriate material color | Material Discolored
material |  | inappropriate material selected | material rigid or stiff
material | 3 | material biocompatibility compromised | biocompatibility issue
material |  | material broken or damaged | material fragmentation
material |  | device contamination with chemical or other material | Foreign Material Present in Device
material | 3 | material deformed or twisted or bent or scratched | device misassembled during manufacturing or shipping
material | 3 | material deviation | material integrity issue
material | 3 | material detached | Material Disintegration
material | 3 | material distorted or torqued or frayed or separation or fragmentation or rupture | Material Distortion; Material Torqued; Material Frayed; Material Separation; Material Frayed; Material Rupture
material | 3 | material erosion | Material Erosion
material | 3 | material extrusion | Extrusion
material | 3 | material flexibility compromised | material rigid or stiff
material | 3 | material integrity failure | material integrity issue
material | 3 | material invagination | Material Invagination
material | 3 | material life short | shelf life exceeded
material | 3 | material opacification | Material Opacification
material | 3 | material perforation or protrusion or puncture | Material Perforation; material protrusion; Material Puncture
material | 3 | material split or torn | torn material
material | 3 | material too rigid or stiff | material rigid or stiff
material | 3 | material too soft or flexible | material rigid or stiff
material | 3 | material washing damaged | device cleaning issue
material | 3 | radiolucent material compromised | material integrity issue
material | 3 | unraveled material | Unraveled Material
material | 3 | using prohibited materials | material integrity issue
matrice | 3
MDU, motor drive unit | 3 | motor drive unit jammed | Mechanical Jam
MDU, motor drive unit | 3 | motor drive unit stalled | Mechanical Jam
measure, measurement, observe | 2 | an erroneous measurement value | incorrect measurement
measure, measurement, observe | 2 | fail to measure | measurement system incompatibility
measure, measurement, observe |  | fail to populate measurement | incorrect measurement; loss of data; patient data issue
measure, measurement, observe |  | fail to pick end point | incorrect measurement; human-device interface issue; use of device issue
measure, measurement, observe |  | fail to disclose measurement accuracy | instruction for use issue
measure, measurement, observe | 2 | inaccurate measurement item | incorrect measurement; measurement system incompatibility | inaccurate measurement
measure, measurement, observe |  | inadeqaute measurement | incorrect measurement
measure, measurement, observe | 2 | incorrect measurement value | incorrect measurement; measurement system incompatibility; patient data issue
measure, measurement, observe |  | inconsistent measurement value | incorrect measurement; measurement system incompatibility; patient data issue
measure, measurement, observe |  | inappropriate measurement unit | incorrect measurement
measure, measurement, observe |  | inappropriate measurement scale | incorrect measurement
measure, measurement, observe |  | inappropriate measurement range | incorrect measurement; measurement system incompatibility
measure, measurement, observe |  | inappropriate measurement algorithm | incorrect measurement; measurement system incompatibility
measure, measurement, observe |  | inconsistent measurement | incorrect measurement; patient data issue
measure, measurement, observe | 2 | loss of measurement | incorrect measurement; measurement system incompatibility; patient data issue
measure, measurement, observe |  | missing measurement value | patient data issue; data issue
measure, measurement, observe |  | low measurement accuracy | imprecision
measure, measurement, observe |  | meassurement changed | application program issue
measure, measurement, observe | 2 | measurement high | incorrect measurement; measurement system incompatibility
measure, measurement, observe | 2 | measurement low | incorrect measurement; measurement system incompatibility
measure, measurement, observe |  | measurement overwritten | patient data issue; data issue
measure, measurement, observe | 2 | measurement range out of specification | incorrect measurement
measure, measurement, observe |  | measurement lists mixed | incorrect measurement
measure, measurement, observe | 2 | sensor measurement interfered | incorrect measurement
measure, measurement, observe |  | fail to label measurement automatically | device markings issue; application program issue
measure, measurement, observe |  | fail to display measurement value | no display or display failure
measure, measurement, observe | 2 | unable to measure | measurement system incompatibility
mechanic, mechanical, mechanism | 3 | crushing surface | Scratched Material
mechanic, mechanical, mechanism | 3 | mechanical problem | mechanical issue
mechanic, mechanical, mechanism |  | loss of mechanical integrity | disassembly; detachment of device component
mechanic, mechanical, mechanism | 3 | mechanics altered | Mechanics Altered
mechanic, mechanical, mechanism | 3 | sharp edge | Material Puncture
MEDDEV | 2 | MEDDEV regulation violated | compliance issue
medication | 1 | fail to nebulization medication | medication error
medication |  | medication clotted | chemical issue
medication | 1 | medication error | medication error
medication |  | medication spilled | fluid leak; medication error
medication | 1 | medication leakage | medication error
membrane | 3 | membrane leak | leak
membrane | 3 | membrane torn | torn material
memory, cache | 2 | data written failed | data issue
memory, cache |  | attention and memory failure | human factors issue
memory, cache |  | lapse of memory | human factors issue
memory, cache | 2 | insufficient memory size | data back-up problem
message |  | error message | device alarm system issue
message |  | exception message error | device alarm system issue
message | 3 | fail to activate message | failure to deploy
message | 3 | fail to auto-reset message | reset issue
message | 3 | fail to clear message | prompts will not clear
message |  | inadequate message display | incorrect display
message | 3 | loss of message | device alarm system issue
message | 3 | unreadable message | false reading from device noncompliance
metal | 3 | metal shedding debris | metal shedding debris
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter | 2 | lack of backup meter | imprecision; no fail-safe mechanism
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter |  | low measurement accuracy | incorrect measurement; imprecision
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter | 2 | meter failure | imprecision
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter | 2 | meter out of range | imprecision
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter |  | meter stuck | physical resistance; sticking
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter |  | incorrect readings | imprecision
meter, oximeter, pyrometer, scintillometer, spectrometer, spirometer, tachometer, thermometer, voltmeter, manometer, bolometer, densitometer, flowmeter, magnetometer, altimeter |  | unstable meter | imprecision
method | 3 | inappropriate method | improper or incorrect procedure or method
method | 3 | incorrect method | improper or incorrect procedure or method
microcircuit | 3
microorganism |  | viable microorganism | bacterial contamination of device
microphone | 3
microscope | 3
migrate, migration, expulsion | 2 | device migrated | migration of device or device component
migrate, migration, expulsion | 2 | implant migrated | migration of device or device component
migrate, migration, expulsion | 2 | unintended part subsidence | migration of device or device component
mirror | 3
misalign, misalignment | 1 | difficult to insert | Difficult to Insert
misalign, misalignment | 1 | fail to insert | Difficult to Insert
misuse | 2 | improper instrument operation | device handling issue
misuse | 2 | incorrect installation | installation-related problem
misuse | 2 | off label use | use of device issue
misuse | 2 | user error | device handling issue
mixer | 3
mobile | 3 | fail to move | failure to deploy
mobile | 3 | overbalancing | unstable
mode, state |  | fail to enter auto mode | human-device interface issue
mode, state | 3 | fail to activate mode | human-device interface issue
mode, state | 3 | fail to deactivate mode | human-device interface issue | exit mode
mode, state |  | fail to select mode | human-device interface issue
mode, state |  | inappropriate mode configuration | human-device interface issue
mode, state | 3 | incorrect mode | human-device interface issue
mode, state |  | incompatible mode | component incompatible
mode, state |  | inadequate therapy modes | human-device interface issue
mode, state |  | fail to adjust independently | human-device interface issue
mode, state |  | unclear state | human-device interface issue
mode, state |  | invisible state | human-device interface issue
mode, state | 3 | mode change slow | human-device interface issue
modem | 3
modify, modification | 2 | fail to implement modification | use of device issue
modify, modification | 2 | lack of change control | improper or incorrect procedure or method
modify, modification | 2 | lack of process | improper or incorrect procedure or method
modify, modification | 2 | lack of validity | improper or incorrect procedure or method
modify, modification | 2 | release modification inappropriately | improper or incorrect procedure or method
modulator, modularity, module |  | fail to integrate | component incompatible
modulator, modularity, module |  | fail to enable module | use of incorrect control settings
modulator, modularity, module |  | inappropriate module selected | use of incorrect control settings; use of device issue
modulator, modularity, module |  | incompatible module | component incompatible
modulator, modularity, module | 3 | defective module | component incompatible; failure to recalibrate
moisture | 3 | component short | electrical shorting; spark
moisture | 3 | dendritic growth | Moisture or Humidity Problem
moisture | 3 | excess moisture | Moisture or Humidity Problem; Moisture Damage
moisture | 3 | liquid ingress | Moisture Damage
moisture | 3 | poor can weld | Solder Joint Failure
monitor | 3 | delayed monitoring | inadequate user interface; improper device output
monitor |  | inadequate display message | inadequate user interface
monitor | 3 | fail to monitor acceleration | failure to read input signal
monitor | 3 | fail to monitor health and status | protective measure issue
monitor | 3 | fail to monitor power event | power conditioning issue
monitor |  | fail to monitor by hardware | failure to read input signal
monitor | 3 | incapable to monitor temperature | temperature issue
monitor | 3 | lack of ambient monitor | environmental control or utility issue
monitor | 3 | lack of health and status monitor | protective measure issue
monitor | 3 | lack of internal temperature monitor | temperature issue
monitor |  | lack of mode monitor | inadequate user interface
monitor |  | inappropriate mode monitor | data issue
monitor |  | blue screen | device inoperable
monitor | 3 | monitor failure | defective component
monitor |  | display to different monitor | incorrect display
monitor | 3 | terminate monitoring with or without alarm | device alarm system issue
morphology | 3 | errors in morphology signal | failure to read input signal
morphology | 3 | fail to cancel morphology | use of device issue
morphology | 3 | lack of morphology evaluation | use of device issue
morphology | 3 | morphology change | use of incorrect control settings
motion | 2 | inelastic motion | unintended system motion
motion | 2 | motion out of range | unintended system motion
motion | 2 | uncontrolled motion | unintended movement
motion | 2 | unintended motion | unintended system motion
motor | 2 | motor failure | defective component
mount | 3 | inappropriate mount | installation-related problem
mount |  | self-dismantling | detachment of device component
mount |  | low mount strength | mechanical issue
mount | 3 | unstable mount | installation-related problem; unstable
move, movement | 2 | device movement in pocket | unintended movement
move, movement | 2 | force on the device due to MRI static field | Electromagnetic Interference (EMI)
move, movement | 2 | hard to initiate movement | unintended movement
move, movement |  | inappropriate movement direction | unintended movement
move, movement | 2 | inadvertent movement | unintended movement
move, movement | 2 | move inelastically within force range | mechanical issue
move, movement | 2 | move inelastically within torque range | defective item
move, movement | 2 | moving distance over limit | device operates differently than expected
move, movement | 2 | moving speed slow | improper device output; imprecision
move, movement |  | fail to provide real time movement | unintended movement
move, movement | 2 | torque on the device due to MRI static field | Electromagnetic Interference (EMI)
move, movement | 2 | uncontrolled movement | unintended movement
move, movement |  | unproportional movement | unintended movement
move, movement | 2 | unwanted movement | unintended movement
MPU, motor protection unit | 3 | fail to detect improper operation | device handling issue
MRI, MR | 2 | battery current drain | electrical issue; Electromagnetic Interference (EMI)
MRI, MR |  | device attracted | Electromagnetic Interference (EMI)
MRI, MR | 2 | device broken or damaged | device damaged prior to use; Electromagnetic Interference (EMI)
MRI, MR | 2 | device malfunction | device operates differently than expected; Electromagnetic Interference (EMI)
MRI, MR | 2 | device reset | reset issue; Electromagnetic Interference (EMI)
MRI, MR | 2 | magnetic fields penetrated | Electromagnetic Interference (EMI)
MRI, MR | 2 | measurement data corrupted | Electromagnetic Interference (EMI)
MRI, MR |  | incorrect MR measurement | data issue
MRI, MR | 2 | MRI induced device vibration | vibration; Electromagnetic Interference (EMI)
MRI, MR | 2 | MRI induced lead voltage | Electromagnetic Interference (EMI)
MRI, MR | 2 | MRI log unavilable | data issue
MRI, MR | 2 | MRI static field strength out of specification | Electromagnetic Interference (EMI)
MRI, MR | 2 | power on reset | Electromagnetic Interference (EMI); reset issue
MRI, MR | 2 | pulse enhanced | Electromagnetic Interference (EMI)
MRI, MR | 2 | RF rectification | Electromagnetic Interference (EMI); Radiofrequency Interference (RFI)
MRI, MR | 2 | software corrupted | Electromagnetic Interference (EMI)
MRI, MR |  | image not in DICOM format | image display error; data issue
MRI, MR | 2 | poor image | poor quality image; Image Resolution Poor; therapeutic or diagnostic output failure
MRI, MR | 2 | unintended episode occur | Electromagnetic Interference (EMI)
MRI, MR | 2 | unintended episode self terminate | Electromagnetic Interference (EMI)
MRI, MR | 2 | unintended stimulation | Electromagnetic Interference (EMI)
MRI, MR | 2 | wrong MRI scan condition | Electromagnetic Interference (EMI); use of incorrect control settings
MTBF, mean time between failure, mean time between failures, MTB | 2 | MTBF out of range | use of device issue
multiplier | 3
muscle | 3 | fail to improve muscle performance | Unexpected Therapeutic Results
muscle | 3 | sore and aching muscle | Unexpected Therapeutic Results
muscle | 3 | stimulate on undesired muscle | Therapy delivered to incorrect body area
mute | 3 | audio mute failed | device difficult to setup or prepare
mute | 3 | audio muted by unauthorized person | unauthorized access to computer system
mute | 3 | sound mute setting failed | device difficult to setup or prepare
N95 | 2 | unqualified bacterial filter | Filtration Issue
needle | 2 | inadvertent puncture | use of device issue; Material Puncture
needle | 2 | needle dull | Dull
needle | 2 | needle extra | component missing |  |  |  |  | how to different from operation/operate
needle | 2 | needle separation | Material Separation
needle | 2 | needle unsheathed | Misassembled by Users
neonate | 3 | inappropriate user | user used incorrect product for intended use
network | 2 | lack of security specification | computer system security issue
network | 2 | loss of connection | connection issue
NMES, neuro muscular electrical stimulation | 2 | fail to improve muscle performance | improper device output
NMES, neuro muscular electrical stimulation | 2 | stimulate on undesired muscle | Therapy delivered to incorrect body area
noise | 2 | ambient noise high | audible noise
noise | 2 | confused about noise detection | inadequate user interface
noise | 2 | data misinterpretation | data issue
noise | 2 | incorrect data entered | data issue
noise |  | noise emitted | audible noise
noise | 2 | noise in power supply | electrical power problem
noise | 2 | noise introduced | Ambient Noise Issue
noise | 2 | noise over specification | audible noise
noise | 2 | noise reversion | audible noise
noise | 2 | oversensing | Over-sensing
noise |  | strange noise | audible noise
non-invasive, noninvasive | 3 | compressed air as the source of air | medical gas supply problem
notify, notifier, notification | 3 | confused patient notification | device alarm system issue
notify, notifier, notification | 3 | excessive notification | device alarm system issue
notify, notifier, notification | 3 | fail to activate alarm | device alarm system issue; device operates differently than expected; improper alarm
notify, notifier, notification | 3 | fail to generate notification | device alarm system issue
notify, notifier, notification | 3 | fail to indicate failure cause | no display or display failure
notify, notifier, notification | 3 | fail to indicate priority | no display or display failure
notify, notifier, notification | 3 | fault notification | false alarm
notify, notifier, notification | 3 | inaudible notification | not audible alarm
notify, notifier, notification |  | lack of read-only prompt | device alarm system issue; computer software issue
notify, notifier, notification | 3 | lack of notification severity | device alarm system issue
notify, notifier, notification | 3 | lack of notification time stamp | imprecision
notify, notifier, notification | 3 | loss of notification | device alarm system issue
nozzle | 3
nut | 3
obturator | 3
occlusion, occlude, obstruct, obstruction | 1 | fail to detect occlusion | occlusion within device; failure to sense; insufficient flow or underinfusion
occlusion, occlude, obstruct, obstruction |  | fail to detect obstruction | occlusion within device
occlusion, occlude, obstruct, obstruction |  | Undetected occlusion | occlusion within device; failure to sense; insufficient flow or underinfusion
occlusion, occlude, obstruct, obstruction | 1 | fail to interrupt infusion | insufficient flow or underinfusion
occlusion, occlude, obstruct, obstruction | 1 | incorrect occlusion | occlusion within device
occlusion, occlude, obstruct, obstruction | 1 | lack of occlusion alarm | occlusion within device; device alarm system issue
occlusion, occlude, obstruct, obstruction |  | occlusion notified | occlusion within device
occlusion, occlude, obstruct, obstruction |  | unspecified occlusion alarm | occlusion within device; device alarm system issue
occlusion, occlude, obstruct, obstruction | 1 | occlusion pressure out of range | occlusion within device; pressure issue
occlusion, occlude, obstruct, obstruction | 1 | under infusion | insufficient flow or underinfusion
oil | 3 | oil leak | leak
onscreen | 3 | fail to pop-up | human-device interface issue; application program issue
onscreen | 3 | inadequate display message | incorrect display
open | 3 | difficult to open | difficult to open or close
operate, operation, operational | 2 | application of bad rule | device handling issue
operate, operation, operational | 2 | device malfunction | device operates differently than expected
operate, operation, operational | 2 | device operates differently than expected | device handling issue
operate, operation, operational | 2 | difficult operation | device handling issue
operate, operation, operational | 2 | inaccessibility | human-device interface issue
operate, operation, operational | 2 | inadequate or improper illumination | Inadequate Lighting
operate, operation, operational | 2 | inadvertent operation | device operates differently than expected
operate, operation, operational | 2 | inappropriate operation environment | improper or incorrect procedure or method
operate, operation, operational | 2 | incorrect operation | device handling issue; device operates differently than expected
operate, operation, operational | 2 | intrusion | material protrusion
operate, operation, operational | 2 | mistiming | imprecision
operate, operation, operational | 2 | no trained operation | Inadequate Training
operate, operation, operational | 2 | omission | device handling issue
operate, operation, operational | 2 | operation early or late | device handling issue; device operates differently than expected
operate, operation, operational | 2 | operation error | device handling issue; device operates differently than expected
operate, operation, operational | 2 | operation out of sequence | device handling issue; device operates differently than expected
operate, operation, operational | 2 | operation speed out of range | device handling issue; device operates differently than expected; imprecision
operate, operation, operational | 2 | operation too briefly | device handling issue; device operates differently than expected
operate, operation, operational | 2 | operation too long | device handling issue; device operates differently than expected
operate, operation, operational | 2 | reversal | device handling issue; device operates differently than expected
operate, operation, operational | 2 | routine violation | device handling issue; device operates differently than expected
operate, operation, operational | 2 | short cut | device handling issue; device operates differently than expected
operate, operation, operational | 2 | unable to operate | device inoperable; device operates differently than expected
operate, operation, operational | 2 | unable to stop operation | device operates differently than expected
operator | 2 | fatigue | use of device issue
operator | 2 | glare | use of device issue
operator | 2 | inadequately trained | Inadequate Training
operator | 2 | misordering | use of device issue
operator | 2 | operator error | use of device issue; device handling issue
operator | 2 | unqualified use | Inadequate Training
optical | 3 | optic edge issue | Material Puncture
optical | 3 | optical decentration | Optical Decentration
optical | 3 | optical discoloration | Optical Discoloration
optical | 3 | optical distortion | Optical Distortion
optical | 3 | optical obstruction | Optical Obstruction
optical | 3 | optical problem | optical issue
optocoupler | 3
orient, orientation | 2 | adverse impacted by orientation | improper device output
orient, orientation | 2 | incorrect device orientation | failure to align; image orientation incorrect
orient, orientation |  | incorrect patient orientation | failure to align; image orientation incorrect
orient, orientation | 2 | incorrect part orientation during assembly | device misassembled during manufacturing or shipping
orient, orientation | 2 | lack of orientation | failure to align; image orientation incorrect
orient, orientation |  | fail to visualize orientation | no display or display failure
orient, orientation |  | fail to auto-orient model | image orientation incorrect
orient, orientation |  | fail to reset orientation | use of device issue; image orientation incorrect
orient, orientation |  | orientation marked incorrectly | failure to align; image orientation incorrect
orient, orientation | 2 | orientation accuracy out of specification | failure to align; image orientation incorrect
oscillator | 3
oscilloscope | 3
osmosis | 3 | reverse osmosis | device operates differently than expected
osseointegrate, osseointegration | 2 | fail to osseointegrate | Failure to Osseointegrate
osseointegrate, osseointegration | 2 | loss of osseointegration | Loss of Osseointegration
output | 3 | false negative output | False Device Output
output | 3 | false output | False Device Output
output | 3 | false positive output | False Device Output
output | 3 | loss of output | no device output
output | 3 | output above specification | output above specifications
output | 3 | output below specification | Output below Specifications
output | 3 | output high | improper device output
output | 3 | output low | improper device output
oven | 3
overheat, overheating | 2 | chemical release | chemical spillage
overheat, overheating | 2 | catch fire | Fire
overheat, overheating |  | inadequate cooling | overheating of device or device component
overheat, overheating |  | insufficient insulation | overheating of device or device component
overheat, overheating | 2 | overheated battery | overheating of device or device component
overheat, overheating | 2 | overheated parts | overheating of device or device component
overheat, overheating | 2 | overheating protection failed | overheating of device or device component
overlay | 3 | confused overlay | inadequate user interface
overlay | 3 | fail to disable overlay | device remains activated
overlay | 3 | fail to display previous data | no display or display failure
overlay | 3 | unable to overlay | difficult to position
overlay | 3 | unable to remove overlay | difficult to remove
overload | 1 | fasten broken or damaged | break
overload | 1 | mechanical damage | mechanical issue
overload | 1 | permanent deformation | mechanical issue
overload | 1 | short circuit | electrical shorting; spark
oversense | 2 | confused diagnostics | inadequate user interface
oversense | 2 | intermittent oversensing | Device Sensing Issue
oversize | 3 | connection broken or damaged | connection issue; break; device damaged prior to use
oversize | 3 | installation difficult | installation-related problem
overtravel | 3 | overtravel out of range | use of device issue
oxidize, oxidation, oxide | 2 | inadequate oxide surface | defective item
oxidize, oxidation, oxide | 2 | insulation failure | electrical shorting
oxidize, oxidation, oxide | 2 | loss of electrical isolation | electrical shorting
oxidize, oxidation, oxide | 2 | part worn out | Naturally Worn
oxygen, O2, medical gas | 3 | fail to cease delivery | medical gas supply problem
oxygen, O2, medical gas | 2 | fail to update oxygen volume | medical gas supply problem
oxygen, O2, medical gas |  | low oxygen concentration | medical gas supply problem; output problem
oxygen, O2, medical gas |  | low oxygen fraction accuracy | medical gas supply problem
oxygen, O2, medical gas |  | oxygen out of specification | medical gas supply problem
oxygen, O2, medical gas |  | fail to enable oxygen flush | failure to deliver
oxygen, O2, medical gas |  | incorrect O2 setting | medical gas supply problem
oxygen, O2, medical gas |  | incorrect supply of medical gas | medical gas supply problem
oxygen, O2, medical gas |  | inappropriate measurement range | incorrect measurement
oxygen, O2, medical gas |  | fail to compatible with Oxygen | Incompatibility Problem
oxygen, O2, medical gas |  | O2 ball stuck | sticking
oxygen, O2, medical gas | 2 | over oxygen concentration | medical gas supply problem; output problem
oxygen, O2, medical gas | 3 | oxygen delivery compromised | medical gas supply problem; use of incorrect control settings
oxygen, O2, medical gas |  | oxygen leakage | air leak
oxygen, O2, medical gas | 3 | oxygen pressure low | medical gas supply problem
oxygen, O2, medical gas | 3 | setup incorrectly | medical gas supply problem
pace, pacing | 2 | apical pacing | failure to pace or properly pace
pace, pacing | 2 | asynchronous pacing | failure to pace or properly pace
pace, pacing | 2 | backup pacing | failure to pace or properly pace
pace, pacing | 2 | competitive pacing | failure to pace or properly pace
pace, pacing | 2 | device explant | device handling issue
pace, pacing | 2 | double counting pace | failure to pace or properly pace
pace, pacing | 2 | inappropriate elevated heart rate | failure to pace or properly pace
pace, pacing | 2 | inappropriate pacing voltage | failure to pace or properly pace
pace, pacing | 2 | induction of an arrhythmia | Device Sensing Issue
pace, pacing | 2 | insufficient pace neutrality | failure to pace or properly pace
pace, pacing | 2 | loss of A-V synchrony | failure to pace or properly pace
pace, pacing | 2 | loss of pacing | No Pacing
pace, pacing | 2 | lower pacing rate | Pacing Inadequately
pace, pacing | 2 | overdrive pacing | failure to pace or properly pace
pace, pacing | 2 | paced refractory event | failure to pace or properly pace
pace, pacing | 2 | pacing around a scar tissue | Energy Output to Patient Tissue Incorrect; failure to pace or properly pace
pace, pacing | 2 | pacing asynchronously | Pacing Asynchronously
pace, pacing | 2 | pacing at wrong place | Energy Output to Patient Tissue Incorrect; failure to pace or properly pace
pace, pacing | 2 | pacing delayed | failure to pace or properly pace
pace, pacing | 2 | pacing inadequately | Pacing Inadequately
pace, pacing | 2 | pacing intermittently | Pacing Intermittently
pace, pacing | 2 | pacing problem | failure to pace or properly pace
pace, pacing | 2 | suboptimal pacing support | use of device issue
pace, pacing | 2 | suboptimal rate responsive pacing | Pacing Inadequately
pace, pacing | 2 | subthreshold pacing | Pacing Inadequately
pacemaker | 2 | device rotated | migration of device or device component
pacemaker | 2 | found in back-up mode | improper or incorrect procedure or method
pacemaker | 2 | Twiddler’s Syndrome | failure to pace or properly pace
pack | 3 | out of dimension | defective component
package, pack, packaging | 3 | faulty sealing | Unsealed Device Packaging
package, pack, packaging | 3 | incomplete packaging | Incomplete or Missing Packaging
package, pack, packaging | 3 | incorrect package labeling | device markings issue
package, pack, packaging |  | inappropriate package design | human-device interface issue
package, pack, packaging | 3 | insufficient packaging | Device Packaging Compromised
package, pack, packaging |  | incorrect packaging | packaging issue; Device Packaging Compromised
package, pack, packaging | 3 | low level package type | Device Packaging Compromised
package, pack, packaging |  | missing content description | device markings issue
package, pack, packaging | 3 | missing brand marking | device markings issue
package, pack, packaging | 3 | missing CE marked | device markings issue; Packaging Issue
package, pack, packaging | 3 | missing legal manufacturer | device markings issue; Packaging Issue
package, pack, packaging | 3 | missing legal representative | device markings issue; Packaging Issue
package, pack, packaging | 3 | missing packaging | Incomplete or Missing Packaging
package, pack, packaging | 3 | missing part or document | component missing; Packaging Issue
package, pack, packaging |  | missing patient implant card | device markings issue; Packaging Issue
package, pack, packaging |  | missing Tracking Form | device markings issue; Packaging Issue
package, pack, packaging |  | missing hand-hold area | human-device interface issue
package, pack, packaging |  | fail to label batch code | device markings issue; Packaging Issue
package, pack, packaging |  | fail to label serial number | device markings issue; Packaging Issue
package, pack, packaging |  | fail to label fragile | device markings issue; Packaging Issue
package, pack, packaging |  | fail to label temperature limit | device markings issue; Packaging Issue
package, pack, packaging |  | fail to label humidity limit | device markings issue; Packaging Issue
package, pack, packaging |  | fail to label Atmospheric pressure limit | device markings issue; Packaging Issue
package, pack, packaging | 3 | out of dimension | defective component
package, pack, packaging | 3 | package compromised | Device Packaging Compromised; device contaminated during manufacture or shipping
package, pack, packaging |  | package broken or damaged | Device Packaging Compromised; Packaging Issue
package, pack, packaging | 3 | package time long | Packaging Issue
package, pack, packaging | 3 | tear or rip or hole in packaging | Tear, Rip or Hole in Device Packaging; Packaging Issue
pad | 3 | bond pad broken or damaged | Loss of or Failure to Bond
pad | 3 | pad corrosion | Corrosion
pain | 1 | chronic pain | Unexpected Therapeutic Results
pain | 1 | intractable pain | Unexpected Therapeutic Results
pain | 1 | pain associated with arthritis | Unexpected Therapeutic Results
pain | 1 | pain relief compromised | Unexpected Therapeutic Results
pain | 1 | pain treatment failed | therapeutic or diagnostic output failure
pain | 1 | temporary pain | Unexpected Therapeutic Results
pallet | 3 | pallet break | break
pallet | 3 | pallet cracked | Crack
palpability | 2
panel | 3 | fail to activate viewing panel | application program issue
paper | 3 | dry paper | defective component
paper | 3 | paper folded | Folded
paper | 3 | paper profile large | use of incorrect control settings
paper | 3 | paper torn | torn material
paper | 3 | thick paper | defective component
paper | 3 | thin paper | defective component
paramagnetic | 3
parameter | 2 | incorrect parameter | use of incorrect control settings
parameter |  | low parameter accuracy | incorrect measurement; imprecision
parameter | 2 | parameter calculation error | parameter calculation error due to software problem
parameter | 2 | parameter confused | inadequate user interface
parameter |  | fail to reset default parameters | use of incorrect control settings
parameter |  | fail to manipulate parameters | application program issue
parameter | 2 | parameter misprogrammed | use of incorrect control settings
paravalvular | 3 | paravalvular leak | Peri-valvular leak
part |  | hole in part | hole in material
part |  | fail to inflate evenly | Inflation Issue
part |  | unable to inflate | Inflation Issue
part |  | unable to deflate | deflation issue
part |  | loose part | detachment of device component
part |  | part broken or damaged | break; split; defective component
part |  | part burned | Burn of Device or Device Component; burning odor
part |  | part burst | Burst
part |  | part cracked | Crack
part |  | part crushed | collapse
part |  | part deflated | deflation issue
part |  | part deformed | material deformation
part |  | part detached | connection issue; detachment of device component
part |  | part dropped off | component falling
part |  | part ejected | unintended ejection
part |  | part extended | Stretched
part |  | part falling apart | split
part |  | part kinked | kinked
part |  | part leakage | leak; fluid leak
part |  | part melted | Melted; fire
part |  | part mismatched | component incompatible; fitting problem
part |  | part misoriented | defective item
part |  | part misshaped | defective component; fitting problem
part |  | part missing | component missing
part |  | part ripped | material separation
part |  | part ruptured | material rupture
part |  | part sheared off | break
part |  | part splitted | split; material separation
part |  | part stuck | sticking
part |  | sticky part | sticking
part |  | part thermal damaged | Burn of Device or Device Component
part |  | part torqued | material torqued
part |  | part twisted | material twisted
part |  | part torn | torn material
part |  | part warped | material deformation
particulate, particulation, particle, foreign matter | 3 | particles in the path | electrical shorting; spark
particulate, particulation, particle, foreign matter |  | particles inside fluid path | Particulates
particulate, particulation, particle, foreign matter |  | particles outside fluid path | Particulates
particulate, particulation, particle, foreign matter |  | Microbiological Growth | Particulates
particulate, particulation, particle, foreign matter |  | Microbiological Mold | Particulates
particulate, particulation, particle, foreign matter | 3 | particulate presented | Particulates
particulate, particulation, particle, foreign matter | 3 | particulate size out of range | Particulates
patch | 3 | lack of patch implement procedure | compliance issue
pathway | 3 | inappropriate insulin pathway | medication error
pathway | 3 | incorrect electrical pathway | electrical shorting
pause | 3 | fail to activate pause | computer software issue
pause | 3 | lack of pause function | computer software issue
pediatric | 3 | inappropriate user | user used incorrect product for intended use
PEMS, programmable electrical medical systems | 3 | failure of the IT-NETWORK | application network issue
PEMS, programmable electrical medical systems | 3 | inconsistent input | failure to read input signal
PEMS, programmable electrical medical systems | 3 | incorporate into an IT-NETWORK | application network issue
PEMS, programmable electrical medical systems | 3 | incorrect timing of data | imprecision
PEMS, programmable electrical medical systems | 3 | input originating from electromagnetic interference | Electromagnetic Interference (EMI)
PEMS, programmable electrical medical systems | 3 | lack of architecture specification | compliance issue
PEMS, programmable electrical medical systems | 3 | lack of data security | application security issue
PEMS, programmable electrical medical systems | 3 | lack of integrity of data | loss of data
PEMS, programmable electrical medical systems | 3 | lack of specification | defective component
PEMS, programmable electrical medical systems | 3 | lead to unacceptable risk | device unsafe to use in environment
PEMS, programmable electrical medical systems | 3 | out of range input | failure to read input signal
PEMS, programmable electrical medical systems | 3 | unavailable data | loss of data
PEMS, programmable electrical medical systems, programmable electronic subsystem | 3 | undesired feedback | use of device issue
PEMS, programmable electrical medical systems, programmable electronic subsystem | 3 | unintended interaction | Electromagnetic Interference (EMI)
PEMS, programmable electrical medical systems, programmable electronic subsystem | 3 | unknown third-party software | computer software issue
PEMS, programmable electrical medical systems, programmable electronic subsystem | 3 | unsolicited input | failure to read input signal
percutaneous | 3
perforate, perforation | 1 | atrium perforated | Material Perforation
perforate, perforation | 1 | blood vessel perforated | Material Perforation; use of device issue
perforate, perforation | 1 | cardiac perforation | Material Perforation
perforate, perforation | 1 | labial perforated | Material Perforation
perforate, perforation | 1 | lingual perforated | Material Perforation
perforate, perforation | 1 | mandibular nerve perforated | Material Perforation
perforate, perforation | 1 | maxillary sinus perforated | Material Perforation
perforate, perforation | 1 | tissue perforated | Material Perforation
perforate, perforation | 1 | ventricle perforated | Material Perforation
perform, performance | 3 | fail to perform | no device output
perform, performance | 3 | loss of performance | no device output
period |  | period out of range | use of incorrect control settings
period | 3 | product life short | shelf life exceeded
perivalvular | 3 | perivalvular leak | Peri-valvular leak
permeability |  | inappropriate permeability | fluid leak
PEMS, programmable electrical medical systems, programmable electronic subsystem | 3 | lead to unacceptable risk | device unsafe to use in environment
pH | 3 | high pH | High pH
pH | 3 | low pH | chemical issue
phantom | 3 | phantom problem | image display error
photo, photograph, photography | 3 | fail to take photo | use of device issue
photodetector | 3
photomultiplier | 3
phototransistor | 3
pickup | 3
pin | 3 | bent pin | Bent
pin | 3 | fail to retain pin | device handling issue
pin | 3 | inappropriate pin material | material integrity issue
pin | 3 | pin broken or damaged | break
pin |  | pin dislodged | dislodged or dislocated
pin |  | pin detached | detachment of device component
pin | 3 | pin length out of specification | device misassembled during manufacturing or shipping
pin | 3 | poor elastic pin | device misassembled during manufacturing or shipping
pin | 3 | stuck pin | sticking
pipette | 3 | failure to aspirate | Aspiration Issue
piston | 3 | loose bolt retaining nut | loose or intermittent connection
piston | 3 | expoxy holding failed | misassembled
piston | 3 | fastener failed | defective component
piston | 3 | mechanical failure | mechanical issue
piston | 3 | part broken or damaged | break; defective component
piston | 3 | piston seal worn | failure to adhere or bond
piston | 3 | roller slot | defective component
piston | 3 | sliding surface worn | Naturally Worn
pit | 3
pivot | 3
place, placement | 2 | implant placed at improper location | dislodged or dislocated; Failure to Advance
place, placement | 2 | incorrect placement | Failure to Advance
place, placement |  | unable to place | difficult to advance
place, placement | 2 | misplacement | Failure to Advance
plate | 3 | plate broken or damaged | break
plot | 3 | fail to reset plot | failure to reset
plot | 3 | incorrect plot | defective component
plotter | 3
plug, pluggable | 3 | fit more power supply cords | misconnection; fitting problem
plug, pluggable | 3 | plug mismatched | component incompatible
plug, pluggable |  | unplug accidently | device handling issue
plunge | 3
plunger | 3 | plunger not aligned properly | failure to align
plunger | 3 | plunger override | improper device output
PN, part number, ID |  | missing site ID | data issue
PN, part number, ID |  | inconsistent number | data issue
PN, part number, ID | 3 | missing part number | device markings issue; data issue
pneumatic | 3 | implosion | break
pneumatic | 3 | Inadequate pneumatic power | loss of power
pneumatic | 3 | loss of pressure | no pressure
pneumatic | 3 | loss of vacuum | suction issue
pneumatic | 3 | miscalibrated relief device | Failure to Calibrate
pneumatic | 3 | mislocated relief device | dislodged or dislocated
pneumatic | 3 | over pressure | Increase in Pressure
pneumatic | 3 | pipe or vessel rupture | Material Rupture; use of device issue
pocket | 2 | bleeding in pocket | Bleed Back
pocket | 2 | device erosion | Material Erosion
pocket | 2 | device exposed | incorrect display
pocket | 2 | device movement | unintended movement
pocket | 2 | fluid in pocket | contamination during use
pocket | 2 | loose pocket | detachment of device component
pocket | 2 | pocket erosion | Material Erosion
pocket | 1 | pocket infection | bacterial contamination of device
pocket | 2 | pocket tissue injury | mechanical issue
pointer, point | 3 | sharp point | human-device interface issue
polarize, polar, polarity, polarization | 3 | fail to sense polarization | electrical issue
polarize, polar, polarity, polarization | 3 | incorrect electrode polarization | electrical issue
polarize, polar, polarity, polarization | 3 | incorrect polarity connection | misconnection
polarizer | 3
porosity | 3 | porosity out of specification | material integrity issue
port, injection port, RS-232 port | 2 | Port blocked | blocked connection
port, injection port, RS-232 port |  | fail to label symbol | device markings issue
port, injection port, RS-232 port |  | fail to label text | device markings issue
port, injection port, RS-232 port |  | port broken or damaged | break
port, injection port, RS-232 port |  | port depressed | fluid leak; Dent in Material
port, injection port, RS-232 port |  | port less durable | material integrity issue
port, injection port, RS-232 port |  | port mismatched | component incompatible; fitting problem
port, injection port, RS-232 port |  | port ruptured | fluid leak
port, injection port, RS-232 port |  | port cracked | Crack
port, injection port, RS-232 port | 2 | port leakage | leak; fluid leak
portable | 3 | difficult to move | difficult to remove
position, positionable | 2 | difficult to position | difficult to position
position, positionable | 2 | improper device position | Malposition of Device; instruction for use issue
position, positionable | 2 | incorrect installation | installation-related problem
position, positionable |  | incorrect patient position | positioning issue
position, positionable | 2 | inefficient position | difficult to position
position, positionable | 2 | lack of position indicator | no display or display failure
position, positionable | 2 | loss of position data | loss of data
position, positionable |  | fail to advice position | instruction for use issue
position, positionable |  | fail to reset position | positioning issue; reset issue
position, positionable | 2 | position out of range | positioning issue
position, positionable | 2 | position repeatability out of specification | positioning issue
position, positionable | 2 | positioning delayed | positioning issue
pouch |  | pouch torn | Tear, Rip or Hole in Device Packaging
powder |  | foreign material inside | Item Contaminated during Manufacturing or Shipping; Packaging Issue
power, unpower | 2 | complete loss of power | loss of power
power, unpower | 2 | fail to power up | failure to power-up
power, unpower | 2 | Insufficient power capacity | Device-Device Incompatibility; Power Source Issue
power, unpower | 2 | intermittent loss of power | loss of power
power, unpower |  | incorrect line power | loss of power
power, unpower | 2 | loss of alarm | device alarm system issue
power, unpower | 2 | power source detection issue | Power Source Issue
power, unpower | 2 | loss of power indication | no display or display failure
power, unpower | 2 | power calculation error | Power Calculation Error due to Software Problem
power, unpower | 2 | power consumption out of range | electrical power problem
power, unpower | 2 | power control failed | power conditioning issue
power, unpower | 2 | power down unexpected | loss of power; data issue
power, unpower |  | power fault displayed | device displays error message
power, unpower | 2 | power frequency out of specification | Power Source Issue
power, unpower | 2 | power interrupted | loss of power
power, unpower | 2 | power on or off switch failed | loss of power
power, unpower | 2 | power on reset | power conditioning issue; reset issue
power, unpower | 2 | power out of range | power conditioning issue
power, unpower | 2 | power restoration | power conditioning issue
power, unpower | 2 | short to ground | electrical shorting; spark
power, unpower | 2 | unintended power up | power conditioning issue
precedence | 3 | inappropriate precedence setting | use of incorrect control settings
precipitate, precipitation | 3 | precipitate in device or device ingredient | sediment, precipitate or deposit in device or device ingredient
precision | 2 | cumulative precision out of range | imprecision; calibration issue
precision |  | loss of precision | data issue
precision | 2 | imprecision | imprecision
precision | 2 | precision out of specification | imprecision
preclotting, preclot
prepare, prep, preparation | 3 | fail to prepare | use of device issue
prescaler | 3
prescribe, prescription, E-prescription, prescriber | 2 | fail to retain prescription | device difficult to maintain
prescribe, prescription, E-prescription, prescriber |  | fail to receive prescription | communication or transmission issue
prescribe, prescription, E-prescription, prescriber | 2 | inadequate labeling | device markings issue
prescribe, prescription, E-prescription, prescriber | 2 | incorrect prescription information | no information
preselector | 3
present, presence | 3 | detect presence failed | failure to sense
pressure | 3 | fail to adjust pressure | pressure issue; device handling issue
pressure |  | fail to hold pressure | pressure issue; device handling issue
pressure |  | fail to label pressure range | device markings issue
pressure |  | fail to disclose pressure changing | instruction for use issue
pressure |  | fail to relieve pressure properly | pressure issue
pressure | 3 | fail to produce pressure | pressure issue; device handling issue
pressure |  | fail to provide sufficient pressure | pressure issue; insufficient flow or underinfusion
pressure |  | fail to withstand pressure | mechanical issue; pressure issue
pressure | 3 | fail to update pressure | pressure issue; device handling issue
pressure | 3 | fail to use default pressure setting | reset issue
pressure | 3 | fluid jet | fluid leak; pressure issue
pressure |  | set leakage | pressure issue
pressure |  | stored pressure | fluid injection
pressure |  | high pressure alarm | device alarm system issue
pressure | 3 | inappropriate pressure change | pressure issue
pressure | 3 | inappropriate pressure setting | pressure issue; use of incorrect control settings; failure to cycle
pressure | 3 | inaccurate pressure reading | pressure issue; Failure to Calibrate; imprecision; failure to recalibrate
pressure |  | inappropriate pressure step | pressure issue; use of incorrect control settings
pressure | 3 | incorrect delivery | inaccurate delivery
pressure |  | incorrect pressure calculation | pressure issue
pressure | 3 | insufficient pressure | Decrease in Pressure
pressure | 3 | lack of high pressure alarm | device alarm system issue
pressure |  | lack of pressure indication | device alarm system issue
pressure | 3 | loss of alarm | device alarm system issue
pressure | 3 | loss of pressure | no pressure
pressure | 3 | loss of vacuum | suction issue
pressure | 3 | low pressure | Decrease in Pressure; device handling issue; improper flow or infusion
pressure |  | low pressure alarm | Decrease in Pressure
pressure |  | low measurement accuracy | pressure issue; imprecision; incorrect measurement
pressure |  | pressure measurement out of range | pressure issue | measurement out of range
pressure |  | fail to update pressure | pressure issue
pressure |  | low pressure accuracy | pressure issue; imprecision; incorrect measurement
pressure |  | negative pressure out of range | pressure issue
pressure | 3 | no pressure | no pressure
pressure | 3 | over pressure | Increase in Pressure
pressure | 3 | pressure control failed | pressure issue
pressure | 3 | pressure drop out of range | pressure issue; no device output
pressure |  | pressure increased | Increase in Pressure
pressure |  | pressure dropped | Decrease in Pressure
pressure |  | pressure over upper limit | pressure issue
pressure | 3 | pressure out of range | pressure issue; fluid leak
prime | 2 | delay delivery | delivery system failure
prime |  | difficult to prime | Failure to Prime
prime | 2 | fail to cancel prime | filling problem
prime | 2 | fail to incorporate step sequence | failure to prime
prime | 2 | fail to indicate prime completion | no display or display failure
prime | 2 | fail to prime | Failure to Prime
prime | 2 | fail to suspend delivery | delivery system failure
prime | 2 | inadequate priming | Short Fill
prime | 2 | incomplete priming | Short Fill
prime | 2 | incorrect prime calculation | dose calculation error due to software problem
prime | 2 | lack of fixed priming | Failure to Prime
prime | 2 | lack of manual priming | Failure to Prime
prime | 2 | no prime | Failure to Prime
prime |  | back prime | Failure to Prime
prime | 2 | prime incorrectly | volume accuracy issue
prime | 2 | prime slow | filling problem
print |  | print missing | device markings issue
printer | 3
probe |  | probe deteched | detachment of device component
probe |  | probe locked | failure to sense; failure to read input signal
probe | 2 | defective probe | failure to sense; failure to read input signal
problem | 2 | fail to address problem | use of device issue
problem | 2 | fail to execution program | use of device issue
problem | 2 | fail to identify potential cause | use of device issue
problem | 2 | fail to resolve problem | use of device issue
problem | 2 | introduce additional problem | protective measure issue
problem | 2 | lack of document | lack of maintenance documentation or guidelines
problem | 2 | lack of problem advice | use of device issue
problem | 2 | lack of problem correction | use of device issue
problem | 2 | lack of problem evaluation | use of device issue
problem | 2 | lack of problem investigation | use of device issue
problem | 2 | program execution problem | application program issue
procedure | 3 | improper procedure | improper or incorrect procedure or method
procedure | 3 | incorrect procedure | improper or incorrect procedure or method
procedure | 3 | lack of procedure | improper or incorrect procedure or method
process | 3 | fail to follow process | improper or incorrect procedure or method
process | 3 | inadequate process | improper or incorrect procedure or method
process | 3 | incorrect process | improper or incorrect procedure or method
process | 3 | redundant process | improper or incorrect procedure or method
processor, microprocessor, risc-processor, CPU, NetDCU |  | processor overheat | overheating of device or device component
processor, microprocessor, risc-processor, CPU, NetDCU |  | low calculation speed | electronic property issue; imprecision
processor, microprocessor, risc-processor, CPU, NetDCU |  | low memory space | electronic property issue
processor, microprocessor, risc-processor, CPU, NetDCU |  | loss of communication | Communication or Transmission Issue
processor, microprocessor, risc-processor, CPU, NetDCU |  | fail to switch redundant processor | protective measure issue
processor, microprocessor, risc-processor, CPU, NetDCU | 2 | loss of function | no device output
profile | 3 | inadequate profile | use of incorrect control settings
profile | 3 | incorrect profile for patient care | use of incorrect control settings
program | 2 | confused program interface | inadequate user interface
program | 2 | inappropriate program selected | application program issue
program | 2 | incorrect program parameter | use of incorrect control settings
program | 2 | incorrect program version | application program version or upgrade problem
program | 2 | intermittent program execution | application program issue
program | 2 | long programming time | programming issue
program | 2 | nonfunctional | no device output
program | 2 | parameter misprogrammed | use of incorrect control settings
program | 2 | program end failed | application program issue
program | 2 | program frozen | application program issue
program | 2 | program upgrade failed | application program version or upgrade problem
program |  | fail to launch program | application program issue
program | 2 | unable to program | Difficult to Program or Calibrate
program | 2 | unintended program shut down | application program issue
programmer |  | interference with magnetic | Electromagnetic Compatibility Issue
programmer | 3 | parameter setting error | use of incorrect control settings
prompt | 3 | prompt failed | Inappropriate Prompts
prosthesis, prosthetic | 3 | prosthesis deterioration | degraded
prosthesis, prosthetic | 3 | short prosthetic life | shelf life exceeded
protector | 3
protrude, protrusion | 1 | antenna protruding | material protrusion
protrude, protrusion | 1 | cable protrusion | material protrusion
protrude, protrusion | 1 | protruding septum | material protrusion
PTFE | 3 | insufficient fit of PTFE liner | fitting problem
PTFE | 3 | PTFE kinked | Kinked
PTFE | 3 | PTFE wrinkled | Wrinkled
pulley | 3
pulse, impulse | 3 | fail to deliver pulse | electronic property issue
pulse, impulse | 3 | inadequate pulse amplitude | electronic property issue
pulse, impulse |  | inappropriate pulse delivered | electronic property issue
pulse, impulse | 3 | inappropriate pulse enhanced | electronic property issue
pulse, impulse | 3 | insufficient pace neutrality | Pacing Inadequately
pulse, impulse | 3 | overdrive pulse | electronic property issue
pulse, impulse | 3 | pulse cancellation | electronic property issue
pulse, impulse | 3 | pulse causing noise | electronic property issue
pulse, impulse | 3 | pulse delayed | electronic property issue
pulse, impulse | 3 | undesired changes in pulse amplitude | electronic property issue
pulser | 3 | fail to initiate pulse | electronic property issue
pump | 2 | altitude accommodation slow | environmental control or utility issue
pump | 2 | decreased pump speed | Decreased Pump Speed
pump | 2 | fail to add bolus over basal rate | inaccurate delivery
pump | 2 | fail to drive motor forward | failure to deploy
pump | 2 | fail to drive motor in reverse | failure to deploy
pump | 2 | fail to force motor rewind | mechanical issue
pump | 2 | fail to generate indication | no display or display failure
pump | 2 | fail to pump | Failure to Pump
pump | 2 | fail to resume basal delivery | reset issue
pump | 2 | fail to setup from PC | device difficult to setup or prepare
pump | 2 | fail to stop delivery | Free or Unrestricted Flow
pump | 2 | fail to upgrade software locally | Operating System Version or Upgrade Problem
pump | 2 | increased pump speed | Increased Pump Speed
pump | 2 | lack of feedback | use of device issue
pump | 2 | liquid ingress | fluid leak
pump | 2 | low delivery accuracy | inaccurate delivery
pump | 2 | low displacement accuracy | dislodged or dislocated
pump | 2 | no delivery | delivery system failure
pump | 2 | over delivery | inaccurate delivery
pump | 2 | post-pumping problem | Pre- or Post-Pumping Problem
pump | 2 | pre-pumping problem | Pre- or Post-Pumping Problem
pump | 2 | pump broken or damaged | break
pump | 2 | pump jam | Pumping Stopped
pump |  | pump detached | disconnection
pump | 2 | pump leakage | leak; fluid leak
pump | 2 | pumping stopped | Pumping Stopped
pump | 2 | under delivery | inaccurate delivery
puncture | 2 | blood vessel punctured | use of device issue; Material Puncture
puncture | 2 | heart punctured | use of device issue; Material Puncture
puncture | 2 | loss of electrical isolation | electrical shorting
puncture |  | fail to puncture | Difficult to Insert
puncture | 2 | low puncture strength | mechanical issue
puncture | 2 | lung punctured | use of device issue; Material Puncture
puncture | 2 | valve punctured | use of device issue; Material Puncture
purge | 2 | fail to purge | device cleaning issue
purge |  | inappropriate purge flow | device cleaning issue
purge | 2 | incomplete purge | device cleaning issue
pusher | 3
radiation | 1 | device broken or damaged | device damaged prior to use
radiation | 1 | radiation leak | Radiation leak
radiation | 1 | radiation output failure | Radiation underexposure
radiation | 1 | radiation overexposure | Radiation overexposure
radiation | 1 | radiation underexposure | Radiation underexposure
radiation | 1 | unexpected radiation output | improper device output
radiation | 1 | unintended radiation exposure | Radiation overexposure
radiation | 1 | unintended radiation output | improper device output
radiolucent, radio-lucent | 3 | radiolucent area compromised | Material Opacification
radiolucent, radio-lucent | 3 | radiolucent material compromised | Material Opacification
rail | 3 | incorrect rail configuration | use of incorrect control settings
ram | 2 | data corrupted | computer software issue
ram | 2 | loss of software application | application program issue
ram | 2 | ram broken or damaged | break
RAM, random access memory | 2 | incorrect revision | Operating System Version or Upgrade Problem
RAM, random access memory | 2 | RAM ROM reversion | positioning issue
ramp | 3 | fail to activate ramp | use of device issue
ramp | 3 | fail to cancel ramp | use of device issue
ramp | 3 | fail to end ramp | use of device issue
ramp | 3 | ramp pressure high | Increase in Pressure
ramp | 3 | ramp pressure low | Decrease in Pressure
range | 3 | out of range | defective item; improper device output
rate | 3 | inadequate rate update | Output Issue
rate | 3 | inappropriate rate decrease | Output Issue
rate | 3 | inappropriate rate increase | Output Issue
rate |  | fail to increase rate | Output Issue
rate |  | high rate | output above specifications
rate |  | low rate | Output below Specifications
rate |  | low rate accuracy | inaccurate flowrate; imprecision
rate | 3 | rate out of range | improper device output
rationale | 3 | lack of rationale | defective item
react, reaction, reactor | 3
read, reader, reading | 2 | reading out of specification | High Readings; Low Readings
read, reader, reading |  | High Readings | High Readings; Incorrect or Inadequate Result
read, reader, reading | 2 | imprecise readings | imprecision
read, reader, reading | 2 | inadequate readings | imprecision
read, reader, reading | 2 | incorrect readings | imprecision
read, reader, reading | 2 | invalid readings | imprecision
read, reader, reading | 2 | low readings | Low Readings; Incorrect or Inadequate Result
read, reader, reading | 2 | no readings | Low Readings; Incorrect or Inadequate Result
read, reader, reading | 2 | unable to obtain readings | Unable to Obtain Readings
read, reader, reading | 2 | unable to read | Unable to Obtain Readings
reattach, reattachable, reattachment | 3 | difficult to be reattached | detachment of device component
reattach, reattachable, reattachment | 3 | fail to reattach | detachment of device component
reattach, reattachable, reattachment | 3 | reattachment failed | detachment of device component
reboot, boot |  | fail to boot up | device inoperable; computer software issue
reboot, boot |  | inappropriate reboot | device inoperable; computer software issue
receive | 3 | loss of communication | Communication or Transmission Issue
receive | 3 | received data incorrectly | data issue
receive | 3 | received data late | data issue
receiver | 3 | receiver failure | failure to read input signal
reconfigure, reconfiguration, reconfigurable | 3 | reconfiguration failed | use of incorrect control settings
record, recorder | 3 | fail to record operative data | loss of data; data back-up problem
record, recorder |  | fail to record monitoring data | loss of data; data back-up problem
record, recorder |  | fail to record diagnosis data | loss of data; data back-up problem
record, recorder |  | fail to record repair data | loss of data; data back-up problem
record, recorder |  | fail to record operational events | loss of data; data back-up problem
record, recorder | 3 | fail to transmit record | Failure to Transmit Record
record, recorder | 3 | invalid record | data back-up problem
record, recorder | 3 | lack of record | data back-up problem
record, recorder | 3 | loss of record | data back-up problem
record, recorder | 3 | loss of traceability | data issue
record, recorder | 3 | unable to record time | data back-up problem
recover | 3 | fail to recover sensor | failure to reset
recover | 3 | recover delayed | reset issue
rectifier | 3
rectify, rectification, rectifications | 3 | fail to rectify situation | use of device issue
rectify, rectification, rectifications | 3 | pulse cancellation | electronic property issue
rectify, rectification, rectifications | 3 | unintended stimulation | Unexpected Therapeutic Results
reed | 3 | fail to electrically open | difficult to open or close
refill | 3 | inability to refill medicine | filling problem
refill |  | unable to refill | filling problem
reflux | 3
register, registration | 2 | incorrect product registration | compliance issue
register, registration |  | fail to obtain registration | compliance issue
register, registration | 3 | register inappropriately | compliance issue
regulate, regulation, regulator, regulatory | 2 | fail to inform regulator on change | Communication or Transmission Issue
regulate, regulation, regulator, regulatory | 2 | fail to inform regulator on problem | Communication or Transmission Issue
regulate, regulation, regulator, regulatory |  | fail to regulate flow to sesired level | improper flow or infusion
regulate, regulation, regulator, regulatory |  | fail to obtain regulatory clerance | compliance issue
regulate, regulation, regulator, regulatory | 2 | regulation failed | compliance issue | may be replaced by "position"
rehabilitation |  | device malfunction | device operates differently than expected
rehabilitation | 2 | rehabilitation delayed | reset issue
reimburse, reimbursement | 3 | loss of reimbursement | Market Issue
relay | 3 | reduced life cycle | shelf life exceeded
release, unrelease | 3 | fail to release | device operates differently than expected
release, unrelease | 3 | inappropriate release | device operates differently than expected
release, unrelease | 3 | lack of local language | Instruction for Use Issue; inadequate user interface
release, unrelease | 3 | release delayed | device operates differently than expected
release, unrelease | 3 | release out of control | device operates differently than expected
reliability | 2 | low confidence level | material integrity issue
reliability | 2 | low reliability level | material integrity issue
remove, removal, removable | 3 | difficult to remove | difficult to remove; defective component
remove, removal, removable | 3 | excessive removal | device handling issue
remove, removal, removable | 3 | unable onehand to remove | difficult to remove
remove, removal, removable | 3 | remove force out of specification | difficult to remove
repeat, repeatability | 2 | repeatability out of specification | device operates differently than expected
repeater | 3
replace, replacement, replaceable | 2 | early device replacement | material integrity issue
replace, replacement, replaceable | 3 | fail to replace component | device maintenance issue
replace, replacement, replaceable | 3 | fail to replace manually | device maintenance issue
replace, replacement, replaceable | 3 | missing replacement tool | component missing
replace, replacement, replaceable | 3 | replacement time long | device difficult to setup or prepare
replace, replacement, replaceable | 3 | require special tool undesired | device difficult to setup or prepare
reposition, repositionable | 2 | fail to reposition | positioning issue
reposition, repositionable | 2 | inefficient reposition | difficult to position
reposition, repositionable | 2 | reposition relied on power | positioning issue; loss of power
reposition, repositionable | 2 | reposition slow | positioning issue
reprocess | 3 | incorrect reprocess | improper or incorrect procedure or method
reproducibility |  | reproducibility out of specification | device operates differently than expected
request | 3 | fail to cancel request | use of device issue
request | 3 | fail to generate periodic request | use of device issue
request | 3 | fail to request | use of device issue
require, requirement | 2 | ambiguity requirement | use of device issue
require, requirement | 2 | contradicting requirement | use of device issue
require, requirement | 2 | duplicated requirement | use of device issue
require, requirement | 2 | fail to establish test criteria | incorrect or inadequate test results
require, requirement | 2 | inadequate requirement | use of device issue
require, requirement | 2 | lack interface requirement | human-device interface issue
require, requirement | 2 | lack of acceptance requirement | use of device issue
require, requirement | 2 | lack of alarm or warning or operator message requirement | device alarm system issue
require, requirement | 2 | lack of capability requirement | use of device issue
require, requirement | 2 | lack of data definition and database requirement | data issue
require, requirement | 2 | lack of functional requirement | use of device issue
require, requirement | 2 | lack of input and output requirement | use of device issue
require, requirement | 2 | lack of installation requirement | installation-related problem
require, requirement | 2 | lack of IT-network requirement | application network issue
require, requirement | 2 | lack of maintenance requirement | lack of maintenance documentation or guidelines
require, requirement | 2 | lack of operation requirement | use of device issue
require, requirement | 2 | lack of regulatory requirement | compliance issue
require, requirement | 2 | lack of security requirement | computer system security issue
require, requirement | 2 | lack of traceability | data issue
require, requirement | 2 | misalignment with the other requirement | use of device issue
reservoir | 3 | bubbles in reservoir | improper gas output
reservoir | 3 | empty reservoir | infusion or flow issue
reservoir | 3 | fail to detect reservoir | failure to sense
reservoir | 3 | fail to reset reservoir | failure to reset
reservoir | 3 | fail to update time remain | imprecision
reservoir | 3 | fail to update volume remain | Failure to Service
reservoir | 3 | flow back reservoir | improper flow or infusion; leak
reservoir | 3 | inaccuracy time remain | imprecision
reservoir | 3 | inaccuracy volume remain | volume accuracy issue
reservoir | 3 | incorrect time remain | imprecision
reservoir | 3 | incorrect volume calculation | parameter calculation error due to software problem
reservoir | 3 | incorrect volume remain | improper device output
reservoir | 3 | low reservoir | infusion or flow issue
reservoir | 3 | reservoir leakage | leak
reset | 3 | fail to reset | failure to reset
reset | 3 | fail to reset sensor | failure to reset
reset | 3 | inappropriate reset | reset issue
reset | 3 | unexpected reset | reset issue
reside | 3 | reside out of area | positioning issue
residue, residual, manufacturing residue | 2 | residue after decontamination | Residue after Decontamination
residue, residual, manufacturing residue |  | foreign matter | Particulates
residue, residual, manufacturing residue |  | fluid residue | Particulates
residue, residual, manufacturing residue | 2 | tissue irritation | Disinfection or Sterilization Issue at User Location
resistance | 3 | high resistance | physical resistance
resistance |  | increased resistance | physical resistance
resistance |  | resistance out of spec | physical resistance
resistance |  | fail to disclose resistance increasion | instruction for use issue
resistance | 3 | loss of resistance | physical resistance
resistance |  | low resistance accuracy | incorrect measurement; imprecision
resistance |  | inadeqaute resistance measurement | physical resistance
resistance |  | inadeqaute temperature resistance | temperature issue
resistance | 3 | resistance inadequate | physical resistance
resistor | 3
resolution | 3 | low motion resolution | imprecision
resolution | 3 | low resolution | Image Resolution Poor
resolution | 3 | low visual resolution | Image Resolution Poor
resolution |  | fail to detect low resolution | Image Resolution Poor
resolution | 3 | resolution out of specification | Image Resolution Poor
resolver | 3
resonator | 3
respiratory |  | respiratory distress | no known device problem
resterilize, resterilizatation | 2 | resterilization inefficacy | Failure to Sterilize
restore |  | fail to restore previous adjustments | use of incorrect control settings; reset issue
restore | 3 | fail to restore factory defaults | use of incorrect control settings; reset issue
restrict, restriction | 2 | movement unrestricted | unintended movement
restrict, restriction | 2 | patient not compliant with usage restriction | improper or incorrect procedure or method
resume | 3 | fail to resume | reset issue
resume | 3 | resume time long | reset issue
resuscitator |  | device works intermittently | Device Stops Intermittently
retain |  | detachment of device component | detachment of device component
retain |  | fail to retain force | device handling issue
retain | 3 | fail to retain insertion | device handling issue
retain |  | fail to retain position | defective component; device handling issue
retain | 3 | fail to retain prescription | device handling issue
retain | 3 | fail to retain settings | device handling issue
retain |  | retain excessively | device handling issue
retainer | 3 | fail to retain position | device handling issue
reticle | 3
retract, retraction, retractable |  | excessive force transferred | device handling issue
retract, retraction, retractable | 2 | insufficient retraction | Retraction Problem
retract, retraction, retractable | 2 | loss of retraction | Retraction Problem
retract, retraction, retractable | 2 | not fully retracted during implant | Retraction Problem
retract, retraction, retractable | 2 | retraction function compromised | Retraction Problem
reuse, re-use, re-usable | 3 | fail to reuse | use of device issue
reuse, re-use, re-usable |  | cross contamination | bacterial contamination of device
reuse, re-use, re-usable |  | missing reuse label | device markings issue
reuse, re-use, re-usable |  | cross infection | bacterial contamination of device
reuse, re-use, re-usable |  | inappropriate reuse | material integrity issue; device disinfection or sterilization issue
revalidate, revalidation | 3 | revalidation failed | protective measure issue
review | 3 | lack of review function | inadequate user interface
review | 3 | reviewed inappropriately | protective measure issue; data issue
rewind | 3 | fail to drive motor in reverse | failure to deploy
rewind | 3 | fail to rewind in reverse | mechanical issue
rewind | 3 | fail to rewind pump | Failure to Pump
rewind | 3 | fail to stop delivery | delivery system failure
rewind | 3 | fail to zero encoder count | Failure to Zero
rework | 3 | rework damage | break
RF, radio frequency | 2 | RF field induced | Radiofrequency Interference (RFI)
RF, radio frequency | 2 | RF interference | Radiofrequency Interference (RFI)
RF, radio frequency | 2 | RF module short | electrical shorting; spark
RF, radio frequency | 2 | RF-induced rectification | Radiofrequency Interference (RFI)
RF, radio frequency | 2 | unintended stimulation | Unexpected Therapeutic Results
rheostat | 3
rhythm | 1 | fail to convert rhythm | Failure to Convert Rhythm
rhythm | 1 | fail to diagnose rhythm | therapeutic or diagnostic output failure; failure to analyze rhythm
rhythm | 1 | fail to override rhythm | Failure to Convert Rhythm
rhythm | 1 | irregular rhythm | Failure to Convert Rhythm
rhythm | 1 | rhythm degrade | Failure to Convert Rhythm; degraded
rhythm | 1 | transient rhythm disturbance | Failure to Convert Rhythm
ribbon | 3
rigid, rigidity | 3 | low rigidity body | material integrity issue
rigid, rigidity | 3 | rigidity out of specification | material rigid or stiff
ring | 3 | loose ring | detachment of device component
ringnet | 2 | loss of communication | Communication or Transmission Issue
rinse | 3 | fail to rinse with sodium hydroxide | device cleaning issue
rinse | 3 | liquid ingress | fluid leak
risk | 2 | fail to analyze risk | Fail-Safe Design Failure
risk | 2 | fail to control risk | no fail-safe mechanism
risk | 2 | fail to evaluate risk | Fail-Safe Design Failure
risk | 2 | fail to identify risk | Fail-Safe Design Failure
risk | 2 | fail to reduce risk | Fail-Safe Design Failure
risk | 2 | lead to new risk | no fail-safe mechanism
risk | 2 | risk from use error | use of device issue
rivet | 3
robot | 3
rod | 3
rohs | 2 | RoHS violated | compliance issue; environmental control or utility issue
ROM, EPROM, read only memory, erasable programmable read only memory | 2 | inappropriate memory size | data back-up problem
ROM, EPROM, read only memory, erasable programmable read only memory | 2 | incorrect revision | Operating System Version or Upgrade Problem
rotate, rotation, rotational | 3 | low rotational stiffness | material rigid or stiff
rotate, rotation, rotational |  | invisible rotation center | incorrect display
rotate, rotation, rotational | 3 | rotation accuracy out of range | imprecision
rotate, rotation, rotational | 3 | rotation failed | defective item; application program issue
rotate, rotation, rotational |  | unable to rotate image | application program issue
rotate, rotation, rotational |  | unable to rotate 3D view | application program issue
rotate, rotation, rotational | 3 | rotation out of range | defective item
route | 3 | missing route | component missing
route |  | Wrong route | misconnection
route |  | fail to route | device alarm system issue
route | 3 | over route | connection issue
route | 3 | router misaligned | failure to align
route | 3 | under route | connection issue
saline | 3
sample | 3 | fail to obtain sample | Failure to Obtain Samples
sample | 3 | incorrect sample size | use of incorrect control settings
scale | 3 | inappropriate scale set | component or accessory incompatibility
scaler | 3
scanner, scan |  | scan slow | device operates differently than expected
scanner, scan | 3 | scan without contrast | poor quality image
scenario | 3 | fail to identify hazard related use scenario | use of device issue
scenario | 3 | inadequate hazard related use scenario | use of device issue
scenario | 3 | inappropriate use scenario selected | use of device issue
scenario | 3 | lack of hazard related use scenario | use of device issue
scenario | 3 | lack of high severity use scenario | use of device issue
scenario | 3 | lack of task | use of device issue
scenario | 3 | lack of task sequence | improper or incorrect procedure or method
sciatica | 2 | pain treatment failed | therapeutic or diagnostic output failure
scrambler | 3
scratch, scrap, scrapped | 2 | coating scratched | Scratched Material
scratch, scrap, scrapped | 2 | loss of electrical connection | intermittent continuity
scratch, scrap, scrapped | 2 | part scratched | Scratched Material; defective component
scratch, scrap, scrapped | 2 | scratches too long | Scratched Material
scratch, scrap, scrapped | 2 | too many scratches | Scratched Material
screen | 3 | blank screen | incorrect display
screen |  | frozen screen | incorrect display
screen |  | fail to freeze screen | incorrect display
screen | 3 | incorrect message display | incorrect display
screen |  | incorrect screen display | incorrect display
screen |  | lack of locked state indication | incorrect display
screen |  | fail to unlock screen | no display or display failure
screen | 3 | unable to lock screen | incorrect display
screw | 3 | screw deformed | material twisted; material distortion
screw |  | low screw strength | mechanical issue
screw | 3 | inappropriate fit | fitting problem
screw | 3 | incorrect screw | component incompatible
screw | 3 | loose screw | detachment of device component
screw | 3 | screw broken or damaged | break
screw | 3 | screw dislocated | dislodged or dislocated
screw | 3 | screw shear failed | mechanical issue
screw | 3 | split nut | Split
screw | 3 | screw stripped | detachment of device component; loose or intermittent connection
scriber | 3
seal | 2 | fail to maintain seal | failure to adhere or bond; bacterial contamination of device
seal |  | fail to seal | failure to adhere or bond; bacterial contamination of device
seal |  | inadequate seal strength | failure to adhere or bond; bacterial contamination of device
seal | 2 | incorrect seal | failure to adhere or bond; deflation issue; Misassembled by Users; bacterial contamination of device
seal |  | incomplete seal | failure to adhere or bond; leak; bacterial contamination of device
seal | 2 | insufficient sealing length | failure to adhere or bond; bacterial contamination of device
seal | 2 | leakage | leak; bacterial contamination of device
seal |  | seal not fully seated | failure to adhere or bond; bacterial contamination of device
seal |  | seal shift | failure to adhere or bond; bacterial contamination of device
seal | 2 | seal broken or damaged | failure to adhere or bond; bacterial contamination of device
seal |  | seal degraded | failure to adhere or bond; material integrity issue; bacterial contamination of device
seal |  | seal missing | failure to adhere or bond; bacterial contamination of device
seal |  | seal torn | material torn; failure to adhere or bond; bacterial contamination of device
seal |  | seal discontinuity | failure to adhere or bond; bacterial contamination of device
secure, securement | 2 | loose attachment | detachment of device component
secure, securement | 2 | fail to secure to bed-mount | migration of device or device component
secure, securement | 2 | fail to secure to cart | migration of device or device component
secure, securement | 2 | fail to secure to pendant-mount | migration of device or device component
sediment | 3 | sediment in device or device ingredient | sediment, precipitate or deposit in device or device ingredient
segment, segmentation | 3 | network segmentation failed | application network issue
segment, segmentation |  | fail to view as solid opacity | incorrect display
segment, segmentation |  | fail to view as outline | incorrect display
segment, segmentation |  | fail to hide segmentation | incorrect display
segment, segmentation |  | fail to show segmentation | incorrect display
segment, segmentation | 3 | segment out of range | defective item
segregate, segregation | 2 | fail to identify segregation | failure to sense
segregate, segregation | 2 | ineffective segregation | protective measure issue
select | 3 | incorrect selecting | use of device issue
select |  | incorrect deselecting | use of device issue
select | 3 | selecting confused | inadequate user interface; use of device issue
select |  | fail to select | application program issue
select |  | fail to deselect | application program issue
select | 3 | selection failed | use of device issue
send | 3 | fail to send out software or data | data issue
sense, sensing | 2 | fail to sense | failure to sense
sense, sensing | 2 | high sensing threshold | high sensing threshold
sense, sensing | 2 | inappropriate sensing | Device Sensing Issue
sense, sensing | 2 | intermittent sensing | Device Sensing Issue
sense, sensing | 2 | invalid sensing | invalid sensing
sense, sensing | 2 | loss of A-V synchrony | failure to sense
sense, sensing | 2 | loss of sensing | failure to sense
sense, sensing | 2 | loss of sensing data | loss of data
sense, sensing | 2 | over sensing | Over-sensing
sense, sensing | 2 | sensing deteriorated | Device Sensing Issue
sense, sensing | 2 | under sensing | under-sensing
sensitivity | 2 | decreased sensitivity | decreased sensitivity
sensitivity | 2 | inappropriate sensitivity setting | use of incorrect control settings
sensitivity | 2 | increased sensitivity | Increased Sensitivity
sensitivity | 2 | low sensitivity | decreased sensitivity
sensitivity | 2 | sensitivity setting out of range | decreased sensitivity; Increased Sensitivity
sensor | 2 | fail to recover sensor | failure to reset
sensor |  | fail to convert to back-up | failure to convert to back-up
sensor |  | fail to disable sensor | device remains activated
sensor |  | fail to detect sensor error | failure to sense
sensor |  | fail to maintain resistance | physical resistance
sensor | 2 | inaccurate sensor measurement data | incorrect measurement
sensor | 2 | incorrect sensor orientation | failure to align
sensor | 2 | insensitive sensing | decreased sensitivity; Device Sensing Issue
sensor | 2 | loss of sensor data | loss of data
sensor | 2 | sensor broken or damaged | break; failure to read input signal; imprecision
sensor |  | sensor falling down | slippage of device or device component; component falling
sensor | 2 | sensor data breached | application security issue
sensor |  | out of dimension | defective component
sensor |  | sensor failed | failure to sense | sensor failure
sensor |  | sensor temperature dropped | temperature issue
sensor | 2 | sensor measurement interfered | incorrect measurement
sensor |  | sensor oversized | physical property issue
sensor | 2 | sensor reads high | High Readings
sensor | 2 | sensor reads low | Low Readings
sensor |  | inappropriate sensor type | component incompatible
sensor |  | invalid sensor | component incompatible
sensor | 2 | sensor settings out of range | use of incorrect control settings
separate, separation | 3 | crimp separation | Material Separation
separate, separation | 3 | difficult to separation | Failure to Separate
separate, separation | 3 | loss of electrical continuity | intermittent continuity
separate, separation | 3 | loss of electrical isolation | electrical shorting
separate, separation | 3 | premature separation | Material Separation
separate, separation | 3 | separation delayed | Failure to Separate
separate, separation | 3 | separation of components | Material Separation
separate, separation | 3 | separation of foreign material from weld site | Material Separation
separator | 3 | separator punctured | Material Puncture
separator | 3 | separator torn | torn material
separator | 3 | thick separator | defective component
serialize, serial, serialization | 2 | incorrect serial number | device markings issue
serialize, serial, serialization | 2 | unique number violated | device markings issue
service, serviceability | 3 | fail to service | Failure to Service; Incorrect or Inadequate Result
service, serviceability |  | service required | calibration issue; Incorrect or Inadequate Result
service, serviceability | 3 | inadequate service | Inadequate Service; Incorrect or Inadequate Result
service, serviceability | 3 | service difficult | device maintenance issue
session | 3 | battery depletion | Premature Discharge of Battery
session | 3 | long session | improper or incorrect procedure or method
set, settable, setup, autoset, setting |  | autoset expired | use of incorrect control settings
set, settable, setup, autoset, setting | 3 | clinical setting changed | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | confused settings | data issue
set, settable, setup, autoset, setting |  | ambiguous settings | data issue
set, settable, setup, autoset, setting |  | unclear settings | data issue
set, settable, setup, autoset, setting | 3 | difficult to setup | device difficult to setup or prepare
set, settable, setup, autoset, setting |  | setup delayed | device difficult to setup or prepare
set, settable, setup, autoset, setting | 3 | fail to provide default setting | data issue
set, settable, setup, autoset, setting |  | fail to retrieve default setting | reset issue; data issue
set, settable, setup, autoset, setting | 3 | fail to set value correctly | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | inadequate setting option | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | incorrect setting format | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | incorrect settings | use of incorrect control settings; data issue
set, settable, setup, autoset, setting |  | inaccurate settings | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | invisible settings | no display or display failure; data issue
set, settable, setup, autoset, setting |  | fail to trigger to pre-determined settings | use of incorrect control settings
set, settable, setup, autoset, setting |  | fail to trigger to backup settings | failure to convert to back-up; data issue
set, settable, setup, autoset, setting |  | fail to export settings | failure to convert to back-up; data issue
set, settable, setup, autoset, setting | 3 | lack of setting | use of incorrect control settings; data issue
set, settable, setup, autoset, setting |  | set broken or damaged | break
set, settable, setup, autoset, setting |  | set distorted | Material Distortion
set, settable, setup, autoset, setting |  | set malfunction | Failure to Infuse
set, settable, setup, autoset, setting | 3 | set date and time failed | reset issue; data issue
set, settable, setup, autoset, setting | 3 | set has a quality issue | use of incorrect control settings
set, settable, setup, autoset, setting | 3 | setup incorrectly | use of incorrect control settings
set, settable, setup, autoset, setting | 3 | setup time long | reset issue
set, settable, setup, autoset, setting | 3 | unclear setup procedures | improper or incorrect procedure or method
set, settable, setup, autoset, setting | 3 | undesired settings changed | use of incorrect control settings; data issue
set, settable, setup, autoset, setting | 3 | user forgets to check the settings | use of device issue
sevo, servo, servo control | 3 | device malfunction | device operates differently than expected
shaft | 3 | shaft break | break
shaft | 3 | shaft split | Split
shape | 3 | inadequacy shape | defective item
sharp | 2 | insulation broken or damaged | electrical shorting
sharp | 2 | sharp edge | Material Puncture
sharp |  | sharp pinch point | Material Puncture
sharp | 2 | tissue or vein damage | mechanical issue
sheath | 2 | sheath surface compromised | Scratched Material
shelf, shelves | 3 | inadequate size | size incorrect for patient
shield | 2 | inadequate shielding | Shielding Failure
shield | 2 | shielding failure | Shielding Failure
shifter | 3
ship, shipment, shipping | 3 | case is physically damaged during transport | Shipping Damage or Problem
ship, shipment, shipping | 3 | contents not held securely during transport | Manufacturing or Shipping Issue Associated with Device
ship, shipment, shipping | 3 | inappropriate shipping environment or condition | environmental control or utility issue; Shipping Damage or Problem
ship, shipment, shipping | 3 | shipping damage | Shipping Damage or Problem
ship, shipment, shipping | 3 | wear debris from shipping tray | metal shedding debris
shock | 2 | device broken or damaged | device damaged prior to use
shock | 2 | device malfunction | device operates differently than expected
shock | 2 | electrical shock | Inappropriate Shock
shock | 2 | excessive leakage current | electrical issue
shock | 2 | fail to deliver shock | Failure to Shock or Properly Shock
shock | 2 | fail to suspend inappropriate shock | Inappropriate Shock
shock | 2 | failure to advise shock | Failure to Shock or Properly Shock
shock | 2 | inappropriate shock | Inappropriate Shock
shock | 2 | inappropriate shock delivery schedule | Inappropriate Shock; improper or incorrect procedure or method
shock | 2 | intermittent shock | Intermittent Shock
shock | 2 | shock to clinician | Inappropriate Shock
shock | 2 | shock to wrong place | Inappropriate Shock
shock | 2 | shock vector changed | Inappropriate Shock
shock | 2 | unintended electrical shock | Inappropriate Shock
shock | 2 | unnecessary shock | Inappropriate Shock
shrinkage | 3 | shrinkage out of specification | defective component
shutter | 3 | shutter failure | defective component
sideport | 3 | fail to introduce pipette | Failure to Advance
sideport | 3 | fail to introduce refill needle | filling problem
sigh | 3 | high peak flow | improper flow or infusion
sigh | 3 | high volume | output above specifications
sigh | 3 | inappropriate sigh delivery | delivery system failure
sigh | 3 | lack of sigh | failure to deliver
sigh | 3 | low volume | Output below Specifications
signal, signage | 2 | fail to analyze signal | failure to read input signal
signal, signage | 2 | fail to read input signal | failure to read input signal
signal, signage | 2 | fail to select signal | Failure to Select Signal
signal, signage | 2 | inaccurate frequency | electronic property issue
signal, signage |  | speed signal out of range | electronic property issue
signal, signage | 2 | incorrect interpretation of signal | failure to read input signal
signal, signage | 2 | invalid signal | failure to read input signal; improper device output
signal, signage | 2 | no signal | no device output
signal, signage | 2 | noise signal | failure to read input signal; improper device output
signal, signage | 2 | signal artifact | artifact
signal, signage | 2 | signal deteriorated | failure to read input signal; improper device output
signal, signage | 2 | signal wandering | failure to read input signal
size | 3 | inadequate size | size incorrect for patient
size | 3 | inadequate size range to meet market demand | size incorrect for patient
size |  | inappropriate part size | defective component
size |  | inconsistent size | size incorrect for patient
size | 3 | mis-sized implant | size incorrect for patient
size | 3 | out of dimension | defective component
size |  | fail to define step size | use of incorrect control settings
size | 3 | size deviation | size incorrect for patient
skin | 1 | clean skin insufficiently | device cleaning issue
skin | 1 | skin allergic reaction | biocompatibility issue
skin | 1 | skin scratched | device handling issue; human-device interface issue
skip | 3 | skip failed | use of device issue
sleep | 3 | unusual sleep schedule | improper or incorrect procedure or method
slide | 3 | low force provided | mechanical issue
slide | 3 | fail to drag slider | human-device interface issue; application program issue
smoke, smoking | 2 | device burned | smoking
snare
socket | 3 | out of dimension | defective component
socket | 3 | socket mismatches to plug | fitting problem; component incompatible
software, SaMD, software as a medical device | 3 | date or time-related software problem | Date-related software issue
software, SaMD, software as a medical device | 3 | fail to update software | application program version or upgrade problem
software, SaMD, software as a medical device |  | fail to deploy software on hardware | computer software issue
software, SaMD, software as a medical device | 3 | inappropriate software release | application program version or upgrade problem
software, SaMD, software as a medical device | 3 | incompatible software versions | application program version or upgrade problem; Operating System Version or Upgrade Problem
software, SaMD, software as a medical device | 3 | incorrect software reference number | application program version or upgrade problem; device markings issue
software, SaMD, software as a medical device | 3 | incorrect software settings | use of incorrect control settings
software, SaMD, software as a medical device |  | incorrect position calculation | computer software issue
software, SaMD, software as a medical device | 3 | lack of software version | application program version or upgrade problem
software, SaMD, software as a medical device | 3 | loss of software application | application program issue
software, SaMD, software as a medical device |  | short software life cycle | computer software issue
software, SaMD, software as a medical device |  | fail to compatible with operating system | Incompatibility Problem
software, SaMD, software as a medical device | 3 | obsolete software used | application program version or upgrade problem
software, SaMD, software as a medical device | 3 | software anomaly | computer software issue
software, SaMD, software as a medical device |  | software defect | therapeutic or diagnostic output failure
software, SaMD, software as a medical device |  | software corrupted | computer software issue
software, SaMD, software as a medical device | 3 | software can not recover from reset | reset issue
software, SaMD, software as a medical device | 3 | software installation problem | problem with software installation
software, SaMD, software as a medical device |  | software download failed | problem with software installation
software, SaMD, software as a medical device |  | software task fault displayed | device displays error message
software, SaMD, software as a medical device | 3 | software version incorrect | application program version or upgrade problem
solder | 3 | excess solder | Solder Joint Failure
solder | 3 | inappropriate soldering | Solder Joint Failure
solder | 3 | insufficient solder | Solder Joint Failure
solder | 3 | reflowed connector solder | improper or incorrect procedure or method
solder | 3 | solder ball | Solder Joint Failure
solder | 3 | solder bridge | electrical shorting
solder | 3 | solder broken or damaged | Solder Joint Failure
solder | 3 | solder cracked | Solder Joint Failure
solder | 3 | solder joint broken or damaged | Solder Joint Failure; break
solenoid, electromagnetic solenoid | 2 | solenoid failed to energize | failure to deliver energy; output issue
solenoid, electromagnetic solenoid | 2 | solenoid stuck in energized position | unintended energization
solvent |  | excess solvent | volume accuracy issue
solvent |  | insufficient solvent | volume accuracy issue
solvent | 3 | incompatible solvent | failure to adhere or bond
sort | 2 | fail to sort alarm message | device alarm system issue
sort |  | fail to sort patient interest list | application program issue
sort |  | inappropriate sort order | data issue; application program issue
sort | 2 | fail to sort data | data issue
sound | 3 | fail to sound audible tone | use of device issue; not audible alarm
sound |  | fail to generate sound | electronic property issue
sound | 3 | incorrect sound pattern | use of incorrect control settings; electronic property issue
sound | 3 | indistinct sound | improper device output; electronic property issue
sound | 3 | sound mute failed | use of device issue
sound | 3 | sound setting failed | use of incorrect control settings
sound | 3 | synchronization sound mute failed | use of device issue
SOUP, software of unknown provenance | 3 | fail to identify hazardous situation | computer software issue
SOUP, software of unknown provenance | 3 | lack of bug fix procedure | computer software issue
SOUP, software of unknown provenance | 3 | lack of functional requirement | computer software issue
SOUP, software of unknown provenance | 3 | lack of hardware requirement | use of device issue
SOUP, software of unknown provenance | 3 | lack of obsolescence procedure | application program version or upgrade problem
SOUP, software of unknown provenance | 3 | lack of patch implement procedure | computer software issue
SOUP, software of unknown provenance | 3 | lack of performance requirement | computer software issue
SOUP, software of unknown provenance | 3 | lack of software requirement | computer software issue
SOUP, software of unknown provenance | 3 | lack of upgrade procedure | application program version or upgrade problem
source | 3 | incorrect source | use of device issue; medical gas supply problem
source | 3 | source detachment | detachment of device component
source | 3 | source leak | leak; medical gas supply problem
source | 3 | source unrestricted | use of device issue; medical gas supply problem
space | 2 | liquid ingress | fluid leak
space | 2 | short circuit | electrical shorting; spark
space | 2 | space gap out of specification | device misassembled during manufacturing or shipping
spacer | 3
spark | 3
speaker, buzzer, magnetic speaker | 3 | alarm audio stuck off | device alarm system issue
speaker, buzzer, magnetic speaker | 3 | alarm audio stuck on | device alarm system issue
speaker, buzzer, magnetic speaker | 3 | frequency out of specification | electronic property issue
speaker, buzzer, magnetic speaker | 3 | inadequate speaker; low volume speaker | Inaudible Voice Prompts
speaker, buzzer, magnetic speaker | 3 | incorrect sound pattern | use of incorrect control settings; electronic property issue
speaker, buzzer, magnetic speaker | 3 | loss of audio notification | device alarm system issue
speaker, buzzer, magnetic speaker | 3 | low audio volume | Inaudible Voice Prompts
speaker, buzzer, magnetic speaker | 3 | missing speaker | component missing
speaker, buzzer, magnetic speaker |  | speaker broken or damaged | break
speaker, buzzer, magnetic speaker | 3 | sound pressure out of specification | pressure issue
speaker, buzzer, magnetic speaker | 3 | volume out of specification | use of device issue
specification | 3 | absence of essential performance | physical property issue
specification | 3 | inadequate specification | physical property issue
specification |  | ambiguous specification | physical property issue
specification |  | untestable specification | physical property issue
specification |  | conflicting specification | physical property issue
specification | 2 | lack of architecture specification | physical property issue
specification | 3 | lack of specification | physical property issue
specification |  | fail to disclose specification | instruction for use issue
specification | 3 | obsolete specification | physical property issue
speed |  | inconsistent speed | imprecision
speed | 3 | speed out of range | imprecision
spill, spillage | 2 | accidental spillage | chemical spillage
splicer | 3
splint | 3 | adjustment difficult | device difficult to setup or prepare
splint | 3 | difficult to remove | difficult to remove
splint | 3 | inadequate size | size incorrect for patient
splint | 3 | lack of ductile | material rigid or stiff
splint | 3 | low stiffness | material rigid or stiff
splint | 3 | missing soft goods | component missing
splint | 3 | missing splint | component missing
splint | 3 | over bending | Bent
splint | 3 | splint dislodgement | dislodged or dislocated
splint | 3 | user confused | inadequate user interface
splint | 3 | user uncomfortable | human-device interface issue
spooler | 3
spring |  | spring missing | component missing
spring |  | spring broken or damaged | break
spring |  | spring torn | torn material
spring |  | Stored Energy | mechanical issue
spring |  | length too large | defective component
spring | 3 | spring malfunction | defective component; Product Quality Issue
stability | 2 | downgrade stability | unstable
stability | 2 | inappropriate stability change | unstable
stability | 2 | insufficient stability | unstable
stabilizer | 3 | loss of function | device inoperable
stack | 3 | excessive stack | unstable; Packaging Issue
stand | 3
staple | 3 | fail to form staple | Failure to form staple
staple | 3
stator | 3 | lead wire solder failed | Solder Joint Failure
stator | 3 | solder joint failed | Solder Joint Failure
stator | 3 | winding open failed | difficult to open or close
stator | 3 | winding short | defective component
stator | 3 | winding system overheat | overheating of device or device component
steam |  | inadequate steam | Residue after Decontamination
stent | 2 | improper stent implanted | Incompatibility Problem
stent |  | invisible under fluoroscopy | poor quality image
sterilize, sterility, sterile, sterilization |  | component cracked | Crack
sterilize, sterility, sterile, sterilization |  | clinical workflow interrupted | use of device issue; disinfection or sterilization issue at user location
sterilize, sterility, sterile, sterilization | 1 | delivered as unsterile product | Delivered As Unsterile Product
sterilize, sterility, sterile, sterilization |  | fail to reduce viable microorganism | Failure to Sterilize
sterilize, sterility, sterile, sterilization |  | false positive | Delivered As Unsterile Product
sterilize, sterility, sterile, sterilization |  | Inappropriate sterility assurance level (SAL) | Delivered As Unsterile Product
sterilize, sterility, sterile, sterilization | 1 | inappropriate sterilization | Failure to Sterilize
sterilize, sterility, sterile, sterilization |  | Inadequate Sterilization | Delivered As Unsterile Product; Failure to Sterilize
sterilize, sterility, sterile, sterilization | 1 | loss of sterility | Delivered As Unsterile Product; Failure to Sterilize
sterilize, sterility, sterile, sterilization |  | material discolored | Material Discolored
sterilize, sterility, sterile, sterilization | 1 | sterility compromising | Delivered As Unsterile Product; Failure to Sterilize
sterilize, sterility, sterile, sterilization | 1 | sterilization cycle failed | device disinfection or sterilization issue
sterilize, sterility, sterile, sterilization | 1 | sterilization inefficacy | Failure to Sterilize
stick | 2 | less sticky | detachment of device component
stiff, stiffness, stiffener | 2 | fatigue fracture | fracture; mechanical issue
stiff, stiffness, stiffener | 2 | perforated atrium | Material Perforation
stiff, stiffness, stiffener | 2 | physician inconvenience | inadequate user interface
stiff, stiffness, stiffener | 2 | punctured blood vessel or valve | Material Puncture; use of device issue
stiff, stiffness, stiffener |  | stiffness increased | material rigid or stiff; disconnection
stiff, stiffness, stiffener |  | stiffness too high | material rigid or stiff; disconnection
stiff, stiffness, stiffener | 2 | stiffness too low | material rigid or stiff
stiff, stiffness, stiffener | 2 | unacceptable tip stiffness | material rigid or stiff; disconnection
stimulate, stimulation | 2 | gradient induced stimulation | electronic property issue
stimulate, stimulation | 2 | high stimulation rate | electronic property issue
stimulate, stimulation | 2 | stimulate on undesired muscle | Therapy delivered to incorrect body area
stimulate, stimulation | 2 | stimulation fails to stop | Unexpected Therapeutic Results
stimulate, stimulation | 2 | unintended diaphragmatic stimulation | Unexpected Therapeutic Results
stimulate, stimulation | 2 | unintended nerve stimulation | Unexpected Therapeutic Results
stimulate, stimulation | 2 | unintended phrenic nerve stimulation | Unexpected Therapeutic Results
stimulate, stimulation | 2 | unintended stimulation | Unexpected Therapeutic Results
stitch | 3 | unraveled stitching | Tear, Rip or Hole in Device Packaging
stop | 2 | stop failed | Failure to Auto Stop
stop | 2 | stop intermittently | Device Stops Intermittently
stop | 2 | stop out of range | defective item
stop | 2 | stopping distance over limit | Fail-Safe Mechanism Issue
stopcock, stopper | 3 | air ingress | air leak
stopcock, stopper |  | stopper puckered | leak
stopcock, stopper | 3 | fail to retain position | device handling issue
stopcock, stopper | 3 | inadequate seal | failure to adhere or bond
stopcock, stopper | 3 | stockcock released inadvertently | device operates differently than expected
stopcock, stopper | 3 | stockcock slipped inadvertently | slippage of device or device component
store, storage, network attached storage | 3 | fail to store data | data issue
store, storage, network attached storage |  | storage hardware failure | device inoperable
store, storage, network attached storage | 3 | inadequate storage | inadequate storage
store, storage, network attached storage | 3 | inappropriate storage environment or condition | environmental control or utility issue
store, storage, network attached storage | 3 | insufficient storage location | inadequate storage
store, storage, network attached storage | 3 | insufficient storage space | inadequate storage
store, storage, network attached storage |  | loss of patient data | loss of data; patient data issue
store, storage, network attached storage | 3 | stored inappropriately | device handling issue
strain | 2 | strain relief | defective item
strap | 3 | adding stability failed | unstable
strap | 3 | adjustment difficult | device difficult to setup or prepare
strap | 3 | fail to avoid inversion movement | unintended movement
strap | 3 | modular support failed | use of device issue
strap | 3 | position adjustment difficult | difficult to position
strap | 3 | strap material torn | torn material
strap | 3 | strap shift | migration of device or device component
strap | 3 | tightening adjustment difficult | device difficult to setup or prepare
strength | 2 | dissociation strength | mechanical issue
strength | 2 | lack of adequate static or fatigue strength | mechanical issue
strength | 2 | less strength | mechanical issue
stress | 2 | deformation | mechanical issue
stress | 2 | part fatigue fracture | fracture; mechanical issue
stress | 2 | loss of electrical continuity | intermittent continuity
stress |  | Residual Stress | mechanical issue
stress | 2 | over stress | mechanical issue
stretch |  | overstretched | leak
stretch |  | overstretched | Stretched
strut | 3 | strut fractured | fracture; mechanical issue
stylet | 3 | incompatible stylet | component or accessory incompatibility
stylet | 3 | stylet too soft | material rigid or stiff
subassembly | 3
substrate | 3
subthreshold | 2 | fail to capture subthreshold | Failure to Capture
subthreshold | 2 | fail to identify subthreshold | Loss of Threshold
subthreshold | 2 | subthreshold increased | capturing issue
suction | 3 | decrease in suction | Decrease in Suction
suction | 2 | inappropriate suction change | Decrease in Suction; Increase in Suction
suction | 2 | inappropriate suction setting | use of incorrect control settings
suction | 2 | inappropriate suction value | suction issue
suction | 2 | lack of low suction alarm | device alarm system issue
suction | 2 | loss of suction | suction issue
suction | 2 | loss of vacuum | suction issue
suction | 2 | low suction | Decrease in Suction
suction | 2 | over suction | Increase in Suction
suction | 2 | suction control failed | suction issue
suction | 2 | suction out of range | suction issue
suit, suitable | 3 | fail to suit | Incompatibility Problem
sunlamp | 3
support | 3 | incapable support | use of device issue
support | 3 | support compromised | use of device issue
support | 3 | support failed | use of device issue
suppressor | 3
surgical, surgery | 2 | surgery extended | use of device issue
surgical, surgery |  | incorrect surgical decision | image display error; use of device issue; application program issue
surgical, surgery | 2 | surgical contamination | contamination during use
suture | 2 | insufficient suture sleeve tie-down force | mechanical issue
suture | 2 | loss of suture | Detachment of Device or Device Component
suture | 2 | loss of tie down strength | mechanical issue
suture | 2 | suture breakage | break
suture | 2 | suture line disruption | Detachment of Device or Device Component
suture | 2 | suture line separation | Detachment of Device or Device Component
suture | 2 | tight suture sleeve | size incorrect for patient
switch | 3 | fail to manage switch | use of device issue
switch |  | fail to toggle visibility | incorrect display
switch | 3 | incorrect incorporated | defective component
switch | 3 | lack of debounce button design | inadequate user interface
switch | 3 | switch back and forth | improper or incorrect procedure or method
switch |  | switch broken or damaged | break
switch | 3 | switch invoked inappropriately | improper or incorrect procedure or method
switch |  | switch on/off issue | defective component; incorrect display
switch | 3 | unable to switch | use of device issue
switch |  | unqualified switch | electrical issue; defective component
switch | 3 | user switching slow | use of device issue
swivel | 3 | low rotational stiffness | material rigid or stiff
swivel | 3 | rotation accuracy out of range | imprecision
swivel | 3 | rotation failed | defective item
swivel | 3 | rotation out of range | defective item
symbol, symbology | 3 | inappropriate symbol | device markings issue
symbol, symbology | 3 | missing symbol | device markings issue
symbol, symbology | 3 | symbol loss of explanation | device markings issue; instruction for use issue
symbol, symbology |  | unstandardized symbol | device markings issue
symbol, symbology | 3 | unclear symbol | device markings issue
synchronize, synchronization, synchronizer | 2 | inaccurate synchronization | Inaccurate Synchronization
synthesizer | 3
syringe | 2 | hole in syringe | hole in material
syringe |  | air in syringe | infusion or flow issue
syringe | 2 | inaccurate syringe marking | device markings issue
syringe |  | check syringe alarm | device alarm system issue
syringe |  | syringe leakage | fluid leak
syringe | 2 | lack of locking ring | component missing
syringe | 2 | syringe broken or damaged | device damaged prior to use
syringe | 2 | syringe does not match with administration sets | component or accessory incompatibility
syringe | 2 | user uses wrong syringe | user used incorrect product for intended use
system, subsystem | 3 | fail to activate system | failure to deploy
system, subsystem |  | inadequate system design | component missing
system, subsystem |  | integration failed | Material Disintegration; incompatibility problem
system, subsystem | 3 | firmware error | problem with software installation
system, subsystem | 3 | hardware error | defective component
system, subsystem |  | system fault displayed | device displays error message
system, subsystem |  | system locked up | device inoperable; application security issue
system, subsystem |  | system frozen | device inoperable; application program issue
system, subsystem |  | unintended system shut down | application program issue
system, subsystem |  | fail to boot up | application program issue
system, subsystem | 3 | unintended system motion | unintended system motion
table | 3 | table break | break
table | 3 | table crushed | collapse
table | 3 | table incrementation incorrect | use of device issue
table | 3 | table tilting | use of device issue
table | 3 | unintended table top motion | unintended system motion
tachycardia | 2 | reentrant ventricular tachycardia | failure to pace or properly pace
talus | 3 | undesired talus movement | unintended movement
tamper | 2 | data abused | application security issue
tamper | 2 | exposure to water | environmental particulates; biological environmental factor
tamper | 2 | non-qualified personnel access | unauthorized access to computer system
tamper | 2 | part broken or damaged | break; defective component
tamper |  | Susceptibility to Tampering | unauthorized access to computer system
tamper | 2 | unauthorized access | unauthorized access to computer system
tape |  | unavailable tape | component missing
tape | 3 | tape out of specification | defective component
task | 3 | fail to conduct task analysis | use of device issue
task | 3 | incomplete task | use of device issue
task | 3 | lack of task | use of device issue
task | 3 | lack of task sequence | improper or incorrect procedure or method
telemetry | 2 | communication data corrupted | Communication or Transmission Issue
telemetry | 2 | fail to tune telemetry | radio signal problem
telemetry | 2 | loss of communication | Communication or Transmission Issue
telemetry | 2 | loss of telemetry | radio signal problem
telemetry | 2 | no telemetry | radio signal problem
telemetry | 2 | poor telemetry | radio signal problem
telemetry | 2 | telemetry discrepancy | Telemetry Discrepancy
telemetry | 2 | telemetry distance short | Telemetry Discrepancy
telemetry | 2 | telemetry interruption | Telemetry Discrepancy
telemetry | 2 | telemetry message slow | Telemetry Discrepancy
telemetry | 2 | telemetry orientation dependently | Telemetry Discrepancy
telemetry | 2 | unique telemetry ID violated | radio signal problem
telemetry | 2 | weak RF telemetry | radio signal problem
teleoperation | 3 | loss of teleoperation | device inoperable
teleoperation | 3 | undesired teleoperation | device operates differently than expected
telephone | 3
teleprinter | 3
teletypewriter | 3
temperature | 2 | device malfunction | device operates differently than expected
temperature | 2 | high temperature | temperature issue
temperature |  | low temperature | temperature issue
temperature |  | temperature dropped | temperature issue
temperature | 2 | inadequate shutdown temperature | temperature issue
temperature | 2 | internal battery short | electrical shorting; spark
temperature | 2 | premature battery depletion | Premature Discharge of Battery
temperature | 2 | temperature out of range | temperature issue
temperature | 2 | warping or dimensional changes | device misassembled during manufacturing or shipping
tens, transcutaneous electrical nerve stimulation | 2 | loss of therapy | therapeutic or diagnostic output failure
tens, transcutaneous electrical nerve stimulation | 2 | pain relief compromised | therapeutic or diagnostic output failure
tens, transcutaneous electrical nerve stimulation | 2 | pain treatment failed | therapeutic or diagnostic output failure
tens, transcutaneous electrical nerve stimulation | 2 | TENS accessory incompatible | accessory incompatible
tensile, tension | 3 | loss of continuity | failure to disconnect
tensile, tension | 3 | loss of electrical isolation | electrical shorting
tensile, tension |  | Device Tension | Retraction Problem
tensile, tension | 3 | low tensile strength | mechanical issue
terminal | 3 | incorrect to use the earth terminal for others | improper or incorrect procedure or method
terminate, termination, terminator | 3 | accidental disconnection | disconnection
terminate, termination, terminator | 3 | attached or detached by tool | detachment of device component
terminate, termination, terminator | 3 | induce self-terminates | electronic property issue
terminate, termination, terminator |  | long termination time | Failure to Auto Stop
terminate, termination, terminator |  | fail to terminate breath | improper device output; device remains activated
terminate, termination, terminator |  | fail to terminate function | improper device output; device remains activated
terminate, termination, terminator |  | fail to disclose terminate device | instruction for use issue
terminate, termination, terminator |  | low control accuracy | imprecision
terminate, termination, terminator | 3 | missing symbol | device markings issue
test, tester | 2 | fail to abort test | failure of device to self-test
test, tester | 2 | fail to cover requirement | Missing Test Results
test, tester | 2 | fail to initiate test | failure of device to self-test
test, tester |  | fail to pass self-test | failure of device to self-test
test, tester |  | fail to pass test | failure of device to self-test; failure to calibrate; failure to recalibrate
test, tester | 2 | fail to record anomaly | Missing Test Results
test, tester |  | fail to unknown test | No Code Available; no known device problem
test, tester | 2 | inadequate test contents | incorrect or inadequate test results
test, tester | 2 | inadequate test procedure | incorrect or inadequate test results
test, tester | 2 | inadequate test record | incorrect or inadequate test results
test, tester |  | inappropriate test configuration | incorrect or inadequate test results
test, tester | 2 | lack of acceptance criteria | incorrect or inadequate test results
test, tester | 2 | lack of responsible tester | incorrect or inadequate test results
test, tester | 2 | lack of test configuration | incorrect or inadequate test results
test, tester | 2 | lack of test date | Missing Test Results
test, tester | 2 | lack of test result | Missing Test Results
test, tester | 2 | lack of test tool | incorrect or inadequate test results
test, tester | 2 | test timeout | Missing Test Results
therapy, therapeutic | 2 | unexpected therapeutic result | Unexpected Therapeutic Results
therapy, therapeutic |  | inadequate therapy modes | Unexpected Therapeutic Results
therapy, therapeutic | 2 | aborting therapy | failure to deliver
therapy, therapeutic | 2 | fail to deliver therapy | failure to deliver
therapy, therapeutic | 2 | fail to suspend therapy | Overfill
therapy, therapeutic | 2 | inappropriate therapy | therapeutic or diagnostic output failure
therapy, therapeutic |  | therapy interrupted | failure to deliver
therapy, therapeutic | 2 | loss of therapy | failure to deliver
therapy, therapeutic | 2 | therapy delayed | failure to deliver
therapy, therapeutic | 2 | therapy delivered to incorrect body area | Therapy delivered to incorrect body area
therapy, therapeutic | 2 | therapy withhold | failure to deliver
thermal | 2 | device malfunction | device operates differently than expected
thermal | 2 | poor thermal expansion | overheating of device or device component
thermal | 2 | thermal damage | overheating of device or device component
thermal | 2 | thermal decomposition | overheating of device or device component
thermistor | 3
thermocouple | 3
thermostat | 3
thick, thickness | 3 | thickness out of specification | size incorrect for patient; defective component
thread |  | Thread Stripped | peeled
thread |  | thread broken or damaged | break
thread | 3 | thread overriding | disconnection
threader | 3
threshold | 2 | fail to capture threshold | capturing issue
threshold | 2 | fail to identify threshold | capturing issue
threshold | 2 | inappropriate threshold setting | capturing issue
threshold | 2 | loss of threshold | Loss of Threshold
threshold | 2 | threshold increased | High Capture Threshold
threshold | 2 | unstable threshold | unstable Capture Threshold
thyristor | 3
tilt | 3 | tilt out of range | defective item
time | 3 | fail to enter time | use of device issue
time | 3 | inappropriate time format | incorrect display
time | 3 | lack of default time | reset issue
time | 3 | lack of start time | reset issue
time |  | time out reported | application program issue
time | 3 | time out of range | imprecision
time |  | lack of time indication | no display or display failure
time |  | incorrect time display range | incorrect display
time |  | incorrect time configuration | use of incorrect control settings
time |  | incorrect time | reset issue
time |  | low time accuracy | imprecision
time | 3 | timing variation | imprecision
timeout | 3 | inappropriate timeout setting | reset issue
timer, timing | 3 | fail to expire timer | device operates differently than expected
timer, timing | 3 | fail to start timer | failure to reset
timer, timing |  | fail to reset timer | failure to reset
timer, timing | 3 | fail to update timer | device operates differently than expected
timer, timing | 3 | timer malfunction | device operates differently than expected
timer, timing |  | timing malfunction | device operates differently than expected
timestamp, time-stamp | 3 | incomplete timestamp | imprecision
timestamp, time-stamp | 3 | missing timestamp | imprecision
tip |  | tip broken or damaged | break
tip |  | tip detached | detachment of device component
tip |  | tip dulls | Dull
tip |  | tip blunt | Dull
tip |  | tip missing | component missing
tip |  | tip over | tip-over
tissue | 2 | insufficient tissue surface | Energy Output to Patient Tissue Incorrect
tissue |  | excessive charring on tissue | Energy Output to Patient Tissue Incorrect
tissue |  | uncoagulated tissue | Energy Output to Patient Tissue Incorrect
tissue |  | inadequate hemostasis | Energy Output to Patient Tissue Incorrect
tissue |  | excessive muscle stimulation | Energy Output to Patient Tissue Incorrect
tissue |  | incorrect mechanical force | use of device issue; mechanical issue; Energy Output to Patient Tissue Incorrect
tissue |  | fail to specify applied tissue | instruction for use issue; Energy Output to Patient Tissue Incorrect
tissue | 2 | thermal damaged tissue | Energy Output to Patient Tissue Incorrect
tissue | 2 | tissue burned | Energy Output to Patient Tissue Incorrect
tissue | 2 | tissue damaged | Energy Output to Patient Tissue Incorrect
tissue | 2 | tissue perforated | Energy Output to Patient Tissue Incorrect
tissue | 2 | vein damaged | Energy Output to Patient Tissue Incorrect
tolerance | 2 | out of tolerance | defective item
tomographic | 2
tool | 3 | require special tool undesired | device maintenance issue
tooth, teeth |  | teeth shift | no known device problem
torque, torsion | 2 | part deformed | defective component
torque, torsion | 2 | induced rotational torque | Electromagnetic Interference (EMI)
torque, torsion | 2 | insufficient torque | defective item
torque, torsion | 2 | locking fragments or breaks | break; Material Fragmentation
torque, torsion | 2 | torque out of range | device handling issue
touch | 2 | antenna touches other components causing short condition | electrical shorting; spark
touch | 2 | component broken or damaged | break
touch |  | unresponsive touch | failure to deploy
touch | 2 | user injury | use of device issue
trace, traceability | 2 | loss of traceability | use of device issue
trace, traceability | 2 | tracking error | use of device issue
trace, traceability | 2 | tracking volume out of specification | use of device issue
track, tracker | 2 | loss of traceability | use of device issue
track, tracker | 2 | tracking error | use of device issue
track, tracker | 2 | tracking volume out of specification | use of device issue
trackball | 3
train, training | 3 | inadequate training | Inadequate Training
train, training | 3 | inadequate training material | Instruction for Use Issue
train, training | 3 | insufficient training | Inadequate Training
train, training | 3 | lack of training material | Inadequate Training
train, training | 3 | no training | Inadequate Training
train, training | 3 | training overwhelmed | Inadequate Training
train, training | 3 | user does not follow the instructions | improper or incorrect procedure or method
transceiver | 3
transducer | 3 | transducer failure | defective component
transducer | 3 | transducer overheating | overheating of device or device component
transfer |  | difficult to transfer through | device handling issue
transfer |  | excessive force transferred | device handling issue
transfer | 3 | fail to transfer data | data issue
transfer |  | fail to transfer force | device handling issue
transfer |  | fail to transfer torque | device handling issue
transfer |  | transfer through inappropriately | device handling issue
transform | 2 | transform failed | device operates differently than expected
transform |  | voltage out of range | electrical issue
transform | 2 | transform slow | device operates differently than expected
Transformer | 3 | overheating | overheating of device or device component
Transformer | 3 | overload of output winding | Premature Deployment
Transformer | 3 | short circuit | electrical shorting; spark
transistor | 3
transit, transition, transmittance | 3 | inappropriate transition point | use of device issue
transit, transition, transmittance |  | contamination | item contaminated during manufacturing or shipping; shipping damage or problem
transit, transition, transmittance | 3 | low transmittance | use of device issue
transit, transition, transmittance | 3 | transition failed | use of device issue
translate, translation | 3 | low translation resolution | use of device issue
translate, translation | 3 | low translation speed | imprecision
translate, translation | 3 | translation error | use of device issue
translator | 3
transmit, transmission, transmitter | 3 | inappropriate operation | device handling issue
transmit, transmission, transmitter | 2 | incorrect transmission | Communication or Transmission Issue
transmit, transmission, transmitter | 2 | information transmission failed | Communication or Transmission Issue
transmit, transmission, transmitter | 3 | loss of communication | Communication or Transmission Issue
transmit, transmission, transmitter | 3 | loss of function | Communication or Transmission Issue
transmit, transmission, transmitter | 2 | no transmission | Communication or Transmission Issue
transmit, transmission, transmitter | 2 | transmitter failure | defective component
transmit, transmission, transmitter | 3 | unsecured data | application security issue
transponder | 3
transport, transportation | 3 | damage during transport | Shipping Damage or Problem
transport, transportation | 3 | fail to transport | Manufacturing or Shipping Issue Associated with Device; device handling issue
transport, transportation | 3 | inappropriate transport pose | Manufacturing or Shipping Issue Associated with Device; device handling issue
transport, transportation | 3 | inappropriate transportation configuration | Manufacturing or Shipping Issue Associated with Device; device handling issue
transport, transportation | 3 | inappropriate transportation environment or condition | Manufacturing or Shipping Issue Associated with Device; device handling issue
trap | 3 | miss guards or protective measures | no fail-safe mechanism
treatment | 2 | confused treatment option | inadequate user interface; instruction for use issue
treatment | 2 | inappropriate treatment | therapeutic or diagnostic output failure
treatment |  | Ill-Treatment | therapeutic or diagnostic output failure
treatment | 2 | ineffective treatment | therapeutic or diagnostic output failure
treatment | 2 | insufficient treatment definition | therapeutic or diagnostic output failure; instruction for use issue
treatment | 2 | treatment with adjunctive condition | therapeutic or diagnostic output failure; instruction for use issue
trend | 2 | fail to trend data or parameter | data issue; use of incorrect control settings
trend | 2 | inaccurate trend result | incorrect or inadequate test results
trend | 2 | lack of trend analysis | use of device issue
trend | 2 | lead to adverse trend | device unsafe to use in environment
trend | 2 | undetected trend | use of device issue
triac | 3
trigger | 3 | fail to initiate trigger | decreased sensitivity
trigger |  | invalid trigger value | use of incorrect control settings; data issue
trigger |  | invalid trigger mode | use of incorrect control settings
trigger |  | invalid trigger control | use of incorrect control settings
trigger |  | time trigger failed | decreased sensitivity
trigger |  | patient trigger failed | use of device issue
trigger |  | pressure trigger failed | decreased sensitivity; pressure issue
trigger |  | flow trigger failed | decreased sensitivity
trigger |  | volume trigger failed | decreased sensitivity
trigger |  | inappropriate trigger type | use of incorrect control settings
trigger |  | incorrect trigger | decreased sensitivity
trigger | 3 | request trigger failed | decreased sensitivity
trigger | 3 | trigger sensitivity low | decreased sensitivity
trim | 3 | incorrect trim | use of device issue
trim | 3 | incorrect trim value | use of device issue; data issue
trocar | 2 | trocar and fiber incompatibility | component or accessory incompatibility
trocar | 2 | trocar and instrument incompatibility | component or accessory incompatibility
trocar | 2 | trocar failure | defective component
trocar | 3
tube, tubing |  | embedded debris | Particulates
tube, tubing |  | air in line | infusion or flow issue; improper flow or infusion
tube, tubing |  | bubbles in line | infusion or flow issue
tube, tubing |  | impurity presented | Particulates
tube, tubing |  | fail to retain position | failure to adhere or bond; use of device issue
tube, tubing |  | fail to deflate | deflation issue
tube, tubing |  | fail to inflate | Inflation Issue
tube, tubing |  | fail to recommend heated tubing | instruction for use issue
tube, tubing |  | flat spot | Pitted
tube, tubing | 3 | incorrect tubing placement | use of device issue
tube, tubing |  | tube missing | component missing
tube, tubing |  | tubing bent | bent
tube, tubing | 3 | tubing blocked | Blockage within Device or Device Component; obstruction within device
tube, tubing |  | tubing broken or damaged | break
tube, tubing | 3 | tubing buckled | Buckled Material
tube, tubing |  | tubing cracked | air leak; leak; Crack; fluid leak
tube, tubing |  | tubing crushed | collapse
tube, tubing |  | tubing clotted | Blockage within Device or Device Component
tube, tubing |  | tubing connection broken or damaged | break
tube, tubing |  | tubing deflated | deflation issue; air leak
tube, tubing |  | tubing detached | detachment of device component; disconnection; fluid leak; material separation
tube, tubing |  | tubing disconnected | disconnection
tube, tubing | 3 | tubing exploded | Burst
tube, tubing |  | tubing flattened | deflation issue
tube, tubing |  | tubing inserted insufficiently | Difficult to Insert
tube, tubing |  | tubing incompatible with liquid | Incompatibility Problem
tube, tubing |  | tubing kinked | kinked; fluid leak; improper flow or infusion
tube, tubing | 3 | tubing leakage | leak; air leak; fluid leak
tube, tubing |  | tubing packaged curly | Packaging Issue
tube, tubing |  | tubing ripped | Split
tube, tubing | 3 | tubing shattered | Burst
tube, tubing | 3 | tubing splitted | Split; fluid leak
tube, tubing |  | tubing stretched | device damaged prior to use; Stretched; use of device issue
tube, tubing |  | tubing stuck | sticking; difficult to open or close; mechanical jam
tube, tubing |  | part torqued | material torqued
tube, tubing |  | tubing twisted | material twisted
tube, tubing |  | tubing mismatched | component incompatible
tube, tubing |  | tubing torn | torn material
tuner | 3
UDI | 3 | illegible UDI | device markings issue
UDI |  | fail to label UDI | device markings issue
UDI | 3 | invalid UDI | device markings issue
ultraviolet | 3 | ultraviolet absorbing | Shielding Failure
ultraviolet | 3
underfill | 3 | infufficient underfill | Short Fill
underfill | 3 | too much underfill | Short Fill
undersense, undersensing | 2 | intermittent undersensing | Device Sensing Issue
undock | 3 | unable onehand to undock | failure to deploy
unfold | 3 | difficult to unfold | Difficult to Fold or unfold
unfold | 3 | fail to unfold | Failure to Unfold or Unwrap
uninstall, uninstallation | 3 | fail to uninstall | installation-related problem
uninstall, uninstallation | 3 | unable onehand to uninstall | installation-related problem
uninstall, uninstallation | 3 | single member uninstallation failed | installation-related problem
uninstall, uninstallation | 3 | uninstall force out of specification | installation-related problem
uninstall, uninstallation | 3 | uninstallation time long | installation-related problem
unit | 3 | inappropriate measurement unit | incorrect measurement; patient data issue
unit |  | incorrect unit format | data issue; patient data issue
unit |  | unit failed | defective component
unit | 3 | unit inactivated | failure to deploy
unlatch | 3 | fail to unlatch | difficult to open or close
unlatch | 3 | unlatch force out of specification | mechanical issue
unlock | 2 | damaged unlocking mechanism | protective measure issue
unlock |  | fail to unlock | protective measure issue
unlock | 2 | unlocking mechanism failed | protective measure issue
unmute | 3 | audio unmute failed | use of device issue
unpack, unpackage | 3 | unpack time long | Packaging Issue
power, unpower | 2 | loss of detection | failure to sense
power, unpower | 2 | loss of power | loss of power
power, unpower | 2 | loss of power indication | no display or display failure
power, unpower | 2 | power consumption out of range | electrical power problem
power, unpower | 2 | power control failed | power conditioning issue
power, unpower | 2 | power frequency out of specification | Output Issue; power conditioning issue
power, unpower | 2 | power interrupted | power conditioning issue
power, unpower | 2 | power on or off switch failed | failure to power-up
power, unpower |  | output power out of range | output energy incorrect
power, unpower | 2 | power out of range | electrical power problem
unrestrict, unrestriction | 2 | movement unrestricted | unintended movement
unrestrict, unrestriction | 2 | patient not compliant with usage restriction | improper or incorrect procedure or method
unsanitary | 3 | bio-contamination | device contamination with biological material; delivered as unsterile product
unwrap | 3 | difficult to unwrap | Failure to Unwrap
unwrap | 3 | fail to unwrap | Failure to Unwrap
update | 2 | document out of date | data issue; patient data issue
update | 2 | fail to update | data issue; patient data issue
update | 2 | fail to update time | data issue; patient data issue
update | 2 | lack of update | data issue; patient data issue
update | 2 | lead to new risk | no fail-safe mechanism
upgrade | 2 | device is restored to the previous version | use of incorrect control settings
upgrade | 2 | diagnostic information is deleted | Failure to Transmit Record
upgrade | 2 | fail to follow process | improper or incorrect procedure or method
upgrade | 2 | fail to upgrade software | application program version or upgrade problem
upgrade | 2 | lack of upgrade procedure | compliance issue
upgrade | 2 | lead to new risk | no fail-safe mechanism
ups | 2 | fail to communicate with driver | Communication or Transmission Issue
ups | 2 | fail to supply power | electrical power problem
ups | 2 | Insufficient capacity | Device-Device Incompatibility
ups | 2 | lack of over charge module | charging issue
ups | 2 | lack of over current module | no fail-safe mechanism
ups | 2 | lack of over heat module | overheating of device or device component
ups | 2 | lack of over voltage module | electronic property issue
ups | 2 | lack of protection circuit | circuit failure
ups | 2 | lack of short circuit module | electrical shorting; spark
ups | 2 | ups switch failed | loss of power
usability | 2 | inadequate usability | inadequate user interface
usability | 2 | lack of acceptance criteria | use of device issue
usability | 2 | poor usability | inadequate user interface
use, misuse |  | inappropriate usage | use of device issue; user used incorrect product for intended use
use, misuse |  | Improper Use | use of device issue; user used incorrect product for intended use
use, misuse | 2 | inappropriate usage environment | environmental control or utility issue
use, misuse | 2 | inappropriate user | user used incorrect product for intended use
use, misuse |  | inadequate capacity | Output Issue
use, misuse | 2 | incorrect usage | user used incorrect product for intended use; component or accessory incompatibility
use, misuse | 2 | off-label use | use of device issue
use, misuse | 2 | over utilization | use of device issue
use, misuse | 2 | unable to use | use of device issue
use, misuse | 2 | usage confused | inadequate user interface
use, misuse |  | part mismatched | component or accessory incompatibility
use, misuse | 2 | use incorrect control settings | use of incorrect control settings
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child | 3 | fail to configure patient information | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child | 3 | inappropriate user | user used incorrect product for intended use; size incorrect for patient; Unexpected Therapeutic Results; patient-device incompatibility
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | unique patient ID violated | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | incorrect patient ID | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | different patient | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | same patient | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child | 3 | incorrect patient information | patient data issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient electrocuted | no information; human-device interface issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | moving patient incorrectly | human-device interface issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | positioning patient incorrectly | human-device interface issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient experienced adverse event | no known device problem; fitting problem; human-device interface issue; size incorrect for patient
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient expired | no information
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient hospitalized | no known device problem
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient incident | no information; human-device interface issue; size incorrect for patient; use of device issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | operator error | device handling issue
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | typo error | inadequate training
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | unskilled user | inadequate training
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | untrained user | inadequate training
user, practitioner, patient, person, personnel, population, disabled, infant, newborn, children, pediatric, pedestal, pendant, new-born, child |  | patient thermal injured | overheating of device or device component; human-device interface issue
utilize, utilization, usage | 3 | over utilization | use of device issue
utilize, utilization, usage | 3 | reduced service life | shelf life exceeded; improper device output
vacuum | 3 | fail to adjust vacuum | suction issue; use of incorrect control settings
vacuum | 3 | incapable to support clinically stable patient | use of device issue
vacuum | 3 | loss of vacuum | suction issue
validate, validation, validity | 2 | fail to follow plan | improper or incorrect procedure or method
validate, validation, validity | 2 | lack of document | device markings issue
validate, validation, validity | 2 | lack of reference | device markings issue
validate, validation, validity | 2 | lack of validation | protective measure issue
validate, validation, validity | 2 | lack of validation plan | protective measure issue
validate, validation, validity | 2 | validation failed | protective measure issue
valve, duckbill valve, duckbill, check valve | 2 | back flow | improper flow or infusion
valve, duckbill valve, duckbill, check valve |  | faulty valve system | No Code Available; Blockage within Device or Device Component
valve, duckbill valve, duckbill, check valve |  | fail to setup valve correctly | device difficult to setup or prepare
valve, duckbill valve, duckbill, check valve |  | fail to control valve | improper flow or infusion
valve, duckbill valve, duckbill, check valve |  | fail to close valve | difficult to open or close
valve, duckbill valve, duckbill, check valve |  | fail to open valve | difficult to open or close
valve, duckbill valve, duckbill, check valve |  | fail to prevent flow | leak; improper flow or infusion
valve, duckbill valve, duckbill, check valve |  | fail to withstand pressure | mechanical issue
valve, duckbill valve, duckbill, check valve | 2 | low life cycle | shelf life exceeded
valve, duckbill valve, duckbill, check valve |  | inappropriate valve location | positioning issue
valve, duckbill valve, duckbill, check valve |  | noisy valve | audible noise
valve, duckbill valve, duckbill, check valve | 2 | undesired valve activation | deployment issue
valve, duckbill valve, duckbill, check valve | 2 | valve broken or damaged | break
valve, duckbill valve, duckbill, check valve |  | valve detached | detachment of device component; disconnection
valve, duckbill valve, duckbill, check valve |  | valve migrated | migration of device or device component
valve, duckbill valve, duckbill, check valve |  | valve cracked | crack; fluid leak
valve, duckbill valve, duckbill, check valve | 2 | valve leakage | leak; improper flow or infusion; fluid leak
valve, duckbill valve, duckbill, check valve |  | valve jammed | Mechanical Jam
valve, duckbill valve, duckbill, check valve | 2 | valve out of control | deployment issue
valve, duckbill valve, duckbill, check valve | 2 | valve partially open | difficult to open or close
valve, duckbill valve, duckbill, check valve | 2 | valve punctured | Material Puncture
valve, duckbill valve, duckbill, check valve | 2 | valve stuck | sticking; difficult to open or close; mechanical jam; fluid leak
valve, duckbill valve, duckbill, check valve | 2 | valve stuck closed | sticking; difficult to open or close; mechanical jam; no flow; fluid leak
valve, duckbill valve, duckbill, check valve | 2 | valve stuck open | sticking; difficult to open or close; mechanical jam
vaporizer | 3
varistor | 3
vasculature |  | vein punctured | use of device issue
vein, venous | 2 | vein punctured | use of device issue; Material Puncture
vein, venous |  | Vein clotted off | human-device interface issue
vein, venous | 2 | vein thrombosis | use of device issue
vein, venous | 2 | venous injury | use of device issue
Velcro | 3 | hook and loops are not holding properly | Misassembled
Velcro | 3 | Velcro assembling breakage | break
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | air delivery compromised | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | all-patient ventilator compromised | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to adjust volume | use of incorrect control settings
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to associate patient | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to compensate pressure | pressure issue
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to conjunct with high pressure oxygen source | medical gas supply problem
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to conjunct with Nitro Oxide | medical gas supply problem
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to control ventilator | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to inline with humidification device | component or accessory incompatibility
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to initiate ventilation | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to compatible with oxygen supply | Device-Device Incompatibility
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to disable ventilation | device remains activated
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to nebulization medication | medication error
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to provide diagnostic test | incorrect or inadequate test results
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to power up | device inoperable
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to roll ventilator in hospital | device handling issue
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to start with preset value | reset issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to stop cycle | reset issue
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to support non-spontaneous breathing patient | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to support spontaneous breathing patient | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | fail to transport patient | device handling issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to vent air out | obstruction within device
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | ventilation failed | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | inadequate ventilation | failure to cycle; improper device output
ventilate, ventilation, ventilatory, ventilator, respirator |  | confused ventilation mode | use of device issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to compatible with accessory | Device-Device Incompatibility
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | inappropriate ventilation support | failure to cycle; improper device output
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to maintain ventilation mode | failure to cycle; improper device output
ventilate, ventilation, ventilatory, ventilator, respirator |  | inadequate ventilation mode | failure to cycle; improper device output
ventilate, ventilation, ventilatory, ventilator, respirator |  | flow rate overspeed | excess flow or over-infusion
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | incorrect ventilation mode | failure to cycle; no flow
ventilate, ventilation, ventilatory, ventilator, respirator |  | incorrect ventilation pattern | improper device output
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | Insufficient power capacity | Device-Device Incompatibility
ventilate, ventilation, ventilatory, ventilator, respirator |  | intermittent ventilation | failure to cycle; intermittent continuity
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | lack of invasive support | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | lack of non-invasive support | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator |  | fail to advice backup ventilation | instruction for use issue; no fail-safe mechanism
ventilate, ventilation, ventilatory, ventilator, respirator |  | lack of backup gas | medical gas supply problem; fail-safe mechanism issue
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | less oxygen supply | medical gas supply problem
ventilate, ventilation, ventilatory, ventilator, respirator |  | lock to incorrect pressure | pressure issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | ventilator malfunction | device operates differently than expected; Product Quality Issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | ventilator turns on by itself | device operates differently than expected; Product Quality Issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | no air delivery | failure to cycle; pressure issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | respiratory failure | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | setup time long | reset issue
ventilate, ventilation, ventilatory, ventilator, respirator |  | thermal event | pressure issue; electronic property issue; fire
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | unable auto ventilation | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator |  | ventilation blocked | Blockage within Device or Device Component
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | ventilation interrupted | failure to cycle; no pressure; device inoperable
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | ventilation loss of control | failure to cycle
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | ventilation stopped | failure to cycle; no pressure; device inoperable
ventilate, ventilation, ventilatory, ventilator, respirator | 2 | ventilation support compromised | failure to cycle
ventricle | 2 | deliver subthreshold | inaccurate delivery
ventricle | 2 | fail to drive full eject | unintended ejection; expulsion
ventricle | 2 | fail to drive partial fill | filling problem
ventricle | 2 | fail to overcome afterload | failure to deploy
ventricle | 2 | inappropriate ventricle depolarization | electrical issue
ventricular | 2 | inappropriate ventricular depolarization | electrical issue
ventricular | 2 | intermittent ventricular pause | Device Stops Intermittently
ventricular | 2 | oversensing | Over-sensing
ventricular | 2 | premature ventricular contraction | Retraction Problem
ventricular | 2 | slow ventricular rate | capturing issue
ventricular | 2 | under sensing | under-sensing
ventricular | 2 | ventricular evoked response small | Failure to Interrogate; Device Sensing Issue
ventricular | 2 | ventricular inhibition | No Pacing
ventricular | 2 | ventricular undersensing | under-sensing
verify, verification | 2 | fail to follow plan | improper or incorrect procedure or method
verify, verification | 2 | inadequate test procedure | incorrect or inadequate test results
verify, verification | 2 | inadequate verification plan | improper or incorrect procedure or method
verify, verification | 2 | inadequate verification strategy | protective measure issue
verify, verification | 2 | incomplete verification | protective measure issue
verify, verification | 2 | lack of acceptance criteria | use of device issue
verify, verification | 2 | lack of deliverable | protective measure issue
verify, verification | 2 | lack of document | device markings issue
verify, verification | 2 | lack of milestone | device markings issue
verify, verification | 2 | lack of task | protective measure issue
verify, verification | 2 | lack of verification | protective measure issue
vessel | 3 | burst vessel | Burst; use of device issue
vessel | 3 | vessel punctured | use of device issue; Material Puncture
vibrate, vibration, vibrator | 2 | device broken or damaged | device damaged prior to use
vibrate, vibration, vibrator | 2 | induced device vibration | vibration
vibrate, vibration, vibrator |  | unwanted tissue vibration | vibration
vibrate, vibration, vibrator | 2 | interaction with gradient & static fields during MRI | Electromagnetic Interference (EMI)
video | 3 | lack of video | inadequate user interface
view |  | fail to hide view | incorrect display
view |  | fail to save view | loss of data; data back-up problem
view |  | inadequate panel view | incorrect display
view |  | unable to rename view | device markings issue; data issue; use of device issue
view |  | unable to display multiple views | incorrect display
view |  | unable to display multiple images | incorrect display
view |  | contradictory view | incorrect display
view |  | switch view | incorrect display
view | 3 | fail to view | no display or display failure
visibility, visible | 3 | image brightness out of specification | poor quality image
visibility, visible | 3 | image color accuracy out of specification | poor quality image
visibility, visible |  | fail to toggle visibility | no display or display failure; incorrect display
visibility, visible |  | fail to indicate visibility | no display or display failure
visibility, visible |  | fail to reset default visibility | reset issue
visibility, visible |  | inadequate visibility design | human-device interface issue
visibility, visible | 3 | loss of visual | human-device interface issue
visibility, visible | 3 | part visible undesired | human-device interface issue
visualize, visual, visualization | 3 | image brightness out of specification | poor quality image
visualize, visual, visualization | 3 | image color accuracy out of specification | poor quality image
visualize, visual, visualization |  | fail to manipulate visualization parameters | optical issue; use of incorrect control settings; poor quality image
visualize, visual, visualization | 3 | loss of visual | human-device interface issue; no display or display failure
void
voltage, VDC, overvoltage | 3 | incorrect voltage value | electronic property issue; electrical issue
voltage, VDC, overvoltage | 3 | induced voltage | electronic property issue; electrical issue
voltage, VDC, overvoltage | 3 | short to ground | electrical shorting; spark
voltage, VDC, overvoltage |  | voltage dropped | electronic property issue; electrical issue
voltage, VDC, overvoltage |  | output voltage out of range | output issue
voltage, VDC, overvoltage | 3 | voltage out of range | electronic property issue; electrical issue
volume | 3 | cumulative volume out of range | output above specifications; output below specifications
volume | 3 | fail to adjust volume | use of incorrect control settings
volume |  | fail to calculate volume | incorrect software programming calculations
volume |  | fail to visualize volume | no display or display failure
volume |  | fail to verify volume | use of device issue
volume | 3 | high volume | output above specifications
volume |  | over solution volume | Overfill
volume | 3 | inaccurate volume | volume accuracy issue
volume |  | misread volume | use of device issue; data issue
volume | 3 | incorrect medicine volume | use of device issue; volume accuracy issue
volume | 3 | incorrect volume calculation | parameter calculation error due to software problem
volume |  | inappropriate volume unit | incorrect measurement
volume | 3 | low production volume | Output below Specifications
volume | 3 | low volume | Output below Specifications
volume |  | Insufficient solution volume | Short Fill
volume |  | unitend volume dropped | Short Fill
volume | 3 | volume control failed | use of device issue; volume accuracy issue
volume |  | low volume accuracy | volume accuracy issue
volume | 3 | volume out of range | use of device issue; volume accuracy issue | volume out of specification
warn, warning | 2 | confused warning | inadequate user interface
warn, warning | 2 | fail to generate warning | device alarm system issue
warn, warning | 2 | inadequate warning | device alarm system issue
warn, warning | 2 | incorrect warning | device alarm system issue
warn, warning | 2 | loss of warning | device alarm system issue
warn, warning |  | warning displayed | device displays error message
warranty | 3 | void warranty | Market Issue
wash, washable, washer |  | loose washer | detachment of device component
wash, washable, washer | 2 | part torn | torn material
wash, washable, washer | 2 | part washing damaged | device cleaning issue
water |  | water damage | fluid leak
water | 1 | water ingress | fluid leak; contamination during use
waveform | 3 | inappropriate waveform | electronic property issue
waveform |  | inappropriate output waveform | electronic property issue; output issue
waveform |  | inappropriate preset value | failure to reset
waveform |  | incorrect derived value | electronic property issue; data issue
waveform |  | frequency out of range | electronic property issue
waveform |  | incorrect repetition rate | electronic property issue
waveform |  | incorrect pulse numbers | electronic property issue
waveform |  | abnormal waveform | electronic property issue
waveform | 3 | incorrect waveform shape | electronic property issue
waveguide | 3
wavelength | 3 | incorrect wavelength | defective component
wear | 3 | part broken or damaged | break; defective component
wear | 3 | remove difficult | difficult to remove
wear | 3 | uncomfortable | human-device interface issue
wear | 3 | wear for prolong periods | use of device issue
wedge | 3 | difficult to wedge | Difficult to Insert
WEEE, Waste Electrical and Electronic Equipment |  | missing disposal mark WEEE | device markings issue
WEEE, Waste Electrical and Electronic Equipment |  | inproper disposal | environmental control or utility issue
WEEE, Waste Electrical and Electronic Equipment | 2 | WEEE violated | compliance issue
weigh, weight | 3 | low weight capacity | patient-device incompatibility
weigh, weight | 3 | over weighing | patient-device incompatibility
weigh, weight | 3 | weight deviation | patient-device incompatibility
weigh, weight |  | null weight value | data issue
weigh, weight | 3 | weight out of range | patient-device incompatibility
weld, weldability, welding | 3 | blown weld | Solder Joint Failure
weld, weldability, welding | 3 | inadequate cleaning from welding | device cleaning issue
weld, weldability, welding | 3 | inadequate weld strength | mechanical issue
weld, weldability, welding | 3 | incomplete welding | Solder Joint Failure
weld, weldability, welding | 3 | limited laser weldability | Solder Joint Failure
weld, weldability, welding | 3 | loss of electrical continuity | intermittent continuity
weld, weldability, welding | 3 | open weld | defective component
weld, weldability, welding |  | over energy | output energy incorrect
weld, weldability, welding | 3 | poor seam weld | Solder Joint Failure
weld, weldability, welding | 3 | poor weld | Solder Joint Failure
weld, weldability, welding | 3 | weld breaking | break
weld, weldability, welding | 3 | weld cracking | Crack
weld, weldability, welding | 3 | weld in wrong location | positioning issue
wetting | 3 | insufficient wetting | Moisture or Humidity Problem
wheelchair | 3
whitelist | 3 | data breached | application security issue
wide, width, widths | 3 | width out of specification | defective component
WIFI | 3 | loss of WIFI connection | connection issue
winding | 3 | high winding temperature | overheating of device or device component
winding | 3 | inadequate insulation | electrical shorting
window, fullscreen | 3 | inappropriate window width | application program issue
wire | 3 | extra wire | device misassembled during manufacturing or shipping
wire | 3 | foreign material inside | Item Contaminated during Manufacturing or Shipping; Packaging Issue
wire | 3 | inadequate cross-sectional area | defective item
wire | 3 | inappropriate route | device misassembled during manufacturing or shipping
wire | 3 | loose wire | device misassembled during manufacturing or shipping
wire | 3 | loss of connection | connection issue
wire | 3 | low wire purity | material integrity issue
wire | 3 | wire bent | device misassembled during manufacturing or shipping
wire | 3 | wire breakage | device misassembled during manufacturing or shipping
wire | 3 | wire defective | defective component
wire | 3 | wire exposed | device misassembled during manufacturing or shipping
wire | 3 | wire rough | Scratched Material
wire | 3 | wire sagging | device misassembled during manufacturing or shipping
wire | 3 | wire shorting | device misassembled during manufacturing or shipping
wire | 3 | wire stress | device misassembled during manufacturing or shipping
wire | 3 | wire tail excess | device misassembled during manufacturing or shipping
wire | 3 | tight wire | device misassembled during manufacturing or shipping
wire |  | wire torn | torn material
wire | 3 | wiring broken or damaged | break
wire | 3 | wiring rubbed | device abrasion from instrument or another object
wireless | 2 | loss of wireless connection | wireless communication issue
withdraw | 3 | connection broken or damaged | connection issue; break; device damaged prior to use
withdraw | 3 | seal damaged | failure to adhere or bond
wrap | 3 | fail to wrap | Packaging Issue
wrinkle | 2 | loss of electrical continuity | disconnection; intermittent continuity
wrinkle | 2 | part fractured | fracture; mechanical issue
xenograft, heterograft |  |  | biocompatibility issue
York | 3
zero | 3 | fail to zero | Failure to Zero
zoom |  | unable to zoom image | application program issue
zoom |  | unable to zoom operation | application program issue
zoom |  | unable to zoom in views | application program issue
zoom |  | unable to zoom out views | application program issue
zoom |  | inappropriate zoom step | use of incorrect control settings
zoom |  | inappropriate zoom out limit | use of incorrect control settings
zoom |  | inappropriate zoom in limit | use of incorrect control settings
zoom |  | inappropriate zoom operation | application program issue
zoom |  | inappropriate zoom level | use of incorrect control settings
zoom |  | incorrect zoom factor | use of incorrect control settings; data issue
zoom |  | incorrect zoom magnification value | use of incorrect control settings; data issue
zoom |  | zoom slow | device sensing issue
zoom |  | unstable zoom level | unstable
zoom | 3 | unable to zoom in small | inadequate user interface
blender |  | blender leakage | leak
blender |  | device malfunction | device operates differently than expected
electronic |  | electronics failure | electronic property issue; defective component
device |  | thermal event | electronic property issue; fire
device |  | fail to boot up | application program issue; device inoperable
device |  | catch fire | fire
device |  | fire involved | fire
tube, tubing |  | hole in tubing | hole in material; fluid leak
restart |  | unexpected restart | inappropriate or unexpected reset
inflate, inflation |  | fail to inflate balloon evenly | Inflation Issue
airway |  | airway obstructed | obstruction within device
tube, tubing |  | back flow | improper flow or infusion
tube, tubing |  | tubing wall thickness out of specification | defective item
tube, tubing |  | tubing unclamped | Free or Unrestricted Flow
tube, tubing |  | tubing malformed | defective component
device |  | inadequate size | size incorrect for patient; patient-device incompatibility
battery |  | battery melted | Melted; battery issue; fire
resuscitator |  | device malfunction | device operates differently than expected; Product Quality Issue
resuscitator |  | fail to hold pressure | pressure issue
cord |  | cord burned | Burn of Device or Device Component; burning odor
poppet |  | poppet not seated properly | failure to align
poppet |  | poppet missing | component missing
poppet |  | poppet broken or damaged | break
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap leakage | fluid leak
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap cracked | fluid leak; cracked
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap broken or damaged | break
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap remains openned | detachment of device or device component
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap adjustment difficult | device difficult to setup or prepare
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap detached | fluid leak; detachment of device or device component; Delivered As Unsterile Product
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | loose cap | loose or intermittent connection; Failure to Prime
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to compatible with ISO luer lock | nonstandard device or device component
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to prime with cap | detachment of device component
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to retain closed | positioning issue
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap tighten during sterilization | device disinfection or sterilization issue
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap tighten during packaging and shipping | device misassembled during manufacturing or shipping
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to remove | difficult to remove
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | difficult to remove | difficult to remove
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to follow clinical request | use of device issue
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap disinfecting over 7 days | Disinfection or Sterilization Issue at User Location
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | fail to expel air | Blockage within Device or Device Component
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | unsecured cap | installation-related problem
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap blocked | Blockage within Device or Device Component
cap, clave cap, microclave cap, Swabcap, disinfecting cap, connector cap |  | cap missing | component missing
silicone |  | silicone stuck | sticking
silicone |  | silicone cracked | crack; fluid leak
silicone |  | silicone broken or damaged | break
silicone |  | silicone missing | fluid leak
connector |  | fail to lock | protective measure issue; fluid leak; loose or intermittent connection
connector |  | locking mechanism failed | Fail-Safe Mechanism Issue; loose or intermittent connection
connector |  | fail to detach | failure to disconnect
connector |  | fail to remain still | fluid leak
connector |  | unable to connect | disconnection
connector |  | closed male luer disconnected | disconnection
connector |  | closed male luer leakage | fluid leak
connector |  | accidental disconnection | disconnection; use of device issue; fluid leak
chemotherapy |  | chemotherapy spill | fluid leak
chemotherapy |  | chemotherapy leakage | fluid leak
cytotoxic, cytostatic, Toxin, Toxocity, Pyrogen, cytokines, endotoxin |  | DEHP/BPA leached | biocompatibility issue
cytotoxic, cytostatic, Toxin, Toxocity, Pyrogen, cytokines, endotoxin |  | exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors; biological environmental factor
cytotoxic, cytostatic, Toxin, Toxocity, Pyrogen, cytokines, endotoxin |  | material biocompatibility compromised | component or accessory incompatibility
cytotoxic, cytostatic, Toxin, Toxocity, Pyrogen, cytokines, endotoxin |  | cytotoxic medication leakage | biocompatibility issue
cartridge |  | tubing cartridge leakage | fluid leak
clave |  | clave leakage | fluid leak
clave |  | clave blocked | Blockage within Device or Device Component; obstruction within device
clave |  | clave depressed | fluid leak
clave |  | clave broken or damaged | break
clave |  | part mismatched | device-device incompatibility
clave |  | clave detached | fluid leak
clave |  | clave connection cracked | fluid leak
burette |  | burette leakage | fluid leak
crack |  | spider crack | fluid leak
spike |  | bag spike dropped | fluid leak
spike |  | bag spike removed | fluid leak
spike |  | fail to compatible with IV container | Incompatibility Problem
spike |  | spike cracked | crack
spike |  | spike leakage | fluid leak
spike |  | spike blocked | fluid leak; mechanical jam
spike |  | spike detached | fluid leak
spike |  | spike dislodged | fluid leak; dislodged or dislocated
spike |  | spike broken or damaged | fluid leak
opacity, transparency |  | fail to adjust opacity level | human-device interface issue; optical issue
opacity, transparency |  | fail to set transparency level | optical issue
opacity, transparency |  | fail to reset transparency level | optical issue
opacity, transparency |  | fail to provide opacity percentage | optical issue
opacity, transparency |  | fail to provide transparency percentage | optical issue
opacity, transparency |  | fail to reset default opacity level | optical issue
opacity, transparency |  | fail to reset default transparency level | optical issue
opacity, transparency |  | incorrect opacity level | optical issue
opacity, transparency |  | increased opacity level | item contaminated during manufacturing or shipping; material opacification; optical issue
lumen |  | lumen blocked | infusion or flow issue
manifold |  | manifold detached | disconnection
retent, retension |  | retent force out of specification | use of device issue
aesthetic, aesthetical |  | product aesthetic compromised | material integrity issue
leach |  | DEHP leached | biocompatibility issue
environment, environmental |  | Hazardous fluids released into environment | environmental control or utility issue; contamination during use
environment, environmental |  | Hazardous materials released into environment | environmental control or utility issue; contamination during use
environment, environmental |  | Hazardous fluids/fumes released into environment | environmental control or utility issue; contamination during use
environment, environmental |  | Hazardous fumes released into environment | environmental control or utility issue; contamination during use
Ergonomic |  | fail to consider ergonomic factors | human-device interface issue
Ergonomic |  | Ergonomic stress | human-device interface issue
surface |  | abrasive surface | Pitted
inject, injection, injector |  | Check Injector Alarm | device alarm system issue
cosmetic |  | component dirty | environmental control or utility issue
cosmetic |  | component wet | Moisture or Humidity Problem
cosmetic |  | component discolored | Material Discolored
cosmetic |  | component damp | Moisture or Humidity Problem
login |  | incorrect login name displayed | application security issue
login |  | invalid username | application security issue
login |  | invalid password | application security issue
dashboard |  | Inaccurate Dashboard | improper device output
license, certificate |  | incorrect license file | use of device issue
license, certificate |  | incompatible license | application program issue
license, certificate |  | license error | application program issue
license, certificate |  | license expired | use of device issue
license, certificate |  | device certificate issue | use of device issue
server |  | low server configuration | use of incorrect control settings
server |  | fail to connect to server | connection issue
server |  | inaccessible server | data issue
server |  | empty server | data issue; loss of data
server |  | fail to upload to server | data issue; patient data issue
server |  | server stopped | computer operating system issue
export |  | unable to export | data issue
database |  | unable to restart database | device inoperable
database |  | inaccessible database partition | application program issue
database |  | database unavailable | data issue
quantity |  | less than stated quantity | imprecision; data issue
quantity |  | more than stated quantity | imprecision; data issue
mobility |  | fail to transport | use of device issue
scalable |  | fail to integrate | accessory incompatible
technology |  | lead to new risk | no fail-safe mechanism
response |  | delayed response | Communication or Transmission Issue
shutdown, shut down |  | unintentional shutdown | device inoperable
shutdown, shut down |  | slow shut down | failure to shut off; device operates differently than expected
shutdown, shut down |  | incorrect shutdown procedure | device handling issue
REACH | REACH | REACH violated | compliance issue; environmental control or utility issue
ratio |  | incorrect ratio calculation | parameter calculation error due to software problem
ratio |  | low ratio accuracy | imprecision
ratio |  | ratio out of range | use of incorrect control settings
retrieve |  | fail to retrieve data | data issue
body |  | fail to specify applied body | instruction for use issue
touchscreen, touch panel, touch screen |  | unresponsive touch screen | decreased sensitivity; inadequate user interface; device inoperable; no display or display failure
touchscreen, touch panel, touch screen |  | unable to lock screen | no display or display failure
touchscreen, touch panel, touch screen |  | long response time | decreased sensitivity; inadequate user interface
touchscreen, touch panel, touch screen |  | low resolution | Image Resolution Poor
touchscreen, touch panel, touch screen |  | insensitive touch screen | decreased sensitivity; inadequate user interface
touchscreen, touch panel, touch screen |  | incorrect touch location detected | positioning issue; inadequate user interface
touchscreen, touch panel, touch screen |  | fail to detect valid touch | decreased sensitivity; inadequate user interface
touchscreen, touch panel, touch screen | usability | constant touch undetected | decreased sensitivity; inadequate user interface
watchdog, watch dog, watchdog test |  | stuck deasserted | sticking; protective measure issue
watchdog, watch dog, watchdog test |  | stuck asserted | sticking; protective measure issue
watchdog, watch dog, watchdog test |  | falsely detect error | Device Sensing Issue
watchdog, watch dog, watchdog test |  | fail to initiate test | failure of device to self-test
watchdog, watch dog, watchdog test |  | fail to detect error | failure to sense; protective measure issue
watchdog, watch dog, watchdog test |  | fail to activate protective actions | fail-safe mechanism issue; protective measure issue
watchdog, watch dog, watchdog test |  | fail to activate alarm | device alarm system issue; protective measure issue
Monitor PIC, Control PIC, PIC |  | long response time | decreased sensitivity
Monitor PIC, Control PIC, PIC |  | loss of connection | connection issue
metadata, meta data
annotation, note, textbox, comment, text |  | inappropriate annotation tool | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | inappropriate annotation format | application program issue; data issue
annotation, note, textbox, comment, text |  | incorrect image annotation | image display error
annotation, note, textbox, comment, text |  | incorrect text direction | data issue
annotation, note, textbox, comment, text |  | fail to add any notes | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to list all annotations | application program issue
annotation, note, textbox, comment, text |  | fail to number all annotations | application program issue
annotation, note, textbox, comment, text |  | fail to select an annotation | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to delete an annotation | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to place annotation on image | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to place annotation on 3D model | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to place annotation | application program issue; use of device issue
annotation, note, textbox, comment, text |  | fail to place arrow annotation | application program issue; use of device issue
annotation, note, textbox, comment, text |  | fail to save note | application program issue; use of device issue; loss of data
annotation, note, textbox, comment, text |  | fail to save annotation | application program issue; use of device issue; data issue; loss of data
annotation, note, textbox, comment, text |  | fail to reset annotation | failure to reset
annotation, note, textbox, comment, text |  | fail to display annotation | application program issue; no display or display failure
annotation, note, textbox, comment, text |  | fail to remove any note | application program issue; human-device interface issue; use of device issue
annotation, note, textbox, comment, text |  | fail to edit note | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | fail to add notes at anywhere | application program issue; human-device interface issue
annotation, note, textbox, comment, text |  | missing annotation | device markings issue; incorrect display
annotation, note, textbox, comment, text |  | fail to transfer comment | loss of data
annotation, note, textbox, comment, text |  | unintend to delete note | use of device issue
annotation, note, textbox, comment, text |  | comment deleted | loss of data
annotation, note, textbox, comment, text |  | lack of annotation tool | application program issue; human-device interface issue
marker
CT |  | image not in DICOM format | image display error; data issue
CT |  | poor image | poor quality image; therapeutic or diagnostic output failure
Anaesthetic |  | incorrect anesthetic supply | medical gas supply problem
Anaesthetic |  | fail to supply range of anesthetic | medical gas supply problem
Anaesthetic |  | inappropriate exposure to anesthetic | use of device issue; environmental control or utility issue
glove |  | glove broken or damaged | bacterial contamination of device
Pyrogenicity |  | pyrogenic reaction | biocompatibility issue
Acid, Alkali |  | chemical leakage | chemical issue
manufacture, manufacturing |  | inadequate manufacturing process | Instruction for Use Issue
malware |  | access malware | unauthorized access to computer system
inject, injection, injector |  | fluid injection | fluid leak; fluid injection
inject, injection, injector |  | high pressure fluid injection | fluid leak; fluid injection
failure |  | Knowledge based failure | use of device issue; human-device interface issue
failure |  | Rule Based failure | use of device issue; human-device interface issue
judge, judgement |  | fault judgement | use of device issue; human-device interface issue
tumor |  | fail to detect tumor spread | poor quality image; image display error; application program issue
tumor |  | fail to detect tumor location | poor quality image; image display error; application program issue
tumor |  | fail to detect tumor size | poor quality image; image display error; application program issue
tumor |  | fail to detect tumor abnormalities | poor quality image; image display error; application program issue
tumor |  | fail to detect tumor defects | poor quality image; image display error; application program issue
tumor |  | fail to visualize color-coded tumor | poor quality image; image display error; application program issue
tumor |  | fail to visualize spatial relationship | poor quality image; image display error; application program issue
tumor |  | fail to remove tumor | use of device issue; human-device interface issue
tumor |  | tumor removed inappropriately | use of device issue; human-device interface issue
cancer |  | fail to visualize cancer location | poor quality image
cancer |  | fail to identify cancer invaded outside wall | poor quality image
cancer |  | overestatement effectiveness of cancer treatment | therapeutic or diagnostic output failure
cancer |  | underestatement effectiveness of cancer treatment | therapeutic or diagnostic output failure
cancer |  | fail to identify structure invaded | poor quality image
anatomy |  | poor anatomy image | optical issue; poor quality image
anatomy | anatomy | fail to identify anatomy variety | optical issue; poor quality image
anatomy |  | invisible anatomy | optical issue; poor quality image
anatomy |  | fail to identify anatomical variation | optical issue; poor quality image
pre-op, pre-operative, preoperative |  | fail to provide virtual clip-applying tool | human-device interface issue
pre-op, pre-operative, preoperative |  | fail to record preoperative data | data back-up problem; loss of data
pre-op, pre-operative, preoperative |  | fail to detect tumor location | poor quality image
pre-op, pre-operative, preoperative |  | fail to provide real time organ calculation | image resolution poor; poor quality image
software, SaMD, software as a medical device |  | IEC 62304 violated | compliance issue; computer software issue
MDR, EU Medical Device Regulation, MDR certificate |  | Eu MDR violated | compliance issue
MDR, EU Medical Device Regulation, MDR certificate |  | MDR submission delayed | compliance issue
MDR, EU Medical Device Regulation, MDR certificate |  | prolonged MDR approval | compliance issue
MDR, EU Medical Device Regulation, MDR certificate |  | inappropriate certificate scope | compliance issue
MDR, EU Medical Device Regulation, MDR certificate |  | insufficient clinical evidence | data issue
MDR, EU Medical Device Regulation, MDR certificate |  | incorrect clinical data | data issue
MDR, EU Medical Device Regulation, MDR certificate |  | lack of clinical data from equivalent device | data issue
FDA |  | prolonged FDR acceptance | compliance issue
QMS |  | audit observations | compliance issue
QMS |  | inadequate QMS document lists | compliance issue
organ |  | fail to calculate volume of resected organ | incorrect software programming calculations
menu |  | fail to access menu | computer software issue; human-device interface issue
menu |  | fail to navigate tab | human-device interface issue
menu |  | confused menu | inadequate user interface; human-device interface issue
menu |  | difficult to load patient information | inadequate user interface; human-device interface issue
menu |  | difficult to access settings | inadequate user interface; human-device interface issue
menu |  | difficult to navigate function | inadequate user interface; human-device interface issue
order |  | fail to select order | application program issue; human-device interface issue
filter
toggle |  | confused toggle | inadequate user interface; human-device interface issue
toggle |  | fail to toggle | failure to deploy; device operates differently than expected; use of device issue
scroll |  | scroll image delayed | application network issue
scroll |  | unable to scroll image | application program issue
pan |  | unable to pan viewing panel | application program issue; human-device interface issue; use of device issue
pan |  | inappropriate pan operation | application program issue; human-device interface issue; use of device issue
pan |  | inappropriate pan limits | use of incorrect control settings
pan |  | fail to pan independently | application program issue; human-device interface issue; use of device issue
pan |  | fail to retrieve default setting | use of incorrect control settings
pan |  | fail to reset pan | failure to reset
pan |  | fail to locate image | data issue; patient data issue
pan |  | unable to pan image | application program issue; human-device interface issue; use of device issue
window, fullscreen |  | fail to auto-adjust window width | application program issue
window, fullscreen |  | fail to auto-set window | application program issue
window, fullscreen |  | fail to adjust window manually | application program issue
window, fullscreen |  | fail to retrieve preset window | use of incorrect control settings
window, fullscreen |  | fail to reset window | failure to reset
window, fullscreen |  | fail to fit to window | failure to reset
window, fullscreen |  | unable to adjust window center | application program issue
window, fullscreen |  | fail to define window center | dislodged or dislocated
window, fullscreen |  | fail to define default window center | dislodged or dislocated
window, fullscreen |  | fail to reset window center | failure to reset
window, fullscreen |  | fail to retrieve preset window center | failure to reset
window, fullscreen |  | unable to adjust window width | application program issue
window, fullscreen |  | fail to define window width | application program issue
window, fullscreen |  | fail to define default window width | application program issue
window, fullscreen |  | fail to reset window width | failure to reset
window, fullscreen |  | fail to retrieve preset window width | reset issue
window, fullscreen |  | unable to adjust window level | application program issue
window, fullscreen |  | fail to enter fullscreen | application program issue; use of device issue
window, fullscreen |  | inappropriate window level | application program issue
endpoint |  | fail to drag endpoint | human-device interface issue; application program issue
drag |  | fail to drag endpoint | human-device interface issue; application program issue
drag |  | fail to drag operation | human-device interface issue; application program issue; use of device issue
drag |  | fail to drag line | human-device interface issue; application program issue
drag |  | fail to drag and drop line | human-device interface issue; application program issue
line |  | fail to resize line | human-device interface issue; application program issue
line |  | fail to drag and drop line | human-device interface issue; application program issue
line |  | fail to add line | human-device interface issue; application program issue
line |  | fail to edit line | human-device interface issue; application program issue
line |  | fail to delete line | human-device interface issue; application program issue
arrow |  | fail to add arrow | human-device interface issue; application program issue; use of device issue
arrow |  | inappropriate arrow fashion | device markings issue
arrow |  | fail to rename arrow | human-device interface issue; application program issue
arrow |  | fail to delete arrow | human-device interface issue; application program issue
checkbox |  | fail to select | application program issue; human-device interface issue
checkbox |  | fail to deselect | application program issue; human-device interface issue
mouse, click |  | wrong mouse click | use of device issue
mouse, click |  | fail to click operation | human-device interface issue; use of device issue
mouse, click |  | mouse navigation frozen | application program issue; device operates differently than expected
mouse, click |  | unresponsive mouse click | human-device interface issue; application program issue
clip, scissors | mechanical | clip mishandled | device handling issue
clip, scissors |  | invisible clip | no display or display failure
clip, scissors |  | fail to apply clip | application program issue; device handling issue
clip, scissors |  | fail to deselect clip | device handling issue
clip, scissors |  | missing voice clip | loss of data
clip, scissors |  | inadequate clip info | data issue
clip, scissors | mechanical | broken clip retention | break
undo, redo |  | fail to undo user interaction | application program issue; use of device issue
undo, redo |  | fail to undo prior action | application program issue; use of device issue
undo, redo |  | fail to redo user interaction | application program issue; use of device issue
resect, resection |  | fail to view resected target anatomy | application program issue
resect, resection |  | fail to view resected pathology | application program issue
resect, resection |  | fail to view resected organ segment | application program issue
resect, resection |  | fail to simulate resection | application program issue
icon |  | fail to display icon | no display or display failure; incorrect display
icon |  | unstandardized icon | device markings issue
intra-op, intra-operative, intraoperative |  | fail to record intraoperative data | data back-up problem; loss of data
post-op, post-operative, postoperative |  | fail to record postoperative data | data back-up problem; loss of data
revision, version |  | insufficient documentation | device markings issue; data issue
revision, version |  | incorrect revision | device markings issue
nerve |  | Nerve Injury | improper or incorrect procedure or method
nerve |  | Nerve damaged | improper or incorrect procedure or method
liver |  | Postoperative liver failure | improper or incorrect procedure or method
Bronchial |  | Bronchial Injury | improper or incorrect procedure or method
Vascular |  | Major Vascular Injury | improper or incorrect procedure or method
projection |  | projection error | image display error
grid |  | unexpected grid | application program issue
grid |  | incorrect grid selected | use of device issue
biopsy |  | incorrect biopsy | unexpected therapeutic results
list |  | inappropriate list displayed | incorrect display
pixel |  | inappropriate display of pixel value tool | incorrect display
pixel |  | inappropriate pixel size | image resolution poor; therapeutic or diagnostic output failure
file, report, study, exam, diagnostic report |  | incorrect file applied | data issue; patient data issue
file, report, study, exam, diagnostic report |  | incorrect report displayed | display misread; patient data issue
file, report, study, exam, diagnostic report |  | incorrect file status change | data issue; patient data issue; application security issue
file, report, study, exam, diagnostic report |  | incorrect file status | data issue; patient data issue; application security issue
file, report, study, exam, diagnostic report |  | incorrect procedural verbiage in report | data issue
file, report, study, exam, diagnostic report |  | incorrect report format | data issue
file, report, study, exam, diagnostic report |  | incorrect report content | data issue
file, report, study, exam, diagnostic report |  | incomplete report | display misread; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | missing voice report | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | missing report | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | blank report | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | missing partial reports | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | missing patient indentifier | patient data issue; loss of data
file, report, study, exam, diagnostic report |  | conflicted report | data issue
file, report, study, exam, diagnostic report |  | missing study | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | fail to display report online | no display or display failure
file, report, study, exam, diagnostic report |  | unable to display report | incorrect display
file, report, study, exam, diagnostic report |  | display to different exam | incorrect display
file, report, study, exam, diagnostic report |  | unable to locate study | data issue; patient data issue
file, report, study, exam, diagnostic report |  | missing header | loss of data
file, report, study, exam, diagnostic report |  | missing data in report | loss of data
file, report, study, exam, diagnostic report |  | missing exam | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | fail to upload report | data issue; patient data issue
file, report, study, exam, diagnostic report |  | inconsistent report | data issue
file, report, study, exam, diagnostic report |  | incorrect report date or time stamped | data issue
file, report, study, exam, diagnostic report |  | incorrect image linked to report | data issue
file, report, study, exam, diagnostic report |  | update report with incorrect info | data issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to update report | data issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to save report | loss of data
file, report, study, exam, diagnostic report |  | unable to open report | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | read report delayed | improper or incorrect procedure or method; data issue
file, report, study, exam, diagnostic report |  | fail to read report | data issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to approve report | data issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to generate report | data issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to export report | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | fail to send report | Communication or Transmission Issue; data issue
|  | send report to wrong place | data issue
file, report, study, exam, diagnostic report |  | fail to validate study | protective measure issue; patient data issue
file, report, study, exam, diagnostic report |  | fail to retrieve study | loss of data; patient data issue
file, report, study, exam, diagnostic report |  | fail to identify report received | data issue; computer software issue
file, report, study, exam, diagnostic report |  | log off report without saving | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | index to incorrect patient | data issue; patient data issue
file, report, study, exam, diagnostic report |  | import study slow | data issue; patient data issue; device operates differently than expected
file, report, study, exam, diagnostic report |  | report crash | data issue; patient data issue; loss of data
file, report, study, exam, diagnostic report |  | report locked | data issue; application program issue
file, report, study, exam, diagnostic report |  | file stuck | difficult to open or close; data issue
file, report, study, exam, diagnostic report |  | file corrupted | data issue; patient data issue; loss of data
tab |  | tab changed unexpectedly | application program issue
osteotomy, tooth, dental |  | incorrect implant placement | positioning issue
snap |  | snap function unexpected | application program issue
hide, hidden |  | fail to hide display | incorrect display
render |  | unable to render 2D views | application program issue
render |  | fail to render image data | data issue
render |  | fail to provide volumetric rendering | application program issue
render |  | fail to display volumetric rendering | no display or display failure
render |  | fail to perform on screen rendering | application program issue
render |  | fail to perform off screen rendering | application program issue
render |  | inappropriate render color | use of incorrect control settings
render |  | inappropriate render transparency | use of incorrect control settings
render |  | unable to provide off-screen rendering | application program issue
render |  | unable to render 3D view | application program issue
ruler |  | incorrect readings | incorrect measurement
name, rename |  | unable to rename | device markings issue; use of device issue
name, rename |  | invalid name | patient data issue; use of device issue
name, rename |  | undesired name changes | patient data issue; use of device issue
preset |  | fail to start with preset value | use of incorrect control settings
preset |  | fail to define presets | device difficult to setup or prepare
preset |  | inappropriate preset value | use of incorrect control settings
linux |  | fail to run ubuntu linux | computer operating system issue
linux |  | incorrect linux version | operating system version or upgrade problem
web, browser |  | fail to access through web browser | unauthorized access to computer system; application security issue; computer system security issue
web, browser |  | fail to capture inputs | computer software issue
web, browser |  | fail to transmit inputs | computer software issue
web, browser |  | low browser window resolution | Image Resolution Poor
web, browser |  | fail to support browsers | application program issue
stream |  | fail to stream rendered view | incorrect display
stream |  | fail to display video stream | incorrect display
stream |  | fail to create rendered stream | data issue
stream |  | low rendered resolution | Image Resolution Poor
stream |  | low rendered frames | Image Resolution Poor
stream |  | fail to compress stream | data issue
stream |  | fail to transmit stream | data issue
deploy |  | fail to deploy software on hardware | computer software issue
map |  | incorrect measurement mapping | data issue
map |  | incorrect map displayed | incorrect display
map |  | duplicated measurement mapping | patient data issue
map |  | fail to map linearly | data issue
center |  | fail to center view | reset issue
resize |  | unable to resize viewing panel | application program issue
resize |  | fail to resize image | reset issue
workspace |  | fail to save workspace | loss of data; data back-up problem
workspace |  | fail to restore previous workspace | no display or display failure; incorrect display; loss of data
viewer |  | viewport frozen | application network issue
viewer |  | inaccuracy distance measurement | incorrect measurement
viewer |  | unexpected closure | application program issue
viewer |  | viewer no responding | application program issue; device inoperable
viewer |  | viewer crash | application program issue; device inoperable
viewer |  | incorrect read status | device operates differently than expected
viewer |  | viewed delayed | application network issue
clinical |  | incorrect clinical data | data issue
harddrive, hard drive, drive, Disc drive |  | insufficient storage space | inadequate storage
harddrive, hard drive, drive, Disc drive |  | disc drive failure | loss of data; data back-up problem; device inoperable
harddrive, hard drive, drive, Disc drive |  | hard drive failure | loss of data; data back-up problem; device inoperable
archive |  | wrong file in archive | data issue
archive |  | archive failed | data issue
archive |  | incomplete data archived | loss of data
subsidence |  | unintended part subsidence | migration of device or device component
virus, software virus, wannacry virus |  | ransomeware virus attack | application security issue
virus, software virus, wannacry virus |  | software virus infected | application security issue
decimal |  | decimal place shifted | data issue
discriminator |  | device dropped | component falling
string |  | duplicated strings | data issue
workstation |  | fail to save info on workstation | loss of data; data back-up problem
status |  | incorrect status change | data issue
status |  | incorrect report status | data issue
template |  | fail to create template | application program issue
template |  | incorrect template to measure | incorrect measurement; imprecision
score, scoring |  | incorrect score calculation | data issue; patient data issue; parameter calculation error due to software problem
score, scoring |  | inaccurate score calculation | data issue; patient data issue; parameter calculation error due to software problem
demographic |  | incorrect demographic | data issue
value |  | value changed incorrectly | data issue
value |  | inaccurate value | data issue; imprecision
value |  | incorrect display value | incorrect display
value |  | fail to update accordingly | data issue
watermark |  | report watermarked off early | data issue
nodule |  | incorrect nodule information | data issue
assign |  | assign image to different patient | patient data issue
assign |  | fail to assign study | data issue
navigate, navigation |  | navigation frozen | human-device interface issue; use of device issue
navigate, navigation |  | fail to navigate to menu | human-device interface issue; use of device issue

## Run
run words | category | failure mode | Hazard
lung recruitment, frequency lung recruitment |  | fail to assess | use of device issue
510k, 510(k), premarket approval | regulatory | lack of 510k approval | compliance issue
error code, error message, error status, fault message, system message, 911 message, explanation message | safety | prompt 911 message failed | device alarm system issue
abnormal use | usability | lead to unacceptable risk | device unsafe to use in environment
abnormal use |  | lack of abnormal use consideration | improper or incorrect procedure or method
absorbed dose |  | inappropriate ionizing radiation energy | output energy incorrect
AC inlet filter, inlet filter | electrical | noise passed | failure to read input signal
AC power, AC main power, AC supply, AC/DC power |  | switch to automobile power failed | electrical power problem
AC power, AC main power, AC supply, AC/DC power |  | switch to AC power failed | electrical power problem
AC power, AC main power, AC supply, AC/DC power |  | power out of range | electrical power problem
AC power, AC main power, AC supply, AC/DC power |  | inappropriate rated AC power | electrical power problem
AC power, AC main power, AC supply, AC/DC power |  | inappropriate power cordset | component incompatible
AC power, AC main power, AC supply, AC/DC power |  | missing safety mark | device markings issue
AC power, AC main power, AC supply, AC/DC power |  | fail to indicate AC power | no display or display failure
AC power, AC main power, AC supply, AC/DC power | power | fail to check voltage periodically | failure of device to self-test
acceptable dose |  | over dose | incorrect or inadequate test results
acceptance criteria |  | lack of acceptance criteria | use of device issue
acceptance criteria |  | inappropriate variable value | use of incorrect control settings; data issue
acceptance criteria |  | inappropriate self-diagnostics | failure of device to self-test
acceptance criteria |  | inappropriate resource allocation | use of device issue
acceptance criteria |  | inappropriate memory management | computer software issue
acceptance criteria |  | inappropriate fault handling | use of device issue
acceptance criteria |  | inappropriate data flow | compliance issue
acceptance criteria |  | inappropriate control flow | compliance issue
acceptance criteria |  | inappropriate boundary condition | use of device issue
acceptance criteria |  | inappropriate acceptance criteria | use of device issue
acceptance criteria |  | inadequate acceptance criteria | use of device issue
acceptance criteria | test | fail to meet acceptance criteria | use of device issue
access code, license code |  | invalid license code | human-device interface issue
access code, license code | safety | invalid access code | human-device interface issue
accompanying documentation, accompanying document, technical description |  | lack of transport document | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | fail to provide manufacturer information | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | fail to provide electronical document | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | fail to provide hard copy | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | lack of storage document | lack of maintenance documentation or guidelines
accompanying documentation, accompanying document, technical description |  | lack of operation document | device handling issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | lack of maintenance and repair document | lack of maintenance documentation or guidelines
accompanying documentation, accompanying document, technical description |  | lack of installation document | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description |  | lack of disposal document | device markings issue; instruction for use issue
accompanying documentation, accompanying document, technical description | standard | inadequate document | device markings issue; instruction for use issue
non-active surgical implant, surgical implant, active surgical implant |  | loss of power | loss of power
activity limit | motion | exercise over activity limit | use of device issue
additive port |  | port broken or damaged | break; fluid leak
adjustable volume | usability | fail to adjust volume | use of incorrect control settings
adjustable volume |  | adjustable volume out of specification | use of incorrect control settings
advanced light engine, ALE, ALEs |  | power out of range | electrical power problem
advanced light engine, ALE, ALEs |  | loss of power indication | no display or display failure
advanced light engine, ALE, ALEs |  | loss of functionality within force limit | mechanical issue
advanced light engine, ALE, ALEs |  | loss of connection | disconnection
advanced light engine, ALE, ALEs |  | lack of power control | power conditioning issue
advanced light engine, ALE, ALEs |  | fail to generate 5% illumination | optical issue
advanced light engine, ALE, ALEs | component | fail to enable laser source output | no device output
adverse event | usability | no apparent | use of device issue
air bubble | infusion | unable to interrupt infusion | excess flow or over-infusion
air bubble | infusion | single air bubble size out of range | defective item
air bubble | infusion | fail to activate air-in-line alarm | device alarm system issue
air bubble | infusion | cumulative bubble volume out of range | improper gas output
air cleaner, air eliminator | environment | air eliminator defective | device cleaning issue
air embolism | infusion | fail to stop infusion | excess flow or over-infusion
air embolism | infusion | fail to allow air embolism | device clogged
air intake | environment | dust ingress | leak
air pump, air pump assembly | electrical | clogged | device clogged
air valve, oxygen valve | ventilator | valve stuck | sticking; difficult to open or close
air valve, oxygen valve | ventilator | valve failure | defective component
air vent |  | loose air vent cover | fluid leak
air vent |  | filter leakage | fluid leak
air vent |  | air vent leakage | fluid leak
air vent |  | air vent distal end leakage | fluid leak
alarm indicator | usability | alarm indicator not visible | alarm not visible
alarm limit, limit of alarm |  | out of limits | device alarm system issue
alarm limit, limit of alarm |  | fail to preset limit correctly | device alarm system issue
alarm limit, limit of alarm | safety | fail to set limit correctly | device alarm system issue
alarm log |  | missing data | data issue
alarm log | user | lack of time-stamped alarm | data issue
alarm system |  | loss of alarm | device alarm system issue
alarm system |  | failure of message-service | device alarm system issue; device maintenance issue
alarm system |  | failure of message-leads off | device alarm system issue
alarm system |  | failure of message-check electrode | failure of device to self-test
alarm system |  | failure of message-battery status | device alarm system issue
alarm system |  | failure of gas-leak/loss | device alarm system issue; improper gas output
alarm system | usability | failure of check-catheter | failure of device to self-test
alarm threshold, alarm condition |  | fail to activate alarm | device alarm system issue
alarm threshold, alarm condition |  | fail to configure alarm threshold | device alarm system issue
alarm threshold, alarm condition | safety | high alarm threshold | device alarm system issue
alert type |  | insufficient alert type | improper alarm
alert type | usability | inappropriate alert type | improper alarm
alternating current | electrical | fail to run on AC | failure to run on AC/DC
ambient noise | environment
ambient temperature | environment
angiographic catheter
ankle brace |  | unstable ankle brace | unstable
ankle brace |  | reduced wearing time | material integrity issue
ankle brace |  | reduced washable | device cleaning issue
ankle brace |  | less flexible | material rigid or stiff
ankle brace |  | inflexible ankle brace | material rigid or stiff
ankle brace |  | inappropriate stability change | unstable
ankle brace |  | heavy weight | patient-device incompatibility
ankle brace |  | fail to fit into shoes | fitting problem
ankle brace |  | elastic plastic layer | defective component
ankle brace | brace | anatomical shape failed | patient-device incompatibility
anode wire | electrical | un-insulated anode wire | electrical shorting
App Key, AppKey |  | invalid key | self-activation or keying
App Key, AppKey |  | fail to retrieve a forgotten key | self-activation or keying
App Key, AppKey | security | fail to identify the user | inadequate user interface
application program |  | unintended program shut down | application program issue
application program |  | program upgrade failed | application program version or upgrade problem
application program |  | program frozen | application program issue
application program |  | nonfunctional | no device output
application program | software | incorrect program version | application program version or upgrade problem
architecture specification |  | lack of redundancy function | use of device issue
architecture specification |  | lack of partitioning of functionality | use of device issue
architecture specification |  | lack of fail-safe function | no fail-safe mechanism
architecture specification |  | lack of diversity function | use of device issue
architecture specification |  | lack of defensive design | protective measure issue
architecture specification |  | lack of components with high-integrity characteristics | compliance issue
architecture specification |  | fail to consider the IT-NETWORK specification | application network issue
architecture specification |  | fail to consider test interval duration and diagnostic coverage | incorrect or inadequate test results
architecture specification |  | fail to consider systematic failure | compliance issue
architecture specification |  | fail to consider risk control measures to subsystems and components | protective measure issue
architecture specification |  | fail to consider protection from reasonably foreseeable misuse | use of device issue
architecture specification |  | fail to consider maintainability | device maintenance issue
architecture specification |  | fail to consider failure modes of components and their effects | compliance issue
architecture specification | standard | fail to consider common cause failures | compliance issue
arm motion | motion | unintended arm motion | Unintended Arm Motion
arrhythmia detection | CRM | undetected arrhythmia | failure to sense
arrhythmia detection |  | inappropriate arrhythmia detection | device sensing issue
arterial pressure |  | fail to alarm failure of arterial pressure | pressure issue; device alarm system issue
articulating bar | brace | articulating bar low rigidity | material rigid or stiff
artificial heart |  | inadequate artificial heart type | use of device issue
artificial heart |  | fail to drive partial fill | filling problem
artificial heart | cardiology | fail to drive full eject | failure to deploy
aseptic presentation |  | bio-contamination | device contamination with biological material; delivered as unsterile product
aseptic processing |  | sterile product mishandled | device handling issue; device disinfection or sterilization issue
aseptic processing |  | ineffectively flushing airflow pattern | difficult to flush
aseptic processing |  | cross contamination | bacterial contamination of device
aseptic processing |  | fail to monitor contamination | bacterial contamination of device
aseptic processing |  | fail to maintain sterility | Delivered As Unsterile Product; Failure to Sterilize
aseptic processing |  | fail to maintain positive pressure | Decrease in Pressure
aseptic processing |  | fail to detect contamination | bacterial contamination of device
aseptic processing |  | container mishandled | device handling issue; device disinfection or sterilization issue
aspiration |  | incomplete aspiration | aspiration issue
aspiration | ventilator | excessive aspiration | aspiration issue
atrial backup pulse | CRM | atrial arrhythmia triggered | self-activation or keying
atrial channel | CRM | atrial channel inhibited | communication or transmission issue
atrial inhibition | CRM | inappropriate atrial inhibition | No Pacing
atrial rate | CRM | high atrial rate | atrial rate
atrial threshold, atrial auto threshold | CRM | inappropriate automatic atrial threshold search | capturing issue
atrial threshold, atrial auto threshold | CRM | fail to capture | Failure to Capture
atrial track | CRM | atrial tracking too fast | device operates differently than expected
audible feedback | usability | no audible feedback | not audible alarm
audible feedback | usability | inaudible or unclear audible feedback | prompts will not clear
audible feedback | usability | inappropriate audible feedback | device alarm system issue
audible noise | usability
audible prompt | usability | no audible prompt | No Voice Prompts
audible prompt | usability | inaudible or unclear audible prompt | Inaudible Voice Prompts
audible prompt | usability | inappropriate audible prompt | Inappropriate Prompts
audible tone | usability | inappropriate audible tone | improper alarm
audible tone | usability | fail to adjust volume | use of incorrect control settings
audio alarm, audible alarm,audible alarms, audio notification, audio notifications |  | low audible alarm | low audible alarm
audio alarm, audible alarm,audible alarms, audio notification, audio notifications |  | short alarm period | device alarm system issue
audio alarm, audible alarm,audible alarms, audio notification, audio notifications |  | loss of audible alarm | not audible alarm
audio alarm, audible alarm,audible alarms, audio notification, audio notifications |  | fail to silence alarm | false alarm
audio alarm, audible alarm,audible alarms, audio notification, audio notifications | usability | fail to configure audio alarm | device alarm system issue
audio bolus |  | low audio volume | use of device issue
audio bolus |  | inappropriate increment amount | inaccurate delivery
audio bolus |  | fail to enable audio bolus | failure to deliver
audio bolus |  | fail to disable audio bolus | device remains activated
audio bolus | pump | audio tones fail to indicate the amount programmed | not audible alarm
audio input | usability | loss of audio input | failure to read input signal
audio jack |  | low audio jack capacity | fitting problem; component incompatible
audio jack | usability | audio jack mismatched | fitting problem; component incompatible
audio latency | usability | audio latency over specification | communication or transmission issue
audio output | usability | loss of audio output | no device output
audio system |  | two-way audio communication failed | communication or transmission issue
audio system |  | loss of audio output | no device output
audio system |  | ambient noise high | ambient noise issue
audio system | usability | ambient echo high | audible noise
auditory indicator, auditory feedback, auditory feedbacks | usability | loss of auditory feedback | not audible alarm
auditory indicator, auditory feedback, auditory feedbacks | usability | loss of audio notification | not audible alarm
auto stop | movement | fail to auto stop | Failure to Auto Stop
auto-fill, auto fill | control | inability to auto-fill | Inability to Auto-fill
automatic off | software | unable to automatic off | Failure to Auto Stop
automatic off |  | lack of automatic off feature | Failure to Auto Stop
automobile power | power | switch to automobile power failed | electrical power problem
auto-reset, auto reset |  | fail to reset error message automatically | reset issue
auto-reset, auto reset | software | fail to reset alarm automatically | device alarm system issue; reset issue
back pain | clinical | pain treatment failed | therapeutic or diagnostic output failure
backup mode | CRM | switch to backup mode incorrectly | human-device interface issue
backup mode | CRM | loss of therapy | therapeutic or diagnostic output failure
backup power, back-up power, back-up powers, back-up source, back-up sources, backup source, backup sources |  | low battery | Low Battery
backup power, back-up power, back-up powers, back-up source, back-up sources, backup source, backup sources |  | fail to convert to back-up | failure to convert to back-up
backup power, back-up power, back-up powers, back-up source, back-up sources, backup source, backup sources |  | backup power activation failed | battery issue; Difficult to Deploy
backup power, back-up power, back-up powers, back-up source, back-up sources, backup source, backup sources | power | backup activation failed | failure to deploy
bacteria filter, bacterial filter, bacterial filters | ventilator | replace delayed | device maintenance issue
bacteria filter, bacterial filter, bacterial filters | ventilator | missing filter | component missing
bacteria filter, bacterial filter, bacterial filters | ventilator | air contaminated | device contamination with biological material
bag hanger, hanger |  | unable to withstand bag or bottle | mechanical issue
balance cell, cell balancing | battery | loss of cell balancing | unstable
balance cell, cell balancing |  | cell balancing stuck on | sticking
balloon catheter |  | out of dimension | defective component
balloon catheter |  | inappropriate fit | fitting problem
balloon catheter |  | inappropriate balloon type | deflation issue; Inflation Issue
basal drug volume | drug | out of range | defective item
basal flow rate, basal rate |  | uneven basal rate | inaccurate delivery
basal flow rate, basal rate |  | unable to repeat | inaccurate delivery
basal flow rate, basal rate |  | unable to cancel basal rate | use of device issue
basal flow rate, basal rate |  | rate out of range | defective item
basal flow rate, basal rate |  | incorrect rate calculation | parameter calculation error due to software problem
basal flow rate, basal rate |  | inadequate basal rate period | Output Issue
basal flow rate, basal rate |  | inaccurate rate | inaccurate delivery
basal flow rate, basal rate | pump | fail to generate basal rate | inaccurate delivery
basal pattern | pump | fail to block basal pattern selection | use of device issue; infusion or flow issue
basal profile |  | lack of start time | reset issue
basal profile |  | lack of basal rate | use of device issue
basal profile | pump | inadequate basal profile | use of incorrect control settings
basal stroke | pump | incorrect stroke calculation | parameter calculation error due to software problem
basal type |  | inadequate basal type | inadequate user interface
basal type |  | fail to select basal type | use of device issue; infusion or flow issue
basal type | pump | fail to block basal type selection | use of device issue; infusion or flow issue
basic safety |  | unacceptable risk under single fault condition | device unsafe to use in environment
basic safety |  | unacceptable risk under normal condition | device unsafe to use in environment
basic safety |  | lack of verification | protective measure issue
basic safety |  | degradation of basic safety | device unsafe to use in environment
basic safety | safety | absence of basic safety | device unsafe to use in environment
battery charger, car charger, car chargers |  | output under voltage | electronic property issue
battery charger, car charger, car chargers |  | output over voltage | electronic property issue
battery charger, car charger, car chargers |  | no output current | electronic property issue
battery charger, car charger, car chargers |  | loss of function | no device output
battery charger, car charger, car chargers |  | intermittent connection | loose or intermittent connection
battery charger, car charger, car chargers |  | inadequate capacity | Device-Device Incompatibility
battery charger, car charger, car chargers |  | fail to charge fully depleted battery | Failure to Charge
battery charger, car charger, car chargers |  | fail to charge battery to full | Failure to Charge
battery charger, car charger, car chargers | power | charger malfunction | charging issue; device operates differently than expected; Product Quality Issue
battery housing, housing of battery, housing of batteries, housings of batteries, battery pack, battery latch, battery wells, battery well, battery door, battery doors, battery slot, battery slots |  | incorrect polarity | device misassembled during manufacturing or shipping; battery issue
battery housing, housing of battery, housing of batteries, housings of batteries, battery pack, battery latch, battery wells, battery well, battery door, battery doors, battery slot, battery slots |  | inappropriate fit | fitting problem
battery housing, housing of battery, housing of batteries, housings of batteries, battery pack, battery latch, battery wells, battery well, battery door, battery doors, battery slot, battery slots |  | fail to engage | disconnection
battery housing, housing of battery, housing of batteries, housings of batteries, battery pack, battery latch, battery wells, battery well, battery door, battery slot | battery | battery gas no vent | improper gas output
battery interface | battery | user confused | inadequate user interface
battery level | battery | incorrect battery level | battery issue; human-device interface issue
battery level |  | fail to indicate battery level | no display or display failure
battery pack, battery cell, onboard battery |  | undesired activation of redundant voltage protection | Premature Deployment
battery pack, battery cell, onboard battery |  | overheat | overheating of device or device component
battery pack, battery cell, onboard battery |  | low battery | Low Battery
battery pack, battery cell, onboard battery |  | internal battery short | electrical shorting; battery issue
battery pack, battery cell, onboard battery |  | incorrect communication information | no information
battery pack, battery cell, onboard battery |  | inappropriate storage environment or condition | environmental control or utility issue
battery pack, battery cell, onboard battery |  | fail to provide redundant voltage protection | protective measure issue
battery pack, battery cell, onboard battery |  | fail to communicate with driver | communication or transmission issue
battery pack, battery cell, onboard battery |  | battery pack dropped | component falling
battery pack, battery cell, onboard battery |  | battery open | circuit failure
battery pack, battery cell, onboard battery |  | battery leakage | battery issue
battery pack, battery cell, onboard battery |  | battery fuse open | Fail-Safe Mechanism Issue
battery pack, battery cell, onboard battery | battery | battery corrosion | Corrosion
battery power | battery | unintended power up | power conditioning issue
battery power | battery | reduced battery life cycles | battery issue; material integrity issue
battery power | battery | power restoration | power conditioning issue
battery power | battery | power out of range | power conditioning issue
battery power | battery | power on or off switch failed | loss of power
battery power | battery | power interrupted | loss of power
battery power | battery | power frequency out of range | Power Source Issue | frequency out of specification
battery power | battery | power control failed | power conditioning issue
battery power | battery | power consumption out of range | electrical power problem
battery power | battery | power calculation error | Power Calculation Error due to Software Problem
battery power | battery | overheat | overheating of device or device component
battery power | battery | non-rechargeable battery | battery issue
battery power | battery | low impedance | Low Impedance
battery power | battery | low battery | Low Battery
battery power | battery | loss of power indication | no display or display failure
battery power | battery | loss of power | loss of power
battery power | battery | loss of detection | failure to sense
battery power | battery | loss of battery measurement | battery issue; measurement system incompatibility
battery power | battery | loss of alarm | device alarm system issue
battery power | battery | lack of short circuit module | electrical shorting
battery power | battery | lack of protection circuit | circuit failure
battery power | battery | lack of over voltage module | battery issue; no fail-safe mechanism
battery power | battery | lack of over heat module | overheating of device or device component
battery power | battery | lack of over current module | battery issue; no fail-safe mechanism
battery power | battery | lack of over charge module | charging issue
battery power | battery | lack of CB certified | battery issue
battery power | battery | internal battery short | electrical shorting; battery issue
battery power | battery | intermittent loss of power | loss of power
battery power | battery | Insufficient power capacity | Device-Device Incompatibility
battery power | battery | incorrect ratio of electrolyte salt | chemical issue
battery power | battery | incorrect electrolyte solvent | chemical issue
battery power | battery | incorrect connection polarity | misconnection; battery issue
battery power | battery | Inadequate capacity | Device-Device Incompatibility
battery power | battery | high impedance | High Impedance
battery power | battery | fail to supply power | electrical power problem
battery power | battery | fail to run on battery | battery issue
battery power | battery | fail to provide low battery indicator | no display or display failure
battery power | battery | fail to power up | failure to power-up
battery power | battery | fail to monitor battery temperature | battery issue; temperature issue
battery power | battery | fail to measure battery voltage | battery issue; measurement system incompatibility
battery power | battery | fail to indicate battery charging status | no display or display failure
battery power | battery | fail to communicate with driver | communication or transmission issue
battery power | battery | expired battery used | device expiration issue
battery power | battery | current drain | battery issue
battery power | battery | confused battery status | inadequate user interface
battery power | battery | complete loss of power | loss of power
battery power | battery | chemical release | chemical spillage
battery power | battery | catch fire | Fire
battery power | battery | battery switch failed | defective component; battery issue
battery power | battery | battery open | circuit failure
battery power | battery | battery leakage | battery issue
battery power | battery | battery increment high | battery issue
battery power | battery | battery depletion | Premature Discharge of Battery
battery power | battery | battery corrosion | Corrosion
battery replacement test |  | lack of unloaded battery test | incorrect or inadequate test results
battery replacement test | battery | lack of loaded battery test | incorrect or inadequate test results
battery replacement test |  | fail to activate alarm | device alarm system issue
battery temperature | battery | incapable to monitor battery temperature | battery issue; temperature issue
battery test |  | lack of unloaded battery test | incorrect or inadequate test results
battery test |  | lack of loaded battery test | incorrect or inadequate test results
battery test |  | inadequate readings | False Reading from Device Noncompliance
battery test | battery | fail to initiate test | failure of device to self-test
battery test |  | fail to indicate battery depletion | no display or display failure
battery test |  | fail to generate alarm | device alarm system issue
beam splitter | energy
bearing |  | part wear out | Naturally Worn
beat rate | driver | incapable to support clinically stable patient | use of device issue
beat rate | driver | fail to adjust beat rate | use of incorrect control settings
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | weight out of specification | patient-device incompatibility
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | usage misaligned | use of device issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | unmatched belt set | component or accessory incompatibility
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | strap torn or broken | torn material; break
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | strap shorten | device misassembled during manufacturing or shipping
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | stitching ripped or unraveled | Detachment of Device or Device Component
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | soft good worn | Naturally Worn
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | slip coating | Split
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | seam torn or broken | torn material; break
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | reduced life cycle | shelf life exceeded
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | part mismatched | component incompatible
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | out of dimension | size incorrect for patient
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | missing brand | device markings issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | material frayed or unraveling | Material Frayed, Unraveled Material
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | material degradation by inappropriate cleaning | device cleaning issue; degraded
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | low strength | mechanical issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | loss of support | use of device issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | lack of labeling | device markings issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | fail to resist normal activities | use of device issue
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | difficult to apply device | device difficult to setup or prepare
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | component torn | torn material
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | component detached | detachment of device component
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | clinic infection | bacterial contamination of device
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | belt size out of specification | size incorrect for patient
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | belt mismatched | component incompatible
belt set, traction belt, traction belts, lumbar belt, lumbar belts |  | accessory unattached | detachment of device component
belt set, traction belt, traction belts, lumbar belt, lumbar belts | brace | accessory missing | component missing
BG reminder, blood glucose Reminder, blood glucose Reminders | diabete | out of range | defective item
BG reminder, blood glucose Reminder, blood glucose Reminders | diabete | loss of function | no device output
BG reminder, blood glucose Reminder, blood glucose Reminders | diabete | fail to select option | use of device issue; infusion or flow issue
BG target | diabete | BG out of range | defective item
bifurcated tubular vascular graft |  | pressurized internal diameter out of specification | physical property issue
bifurcated tubular vascular graft |  | main tube usable length out of specification | physical property issue
bifurcated tubular vascular graft |  | internal diameter out of specification | physical property issue
bifurcated tubular vascular graft |  | branch usable length out of specification | physical property issue
bilevel positive airway pressure, BiPAP | ventilator | low pressure | Decrease in Pressure
bio-decontamination |  | inable to remove microbiological contamination | bacterial contamination of device
bio-decontamination |  | fail to reduce microbiological contamination | bacterial contamination of device
biofilm |  | biofilm coating in device | biofilm coating in device
biological indicator |  | fail to provide indicator | no display or display failure
biological material, animal origin, animal tissue, animal derivate |  | incorrect tissue type and site | biocompatibility issue
biological material, animal origin, animal tissue, animal derivate |  | incorrect nature of chemical treatment or modification | biocompatibility issue
biological material, animal origin, animal tissue, animal derivate |  | incorrect material origin | biocompatibility issue
biomechanical test, biomechanics test, biomechanics testing | clinical | biomechanics testing failed | failure of device to self-test
biventricular threshold | CRM | inappropriate pacing sequence | failure to pace or properly pace
bleed flow | physiology | over flow | excess flow or over-infusion
bleed flow |  | insufficient flow | insufficient flow or underinfusion
blood flow |  | reduced blood flow | insufficient flow or underinfusion
blood flow |  | insufficient blood flow | insufficient flow or underinfusion
blood flow |  | inappropriate low blood flow | improper flow or infusion
blood flow | clinical | inappropriate accelerated blood flow | improper flow or infusion
blood sugar | diabete | fail to check blood sugar periodically | failure of device to self-test
body temperature and pressure saturated, BTPS | ventilator | low pressure | Decrease in Pressure
body weight set |  | setup incorrectly | use of incorrect control settings
body weight set | physiology | inappropriate patient height and body weight relationship | patient-device incompatibility
bolus button | usability | lack of button | inadequate user interface
bolus history | pump | unable to access bolus history | unauthorized access to computer system
bolus history |  | lack of delivery time | reset issue
bolus history |  | lack of delivery date | Date-related software issue
bolus history |  | lack of bolus amount | failure to deliver
bolus maximum |  | bolus size out of range | inaccurate delivery
bolus maximum | pump | bolus amount out of range | inaccurate delivery
bolus mechanism | pump | fail to generate bolus | failure to deliver
bolus size | pump | inadequate bolus size | inaccurate delivery
bolus size | pump | bolus size mismatched | inaccurate delivery
bolus wizard |  | incorrect bolus wizard | inaccurate delivery
bolus wizard | pump | bolus amount out of range | inaccurate delivery
brake system | safety | unwanted movement | unintended movement
breath delivery |  | no delivery | delivery system failure
breath delivery | ventilator | delivery out of control | delivery system failure
breath delivery modality | ventilator | ventilation compromised | device operates differently than expected
breath mode, breath control, breath control servo |  | fail to control flow | infusion or flow issue
breath mode, breath control, breath control servo | ventilator | inadequate breathing mode | inadequate user interface
breath period |  | incorrect breath period | improper device output
breath type, type of breath, , type of breaths, , types of breath |  | lack of breath type | improper device output
breath type, type of breath, , type of breaths, , types of breath |  | incorrect breath type | improper device output
breath type, type of breath, , type of breaths, , types of breath |  | inadequate breath type | improper device output
breath type, type of breath, , type of breaths, , types of breath | ventilator | confused breath type | inadequate user interface
breathe atomizer |  | loose washer | detachment of device component
breather vent | ventilator | clogged | device clogged
brightness adjustment |  | fail to adjust settings | use of incorrect control settings
brightness adjustment |  | fail to adjust manually | device difficult to setup or prepare
bubble detector | measure | loss of function | no device output
cable length | dimension | cable length out of range | defective component
calibration check | test | fail to generate periodic check | failure of device to self-test
calibration check |  | fail to activate alert | device alarm system issue; failure to deploy
calibration data | data | loss of data | loss of data
calibration test | test | fail to activate test | failure of device to self-test
cannula guide | lead | no description | inadequate user interface
cannula guide |  | needle insertion difficult | Difficult to Insert
cannula guide |  | difficult to locate | dislodged or dislocated
capability requirement | standard | lack of capability requirement | use of device issue
capacitive coupling | electrical
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | incompatible connection | Device-Device Incompatibility
capture threshold | CRM | incorrect measurement of capture threshold | capturing issue; incorrect measurement
capture threshold | CRM | high capture threshold | High Capture Threshold
capture threshold | CRM | fail to identify capture threshold | capturing issue
capture threshold | CRM | capture threshold misinterpreted | capturing issue
capture threshold | CRM | capture threshold decreased | Low Sensing Threshold
carbohydrate unit | diabete | fail to reset carbohydrate ratio value | failure to reset; data issue
carbon absorption | ventilator
cardiac enzyme evaluation | chemical
cardiac output |  | output low | improper device output
cardiac output | cardiology | incapable to support clinically stable patient | use of device issue
cardiac output |  | inadequate cardiac output | improper device output
cardiac pacing | CRM | loss of pacing | No Pacing
cardiac signal | CRM | loss of signal | failure to read input signal
carrier frequency | communication | fail to penetrate obstacle | Telemetry Discrepancy
caster lock |  | long slide distance | device operates differently than expected
caster lock | safety | lack of caster lock | protective measure issue
catheter mount |  | part cracked | crack
catheter mount |  | part broken or damaged | break
catheter withdrawal inference | catheter
CE Mark |  | missing legal representative | device markings issue
CE Mark |  | missing legal manufacturer | device markings issue
CE Mark |  | fail to comply with CE Mark | compliance issue
CE Mark | regulatory | lack of CE mark | device markings issue
cervical collar | brace | restricted height adjustability | use of incorrect control settings
cervical collar | brace | over weighing | patient-device incompatibility
channel |  | incorrect output | improper device output
channel | communication | faulty channel invisible | no display or display failure
charge cable |  | plug to undesired connector | blocked connection
charge cable | electrical | inappropriate fit | fitting problem
charge path, discharging path, discharge path |  | charge path stuck opened | Failure to Charge
charge path, discharging path, discharge path | battery | charge path stuck closed | Failure to Charge
charge state, state of charge, state of charges |  | incorrect state of charge calculation | parameter calculation error due to software problem
charge state, state of charge, state of charges | battery | fail to display state of charge | no display or display failure
charge time | power | excessive charge time | delayed charge time
charge time | power | charge time extended | delayed charge time
check valve, vacuum break |  | valve blocked | Blockage within Device or Device Component; device clogged; failure to deliver; no flow
check valve, vacuum break |  | particulate contamination holding valve open | Particulates
check valve, vacuum break |  | fail to prevent gas contaminating | protective measure issue; contamination during use
check valve, vacuum break |  | fail to open enough | difficult to open or close
check valve, vacuum break | motor | fail to build vacuum up | suction issue
check valve, vacuum break |  | fail to allow air in and out | difficult to open or close
chemical indicator |  | fail to provide indicator | no display or display failure
chemical reaction | chemical | improper chemical reaction | Improper Chemical Reaction
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | false output | electronic property issue; False Device Output
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | self-dismantling | detachment of device component
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | incompatible connection | fitting problem
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB | electrical | circuit board broken or damaged | circuit failure
circuit breaker | electrical | circuit breaker tripped | component falling
circuit calibration | ventilator | fail to request calibration | calibration issue
cleaning agent |  | unqualified cleaning agent | device cleaning issue
cleaning agent |  | loss of product function | no device output
cleaning agent |  | expired cleaning agent | device expiration issue; device cleaning issue
cleaning agent | sterility | appearance compromised | Improper Chemical Reaction
cleaning cycle |  | fail to reach designated cycles | failure to disinfect
clearance angle | safety | angle out of range | defective item
clinical evaluation |  | incorrect clinical evaluation | use of device issue
clinical evaluation | clinical | inappropriate clinical evaluation | use of device issue
clinical history |  | unable to record event | data back-up problem
clinical history |  | lack of delivery stop event | Failure to Auto Stop
clinical history |  | inadequate clinical history data | data issue
clinical history | clinical | fail to access clinical history | human-device interface issue
clinician station | clinical | fail to retrieve patient treatment data | patient data issue
clinician station |  | fail to retrieve patient diagnostics data | therapeutic or diagnostic output failure; loss of data
clinician station |  | fail to manage data | data issue
clock speed | motion | clock speed out of range | imprecision
clock test | test | fail to issue RTC error | False Device Output
clock test | test | fail to generate periodic test | failure of device to self-test
closure integrity |  | fail to minimize ingress of microorganism | bacterial contamination of device
CO2 monitor, O2 monitor subassembly | measure | delayed monitoring | improper device output
coin cell, button cell, button cells | battery | low battery | Low Battery
colony forming unit | biotech
color accuracy | usability | image color accuracy low | poor quality image
color-blind, color blind | usability | fail to support color-blind | use of device issue
combination product |  | fail to mix as a single entity | Material Disintegration
common receptacle | component | improper fit to common receptacle | fitting problem
communication cable |  | loss of communication | communication or transmission issue
communication cable | component | cable short | defective component
communication interface | communication | loss of communication | communication or transmission issue
component failure | failure | loss of function | no device output
composite prosthesis
compound prosthesis, coated prosthesis
computer hardware | computer | computer hardware error | defective component
computer system | software | unauthorized access to computer system | unauthorized access to computer system
computer system security | software
condition of use, use condition |  | unable to confirm conditions of use | use of device issue
condition of use, use condition |  | lack of social conditions | use of incorrect control settings
condition of use, use condition |  | lack of personnel-specific conditions | use of incorrect control settings
condition of use, use condition | usability | lack of location-specific conditions | environmental control or utility issue
condition of use, use condition |  | inadequate conditions | use of incorrect control settings
conductor coil | electrical
configuration item |  | uncontrolled configuration version | use of incorrect control settings
configuration item |  | non unique configuration item | use of incorrect control settings
configuration item |  | lack of verification | protective measure issue
configuration item |  | lack of configuration item | use of incorrect control settings
configuration item |  | lack of change control | use of device issue
configuration item | software | inadequate configuration | use of incorrect control settings
connecting rod, heater rod | component
connection point |  | fluid residue | Particulates
connection point |  | difficult to close | leak
connector pin | mechanical | insufficient current capacity | Device-Device Incompatibility
connector pin | mechanical | connector pin pulled out | device misassembled during manufacturing or shipping
connector pin | mechanical | connector pin failure | defective component
connector ring |  | leakage pathway between connector rings | electrical issue
connector ring |  | connector ring separated | failure to separate
connector ring | mechanical | connector ring crimp reduced | mechanical issue
connector spring | mechanical | connector spring damaged | break
contact len | vision | problem with trial set | use of incorrect control settings
control measure, risk control, protective measure |  | unperceivable safety information | no information
control measure, risk control, protective measure |  | misunderstandable safety information | no information
control measure, risk control, protective measure |  | lack of verification | protective measure issue
control measure, risk control, protective measure |  | lack of validation | protective measure issue
control measure, risk control, protective measure |  | lack of software safety class | compliance issue
control measure, risk control, protective measure |  | lack of risk control measure requirement | protective measure issue
control measure, risk control, protective measure |  | lack of risk control measure | protective measure issue
control measure, risk control, protective measure |  | lack of protective measure | protective measure issue
control measure, risk control, protective measure |  | lack of inherent safety by design | protective measure issue
control measure, risk control, protective measure |  | lack of information for safety | no information
control measure, risk control, protective measure |  | ineffective risk control measure | protective measure issue
control measure, risk control, protective measure |  | inadequate risk control measure | protective measure issue
control measure, risk control, protective measure |  | fail to support correct use | use of device issue
control measure, risk control, protective measure |  | fail to reduce user interface-related risk | human-device interface issue
control measure, risk control, protective measure |  | fail to implement control measure | protective measure issue
control measure, risk control, protective measure |  | fail to define risk control measure | protective measure issue
control measure, risk control, protective measure | risk | conscious disregard safety information | no information
control switch | electrical
control system |  | sneak software | application program issue
control system |  | sneak circuit | circuit failure
control system |  | confused control system | inadequate user interface
control system |  | complex control system | inadequate user interface
control system | control | power outage | power conditioning issue
control system |  | moisture in system | Moisture or Humidity Problem
control system |  | lightning strike | electrical issue
control system |  | inadvertent activation | deployment issue
control system |  | grounding failure | Grounding Malfunction
control valve, flow control valve, blender flow control valve | component | valve stuck | sticking; difficult to open or close
cooling module |  | incompatible module | component incompatible
cooling module | chemical | loss of function | no device output
cooling system | chemical | loss of function | no device output
cord anchorage | electrical | loose cord anchorage | detachment of device component
cord connection, connection of cord, connection of cords | component | loss of connection | connection issue
correct use | usability | lead to unacceptable risk | no fail-safe mechanism
cost of good, COG, COGs | business | high cost | Market Issue
cost of good, COG, COGs |  | COG out of range | Market Issue
CPU test |  | fail to generate periodic test | failure of device to self-test
CPU test | test | fail to activate alert | device alarm system issue; failure to deploy
crank arm | mechanical
crimp tube | mechanical | insufficient crimp | mechanical issue
cross connection | electrical
cross reactivity | electrical
CRT, cathode ray tube | electrical
CSTD, closed system transfer device, chemolock, chemoclave |  | transfer lock opened | mechanical issue
CSTD, closed system transfer device, chemolock, chemoclave |  | closed system transfer device leakage | fluid leak
CSTD, closed system transfer device, chemolock, chemoclave |  | closed system transfer device detached | fluid leak
CSTD, closed system transfer device, chemolock, chemoclave |  | closed system transfer device cracked | crack
CT scan, CT image, CT map | image | wrong scan condition | Electromagnetic Interference (EMI); use of incorrect control settings
CT scan, CT image, CT map |  | incorrect color map | incorrect display; image display error
CT scan, CT image, CT map | image | unclear image | poor quality image
cup seal | mechanical
current control | control | current out of control | electronic property issue
current limiter | electrical
current source | electrical
cursor function | software | fail to display point data | no display or display failure
cutoff point | mechanical
cylinder assembly |  | sliding surface worn | Naturally Worn
cylinder assembly |  | roller slot | defective component
cylinder assembly |  | part broken | break
cylinder assembly |  | mechanical failure | mechanical issue
cylinder assembly | motor | loose bolt retaining nut | detachment of device component
cylinder assembly |  | fastener failed | defective component
cylinder assembly |  | expoxy holding failed | misassembled
D value |  | inappropriate exposure time required | imprecision
data acquisition unit | data
data cable | cable | loss of connection | connection issue
data communication, communication data | data | unsecured data | application security issue
data communication, communication data | data | loss of communication | communication or transmission issue
data communication, communication data | data | data corrupted | data issue
data connection, data connector | data | loss of connection | connection issue
data interface | data | loss of communication | communication or transmission issue
daylight savings time |  | fail to unresponse daylight savings time | reset issue
daylight savings time | time | fail to response daylight savings time | reset issue
default value | data | incorrect default value | reset issue; data issue
defibrillation therapy | CRM | loss of defibrillation therapy | failure to pace or properly pace
defibrillation therapy | CRM | inappropriate defibrillation therapy | failure to pace or properly pace
defibrillator paddle | CRM
defibrillator subassembly | CRM
delay line | control
delivery pattern | pump | unable to repeat | device operates differently than expected
delivery system |  | low torsional bond strength | mechanical issue
delivery system |  | low tensile bond strength | delivery system failure
delivery system |  | inappropriate retraction | delivery system failure
delivery system |  | inappropriate advancement | delivery system failure
delivery system | equipment | delivery system failure | delivery system failure
demographic information | data | inadequate information | no information
depyrogentation |  | fail to remove endotoxins | biocompatibility issue
detection interval | CRM | inappropriate interval settings | use of incorrect control settings
developable surface | clinical | developable surface invisible | no display or display failure
device explant | CRM | misleading device explant | device handling issue
device information | data | inadequate device information | no information
device output | control | no device output | no device output
device output | control | improper device output | improper device output
DHF, design history file, design history files, deliverable checklist, deliverable checklists | standard | inadequate DHF | device markings issue
diabetes management |  | ineffective algorithm | programming issue
diabetes management |  | incorrect algorithm | programming issue
diabetes management | diabete | fail to require BG value | Unable to Obtain Readings; data issue
diagnostic output | clinical
diagnostic radiation, therapeutic radiation | CRM | transient inhibition of cardiac pacing | No Pacing
diagnostic radiation, therapeutic radiation | CRM | single bit flip | improper device output
diagnostic radiation, therapeutic radiation |  | unintended radiation exposure | Radiation overexposure
diagnostic radiation, therapeutic radiation | CRM | multiple bit flips | improper device output
diagnostic test |  | inadequate diagnostic tests | incorrect or inadequate test results
diagnostic test | service | fail to provide diagnostic test | failure of device to self-test
diagnostic ultrasound | environment | device damaged | device damaged prior to use
directional valve, dump valve, cylinder valve | component | valve stuck | sticking; difficult to open or close
discharge of battery, battery discharge | power | premature discharge of battery | Premature Discharge of Battery
discharge pulse | CRM | inappropriate pulse period | electronic property issue
disinfecting agent |  | unqualified disinfecting agent | failure to disinfect
disinfecting agent |  | loss of product function | no device output
disinfecting agent |  | expired disinfecting agent | failure to disinfect
disinfecting agent | sterility | appearance compromised | material integrity issue
display input | video | invalid input | failure to read input signal
display output |  | low resolution | Image Resolution Poor
display output | video | loss of output | no device output
display port |  | unmatched port | component incompatible
display port | component | display port damaged | no display or display failure
display size | computer | inappropriate display size | incorrect display
distal end, distal tip, patient end |  | rough distal tip | Scratched Material
distal end, distal tip, patient end | lead | loss of electrical connection | intermittent continuity
distal end, distal tip, patient end |  | inadequate weld strength | mechanical issue; Solder Joint Failure
distal end, distal tip, patient end |  | inadequate crimp | mechanical issue
distal end, distal tip, patient end |  | blood clogging | device clogged
distilled water | physiology | contaminated | contamination during use
distribution cycle | test | shipping test failed | failure of device to self-test
docking position, docking location, docking locations, dock location, dock locations |  | inappropriate docking location | positioning issue
docking position, docking location, docking locations, dock location, dock locations | motion | docking position out of specification | positioning issue
don-on, don-off, don on, don off | motion | difficult to apply device | device difficult to setup or prepare
dose mapping
drip chamber |  | inappropriate drop permit | infusion or flow issue
drip chamber |  | drop amount out of spec | infusion or flow issue
drip chamber |  | drip chamber discolored | Material Discolored
drip chamber |  | drip chamber cracked | crack
drip chamber |  | drip chamber cloudy | Material Discolored; Material Opacification
driveline pressure |  | incapable to support clinically stable patient | use of device issue
driveline pressure | driver | fail to adjust driveline pressure | use of incorrect control settings
driver housing |  | static interference | electromagnetic compatibility issue
driver housing |  | misaligned parts | failure to align
driver housing |  | liquid ingress | fluid leak
driver housing |  | inappropriate fit | fitting problem
driver housing |  | inability to full insert | Difficult to Insert
driver housing |  | exposed internal component | Item Contaminated during Manufacturing or Shipping
driver housing |  | electro-magnetic interference (EMI) | Electromagnetic Interference (EMI)
driver housing |  | dust ingress | leak
driver housing |  | detachment of device component | detachment of device component
driver housing | pump | cylinder seal contamination | Item Contaminated during Manufacturing or Shipping
drug administrator |  | users confused | inadequate user interface
drug administrator | usability | hard to use | use of device issue
drug availability | drug | fail to interrupt infusion | excess flow or over-infusion
drug concentration | drug | out of safe limits | protective measure issue
drug data, drug library, drug database |  | unable to edit | human-device interface issue
drug data, drug library, drug database |  | unable to download library | data issue
drug data, drug library, drug database |  | incorrect drug name | medication error
drug data, drug library, drug database |  | incorrect concentration | medication error
drug data, drug library, drug database |  | fail to replicate | use of device issue
drug data, drug library, drug database |  | fail to finalize | use of device issue
drug data, drug library, drug database | drug | database outdated | data issue
drug date | drug | drug expired | device expiration issue; medication error
drug flow |  | out of range | defective item
drug flow | drug | insufficient drug flow | medication error
drug reservoir | drug | low reservoir | infusion or flow issue
drug safe limit |  | exceed drug safe limits | medication error
drug safe limit | drug | drug concentration not in safe range | medication error
drug volume |  | out of range | defective item
drug volume |  | low volume | use of device issue; volume accuracy issue
drug volume | drug | cumulative volume out of range | volume accuracy issue
dual bolus | pump | lack of square wave portion | inaccurate delivery
dual bolus |  | lack of normal portion | inaccurate delivery
dual wave bolus |  | inappropriate second delivery | delivery system failure
dual wave bolus | pump | inappropriate immediate delivery | delivery system failure
dual wave bolus |  | inappropriate delivery duration | inaccurate delivery
dual wave bolus |  | fail to support dual wave bolus | use of device issue
dummy load | mechanical
duty cycle, duty-cycle, duty-cycles |  | fail to reach designated cycles | failure to cycle
duty cycle, duty-cycle, duty-cycles | quality | duty cycle low | device operates differently than expected
earth ground, ground strap, ground post |  | fail to label standard symbol | device markings issue
earth ground, ground strap, ground post | electrical | loss of earth protection | protective measure issue
easy bolus | pump | unintended program remotely | programming issue
easy bolus | pump | fail to set normal bolus | inaccurate delivery; use of incorrect control settings
easy bolus | pump | bolus amount out of range | inaccurate delivery
ECG monitor, EKG monitor | equipment | delayed monitoring | improper device output
ECG subassembly, EKG subassembly | component
effective length |  | effective length out of range | defective component
electrical interface | electronic | loss of electrical connection | intermittent continuity
electrical isolation | electrical | loss of electrical isolation | electrical shorting
electrical overstress | electrical
electrical power |  | power out of range | power conditioning issue
electrical power | power | loss of power | loss of power
electrode contact | electrical
electrolyte reservoir | capacitor | compromised electrolyte reservoir | infusion or flow issue
electromagnetic capability |  | MRI image distorted | image display error
electromagnetic capability |  | device malfunction | device operates differently than expected
electromagnetic capability | magnetic | data corrupted | data issue
Electro-magnetic emission, Electromagnetic emission, electromagnetic energy, electromagnetic radiation, Electric Field |  | excessive electromagnetic energy emitted | Electromagnetic Interference (EMI)
Electro-magnetic emission, Electromagnetic emission, electromagnetic energy, electromagnetic radiation, Electric Field | test | magnetic field out of range | Electromagnetic Interference (EMI)
electronic property | electronic
electro-surgery, electrosurgery | environment | device damaged | device damaged prior to use
emergency button, emergency hold button, emergency hold buttons, emergency stop device, emergency stop devices |  | multiple action required | protective measure issue
emergency button, emergency hold button, emergency hold buttons, emergency stop device, emergency stop devices |  | label missed | device markings issue
emergency button, emergency hold button, emergency hold buttons, emergency stop device, emergency stop devices |  | Inadvertently activating button accidentally | deployment issue
emergency button, emergency hold button, emergency hold buttons, emergency stop device, emergency stop devices | safety | fail to access | human-device interface issue
emergency power | power
emergency stop |  | use for normal operation undesired | device handling issue
emergency stop |  | introduce further hazard | device handling issue
emergency stop |  | inappropriate marking symbol | device markings issue
emergency stop |  | inappropriate color scheme | defective item
emergency stop |  | fail to suit | Incompatibility Problem
emergency stop |  | fail to require one single action | inadequate user interface
emergency stop |  | fail to reduce risk | Fail-Safe Design Failure
emergency stop |  | fail to break full circuit | device inoperable
emergency stop |  | fail to access | human-device interface issue
emergency stop | safety | difficult to maintain | device difficult to maintain
emergency table stop | mechanical
EMI source | image | device malfunction due to TENS field | device operates differently than expected
EMI source | image | device malfunction due to radiation field | device operates differently than expected
EMI source | image | device malfunction due to motor field | device operates differently than expected
EMI source | image | device malfunction due to external defibrillation field | device operates differently than expected
EMI source | image | device malfunction due to electrocautery field | device operates differently than expected
EMI source | image | device malfunction due to electrical tool field | device operates differently than expected
EMI source | image | device malfunction due to diathermy field | device operates differently than expected
empty reservoir |  | fail to indicate empty reservoir | no display or display failure
empty reservoir | pump | fail to activate alarm | device alarm system issue
encoder count, encoder drive count |  | incorrect encoder count | Incorrect Software Programming Calculations
encoder count, encoder drive count | motor | inaccuracy encoder count | imprecision
encrypt data | data | unsecured data | application security issue
endoscope sterilization | sterility | sterility compromising | Delivered As Unsterile Product; Failure to Sterilize
endoscopic accessory | component | endoscopic accessory melt | Melted
endoscopic accessory | component | endoscopic accessory catch fire | Fire
endovascular graft
endovascular implant
endovascular prosthesis |  | sterility compromising | Delivered As Unsterile Product; Failure to Sterilize
endovascular prosthesis |  | reduced life cycle | shelf life exceeded
endovascular prosthesis |  | low migration resistance | migration of device or device component
endovascular prosthesis |  | low kinking resistance | Kinked
endovascular prosthesis |  | loss of long-term structure integrity | material integrity issue
endovascular prosthesis |  | loss of function | device inoperable
endovascular prosthesis |  | leakage at seal zone | failure to adhere or bond
endovascular prosthesis |  | invisible under fluoroscopy | poor quality image
endovascular prosthesis |  | incompatible dimensions | size incorrect for patient
endovascular prosthesis |  | inappropriate interact between components | disassembly
endovascular prosthesis |  | inable to safely use magnetic resonance imaging (MRI) | electromagnetic compatibility issue
endovascular prosthesis |  | inable to maintain adequate integrity | material integrity issue
endovascular prosthesis |  | inable to maintain adequate blood flow through lumen | insufficient flow or underinfusion
endovascular prosthesis |  | inable to isolate the lesion | fluid leak
endovascular prosthesis |  | inable to effective fixation within the vasculature | Detachment of Device or Device Component
endovascular prosthesis |  | corrosion | Corrosion
endovascular system |  | sterility compromising | Delivered As Unsterile Product; Failure to Sterilize
endovascular system |  | out of dimension | size incorrect for patient
endovascular system |  | invisible under fluoroscopy | poor quality image
endovascular system |  | Inappropriate sterility assurance level (SAL) | Delivered As Unsterile Product
endovascular system |  | inappropriate dislodgement force | device handling issue
endovascular system |  | inappropriate deployment force | device handling issue
endovascular system |  | inable to safely withdraw the delivery system | use of device issue
endovascular system |  | inable to minimize blood loss | use of device issue
endovascular system |  | inable to consistently and accurately and safely deploy the endovascular prosthesis | use of device issue
endovascular system |  | inable to consistently and accurately and safely access the intended location | use of device issue
endovascular system |  | fail to label sterile | device markings issue
energy activation | energy | loss of energy | failure to deliver energy
energy activation |  | fail to activate energy | failure to deliver energy
energy delivery, delivery energy, delivery energies |  | over energy delivery | output energy incorrect
energy delivery, delivery energy, delivery energies | energy | no energy delivery | failure to deliver energy
energy delivery, delivery energy, delivery energies |  | incorrect energy delivery status | output energy incorrect
energy generator |  | power on or off switch failed | loss of power
energy generator |  | loss of state (activate or deactivate or alarm or warning) | deployment issue
energy generator |  | loss of output | no device output
energy generator |  | fail to power up | failure to power-up
energy generator |  | fail to deliver energy | failure to deliver energy
energy generator | power | energy activation or deactivation failed | failure to deliver energy
energy output |  | intermittent energy output | output energy incorrect
energy output | power | energy output to patient tissue incorrect | Energy Output to Patient Tissue Incorrect
energy output port, output port, output ports | part | port out of dimension | poor quality image
energy spectrum | power | energy spectrum incorrect | Energy Spectrum Incorrect
energy tool |  | tool activation or deactivation failed | deployment issue
energy tool |  | incorrect tool settings | use of incorrect control settings
energy tool | equipment | incorrect tool identification | use of device issue
engineering file | usability | inadequate file | device markings issue
engineering process, usability engineering |  | misalignment with the other process | improper or incorrect procedure or method
engineering process, usability engineering |  | lack of user interaction process | human-device interface issue
engineering process, usability engineering |  | lack of maintenance process | device maintenance issue
engineering process, usability engineering |  | lack of implementation process | improper or incorrect procedure or method
engineering process, usability engineering |  | lack of establishment process | improper or incorrect procedure or method
engineering process, usability engineering |  | lack of document control procedure | device markings issue
engineering process, usability engineering |  | inappropriate technique | use of device issue
engineering process, usability engineering |  | inappropriate method | improper or incorrect procedure or method
engineering process, usability engineering |  | fail to tailor process | improper or incorrect procedure or method
engineering process, usability engineering | usability | fail to follow standard | compliance issue
entrap, entrapment | movement | device entrapment | Entrapment of Device or Device Component
environment condition, environmental condition | environment | inappropriate environment condition | environmental control or utility issue
enzymatic cleaner | chemical | removal problem | device handling issue
error code, error message, error status, fault message, system message, 911 message, explanation message |  | incorrect error code | Incorrect Error Code
error code, error message, error status, fault message, system message, 911 message, explanation message |  | fail to receive error message | device alarm system issue
error code, error message, error status, fault message, system message, 911 message, explanation message |  | fail to disclose error message | instruction for use issue
error code, error message, error status, fault message, system message, 911 message, explanation message |  | fail to disclose trouble shooting | instruction for use issue
error code, error message, error status, fault message, system message, 911 message, explanation message | safety | fail to auto-reset error message | failure to reset
error code, error message, error status, fault message, system message, 911 message, explanation message |  | error message reported | device alarm system issue
ESD, electrical-static discharge, electrostatic discharge |  | fail to pass ESD immunity test | failure of device to self-test
ESD, electrical-static discharge, electrostatic discharge |  | electrical short | electrical shorting; spark
ESD, electrical-static discharge, electrostatic discharge |  | dielectric breakdown | electrical shorting
ESD, electrical-static discharge, electrostatic discharge |  | device malfunction | device operates differently than expected
ESD, electrical-static discharge, electrostatic discharge |  | device damaged | device damaged prior to use
ESD, electrical-static discharge, electrostatic discharge | safety | component damaged | break
essential performance |  | lack of verification | protective measure issue
essential performance |  | degradation of essential performance | compliance issue
essential performance | clinical | absence of essential performance | compliance issue
ethernet connection | communication | loss of connection | connection issue
ethernet interface | communication | loss of communication | communication or transmission issue
ethernet port | component | port mismatched | component incompatible
EtO, Et0, ethylene oxide | safety | sterilization inefficacy | Failure to Sterilize
EtO, Et0, ethylene oxide |  | residual over limit | Particulates
EtO, Et0, ethylene oxide |  | inappropriate sterilization cycle | Failure to Sterilize
EtO, Et0, ethylene oxide |  | fail to meet Sterility Assurance Level (SAL) | Delivered As Unsterile Product; Failure to Sterilize
excessive current | electrical | catch fire | Fire
exhalation filter | ventilator | difficult to replace | difficult to remove; device maintenance issue
exhalation threshold | ventilator | set exhalation threshold incorrectly | capturing issue
expiratory time, E-time, exhalation time, respiratory time | ventilator | exhalation time short | imprecision
exhalation valve, inhalation valve, inspiratory valve |  | valve stuck | sticking; difficult to open or close
exhalation valve, inhalation valve, inspiratory valve |  | fail to control valve | difficult to open or close
exhalation valve, inhalation valve, inspiratory valve |  | fail to open valve | difficult to open or close; pressure issue
exhalation valve, inhalation valve, inspiratory valve |  | fail to fail-safe open | fail-safe design failure
exhalation valve, inhalation valve, inspiratory valve |  | fail to modulate valve force | mechanical issue
exhalation valve, inhalation valve, inspiratory valve |  | fail to control delivered pressure | pressure issue
exhalation valve, inhalation valve, inspiratory valve |  | fail to provide exit path | fail-safe design failure
exhalation valve, inhalation valve, inspiratory valve |  | fail to adjust gas flow | improper flow or infusion
exhalation valve, inhalation valve, inspiratory valve |  | fail to maintain target flow rate | improper flow or infusion
exhalation valve, inhalation valve, inspiratory valve |  | lack of backup valve | no fail-safe mechanism
exhalation valve, inhalation valve, inspiratory valve | component | undesired valve activation | Premature Deployment
exhalation valve, inhalation valve, inspiratory valve |  | fail to close valve | difficult to open or close
exhalation valve breath, exhalation valve cycle | data | reduced service life | shelf life exceeded
exhalation valve breath, exhalation valve cycle |  | inable to provide breath | delivery system failure
exhalation valve breath, exhalation valve cycle |  | fail to record data | data back-up problem
exhaust Fan, cooling fan, cool fan |  | not enough ventilation | device operates differently than expected
exhaust Fan, cooling fan, cool fan |  | low speed | imprecision
exhaust Fan, cooling fan, cool fan |  | loud fan noise | audible noise
exhaust Fan, cooling fan, cool fan |  | fan detachment | detachment of device component
exhaust Fan, cooling fan, cool fan |  | fan broken or damaged | break
exhaust Fan, cooling fan, cool fan |  | fail to label text | device markings issue
exhaust Fan, cooling fan, cool fan | component | fail to rotate | device inoperable
expiration date, valid date, expiry date |  | missing expiration date | device markings issue
expiration date, valid date, expiry date |  | expired parts used | device expiration issue
expiration date, valid date, expiry date | usage | expiration date error | Expiration Date Error
expiratory hold | ventilator | incorrect measurement | incorrect measurement
expiratory phase | ventilator | stuck in expiratory phase | sticking
expiratory positive airway pressure (EPAP), expiratory pressure, PEEP, positive end expiratory pressure, PEEP control, PIP | ventilator | fail to update pressure | pressure issue
expiratory positive airway pressure (EPAP), expiratory pressure, PEEP, positive end expiratory pressure, PEEP control, PIP |  | expiratory pressure out of range | pressure issue
expiratory positive airway pressure (EPAP), expiratory pressure, PEEP, positive end expiratory pressure, PEEP control, PIP |  | incorrect display range | incorrect display
expiratory positive airway pressure (EPAP), expiratory pressure, PEEP, positive end expiratory pressure, PEEP control, PIP |  | fail to reduce pressure | no pressure
exposed edge, bent wire, wire edge |  | fail to be radiused | defective component
exposed edge, bent wire, wire edge | component | fail to be deburred | Material Puncture
exposed surface |  | surface adverse impacted by cleaning agent | device cleaning issue
exposed surface |  | size out of specification | physical property issue
exposed surface |  | sharp surface | Material Puncture
exposed surface |  | high temperature | temperature issue
exposed surface |  | electrical leakage | electrical issue
exposed surface | component | appearance compromised | material integrity issue
extension plate | component | insufficient extended length | defective component
extension strap |  | out of dimension | size incorrect for patient
extension strap | brace | fail to cover all thigh/calf diameter | device misassembled during manufacturing or shipping
extractor | equipment
eye tracker, eye track, eye tracking, track system, track systems |  | undetected eye gaze off | failure to sense
eye tracker, eye track, eye tracking, track system, track systems |  | gaze volume out of specification | optical issue
eye tracker, eye track, eye tracking, track system, track systems |  | fail to wear prescription glasses | use of device issue
eye tracker, eye track, eye tracking, track system, track systems |  | detect eye gaze position inaccurately | invalid sensing
eye tracker, eye track, eye tracking, track system, track systems |  | detect eye gaze off slow | device sensing issue
eye tracker, eye track, eye tracking, track system, track systems |  | detect eye gaze off failed | failure to sense
eye tracker, eye track, eye tracking, track system, track systems | system | angular error out of specification | physical property issue
factory anastomosis |  | seam-line broken | failure to adhere or bond
factory default | service | fail to restore to factory default state | use of incorrect control settings
factory default | usability | fail to clear data | data issue
fail-safe design, fail-safe system, fail-safe mechanism, fail-safe | safety | loss of function | no device output
fail-safe design, fail-safe system, fail-safe mechanism, fail-safe |  | lack of fail-safe mechanism | no fail-safe mechanism
fail-safe design, fail-safe system, fail-safe mechanism, fail-safe |  | fail to enter safe state | device unsafe to use in environment
fail-safe design, fail-safe system, fail-safe mechanism, fail-safe |  | fail to activate alarm | device alarm system issue
fatigue cycle | mechanical | insufficient fatigue cycle | material integrity issue
fault condition |  | loss of detection | failure to sense
fault condition | safety | fail to stop delivery | Failure to Auto Stop
fault condition |  | fail to display nature of fault condition | no display or display failure
fault status | safety | incorrect fault status | protective measure issue
FCC, Federal Communication Commission | regulatory | FCC regulation violated | compliance issue
FET, field effect transistor | electronic
fiberoptic cladding | optical | material separation | Material Separation
fiberoptic cleaver | optical | failure of fiberoptic cleaver | defective component
film cartridge | optical
film processing | optical | inadequate film processing | improper or incorrect procedure or method
film processor | optical | film processor failure | defective component
filter assembly | mechanical
filter tray |  | filter tray leakage | fluid leak
fitness enthusiast, patient population |  | fail to specify patient population | instruction for use issue
fitness enthusiast, patient population | user | inappropriate user | user used incorrect product for intended use
fixation system |  | system migrated | migration of device or device component
fixed prime |  | incorrect prime calculation | dose calculation error due to software problem
fixed prime |  | fail to suspend delivery | delivery system failure
fixed prime |  | fail to prime | Failure to Prime
fixed prime |  | fail to drive motor forward | failure to deploy
fixed prime |  | fail to cancel prime | filling problem
fixed prime | pump | delivery out of range | defective item
flap valve | component | valve stuck | sticking; difficult to open or close
flex joint | mechanical | flex joint damaged | connection issue
flex trace | mechanical | broken flex trace | break
floor level | environment | inappropriate floor level setting | use of incorrect control settings
flow rate | measure | restricted flow rate | Restricted Flowrate
flow rate | measure | out of range | defective item
flow rate | measure | inaccurate flow rate | inaccurate flowrate
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | sensor zero offset out of range | Failure to Zero
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | sensor measurement interfered | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | sensor failure | defective component; failure to read input signal
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | measurement low | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | measurement high | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | low measurement accuracy | incorrect measurement; imprecision
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to read status data | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to read pressure data | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to regulate pressure condition | pressure issue
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to interface with other device | device-device incompatibility
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to measure temperature | incorrect measurement
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to compensate flow | insufficient flow or underinfusion
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to restrict flow | free or unrestricted flow
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to determine flow | parameter calculation error due to software problem; failure to sense
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to determine direction | failure to sense
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to calibrate oxygen | Failure to Calibrate
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to calibrate air | Failure to Calibrate
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | improper sensor position | malposition of device
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS | ventilator | loss of sensor data | loss of data
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | lack of sensor heating | overheating of device or device component
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | fail to detect sensor error | failure to sense
flow sensor, exhalation flow sensor, flow restrictor, exhalation X-Flow Sensor, proximal flow sensor, hot wire flow sensor, HWFS |  | condensed water vapor | fluid leak
flow valve, proportional valve |  | low proof pressure | pressure issue
flow valve, proportional valve |  | inappropriate oxygen fraction | medical gas supply problem
flow valve, proportional valve | component | valve stuck | sticking; difficult to open or close
Flow-Trak |  | pressure control failed | pressure issue
Flow-Trak |  | low volume | Output below Specifications
Flow-Trak | ventilator | lack of Flow-Trak | infusion or flow issue
flow-trak |  | high volume | output above specifications
Flow-Trak |  | fail to engage | disconnection
Flow-Trak |  | align to incorrect breath type | improper device output
fluid handling system |  | under delivery | inaccurate delivery
fluid handling system | pump | over delivery | inaccurate delivery
fluid handling system |  | clogged fluid delivery line | device clogged
fluid path |  | Open fluid path | fluid leak
fluid path |  | Fluid path occluded | occlusion within device
flux rate | control | high flux rate | use of incorrect control settings
food bolus | diabete | incorrect bolus calculation | dose calculation error due to software problem
foot pedal |  | low load capacity | Device-Device Incompatibility
foot pedal |  | loose foot pedals | Detachment of Device or Device Component
foot pedal | component | insensitive sensing | decreased sensitivity; device sensing issue
foot pedal |  | fail to activate | failure to deploy
foot switch, footswitch, footswitches | component | sticky foot switch | difficult to open or close
foot switch, footswitch, footswitches |  | failure of foot switch | defective component
formative evaluation |  | lack of user interface | inadequate user interface
formative evaluation |  | lack of training | Inadequate Training
formative evaluation |  | lack of timeline | reset issue
formative evaluation |  | lack of formative evaluation | use of device issue
formative evaluation |  | lack of evaluation method | improper or incorrect procedure or method
formative evaluation |  | inappropriate technique | use of device issue
formative evaluation |  | inappropriate method | improper or incorrect procedure or method
formative evaluation |  | inadequate formative evaluation | use of device issue
formative evaluation | usability | formative evaluation delayed | improper or incorrect procedure or method; data issue
fraction of inspired oxygen, FiO2 |  | setup incorrectly | use of incorrect control settings
fraction of inspired oxygen, FiO2 |  | incorrect configuration | use of incorrect control settings
frame rate | motion | rate out of range | defective item
function indicator | safety
function indicator light | safety | failure of function indicator light | no display or display failure
function lid | mechanical
functional requirement | standard | lack of functional requirement | use of device issue
gap analysis | standard | lack of gap analysis | use of device issue
gas delivery, gas delivery system |  | fail to deliver compressed ambient air | delivery system failure
gas delivery, gas delivery system |  | fail to deliver compressed oxygen | delivery system failure
gas delivery, gas delivery system |  | fail to deliver required flow rate | delivery system failure
gas delivery, gas delivery system |  | fail to deliver target pressure | delivery system failure; pressure issue
gas delivery, gas delivery system |  | fail to mix gas | delivery system failure
gas delivery, gas delivery system |  | fail to exhalate gas | delivery system failure
gas delivery, gas delivery system | ventilator | failure of gas delivery system | delivery system failure
gas exhaust outlet, gas outlet, inspiratory gas outlet, outlet port, tube port |  | sparked | spark
gas exhaust outlet, gas outlet, inspiratory gas outlet, outlet port, tube port |  | connection detached | disconnection
gas exhaust outlet, gas outlet, inspiratory gas outlet, outlet port, tube port |  | size out of specification | physical property issue
gas exhaust outlet, gas outlet, inspiratory gas outlet, outlet port, tube port |  | fail to label text | device markings issue
gas exhaust outlet, gas outlet, inspiratory gas outlet, outlet port, tube port | ventilator | close to electrical connector | deployment issue
gas flow |  | inappropriate gas flow | infusion or flow issue
gas flow |  | reversed gas flow | improper flow or infusion
gas flow | ventilator | contamination | contamination during use
gas supply, gas supply system, gas cylinder | ventilator | failure of medical gas supply | medical gas supply problem
gas supply, gas supply system, gas cylinder |  | fail to switch backup gas | medical gas supply problem
gas supply, gas supply system, gas cylinder |  | pressure out of range | pressure issue
gas supply, gas supply system, gas cylinder |  | fail to compatible with oxygen supply | medical gas supply problem
gearhead assembly |  | track bearing retainer ring failed | device misassembled during manufacturing or shipping
gearhead assembly |  | track bearing post dislocated | dislodged or dislocated
gearhead assembly |  | track bearing failed | defective component
gearhead assembly |  | stator extension failed | mechanical issue
gearhead assembly |  | stator bonding failed | failure to adhere or bond
gearhead assembly |  | spacer and wave washer failed | defective component
gearhead assembly |  | spacer and wave washer dislocated | dislodged or dislocated
gearhead assembly |  | shaft shear failed | mechanical issue
gearhead assembly |  | screw sheared | mechanical issue
gearhead assembly |  | screw shear failed | mechanical issue
gearhead assembly |  | screw dislocated | dislodged or dislocated
gearhead assembly |  | ring gear failed | defective component
gearhead assembly |  | planet washer failed | defective component
gearhead assembly |  | planet carrier failed | defective component
gearhead assembly |  | planet bearings failed | defective component
gearhead assembly |  | output plate failed | break
gearhead assembly |  | Input Pinion failed | defective component
gearhead assembly |  | heavy duty retaining ring failed | device misassembled during manufacturing or shipping
gearhead assembly |  | heavy duty retaining ring dislocated | dislodged or dislocated
gearhead assembly |  | grease | misassembled
gearhead assembly |  | gear failed | defective component
gearhead assembly | motor | fail to drive piston | failure to deploy
gearhead assembly |  | dowell pin failed | defective component
gearhead assembly |  | bearing failed | defective component
graft material |  | material split or torn | Split; torn material
graft material |  | inappropriate permeability | fluid leak
graft material |  | hole in materials | hole in material
GUI help button, GUI element, on-screen button, onscreen help |  | lack of button | inadequate user interface
GUI help button, GUI element, on-screen button, onscreen help |  | inadequate message display | incorrect display
GUI help button, GUI element, on-screen button, onscreen help | usability | fail to pop-up | human-device interface issue; application program issue
GUI touch, touch key, touch-key | usability | invalid GUI touch | human-device interface issue
GUI touch, touch key, touch-key |  | short confirmation time | use of device issue
GUI touch, touch key, touch-key |  | fail to detect GUI touch | failure to sense
hand immobilizer | equipment
hand pass |  | pull strength low | mechanical issue
hand pass | brace | Hand pass breakage | break
hand rheumatoid arthritis, hand rheumatism, hand hemiplegia |  | hand retraction | Retraction Problem
hand rheumatoid arthritis, hand rheumatism, hand hemiplegia | brace | brace mismatched | component incompatible
hazardous situation |  | lack of hazardous situation | device unsafe to use in environment
hazardous situation |  | inadequate hazardous situation information | no information
hazardous situation |  | fail to identify non-user interface related hazardous situation | inadequate user interface
hazardous situation |  | fail to identify hazardous situation | compliance issue
hazardous situation |  | fail to analyze new hazardous situation | device unsafe to use in environment
hazardous situation | risk | data may not be stored | data issue
head immobilizer | equipment
header post |  | insufficient header post strength | mechanical issue
header post | lead | break off | break
healing process |  | missing removable rigid part | component missing
healing process | brace | inappropriate healing process | improper or incorrect procedure or method
health data, healthcare data, patient data, Electronic Medical Record |  | undesired patient data empty | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | undesired patient data changes | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | unauthorized access | unauthorized access to computer system
health data, healthcare data, patient data, Electronic Medical Record |  | patient data exposured | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | patient data breached | application security issue; patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | patient data updated incorrectly | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | patient data merged | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | patient order merged | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | loss of patient data | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | lack of data validation | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | incorrect data format | patient data issue; data issue
health data, healthcare data, patient data, Electronic Medical Record |  | incorrect data chart | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | incorrect data | data issue
health data, healthcare data, patient data, Electronic Medical Record |  | inadequate patient data | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | incorrect patient laterality | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | incorrect patient orientation | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | healthcare data breached | application security issue; patient data issue; unauthorized access to computer system
health data, healthcare data, patient data, Electronic Medical Record |  | fail to display data | no display or display failure
health data, healthcare data, patient data, Electronic Medical Record |  | fail to reset patient data | patient data issue
health data, healthcare data, patient data, Electronic Medical Record |  | data measured incorrectly | patient data issue
health data, healthcare data, patient data, Electronic Medical Record | data | data calculated incorrectly | parameter calculation error due to software problem; patient data issue
heart rate | CRM | inappropriate heart rate | failure to pace or properly pace; patient data issue
heart rate | CRM | fail to alarm failure of fast heart rate | device alarm system issue
heart valve cage | component
heart valve leaflet | component
heart valve sewing ring | component
heat and moisture exchanger | ventilator | poor exchange | device operates differently than expected
heat exchanger | equipment
heat sink | equipment
height adjustment, height adjustability |  | unable single hand to adjust | device difficult to setup or prepare
height adjustment, height adjustability |  | restricted height adjustability | use of incorrect control settings
height adjustment, height adjustability |  | lack of position indicator | no display or display failure
height adjustment, height adjustability |  | inaccurate height adjustment | imprecision
height adjustment, height adjustability | motion | adjustment height out of specification | device difficult to setup or prepare
helifix coil | component
helix |  | helix retraction failed | Retraction Problem
helix | lead | helix extension failed | mechanical issue
hemo valve | component | valve stuck | sticking; difficult to open or close
hemofiltration | control
hermetic seal | component | loss of hermetic seal | failure to adhere or bond
high efficiency particulate air filter, HEPA filter |  | particle collection inefficiency | Environmental Particulates
high efficiency particulate air filter, HEPA filter |  | integrity compromised | Material Disintegration
high frequency oscillation, HFO |  | fail to deliver constant flow | improper device output
high frequency oscillation, HFO |  | fail to enable high frequency | improper device output
high frequency oscillation, HFO |  | fail to deliver high frequency pressure | pressure issue
high frequency oscillation, HFO |  | fail to deliver oscillated pressure | pressure issue
high frequency oscillation, HFO |  | fail to deliver | therapeutic or diagnostic output failure
high frequency ventilation | ventilator | ventilation interrupted | failure to cycle
hollow fiber | material
home switch | component | back-off home switch failed | defective component
home use |  | improper or incorrect use of device | use of device issue
home use | clinical | clinical setting environment violated | use of incorrect control settings
homeostasis valve | lead | inappropriate fit | fitting problem
hose line |  | hose line rupture | Material Rupture
hose line | component | hose line occlusion | occlusion within device
hospital database | data | unknown drug | medication error
hospital environment | environment | missing direct medical supervision | Instruction for Use Issue
human factor | user
human-device interface | user
HV therapy, high voltage therapy | CRM | subthreshold pacing | failure to pace or properly pace
HV therapy, high voltage therapy | CRM | loss of HV therapy | therapeutic or diagnostic output failure
HV therapy, high voltage therapy | CRM | lack of HV therapy marker | device markings issue
HV therapy, high voltage therapy | CRM | HV therapy suspended | therapeutic or diagnostic output failure
HV therapy, high voltage therapy | CRM | HV therapy delayed | therapeutic or diagnostic output failure
HV therapy, high voltage therapy | CRM | fail to deliver HV therapy | therapeutic or diagnostic output failure
hybrid circuit | electrical
hydratable intravascular catheter
hydraulic bed system | system | fluid leak | fluid leak
hydraulic system | system | failure of hydraulic system | device inoperable
hydrogen peroxide sterilization | system | explosion | burst
image brightness | image | image brightness out of specification | poor quality image
image color |  | image color out of specification | poor quality image
image color | image | image color accuracy low | poor quality image
impedance converter | equipment
implant catheter |  | reuse implant unit | use of device issue
implant catheter |  | over expiration date | Expiration Date Error
implant catheter | catheter | improper disposal of explanted material | device handling issue
implant mobility NOS | system
implant pump |  | unique pump communicator violated | communication or transmission issue
implant pump |  | reuse implant unit | use of device issue
implant pump |  | over expiration date | Expiration Date Error
implant pump |  | over delivery | inaccurate delivery
implant pump |  | low alarm volume | low audible alarm
implant pump |  | lack of storage mode | inadequate storage
implant pump |  | lack of audible alarm | not audible alarm
implant pump |  | inappropriate insulin pathway | medication error
implant pump |  | improper disposal of explanted material | device handling issue
implant pump |  | failed to be sterilized internally | Failure to Sterilize; Delivered As Unsterile Product
implant pump |  | failed to be sterilized externally | Failure to Sterilize; Delivered As Unsterile Product
implant pump |  | fail to stop pump | pumping issue
implant pump |  | fail to rinse with sodium hydroxide | device cleaning issue
implant pump |  | fail to draw out the rinsing | device cleaning issue
implant pump | pump | fail to communicate with telemetry | radio signal problem
implantable injection chamber, injection chamber |  | chamber leakage | fluid leak
implantable state
incremental dose
information for safety, safety information |  | unperceivable safety information | no information
information for safety, safety information |  | misunderstandable safety information | no information
information for safety, safety information |  | incomplete safety information | no information
information for safety, safety information |  | inadequate safety information | no information
information for safety, safety information |  | fail to support correct use | use of device issue
information for safety, safety information | safety | conscious disregard safety information | no information
infusion mode, mode of Infusion, modes of infusions | infusion | incorrect bolus calculation | dose calculation error due to software problem
infusion mode, mode of Infusion, modes of infusions | infusion | incorrect basal calculation | dose calculation error due to software problem
infusion mode, mode of Infusion, modes of infusions | infusion | inappropriate infusion mode selected | use of device issue
infusion mode, mode of Infusion, modes of infusions | infusion | inadequate infusion mode | inadequate user interface
infusion pump | infusion | unable to self-adminster | failure of device to self-test
infusion pump | infusion | single drug infusion violated | medication error
infusion pump | infusion | liquid ingress | fluid leak
infusion pump | infusion | incorrect prescription information | medication error
infusion pump | infusion | inaccurate delivery | inaccurate delivery
infusion pump | infusion | fail to follow delivery schedule | improper or incorrect procedure or method
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | water mark | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | set malfunction | Failure to Infuse
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | set leakage | fluid leak; pressure issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | set detached | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | medication spilled | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | jacket wet | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | infusion set empty | Failure to Infuse
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | filter leakage | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | fail to withstand pressure | pressure issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset | pump | fail to fill infusion set | infusion or flow issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | fail to compatible with Taxol | Incompatibility Problem; chemical issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | fail to compatible with Nitroglycerin | Incompatibility Problem; chemical issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | dry or wet stain | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | cytotoxic drug leakage | fluid leak
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | chemical resistance failed | chemical issue
infusion set, IV set, plum set, administration set, tubing set, extension set, plumset |  | chamber leakage | fluid leak
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time |  | inhalation time over | imprecision
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time | ventilator | inaccurate inhalation time | imprecision
injection system |  | under infusion | insufficient flow or underinfusion
injection system |  | over infusion | excess flow or over-infusion
injection system |  | no infusion | insufficient flow or underinfusion
injection system |  | incompatible infusion | improper flow or infusion
injection system | system | failure to infuse | Failure to Infuse
AC inlet filter, inlet filter |  | missing filter | component missing
AC inlet filter, inlet filter |  | lost of fire protection | Failure to Fire
AC inlet filter, inlet filter | ventilator | filter plugged | Filtration Issue
inlet port valve, outlet port valve | component | valve stuck | sticking; difficult to open or close
inner coil, proximal coil, distal coil |  | loss of integrity | Material Disintegration
inner coil, proximal coil, distal coil | lead | insufficient fatigue cycle | material integrity issue
inner coil, proximal coil, distal coil |  | inner colil fracture due to over torque | device misassembled during manufacturing or shipping
inner coil, proximal coil, distal coil |  | inner coil separated from shaft | failure to separate
inner coil, proximal coil, distal coil |  | inner coil abraded | device abrasion from instrument or another object
inner coil, proximal coil, distal coil |  | fatigue fracture | fracture; mechanical issue
inner coil, proximal coil, distal coil |  | excessive body fluids in inner coil | contamination during use
inner coil, proximal coil, distal coil |  | excessive blood in inner coil | bleed back
inner coil, proximal coil, distal coil |  | residual torque in inner coil | mechanical issue
inner coil, proximal coil, distal coil |  | over force | device handling issue
inner coil, proximal coil, distal coil |  | inadequate coil weld strength | mechanical issue; Solder Joint Failure
inner coil, proximal coil, distal coil |  | fail to retain position | failure to adhere or bond
inoperative state |  | fail to enter inoperative state | device inoperable
insert force, insertion force | mechanical | insert force out of specification | Difficult to Insert
inspiratory hold |  | invalid to hold inspiratory | misassembled
inspiratory hold | ventilator | fail to hold inspiratory | misassembled
inspiratory phase | ventilator | stuck in inspiratory phase | sticking
inspiratory pressure |  | fail to alarm low inspiratory pressure | device alarm system issue
inspiratory pressure |  | fail to alarm high inspiratory pressure | device alarm system issue
inspiratory pressure |  | fail to alarm failure of low inspiratory pressure | device alarm system issue
inspiratory pressure |  | incorrect trigger | use of incorrect control settings
inspiratory pressure |  | pressure out of range | pressure issue
inspiratory pressure |  | fail to alarm failure of high inspiratory pressure | device alarm system issue
Inspiratory resistance, expiratory resistance, pulmonary reactance |  | inadeqaute resistance measurement | physical resistance
Inspiratory resistance, expiratory resistance, pulmonary reactance |  | inadeqaute reactance measurement | incorrect measurement
Inspiratory resistance, expiratory resistance, pulmonary reactance |  | Inspiratory resistance out of spec | physical resistance
Inspiratory resistance, expiratory resistance, pulmonary reactance |  | expiratory resistance out of spec | physical resistance
Inspiratory resistance, expiratory resistance, pulmonary reactance |  | resistance out of spec | physical resistance
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time | ventilator | inappropriate inspiratory time | use of incorrect control settings
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time |  | low inspiratory duration accuracy | use of incorrect control settings
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time |  | inappropriate cycle breath time | use of incorrect control settings
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time |  | inspiratory time out of range | use of incorrect control settings
inspiratory time, inspiration time, inspiratory duration, I-time, inhalation time |  | inappropriate Auto-Trak time | use of incorrect control settings
insulated conductor |  | over bending | bent
insulated conductor |  | lack of heat-resistant material | material integrity issue
insulated conductor | component | compromised insulation | electrical shorting
insulin concentration | diabete
insulin pump | diebete | lack of bolus function | Inadequate Service
insulin pump |  | lack of basal function | Inadequate Service
insulin sensitivity |  | insulin sensitivity out of range | decreased sensitivity; Increased Sensitivity
insulin sensitivity | diabete | insulin sensitivity changed | device sensing issue
insulin sensitivity |  | incorrect insulin calculation | dose calculation error due to software problem
insulin sensitivity |  | fail to store insulin sensitivity factors | use of incorrect control settings
insulin sensitivity |  | fail to set insulin sensitivity factors | use of incorrect control settings
insulin type |  | incorrect insulin type selected | medication error; infusion or flow issue
insulin type | diabete | incorrect insulin concentration | medication error
insurance coverage | business | void insurance coverage | Market Issue
integral water permeability |  | inappropriate permeability | fluid leak
integrity check |  | thermal damage | overheating of device or device component
integrity check |  | mechanical damage | mechanical issue
integrity check | quality | integrity check failed | failure of device to self-test
intended use, Indications for Use, intended for hospital, heathcare facility use, intended for use |  | inappropriate user | user used incorrect product for intended use
intended use, Indications for Use, intended for hospital, heathcare facility use, intended for use |  | inappropriate usage | use of device issue; user used incorrect product for intended use
intended use, Indications for Use, intended for hospital, heathcare facility use, intended for use |  | Indications for Use misalignment | failure to align
intended use, Indications for Use, intended for hospital, heathcare facility use, intended for use | usability | failure for traditional use | use of device issue
intended use, Indications for Use, intended for hospital, heathcare facility use, intended for use |  | failure for immediate or early load | Premature Deployment
interconnection board intubate | component
interface device, uid, uids |  | unauthorized access | unauthorized access to computer system
interface device, uid, uids |  | over utilization | use of device issue
interface device, uid, uids |  | over noise exposure | audible noise
interface device, uid, uids |  | over magnetic exposure | use of device issue; user used incorrect product for intended use; electromagnetic compatibility issue
interface device, uid, uids |  | missing control function | inadequate user interface
interface device, uid, uids |  | loss of feedback | human-device interface issue
interface device, uid, uids |  | lack of uid | inadequate user interface
interface device, uid, uids |  | lack of internal temperature monitor | temperature issue
interface device, uid, uids |  | interface loss of control | human-device interface issue
interface device, uid, uids |  | information transmission failed | communication or transmission issue
interface device, uid, uids |  | indistinct button | human-device interface issue
interface device, uid, uids | usability | grasp force low | device handling issue
interface evaluation |  | lack of user interface evaluation plan | human-device interface issue
interface evaluation |  | lack of summative evaluation | use of device issue
interface evaluation |  | lack of quantitative analysis | inadequate user interface
interface evaluation |  | lack of qualitative analysis | inadequate user interface
interface evaluation |  | lack of formative evaluation | use of device issue
interface evaluation | usability | inappropriate evaluation environment setting (lab or simulated or actual) | improper or incorrect procedure or method
interface output | electrical | loss of output | no device output
interface specification |  | obsolete interface specification | human-device interface issue
interface specification |  | lack of user interface requirement consideration | human-device interface issue
interface specification |  | lack of use specification consideration | use of device issue
interface specification |  | lack of use specification | use of device issue
interface specification |  | lack of use error consideration | use of device issue
interface specification |  | lack of training indication | Inadequate Training
interface specification |  | lack of hazard related use scenario consideration | use of device issue
interface specification |  | lack of documentation indication | no display or display failure
interface specification |  | inadequate interface specification | human-device interface issue
interface specification | usability | ambiguous interface specification | human-device interface issue
interlumen communication | communication
internal battery |  | fail to warn use of internal battery | instruction for use issue
internal battery | power | undesired battery removal | device handling issue
internal fixation | mechanical
inter-nodal distance |  | distance out of range | positioning issue
intra-cavity procedure | procedure
intravascular catheter |  | radio detectability failed | failure to sense
intravascular catheter |  | low tensile strength | mechanical issue
intravascular catheter |  | low burst pressure | pressure issue; Burst
intravascular catheter |  | leakage at hub or connection | fluid leak
intravascular catheter |  | crushing surface | Scratched Material
intravascular catheter |  | corrosion | Corrosion
intravascular catheter |  | biocompatibility compromised | biocompatibility issue
intravascular catheter |  | adverse impacted by side hole | hole in material
intravenous pole, IV pole, Mini-pole |  | fail to hang bag | disassembly
intravenous pole, IV pole, Mini-pole | physiology | fail to assemble | Misassembled by Users
intravenous pole, IV pole, Mini-pole |  | fail to adjust height | device difficult to setup or prepare
introducer catheter |  | radio detectability failed | failure to sense
introducer catheter |  | low peak tensile force | mechanical issue
introducer catheter |  | effective length out of range | defective component
introducer needle |  | size out of specification | size incorrect for patient
introducer needle |  | needle point hook | Dull
introducer needle |  | needle point burr | Material Puncture
invasive ventilation, invasive ventilator, invasive ventilators |  | lack of invasive support | failure to cycle
invasive ventilation, invasive ventilator, invasive ventilators |  | incorrect ventilation mode | failure to cycle
invasive ventilation, invasive ventilator, invasive ventilators |  | inadequate flow | failure to cycle
invasive ventilation, invasive ventilator, invasive ventilators |  | fail to conjunct with high pressure oxygen source | medical gas supply problem
invasive ventilation, invasive ventilator, invasive ventilators |  | compressed air as the secondary source of air | medical gas supply problem
invasive ventilation, invasive ventilator, invasive ventilators | ventilator | blower as the primary source of air | medical gas supply problem
IP level |  | water resistant failed | fluid leak
IP level |  | water proof failed | fluid leak
IP level | environment | dust ingress | leak
IR link | communication | fail to configure pump communicator | device difficult to setup or prepare
irradiation container
irradiator operator
irritating area | clinical | user inflammation or allergy | biocompatibility issue
it-network, it network, it networks |  | lack of security specification | computer system security issue
it-network, it network, it networks |  | lack of routing | application network issue
it-network, it network, it networks |  | lack of purpose | application network issue
it-network, it network, it networks |  | lack of information flow | no information
it-network, it network, it networks |  | lack of configuration | use of incorrect control settings
it-network, it network, it networks | network | lack of characteristics | application network issue
IV bag |  | bag leakage | fluid leak
IV catheter | component | part detachment | detachment of device component
IV catheter connection, IV catheter port connection
IV stand |  | fail to connect to IV stand | connection issue
IV tubing set connection |  | connection leakage | fluid leak
IV tubing set port connection |  | connection leakage | fluid leak
IV, intravenous therapy |  | infusion of particulates | Particulates; infusion or flow issue
IV, intravenous therapy |  | infiltration | Filtration Issue
IV, intravenous therapy |  | exposure to infectious agent | bacterial contamination of device; environmental particulates; biological environmental factor
IV, intravenous therapy |  | catheter misplacement | use of device issue
IV, intravenous therapy |  | catheter clotting | chemical issue
IV, intravenous therapy |  | blood loss | use of device issue
IV, intravenous therapy |  | allergens | biocompatibility issue
IV, intravenous therapy | infusion | air in line | improper flow or infusion
knee joint | anatomy | knee joint motion out of range | unintended system motion
knee patient | anatomy | Inadequate extension | mechanical issue
laparoscopic sterilization | vision
laser nozzle | laser | laser nozzle separation | Material Separation
lateral hinge, medial hinge, medial hinges |  | unsecured hinge | protective measure issue
lateral hinge, medial hinge, medial hinges | brace | support failed | use of device issue
lateral hinge, medial hinge, medial hinges |  | flexion degree out of range | defective component
lateral hinge, medial hinge, medial hinges |  | flexion and extension stop failed | Failure to Auto Stop
lcd backlight | usability | timeout late | imprecision
lcd backlight |  | Insufficient brightness | optical issue
lcd backlight |  | fail to turn on automatically | device operates differently than expected
lcd screen |  | unreadable display | incorrect display
lcd screen |  | screen broken or damaged | break; no display or display failure
lcd screen |  | no display | no display or display failure
lcd screen |  | Insufficient brightness | optical issue
lcd screen |  | incorrect display | incorrect display
lcd screen | usability | image color accuracy low | poor quality image
lead body | CRM | PVP coating delamination | delamination
lead body | CRM | medical adhesive delamination | delamination; failure to adhere or bond
lead body | CRM | loss of tie down strength | mechanical issue
lead body | CRM | lead body slipped | slippage of device or device component
lead body | CRM | lead body kinked during implant | Kinked
lead body | CRM | lead body insulation damaged | electrical shorting
lead body | CRM | lead body damaged | break
lead body | CRM | insulation material degradation | electrical shorting; degraded
lead body | CRM | insufficient torque transfer | mechanical issue | defective Item
lead body | CRM | insufficient insulation biostability | electrical shorting; unstable
lead body | CRM | flex fatigue fracture | device misassembled during manufacturing or shipping
lead conductor | lead
lead distal, lead tip, lead tips | lead | unacceptable tip stiffness | material rigid or stiff
lead distal, lead tip, lead tips |  | spring jam | mechanical jam
lead distal, lead tip, lead tips |  | soft-tip delaminated | delamination
lead distal, lead tip, lead tips |  | over crimp | mechanical issue
lead distal, lead tip, lead tips |  | loss of soft-tip during implant | component missing
lead distal, lead tip, lead tips |  | lead tip fragments remain in ventricle | Material Fragmentation
lead distal, lead tip, lead tips |  | lead tip breakage at implant | break
lead distal, lead tip, lead tips |  | ineffective opacity of distal marker | device markings issue
lead distal, lead tip, lead tips |  | inadequate crimp | mechanical issue
lead distal, lead tip, lead tips |  | header separated during extraction | failure to separate
lead distal, lead tip, lead tips |  | fatigue fracture | fracture; mechanical issue
lead distal, lead tip, lead tips |  | excessive gap | misassembled
lead distal, lead tip, lead tips |  | distal tip flex fatigue fracture | fracture; mechanical issue
lead distal, lead tip, lead tips |  | distal tip bent | bent
lead distal, lead tip, lead tips |  | diameter of distal tip oversized | defective component
lead electrode | CRM | open circuit | circuit failure
lead electrode | CRM | heating | overheating of device or device component
lead impedance | CRM | incorrect lead impedance measurement | incorrect measurement; Impedance Issue
lead impedance | CRM | inaccurate lead impedance measurement | incorrect measurement; Impedance Issue
lead impedance | CRM | impedance out of range | Impedance Issue
lead implant | lead | excessive lead remove force | difficult to remove
lead implant |  | excessive lead insert force | Difficult to Insert
lead insulation | CRM | lead insulation damaged | electrical shorting
lead insulation |  | lead insulation breached | electrical shorting
lead insulation |  | high current | electrical issue
lead insulation |  | arcing | arcing
lead labeling | lead | lead labeling misinterpreted | device markings issue
lead material | CRM | material biocompatibility compromised | biocompatibility issue
lead noise | CRM | lead noise detected incorrectly | device sensing issue
lead voltage | CRM | gradient induced lead voltage | electronic property issue
leaflet disk | heart | leaflet disc escape | Escape
leak flow | ventilator | fail to update leak flow | improper flow or infusion
leak flow | ventilator | fail to compensate flow | insufficient flow or underinfusion
leakage current, leak current |  | excess pass current through | electrical issue
leakage current, leak current | electrical | leakage current out of specification | electrical issue
legacy software |  | software version incorrect | application program version or upgrade problem
legacy software |  | lack of risk management activity | Fail-Safe Design Failure
legacy software |  | lack of rationale for using legacy software | application program version or upgrade problem
legacy software |  | lack of gap analysis | computer software issue
legacy software |  | integration failed | Material Disintegration
legacy software |  | fail to identify potential cause | compliance issue
legacy software |  | fail to identify hazardous situation | compliance issue
legacy software |  | fail to follow standard | nonstandard device or device component
legacy software |  | fail to define risk control measure | compliance issue
legacy software | software | fail to assess feedback | use of device issue
len aberration | len | distortion | Material Distortion
len disc | len
len implant | len
len toric | len
len vaulting | len
len washer | len | fail to len washer | device cleaning issue
limit switch | safety
line conditioner | equipment
line cord | component
liquid ingress, ingress of liquid, ingress of water, water ingress |  | water proof failed | fluid leak, electrical shorting
liquid ingress, ingress of liquid, ingress of water, water ingress |  | short circuit | electrical shorting
liquid ingress, ingress of liquid, ingress of water, water ingress | safety | inadvertent activation | deployment issue
lithium lodide battery | power
localization configuration |  | inappropriate time format | imprecision
localization configuration |  | inappropriate date format | use of incorrect control settings
localization configuration | software | inadequate language settings | Instruction for Use Issue; inadequate user interface
lock mechanism, closure mechanism, lockout mechanism |  | self-locking mode failed | Fail-Safe Mechanism Issue
lock mechanism, closure mechanism, lockout mechanism |  | part disconnected | failure to disconnect
lock mechanism, closure mechanism, lockout mechanism |  | locking mechanism failure | Fail-Safe Mechanism Issue; Fail-Safe Design Failure
lock mechanism, closure mechanism, lockout mechanism |  | lock mechanism failed | protective measure issue
lock mechanism, closure mechanism, lockout mechanism |  | fail-lock mode failed | improper or incorrect procedure or method
lock mechanism, closure mechanism, lockout mechanism | safety | apply inconveniently | device difficult to setup or prepare
lock sleeve | component
log file |  | log file overwritten | patient data issue
log file |  | log file full | patient data issue
log file | data | fail to update log file | patient data issue
log file |  | inadequate log file | patient data issue
log file |  | fail to create new log file | patient data issue
loss of communication | communication | loss of communication | communication or transmission issue
loss of power | power | loss of power | loss of power
low reservoir limit | pump | lack of unit selection | inadequate user interface
low reservoir limit |  | lack of time selection | reset issue
low reservoir limit |  | fail to generate alert | device alarm system issue; failure to deploy
luer connector, luer slip connector, luer lock connector, Y-connector |  | out of tolerance | defective item
luer connector, luer slip connector, luer lock connector, Y-connector |  | out of dimension | defective component
luer connector, luer slip connector, luer lock connector, Y-connector |  | fail to detach | failure to disconnect
luer connector, luer slip connector, luer lock connector, Y-connector |  | inadequate elasticity in flexure or in tension | material rigid or stiff
luer connector, luer slip connector, luer lock connector, Y-connector |  | thread or lug overriding | disconnection
luer connector, luer slip connector, luer lock connector, Y-connector |  | sub-atmospheric pressure air leakage | fluid leak
luer connector, luer slip connector, luer lock connector, Y-connector |  | stress cracking | Crack
luer connector, luer slip connector, luer lock connector, Y-connector |  | separation from unscrewing | Material Separation
luer connector, luer slip connector, luer lock connector, Y-connector |  | separation from axial load | Material Separation
luer connector, luer slip connector, luer lock connector, Y-connector |  | positive pressure liquid leakage | fluid leak
luer connector, luer slip connector, luer lock connector, Y-connector |  | misconnection | misconnection
luer connector, luer slip connector, luer lock connector, Y-connector |  | excessive leakage by pressure decay | fluid leak
luer valve | component | valve stuck | sticking; difficult to open or close
lung recruitment, frequency lung recruitment |  | fail to raise pressure | pressure issue
lung recruitment, frequency lung recruitment |  | frequency out of range | electronic property issue
lung recruitment, frequency lung recruitment |  | inappropriate technique applied | improper or incorrect procedure or method
LV Capture | CRM | loss of capture | Failure to Capture
magnetic exposure |  | sensor rate maximized | incorrect measurement
magnetic exposure |  | MRI image distorted | image display error
magnetic exposure |  | device overheating | overheating of device or device component
magnetic exposure | magnetic | device movement | unintended movement
main body deployment |  | fail to transfer force | device handling issue
maintenance documentation, maintenance interval, maintenance service interval |  | lack of maintenance interval | lack of maintenance documentation or guidelines
maintenance documentation, maintenance interval, maintenance service interval |  | inappropriate interval settings | maintenance does not comply to manufacturers recommendations
maintenance documentation, maintenance interval, maintenance service interval | standard | lack of maintenance documentation | lack of maintenance documentation or guidelines
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | loss of pressure | no pressure
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | no breath | use of incorrect control settings
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | incorrect breath type setting | use of incorrect control settings
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | incorrect breath type | improper device output
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inappropriate pressure | pressure issue
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inappropriate peak flow | improper flow or infusion
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inappropriate inspiratory time | imprecision
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inappropriate flow volume | improper flow or infusion
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inappropriate flow pattern | improper flow or infusion
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | inadequate breath types | improper device output
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths |  | breath delayed | improper device output
mandatory breath, assisted breath, spontaneous breath, supported breath, assisted breaths, spontaneous breaths, supported breaths | ventilator | air delivery compromised | inaccurate delivery
mandatory inspiration, assisted Inspiration |  | fail to reduce pressure | pressure issue
mandatory inspiration, assisted Inspiration | ventilator | fail to raise pressure | pressure issue
manifold pole mount |  | fail to connect to IV stand | connection issue
manual cleaning
manual infusion |  | inappropriate infusion stop | pumping issue
manual infusion | infusion | inappropriate infusion start | infusion or flow issue
manual prime |  | fail to manual prime | Failure to Prime
manual prime |  | fail to indicate prime completion | no display or display failure
manual prime |  | fail to incorporate step sequence | Failure to Prime
manual prime |  | fail to drive motor forward | failure to deploy
manual prime |  | fail to disable low reservoir detection | device remains activated
manual prime | pump | fail to deliver | failure to deliver
manual prime chirp | usability | loss of audio notification | not audible alarm
manufacturer address, MFG address, MFG addresses |  | obsolete manufacturer address | device markings issue
manufacturer address, MFG address, MFG addresses |  | incorrect manufacturer name | device markings issue
manufacturer address, MFG address, MFG addresses | business | incorrect manufacturer address | device markings issue
measure activity | measure | measure activity failed | incorrect measurement
measurement accuracy | measure | low measurement accuracy | incorrect measurement; imprecision
measurement system, measurement tool |  | incorrect measurement value | incorrect measurement
measurement system, measurement tool | system | measurement system incompatibility | measurement system incompatibility
mechanical hard stop | safety | stop failed | Failure to Auto Stop
mechanical hazard | mechanical | user cutted | device handling issue
mechanical interface, mechanical retention |  | unstable mount | installation-related problem
mechanical interface, mechanical retention | mechanical | weak mechanical interface strength | mechanical issue
mechanical jam | component
mechanical locker | safety | two items plugged simultaneously | device handling issue
mechanical resistance | mechanical | high resistance | physical resistance
mechanical rotation | motion | restrict mechanical rotation | difficult to advance
meter needle | component
microbial barrier |  | fail to minimize ingress of microorganism | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
mineralizate, mineralization | material
cart, mobile cart, mobile stand, mobile configuration | safety | overbalancing | unstable
cart, mobile cart, mobile stand, mobile configuration |  | missing handle | device handling issue
cart, mobile cart, mobile stand, mobile configuration | movement | lead to unsafe condition | device unsafe to use in environment
cart, mobile cart, mobile stand, mobile configuration |  | unsecured stand | migration of device or device component
cart, mobile cart, mobile stand, mobile configuration | ventilator | fail to move | difficult to remove
model number | business | incorrect model number | device markings issue
monitor display alarm | safety | failure of monitor display alarm | device alarm system issue
morphology discriminator | CRM | arrhythmia misdiagnosed | therapeutic or diagnostic output failure
morphology template | CRM | morphology template changed | use of device issue
morphology template | CRM | arrhythmia misdiagnosed | therapeutic or diagnostic output failure
motion detector | measure | motion detector failure | failure to sense
motor controller |  | circuit triggered without software requesting activation | self-activation or keying
motor controller |  | circuit triggered when voltage is too low | self-activation or keying
motor controller | electrical | circuit triggered when voltage is too high | self-activation or keying
motor currect | pump | fail to check motor currect periodically | failure of device to self-test
motor encoder | motor | fail to read motor encoder | unable to obtain readings
motor pulse | motor | fail to drive motor forward | failure to deploy
motor roll pin | component
movement circle, movement cone, movement cones, movement ellipse, movement ellipses | motion | movement out of specification | unintended movement
moving part |  | foreign matter | Particulates
moving part |  | unintended motion | unintended system motion
moving part | motion | device part collision | Unintended Collision
MRI checklist | image | lack of test result | missing test results
MRI checklist | image | inadequate checklist | use of incorrect control settings
MRI condition, MR conditionality, MR Conditional | image | scan with the wrong MRl condition | Electromagnetic Interference (EMI)
MRI condition, MR conditionality, MR Conditional | image | MRl condition misinterpreted | Electromagnetic Interference (EMI)
MRI condition, MR conditionality, MR Conditional | image | MR conditionality confused | inadequate user interface
MRI condition, MR conditionality, MR Conditional | image | magnetic fields penetrated | Electromagnetic Interference (EMI)
MRI condition, MR conditionality, MR Conditional | image | lack of MR conditionality | Electromagnetic Interference (EMI)
MRI image, diagnostic imaging, mammograms |  | poor quality image | therapeutic or diagnostic output failure
MRI image, diagnostic imaging, mammograms | image | MRI image distorted | image display error
MRI log | image | lack of log file | patient data issue
MRI scan | image | wrong MRI scan condition | Electromagnetic Interference (EMI)
MRI scan | image | unintended stimulation | Unexpected Therapeutic Results
MRI scan | image | pacing inhibition | No Pacing
MRI scan | image | MRI scan without MRI settings enabled | use of incorrect control settings
MRI scan | image | MRI scan restriction confused | inadequate user interface
MRI scan | image | inappropriate MRI scan | Electromagnetic Interference (EMI)
MRI scan | image | fail to receive prescribed MRI scan | Electromagnetic Interference (EMI)
MRI setting | image | MRI setting confused | inadequate user interface
MRI setting | image | incorrect MRI setting | use of incorrect control settings
MRI setting | image | fail to enable MRI setting | device difficult to setup or prepare
MRI setting | image | fail to disable MRI setting | device difficult to setup or prepare
MRI test | image | test suspended | incorrect or inadequate test results
muscle fiber |  | inappropriate muscle fiber excitement | Unexpected Therapeutic Results
muscle fiber | physiology | inappropriate muscle fiber contraction | Retraction Problem
navigation button, confirmation button, navigation buttons, confirmation buttons | usability | design lack of intuitively implemented | Human Factors Issue
navigation button, confirmation button, navigation buttons, confirmation buttons |  | confused button | inadequate user interface
nebulizer, nebulization, nebulizer output port |  | nebulization period out of range | Output Issue
nebulizer, nebulization, nebulizer output port |  | fail to label text | device markings issue
nebulizer, nebulization, nebulizer output port | equipment | fail to pause nebulization | improper device output
needle connection
needle hub, needle tube |  | low strength | mechanical issue
negative pressure |  | low peak nagative pressure | pressure issue
net flow | physiology | insufficient flow | insufficient flow or underinfusion
network interface | usability | improper fit between connection | fitting problem; component or accessory incompatibility
network interface |  | fail to support multiple virtual network interfaces | application network issue
network segmentation |  | network segmentation failed | application network issue
network segmentation | safety | network interference | application network issue
network switch | network | loss of power | loss of power
network switch |  | fail to update configuration | device maintenance issue; data issue
network switch |  | fail to support network segmentation | application network issue
network switch |  | fail to support multiple virtual network interfaces | application network issue
network switch |  | fail to support high speed | application network issue
network switch |  | fail to support ethernet small form-factor pluggable (SFP) | use of device issue
network switch |  | fail to support accessory | accessory incompatible
network switch |  | fail to block non-essential communication | application security issue
non-active surgical implant, surgical implant, active surgical implant |  | size out of specification | size incorrect for patient
non-active surgical implant, surgical implant, active surgical implant |  | reduced service life | shelf life exceeded
non-active surgical implant, surgical implant, active surgical implant |  | radio detectability failed | failure to sense
non-active surgical implant, surgical implant, active surgical implant |  | patient-device interaction | human-device interface issue
non-active surgical implant, surgical implant, active surgical implant |  | material degraded | degraded
non-active surgical implant, surgical implant, active surgical implant |  | implant rejection | biocompatibility issue
non-active surgical implant, surgical implant, active surgical implant |  | fluid leak | fluid leak
non-active surgical implant, surgical implant, active surgical implant |  | electro-magnetic interference (EMI) | Electromagnetic Interference (EMI)
non-active surgical implant, surgical implant, active surgical implant |  | difficult to replace | difficult to remove
non-active surgical implant, surgical implant, active surgical implant |  | difficult to remove | difficult to remove
non-active surgical implant, surgical implant, active surgical implant |  | contamination | bacterial contamination of device
non-active surgical implant, surgical implant, active surgical implant |  | ageing | shelf life exceeded
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | lack of non-invasive support | failure to cycle
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | incompatible nasal prongs | accessory incompatible
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | incompatible nasal mask | accessory incompatible
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | incorrect ventilation mode | failure to cycle
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | inadequate flow | insufficient flow or underinfusion
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators |  | inadequate breathing mode | failure to cycle
noninvasive ventilation, non-invasive ventilation, non-invasive ventilator, non-invasive ventilators | ventilator | compressed air as the source of air | medical gas supply problem
nonstandard device | equipment
normal operation, operation period |  | reduced service life | shelf life exceeded
normal operation, continuous operation | user | operation error | device handling issue
normal use | usability | risk from normal use | no fail-safe mechanism
nozzle tip |  | nozzle tip stressed | mechanical issue
nozzle tip |  | nozzle tip split | Split
nozzle tip |  | nozzle tip rough | material rigid or stiff
nozzle tip | component | nozzle tip issue | defective component
nurse call, nurse call system, nurse call connector port |  | fail to send out signal | communication or transmission issue
nurse call, nurse call system, nurse call connector port |  | fail to label alarm symbol | device markings issue
nurse call, nurse call system, nurse call connector port |  | loss of call | use of device issue
nurse call, nurse call system, nurse call connector port | control | fault call | use of device issue
objective evidence |  | inadequate objective evidence | use of device issue
objective evidence | data | contrary objective evidence | use of device issue
occlusion alarm |  | fail to silence alarm | device alarm system issue
occlusion alarm |  | fail to set or reset alarm automatically | device alarm system issue; reset issue
occlusion alarm |  | fail to disable ventilation | device remains activated
occlusion alarm |  | fail to disable gas flow | device remains activated
occlusion alarm | ventilator | fail to activate occlusion alarm | occlusion within device; device alarm system issue
occlusion pressure | pump | occlusion pressure out of range | occlusion within device; pressure issue
occlusion sensitivity | pump | sensitivity setting out of range | decreased sensitivity; Increased Sensitivity
occlusion threshold | pump | inappropriate threshold setting | capturing issue
offline switching power supply | power | fail to shut-down device | failure to shut off
offline switching power supply |  | fail to provide conducted EMI filtering | failure to conduct
offset value | business | incorrect offset value | use of incorrect control settings; data issue
one-way valve | component | valve stuck | sticking; difficult to open or close
one-way valve |  | thread overriding | disconnection
on-screen keyboard | computer | lack of keyboard | inadequate user interface
on-screen keyboard |  | fail to support multiple languages | Instruction for Use Issue; inadequate user interface
on-screen keyboard |  | fail to support alpha-numeric entry | use of device issue
operating humidity | pump | humid out of range | Moisture or Humidity Problem
operating mode, operation mode, operation modes | usability | inadequate operating mode | inadequate user interface
operating pressure | pump | pressure out of range | pressure issue
operating surface | mechanical | size out of specification | physical property issue
operating system, computer operating system |  | version incorrect | operating system version or upgrade problem
operating system, computer operating system |  | upgrade failed | operating system version or upgrade problem
operating system, computer operating system | software | nonfunctional | no device output
operating table, operating room table, operating room tables | clinical | table adjustment failed | device difficult to setup or prepare
operating temperature | pump | temperature out of range | temperature issue
operating voltage | pump | voltage out of range | electrical issue
operative sterile, preoperative sterile, preoperative steriles, pre-operative sterile, pre-operative steriles | sterility | sterility compromised | Delivered As Unsterile Product; Failure to Sterilize
optical cable | component | cable break | break
optical disk | component
optical fiber | material
otc, over-the-counter device | 2 | miss labeling | device markings issue
outer coil | component | outer coil fractured | fracture; mechanical issue
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube | ventilator | tubing blocked | blockage within device or device component
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | tubing pinched | Pitted
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | tubing leakage | fluid leak
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | tubing leakage | leak
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | tubing kinked | Kinked
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | Tubing installed improperly | installation-related problem
outlet tubing, inlet tubing, inlet tubings, sensor tubing, sensor tubings, pressure tubing, pressure tubings, circular tube |  | tubing disconnect | disconnection
out-of-hospital, out of hospital | environment | inappropriate environment | environmental control or utility issue
output interface |  | unstable interface | human-device interface issue
output interface |  | loss of output | no device output
output interface | usability | inappropriate interface design | human-device interface issue
outside diameter
overcurrent device, over-current device, over-current devices | electrical | overcurrent protection failed | protective measure issue
over-infusion Processor, overinfusion Processor, overinfusion Processors |  | over infusion | excess flow or over-infusion
over-infusion Processor, overinfusion Processor, overinfusion Processors | electronic | fail to activate alarm | device alarm system issue
oxygen analyzer | measure | inaccurate measurement | incorrect measurement
oxygen concentration, O2 concentration |  | oxygen out of specification | medical gas supply problem
oxygen concentration, O2 concentration |  | ignition | fire
oxygen concentration, O2 concentration |  | incorrect concentration | medication error
oxygen pressure |  | oxygen pressure low | medical gas supply problem
oxygen sensor, O2 sensor |  | short validation period | protective measure issue
oxygen sensor, O2 sensor |  | sensor failure | defective component; failure to read input signal
oxygen sensor, O2 sensor |  | inappropriate sensor type | component incompatible
oxygen sensor, O2 sensor |  | out of calibration | calibration error
oxygen sensor, O2 sensor |  | fail to compensate for pressure | imprecision
oxygen sensor, O2 sensor |  | fail to compensate for vibration | imprecision
oxygen sensor, O2 sensor |  | fail to compensate for temperature | imprecision
oxygen sensor, O2 sensor |  | fail to monitor fraction of oxygen | incorrect measurement
oxygen sensor, O2 sensor |  | fail to request calibration | calibration issue
oxygen sensor, O2 sensor |  | fail to provide calibration gas and air | Failure to Calibrate
oxygen sensor, O2 sensor | ventilator | fail to calibrate manually | Failure to Calibrate
pacing mode | CRM | inappropriate pacing mode | use of device issue
pacing path | CRM | short circuit | electrical shorting
pacing path | CRM | open circuit | circuit failure
pacing path | CRM | lead fractured | device misassembled during manufacturing or shipping
pacing pulse | CRM | temporary pacing inhibition | No Pacing
pacing pulse | CRM | pulse cancellation | failure to pace or properly pace
pacing pulse | CRM | pacing pulse to wrong place | Energy Output to Patient Tissue Incorrect
pacing pulse | CRM | pacing pulse reduced | Pacing Inadequately
pacing pulse | CRM | pacing pulse energy reduced | Pacing Inadequately
pacing pulse | CRM | inappropriate pacing pulse | failure to pace or properly pace
pacing rate, pace rate | CRM | pacing rate high | failure to pace or properly pace
pacing rate, pace rate | CRM | low pacing rate | Pacing Inadequately
pacing rate, pace rate | CRM | incorrect pacing rate limit selected | failure to pace or properly pace
pacing rate, pace rate | CRM | inappropriate pacing rate | failure to pace or properly pace
pacing sensor | measure | sensor failure | defective component; failure to read input signal
pacing therapy |  | pacing delayed | failure to pace or properly pace
pacing therapy |  | loss of pacing therapy | No Pacing
pacing therapy | CRM | fail to deliver rapid pacing therapy | failure to pace or properly pace
pacing threshold, pacing subthreshold, subthreshold pacing | CRM | loss of AV synchrony | failure to pace or properly pace
pacing threshold, pacing subthreshold, subthreshold pacing | CRM | increased pacing threshold | High Capture Threshold
pacing threshold, pacing subthreshold, subthreshold pacing | CRM | increased pacing subthreshold | High Capture Threshold
pacing threshold, pacing subthreshold, subthreshold pacing | CRM | inappropriate subthreshold pacing | capturing issue
packaging material | user | difficult to open or remove packaging material | Difficult to Open or Remove Packaging Material
patch test | test | abnormal results | improper device output
patient bolus | pump | bolus amount out of range | inaccurate delivery
patient bolus flow rate, flow rate of patient bolus | pump | flow rate out of range | inaccurate flowrate
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | sprayed contents | leak; air leak
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | loss of pressure | no pressure
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | incorrect circuit type | component incompatible
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | inappropriate circuit configuration | component incompatible
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | missing part | component missing
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | component incompatible | fitting problem
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | fail to comply with standards or regulation | nonstandard device or device component
patient circuit, ventilator circuit, breath circuit, breathing circuit | ventilator | disconnected circuit | leak; air leak
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | circuit changed | leak; air leak
patient circuit, ventilator circuit, breath circuit, breathing circuit |  | breathing circuit damaged | leak; air leak
patient device, patient-device | user | patient-device interaction | human-device interface issue
patient device, patient-device |  | patient-device incompatibility | patient-device incompatibility
patient graph, graph of patient, graph of patients, graphs of patients, graphs of patient | usability | loss of image | image display error
patient implant card |  | missing patient implant card | device markings issue
Patient information |  | patient information overwritten | loss of data
Patient information |  | incorrect patient information | patient data issue
Patient information | data | fail to configure Patient information | patient data issue
patient notifier, patient notification |  | weak patient notifier vibration | device alarm system issue; vibration
patient notifier, patient notification |  | patient notifier undetected | failure to sense
patient notifier, patient notification |  | patient notifier malfunction | device operates differently than expected
patient notifier, patient notification |  | patient notifier inoperable | device inoperable
patient notifier, patient notification |  | notification confused | device alarm system issue
patient notifier, patient notification |  | inaudible notification | not audible alarm
patient notifier, patient notification | usability | fail to vibrate | vibration
patient notifier, patient notification |  | fail to deliver notification | device alarm system issue
patient notifier, patient notification |  | annoying notification | device alarm system issue
patient photo, photo of patient, photo of patients, photos of patients, photos of patient | usability | retaking photo failed | use of device issue
patient photo, photo of patient, photo of patients, photos of patients, photos of patient | usability | fail to take photo | use of device issue
patient photo, photo of patient, photo of patients, photos of patients, photos of patient | usability | accepting photo failed | poor quality image
patient position | usability | improper device position | malposition of device
standby mode, patient standby mode, standby control |  | unable auto ventilation | computer software issue
standby mode, patient standby mode, standby control |  | out of calibration | calibration error
standby mode, patient standby mode, standby control |  | incorrect mode | use of device issue
standby mode, patient standby mode, standby control |  | fail to deactivate mode | computer software issue; use of device issue
standby mode, patient standby mode, standby control | software | fail to activate mode | computer software issue; use of device issue
standby mode, patient standby mode, standby control |  | fail to cease delivery | delivery system failure
standby mode, patient standby mode, standby control |  | fail to maintain constant minimum flow | no flow
standby mode, patient standby mode, standby control |  | fail to maintain constant pressure | pressure issue
standby mode, patient standby mode, standby control |  | fail to generate indication | no display or display failure
standby mode, patient standby mode, standby control |  | fail to disable alarm | device alarm system issue
standby mode, patient standby mode, standby control |  | device overheating | overheating of device or device component
patient transport | ventilator | loss of power | loss of power
patient type |  | unclear patient type | patient data issue
patient type |  | inappropriate patient type | use of device issue; Unexpected Therapeutic Results
patient type | usability | fail to select patient type | use of device issue
pause function | software | inappropriate pause period | improper device output
pause function |  | fail to terminate process completely | device operates differently than expected
peak flow | ventilator | high peak flow | improper flow or infusion
PEMS development life-cycle |  | lack of verification method | protective measure issue
PEMS development life-cycle |  | lack of risk management activity | Fail-Safe Design Failure
PEMS development life-cycle |  | lack of problem resolution | compliance issue
PEMS development life-cycle |  | lack of milestone | device markings issue
PEMS development life-cycle |  | lack of document requirement | device markings issue
PEMS development life-cycle |  | lack of document | device markings issue
PEMS development life-cycle |  | lack of activity output | no device output
PEMS development life-cycle |  | lack of activity input | compliance issue
PEMS development life-cycle |  | lack of activity | use of device issue
PEMS development life-cycle |  | fail to tailor process | improper or incorrect procedure or method
PEMS development life-cycle | software | fail to follow process | improper or incorrect procedure or method
PEMS validation |  | lack of validation plan | protective measure issue
PEMS validation |  | lack of reference | compliance issue
PEMS validation |  | lack of document | device markings issue
PEMS validation |  | inappropriate validation personnel | protective measure issue
PEMS validation |  | fail to validate essential performance | protective measure issue
PEMS validation |  | fail to validate basic safety | protective measure issue
PEMS validation |  | fail to follow plan | improper or incorrect procedure or method
PEMS validation | software | fail to document method | use of device issue
periodic notification | usability | fail to generate periodic notification | device alarm system issue
periodic notification |  | fail to clear periodic notification | prompts will not clear
personal event | clinical | unable to record event | data back-up problem
personal event |  | lack of sickness info | patient data issue
personal event |  | lack of meal info | patient data issue
personal event |  | lack of exercise info | use of incorrect control settings
personnel competent |  | inappropriate training | Inadequate Training
personnel competent |  | inappropriate skill | Inadequate Training
personnel competent |  | inappropriate personnel competent | Inadequate Training
personnel competent |  | inappropriate experience | Inadequate Training
personnel competent | human | inappropriate education | Inadequate Training
phone number, phone no. | business | incorrect phone number | patient data issue
photomultiplier tube | component | photomultiplier tube defective | defective component
physical resistance | physical | sticking | sticking
pilot balloon valve | component | valve stuck | sticking; difficult to open or close
pin number |  | reset pin number failed | failure to reset
pin number | data | invalid pin number | device markings issue
piston bearing | mechanical
piston guide | component
piston head | component
piston pad | component
plasma bond |  | poor seam weld | Solder Joint Failure
plasma clean, plasma cleaning | maintenance | Plasma Clean Expired | device expiration issue; device cleaning issue
pneumatic driver | pump | Inadequate pneumatic power | loss of power
pocket stimulation | physiology
Port interface | part | wrong connection | misconnection
position sensing, position sensor |  | position sensing out of specification | positioning issue
position sensing, position sensor |  | loss of position sensing data | loss of data
position sensing, position sensor | measure | fail to indicate valid position | no display or display failure
positive pressure |  | low peak positive pressure | pressure issue
post surgical care | therapy | therapy delayed | Unexpected Therapeutic Results
post-hydration
potential cause |  | SOUP failure | application program issue
potential cause |  | software defect | computer software issue
potential cause |  | reasonably foreseeable misuse | use of device issue
potential cause |  | lack of potential cause | use of device issue
potential cause |  | incorrect specification | physical property issue
potential cause |  | incomplete specification | physical property issue
potential cause |  | inappropriate potential cause | use of device issue
potential cause |  | hardware failure | defective component
potential cause | risk | fail to identify potential cause | compliance issue
potential cause |  | fail to identify potential cause | compliance issue
potentiometer |  | low reading | Low Readings
potentiometer | component | high reading | High Readings
power adaptor |  | power outage | power conditioning issue
power adaptor |  | installation difficult | installation-related problem
power adaptor |  | fail to output current | electronic property issue
power adaptor | power | fail to accept power source | Power Source Issue
power button |  | loss of power | loss of power
power button |  | Inadvertently activating button accidentally | deployment issue
power button | power | button damaged | break
power cable |  | loss of label | device markings issue
power cable |  | loss of connection | connection issue
power cable |  | loss of common receptacle | fitting problem; component or accessory incompatibility
power cable |  | cable short | defective component
power cable | power | cable disconnection | disconnection
power conditioning | energy
power connection |  | loss of power | loss of power
power connection | power | loss of connection label | device markings issue
power cord, power supply cord, power supply cords |  | over bending | bent
power cord, power supply cord, power supply cords |  | less strength | mechanical issue
power cord, power supply cord, power supply cords |  | lack of retention clip | defective component
power cord, power supply cord, power supply cords |  | lack of socket | defective component
power cord, power supply cord, power supply cords |  | less cross-sectional area | defective item
power cord, power supply cord, power supply cords | electrical | downsized power cord size | defective component
power distribution | power | loss of power | loss of power
power draw | power | power draw out of range | electrical power problem
power injection
power injector |  | injector burst | pressure issue; Burst
power input | power | invalid input | failure to read input signal
power lamp | energy
power line | electrical | loss of power | loss of power
power module |  | missing part | component missing
power module | energy | incompatible module | component incompatible
power on self-test, power on self test, power on self tests, POST test, POST tests, POST check, POST checks, POST results, POST result |  | fail to receive self-test results | failure of device to self-test
power on self-test, power on self test, power on self tests, POST test, POST tests, POST check, POST checks, POST results, POST result | test | fail to initiate self-test | failure of device to self-test
power on self-test, power on self test, power on self tests, POST test, POST tests, POST check, POST checks, POST results, POST result | test | fail to detect power status | failure to sense
power output | power | power output out of range | electrical power problem
power output | power | low power output | power conditioning issue
power path | power | loss of power | loss of power
power plug | power | plug mismatches socket | fitting problem; component incompatible
power range | power | power out of range | electrical power problem
power source | power | power source incompatibility | Power Source Issue
power source | power | power interrupted | loss of power
power source | power | lack of independent power source | Power Source Issue; Fail-Safe Design Failure
power source | power | fail to indicate AC or battery power | no display or display failure
power state |  | power state transition failed | device operates differently than expected
power state |  | power state initial failed | device operates differently than expected
power state |  | power state confused | device operates differently than expected
power state | power | incorrect power state | power conditioning issue
power supply |  | power supply malfunction | electrical power problem; device operates differently than expected; power source issue; Product Quality Issue
power supply |  | inappropriate power supply | power source issue
power supply |  | fail to disclose input voltage and frequency | instruction for use issue
power supply |  | power source detection issue | power source issue
power supply |  | power supply burned | power source issue
power supply |  | power interrupted | loss of power
power supply | energy | loss of state indication | no display or display failure
power supply |  | device burned | burn of device or device Component; electrical power problem; smoking
power supply |  | defective power supply unit | defective component; power source issue
power switch, power control |  | power on or off switch failed | loss of power
power switch, power control | power | fail to supply power | failure to power-up
power switch, power control |  | short confirmation time | use of device issue
power switch, power control |  | fail to manage power | power conditioning issue
power switch, power control |  | defective switch | defective component
prescribed exercise |  | untimely prescribed exercise | imprecision
prescribed exercise |  | prescribed exercise confused | improper or incorrect procedure or method
prescribed exercise | therapy | incorrect prescribed exercise | improper or incorrect procedure or method
pressure alarm |  | lack of patient disconnect alarm | device alarm system issue
pressure alarm |  | lack of occlusion disconnect alarm | occlusion within device; device alarm system issue
pressure alarm |  | lack of low mandatory Vt alarm | device alarm system issue
pressure alarm |  | lack of low Inspiratory pressure alarm | device alarm system issue
pressure alarm |  | lack of low Baseline Pressure alarm | device alarm system issue
pressure alarm |  | lack of high mandatory Vt alarm | device alarm system issue
pressure alarm | ventilator | lack of high inspiratory pressure alarm | device alarm system issue
pressure gage |  | component issue | defective component
pressure level |  | pressure level out of specification | pressure issue
pressure level | control | inappropriate pressure level | pressure issue
pressure limit | component | incorrect limit setting | use of incorrect control settings
pressure signal | electronic
pre-use check | quality | inadequate cross-check | failure of device to self-test
preventative maintenance | service | missing maintenance | device maintenance issue
preventative maintenance |  | fail to disclose preventative maintenance | lack of maintenance documentation or guidelines
preventative maintenance |  | fail to disclose service personnel | lack of maintenance documentation or guidelines
preventative maintenance |  | fail to disclose maintenance frequency | lack of maintenance documentation or guidelines
preventative maintenance |  | fail to disclose characterization of exhalation valve | lack of maintenance documentation or guidelines
preventative maintenance |  | fail to disclose battery replacement | lack of maintenance documentation or guidelines
preventative maintenance |  | incorrect preventative maintenance | Maintenance Does not comply to Manufacturers Recommendations
primary operating function |  | fail to identify safety | compliance issue
primary operating function | usability | fail to identify primary operating function | compliance issue
primary packaging
priority alarm |  | fail to raise medium priority alarm | improper alarm
priority alarm |  | fail to raise low priority alarm | improper alarm
priority alarm | usability | fail to raise high priority alarm | improper alarm
priority alarm |  | fail to escalate high priority alarm | improper alarm
privacy policy, privacy regulation |  | fail to follow privacy regulation | compliance issue
privacy policy, privacy regulation | safety | fail to email privacy policy | Market Issue
problem report |  | lack of problem report | use of device issue
problem report |  | lack of critical statement | use of device issue
problem report |  | fail to close problem report | compliance issue
problem report |  | fail to address problem | use of device issue
problem report | problem | absence of basic safety | device unsafe to use in environment
problem resolution system |  | lead to new problem | use of device issue
problem resolution system |  | lack of risk assessment | Fail-Safe Design Failure
problem resolution system |  | lack of mitigation action | use of device issue
problem resolution system |  | lack of document | device markings issue
problem resolution system |  | lack of acceptance criteria | use of device issue
problem resolution system |  | fail to resolve problem | use of device issue
problem resolution system | problem | fail to follow process | improper or incorrect procedure or method
problem resolution system |  | absence of essential performance | compliance issue
problem resolution system |  | absence of basic safety | device unsafe to use in environment
Product variation |  | inadequate size range to meet market demand | size incorrect for patient
Product variation | quality | inadequate language | Instruction for Use Issue; inadequate user interface
profile segment |  | inappropriate profile rate | use of incorrect control settings
profile segment | pump | inadequate profile | use of incorrect control settings
profile segment |  | fail to activate at start time | reset issue
Program parameter |  | insufficient program parameter | programming issue
Program parameter | program | incorrect program parameter | programming issue
programmable read only memory | electronic
protection circuit |  | lack of short circuit module | electrical shorting
protection circuit |  | lack of over voltage module | electronic property issue
protection circuit |  | lack of over heat module | overheating of device or device component
protection circuit |  | lack of over current module | electronic property issue
protection circuit | electrical | lack of over charge module | charging issue
protection system, over-voltage protection |  | lack of protection system | protective measure issue
protection system, over-voltage protection |  | lack of over-voltage protection | protective measure issue
protection system, over-voltage protection | safety | energy deactivation failed | unintended energization
protective device | safety | inadequate breaking capacity | Device-Device Incompatibility
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | protective earth conductor failed | failure to conduct
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | fail to disclose protective earth required | instruction for use issue
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement | safety | missing detachment tool | detachment of device component
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | loss of connection | connection issue
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | electrical shock | Inappropriate Shock
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | loose protective earth conductor | loose or intermittent connection
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | lack of earth pin | device misassembled during manufacturing or shipping
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | lack of colored insulation | protective measure issue
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | incorrect to use the earth terminal for others | improper or incorrect procedure or method
protective earth connection, protective earth conductor, protective earth terminal, protective earth, PE, PE requirement |  | incorrect to use moving part as protective earth element | improper or incorrect procedure or method
control measure, risk control, protective measure | safety | inappropriate protective measure | protective measure issue
protective packaging, protective package, packaging system |  | inadequate protection | Device Packaging Compromised
proximal end, free end |  | metal ion oxidation (MIO) | defective item
proximal end, free end | lead | loss of electrical insulation | electrical shorting
proximal end, free end |  | loss of electrical continuity | intermittent continuity
proximal end, free end |  | insulation breached | electrical shorting
proximal end, free end |  | insufficient electrical continuity | connection issue
proximal end, free end |  | insufficient crimp strength | mechanical issue
proximal end, free end |  | excessive friction | physical resistance
proximal end, free end |  | difficult to wrap | Packaging Issue
proximal end, free end |  | difficult to position | difficult to position
proximity switch | component
pull strength | mechanical | pull strength low | mechanical issue
pump communicator |  | stop delivery undesired | Pumping Stopped
pump communicator |  | stop bolus undesired | Pumping Stopped
pump communicator |  | loss of vibrating alarm | device alarm system issue
pump communicator |  | loss of audio alarm | device alarm system issue
pump communicator |  | initiate delivery undesired | delivery system failure
pump communicator |  | initiate bolus undesired | inaccurate delivery
pump communicator |  | fail to signal low reservoir | Failure to Interrogate
pump communicator | communication | fail to initiate bolus | failure to deliver
pump communicator |  | fail to indicate pump status | no display or display failure
pump communicator |  | fail to deliver square wave bolus | failure to deliver
pump communicator |  | fail to deliver immediate bolus | failure to deliver
pump communicator |  | fail to deliver dual wave bolus | failure to deliver
pump communicator |  | fail to configure pump communicator | device difficult to setup or prepare
pump communicator |  | fail to communicate with pump | communication or transmission issue
pump communicator | pump | data corrupted | data issue
pump communicator |  | component dropped | component falling
pump mechanism |  | reservoir leakage | leak
pump mechanism | pump | propellant leakage | leak
pump motor |  | motor cable failed | defective component
pump motor | pump | fail to drive pump at set speed | Decreased Pump Speed; Increased Pump Speed
pump motor |  | beat rate potentiometer failed | defective component
pump refill |  | unable to fill | filling problem
pump refill | pump | incomplete aspiration | aspiration issue
pump refill |  | contamination | contamination during use
pump stroke |  | unable to drive motor | failure to deploy
pump stroke |  | short interval | defective item
pump stroke |  | motor stop time short | imprecision
pump stroke |  | low accuracy | imprecision
pump stroke |  | lack mode of clinically insignificant amount | improper or incorrect procedure or method
pump stroke |  | incorrect stroke calculation | parameter calculation error due to software problem
pump stroke |  | incorrect encoder count | Incorrect Software Programming Calculations
pump stroke | pump | inaccuracy encoder count | Incorrect Software Programming Calculations
Radio Equipment Directive | regulatory | Directive violated | compliance issue
radio equipment directive, RED, REDs | regulatory | RED regulation violated | compliance issue
radio frequency transceiver, RF transceiver, RF transceivers |  | unable to power off | power conditioning issue
radio frequency transceiver, RF transceiver, RF transceivers | communication | loss of communication | wireless communication issue
radio frequency, RF | electrical | loss of communication | wireless communication issue
radio signal | electronic | radio signal interference | environmental control or utility issue
radiofrequency interference, radio-frequency interference, RFI | electrical
radio-opaque stripe, radiopaque stripe | component | lack of radiopaque strip | device markings issue
RAM test | test | fail to generate periodic test | incorrect or inadequate test results; electronic property issue
ramp time, ramp operation | time | fail to activate ramp time | imprecision
range limit | tolerance | out of specification | physical property issue
range of motion, motion range, motion ranges | brace | motion out of range | unintended system motion
read only memory, ROM | electronic
ready validation, ready check, ready checks |  | validation failed | protective measure issue
ready validation, ready check, ready checks |  | lack of validation | protective measure issue
ready validation, ready check, ready checks | infusion | fail to send out command | communication or transmission issue
receiver stimulator unit | component | receiver stimulator unit failure | defective component
re-chargeable battery, rechargeable battery, rechargeable batteries | electrical | charge cycles out of specification | charging issue
reconstructive knee | brace | reconstructive knee delayed | use of device issue
reconstructive knee |  | device malfunction | device operates differently than expected; Product Quality Issue
recovery program |  | lack of various preloaded muscle training | Inadequate Training
recovery program |  | inappropriate program selected | programming issue
recovery program |  | inadequate muscle training | Inadequate Training
recovery program |  | fail to clear remaining lactic acid | prompts will not clear; chemical spillage
recovery program | program | fail to bring body back to resting phase | device operates differently than expected
recovery status |  | untimely recovery status | imprecision
recovery status | therapy | recovery status confused | inadequate user interface
reference connector |  | rough surface | Scratched Material
reference connector |  | corrosion | Corrosion
reference guide |  | missing reference guide | device markings issue
reference guide |  | incorrect reference guide | device markings issue
reference guide | usability | inadequate reference guide | device markings issue
refill kit |  | fail to label single-use only | improper or incorrect procedure or method
refill kit | infusion | fail to be sterilzed | Failure to Sterilize; Delivered As Unsterile Product
refill needle |  | single use violated | use of device issue
refill needle |  | out of dimension | defective component
refill needle |  | inappropriate lubricant | component incompatible
refill needle |  | flow blocked | no flow
refill needle |  | fail to rounded tip | Failure to Fold
refill needle | infusion | fail to be sterilzed | Failure to Sterilize; Delivered As Unsterile Product
refill syringe | service | inappropriate syringe barrel | component incompatible
refill syringe |  | inappropriate plunger | component incompatible
refractory period | CRM | atrial pulses falling in refractory period | electronic property issue
regulatory requirement | standard | lack of regulatory requirement | compliance issue
relief valve, pressure relief valve, passive pressure relief valve |  | fail to trigger valve to reduce pressure | fail-safe design failure; protective measure issue
relief valve, pressure relief valve, passive pressure relief valve | component | valve stuck | sticking; difficult to open or close
remaining battery, battery remaining, battery remainings | power | fail to identify remaining battery | battery issue
remaining insulin, insulin remaining, insulin remainings |  | inaccuracy volume remaining | use of device issue
remaining insulin, insulin remaining, insulin remainings |  | inaccuracy remaining insulin | volume accuracy issue
remaining insulin, insulin remaining, insulin remainings |  | fail to display insulin volume remaining | no display or display failure
remaining insulin, insulin remaining, insulin remainings |  | fail to display insulin time remaining | no display or display failure
remaining insulin, insulin remaining, insulin remainings | diabete | fail to activate alarm | device alarm system issue
remaining volume, volume remaining, volume remainings |  | inaccuracy volume remaining | volume accuracy issue
remaining volume, volume remaining, volume remainings | infusion | fail to activate alarm | device alarm system issue
remote control | control | loss of control | deployment issue
remote transmitter |  | unsecured data | application security issue
remote transmitter |  | loss of function | no device output
remote transmitter |  | loss of communication | wireless communication issue
remote transmitter | communication | inappropriate operation | device handling issue
replaceable unit, replacement part, elective replacement |  | fail to disclose installation instruction | installation-related problem
replaceable unit, replacement part, elective replacement | component | replaceable unit damaged | break
residual risk | risk | fail to evaluate residual risk | use of device issue
residual risk |  | fail to determine residual risk | use of device issue
respiratory rate |  | slow respiratory rate | output problem
response time | software | long response time | imprecision
retaining pin, Equalization Pin | component | pin detached | protective measure issue
retention mechanism connection
reverse delivery | infusion | fail to stop infusion | pumping issue
reverse delivery | infusion | fail to allow reverse delivery | delivery system failure
reverse delivery | infusion | fail to activate alarm | device alarm system issue
RF communication | communication | loss of RF communication | wireless communication issue
RF communication |  | disabled RF communication | wireless communication issue
RF communication |  | degraded RF communication | wireless communication issue; degraded
RF rectification |  | unintended stimulation | Unexpected Therapeutic Results
RF rectification | communication | RF rectification during MRI scan | use of device issue
rigid part | brace | covering ankle undesired | Human Factors Issue
ring electrode |  | ring electrode separated | failure to separate
ring electrode | lead | insufficient surface area | defective item
rise time, rise time control, rise-time |  | rise time control failed | device operates differently than expected
rise time, rise time control, rise-time |  | low rise time accuracy | incorrect measurement
rise time, rise time control, rise-time |  | fail to activate control | device inoperable
rise time, rise time control, rise-time |  | fail to deactivate control | device remains activated
rise time, rise time control, rise-time |  | rise time out of range | imprecision; use of incorrect control settings
risk analysis | risk | inadequate risk analysis | Fail-Safe Design Failure
risk management activity | risk | lack of risk management activity | Fail-Safe Design Failure
RO marker, radiopaque marker |  | invisible under fluoroscopy | poor quality image
root cause | problem | fail to identify root cause | compliance issue
rotor assembly |  | loss of bond | failure to adhere or bond
rotor assembly | 3 | looss of bond | Loss of or Failure to Bond
rotor assembly |  | loose rotor | detachment of device component
rotor assembly | 3 | loose rotor | detachment of device component
rotor assembly | pump | loose magnet | detachment of device component
rotor assembly | 3 | loose magnet | detachment of device component
rough handle | motion | part damaged | material integrity issue
rough surface |  | tissue or vein damage | use of device issue; mechanical issue; Energy Output to Patient Tissue Incorrect
rough surface |  | loss of electrical isolation | electrical shorting
rough surface | mechanical | insulation damage | electrical shorting
safe distance | safety | safe distance out of specification | Fail-Safe Mechanism Issue
safe limit | safety | temperature out of range | temperature issue
safe limit | safety | pressure out of range | pressure issue
safe limit | safety | humidity out of range | Moisture or Humidity Problem
safe limit | safety | drug concentration not in safe range | medication error
safe range | drug
safety check | safety | safety check failed | failure of device to self-test
safety check |  | invalid limit applied | device handling issue
safety check |  | inadequate check contents | failure of device to self-test
safety factor | safety | low safety factor | device unsafe to use in environment
safety factor | safety | inappropriate safety factor | device unsafe to use in environment
safety interlock | safety | safety interlock inadequate | device unsafe to use in environment
safety margin | quality | inappropriate safety margin | device unsafe to use in environment
safety valve, solenoid valve, discharge solenoid valve, pressure discharge solenoid valve, bleed valve |  | fail to switch backup valve | fail-safe design failure
safety valve, solenoid valve, discharge solenoid valve, pressure discharge solenoid valve, bleed valve |  | fail to trigger valve to reduce pressure | protective measure issue
safety valve, solenoid valve, discharge solenoid valve, pressure discharge solenoid valve, bleed valve |  | fail to trigger calibration | Failure to Calibrate
safety valve, solenoid valve, discharge solenoid valve, pressure discharge solenoid valve, bleed valve |  | valve stuck open | sticking; difficult to open or close
safety valve, solenoid valve, discharge solenoid valve, pressure discharge solenoid valve, bleed valve | component | valve stuck close | sticking; difficult to open or close
salt tablet | chemical
sample size | quality | insufficient sample size | use of incorrect control settings
scar tissue | physiology | sensing deteriorated | device sensing issue
scar tissue | physiology | pacing around a scar tissue | failure to pace or properly pace; Energy Output to Patient Tissue Incorrect
screen graph | image | inadequate screen graph | image display error
screw head | component | incorrect screw head | component incompatible
screw tapper | component | screw tapper damaged | break
seal integrity
seal ring | mechanical | sealing ring damaged | failure to adhere or bond
seal strength
seal tracheal | component
secondary packaging
security code |  | security code breached | application security issue
sediment filter | equipment
selector valve | component | valve stuck | sticking; difficult to open or close
self-checking, self checking, self checkings |  | fail to generate alarm | device alarm system issue
self-checking, self checking, self checkings | test | fail to detect errors | device sensing issue
self-inspection, self inspection, self inspections | test | fail to initiate self-inspection | failure of device to self-test
self-test, self test, self tests |  | system performance impacted | improper device output
self-test, self test, self tests |  | inadequate test contents | incorrect or inadequate test results
self-test, self test, self tests |  | fail to test watchdog | incorrect or inadequate test results
self-test, self test, self tests |  | fail to test memory address | incorrect or inadequate test results
self-test, self test, self tests |  | fail to test data buses | failure of device to self-test
self-test, self test, self tests |  | fail to test current | incorrect or inadequate test results
self-test, self test, self tests | test | fail to test alarm | failure of device to self-test
self-test, self test, self tests |  | fail to pass self-test | failure of device to self-test; failure to calibrate
self-test, self test, self tests | test | fail to initiate self-test | failure of device to self-test; failure to calibrate
self-test, self test, self tests |  | fail to generate alarm | device alarm system issue
self-test, self test, self tests |  | fail to abort remaining test steps if failed | incorrect or inadequate test results
sense threshold | measure | low sensing threshold | Low Sensing Threshold
sensing switch, position sensing switch, position sensing switches | motor | wiring connection broken | break; disconnection
sensing switch, position sensing switch, position sensing switches |  | sensor magnet ripped from piston | incorrect measurement
sensing switch, position sensing switch, position sensing switches |  | no magnet field to trip switch | defective component
sensitivity segment |  | segment out of range | defective item
sensitivity segment |  | lack of start time | reset issue
sensitivity segment | pump | incorrect start time | reset issue
sensor calibration | quality | out of calibration | calibration error
sensor calibration |  | inaccurate sensor measurement data | incorrect measurement
sensor fitting | motor
separative device |  | fail to create assured separation level | failure to separate
serial number |  | unique number violated | compliance issue
serial number |  | missing product serial number | device markings issue
serial number |  | loss of data | loss of data
serial number | data | incorrect serial number | device markings issue
service interval, service life, operating life, service date | service | reduced service interval | device maintenance issue
service interval, service life, operating life, service date |  | fail to set next service date | failure to service
service interval, service life, operating life, service date | service | reduced service life | shelf life exceeded; failure to service
service manual, service procedure, service procedures | service | inadequate information | no information; failure to service
service mode |  | inadequate diagnostic features | use of device issue; failure to service
service mode |  | fail to access | human-device interface issue; failure to service
service personnel, service professional, service professionals | service | misconduct service activities | Maintenance Does not comply to Manufacturers Recommendations
service personnel, service professional, service professionals |  | incorrect settings | use of incorrect control settings
service personnel, service professional, service professionals |  | fail to upgrade software | application program version or upgrade problem; failure to service
service personnel, service professional, service professionals |  | fail to transfer data | data issue; failure to service
service personnel, service professional, service professionals |  | fail to restore to factory default state | use of incorrect control settings; failure to service
service personnel, service professional, service professionals |  | fail to replace component | device maintenance issue
service personnel, service professional, service professionals |  | fail to perform | no device output
service personnel, service professional, service professionals |  | fail to access | human-device interface issue; failure to service
service personnel, service professional, service professionals |  | difficult to remove | difficult to remove; failure to service
set time | service | set time failed | reset issue
set up, set-up |  | user setups incorrectly | use of incorrect control settings
set up, set-up |  | unclear setup procedures | improper or incorrect procedure or method
set up, set-up |  | setup time long | reset issue
set up, set-up |  | setup time failed | reset issue
set up, set-up |  | setup incorrectly | use of incorrect control settings
set up, set-up | maintenance | difficult to setup | device difficult to setup or prepare
setscrew, set screw |  | unable to tighten setscrew | detachment of device component
setscrew, set screw |  | unable to loose setscrew | detachment of device component
setscrew, set screw |  | loose setscrew | detachment of device component
setscrew, set screw | mechanical | fluid inside setscrew bore | fluid leak
setup test | test | fail to activate test | failure of device to self-test
shared power resource | power | loss of power | loss of power
sharp corner |  | tissue or vein damage | use of device issue; mechanical issue
sharp corner |  | loss of electrical isolation | electrical shorting
sharp corner | safety | insulation damage | electrical shorting
sharp edge |  | tissue or vein damage | use of device issue; mechanical issue
sharp edge |  | skin scratched | device handling issue
sharp edge |  | loss of electrical isolation | electrical shorting
sharp edge | safety | insulation damage | electrical shorting
shear strength |  | part torn | torn material
shear strength | mechanical | less strength | mechanical issue
sheath introducer |  | tip rollback | tip-over
sheath introducer |  | tip cracked | Crack
sheath introducer |  | radio detectability failed | failure to sense
sheath introducer |  | low peak tensile force | mechanical issue
sheath introducer |  | less flexible | material rigid or stiff
sheath introducer |  | fluid leak | leak
sheath introducer |  | effective length out of range | defective component
shelf life, shelf-life, storage life |  | shelf life short | shelf life exceeded
shelf life, shelf-life, storage life |  | shelf life exceeded | shelf life exceeded
shelf life, shelf-life, storage life | service | parts expired | device expiration issue
shift register | electronic
shipping container, shipping carton, shipping parcel, shipping package |  | shipping test failed | failure of device to self-test
shipping container, shipping carton, shipping parcel, shipping package |  | shipping damage | break
shipping container, shipping carton, shipping parcel, shipping package |  | inadequate labeling | device markings issue
shipping container, shipping carton, shipping parcel, shipping package |  | fail to label UDI | device markings issue
shipping container, shipping carton, shipping parcel, shipping package |  | excessive stack | unstable
shipping container, shipping carton, shipping parcel, shipping package | ship | part damaged | break
shock capacitor, shocking capacitor | component | inappropriate capacitor charging | charging issue; Device-Device Incompatibility
shock capacitor, shocking capacitor | component | capacitor damaged | break
shock coil | lead
shock counter | component
shock energy | CRM | shock to wrong place | Inappropriate Shock; use of device issue
shock test, vibration test, vibration tests, free fall test, free fall tests | electrical | fail to be tested | failure of device to self-test
short circuit | electrical | catch fire | Fire
shoulder bag |  | zipper stuck | sticking
shoulder bag |  | zipper broken | break
shoulder bag |  | unable to view display | no display or display failure
shoulder bag |  | unable to connect power | loss of power
shoulder bag |  | unable to carry in bag | size incorrect for patient
shoulder bag |  | strap broken | break
shoulder bag |  | stitching broken | break
shoulder bag |  | inadequate padding in strap | material integrity issue
shoulder bag |  | inability to hear alarm | low audible alarm
shoulder bag | accessory | bag improperly sized | size incorrect for patient
shoulder rest | component | shoulder rest detachment | detachment of device component
shoulder support | component | shoulder support detachment | detachment of device component
shut off | electrical | fail to shut off | failure to shut off
shutdown, shut down | control | unexpected shutdown | False Device Output; device inoperable
sigh breath |  | low volume | Output below Specifications
sigh breath |  | lack of sigh | failure to deliver
sigh breath |  | inappropriate sigh delivery | inaccurate delivery
sigh breath |  | high volume | output above specifications
sigh breath | ventilator | high peak flow | excess flow or over-infusion
signal conditioner | equipment
signal transmission | electrical | loss of transmission | communication or transmission issue
silicone lubrication |  | lack of lubrication | component missing
silicone lubrication |  | insufficient lubrication | component missing
silicone lubrication |  | excessive lubrication | defective component
Silicone Seal |  | inadequate seal | failure to adhere or bond
silicone sleeve |  | silicone sleeve stuck down | sticking
simulator | equipment
single use, single-use, single patient, single patient use |  | fail to label single-use only | device markings issue
single use, single-use, single patient, single patient use |  | fail to disclose single-use only | instruction for use issue
single use, single-use, single patient, single patient use |  | fail to label DO NOT REUSE | device markings issue
single use, single-use, single patient, single patient use | usability | single use violated | failure to sterilize; disinfection or sterilization issue at user location
size adjustment | motion | fail to adjust size | device difficult to setup or prepare; defective component
size chart |  | size chart confused | inadequate user interface
size chart | measure | less size | size incorrect for patient
skin contact | safety | skin allergic reaction | biocompatibility issue
slip ring | component
slotter line | component
snap electrode |  | metal residual existed | Particulates
snap electrode |  | material biocompatibility compromised | component or accessory incompatibility
snap electrode |  | loss of adhesive | failure to adhere or bond
snap electrode |  | lack of expiration date | Expiration Date Error
snap electrode |  | incorrect snap size | size incorrect for patient
snap electrode | physiology | inadequate hydrogel | material integrity issue
socket adaptor | component
soft part | brace | breathable failed | use of device issue
softgood, soft good, soft goods | material | uncomfortable | human-device interface issue
softgood, soft good, soft goods |  | inappropriate size | size incorrect for patient
softgood, soft good, soft goods |  | fail to hand wash | device cleaning issue
softgood, soft good, soft goods |  | fail to air dry | environmental control or utility issue
software configuration |  | software version incorrect | application program version or upgrade problem
software configuration |  | software firmware not compatible | application program version or upgrade problem
software configuration |  | lack of tool | computer software issue
software configuration |  | lack of time control | reset issue
software configuration |  | lack of task | use of device issue
software configuration |  | lack of setting | use of incorrect control settings
software configuration |  | lack of relationship | computer software issue
software configuration |  | lack of problem resolution | compliance issue
software configuration |  | lack of organization responsible | computer software issue
software configuration |  | incorrect configuration | use of incorrect control settings
software configuration | software | inadequate configuration | use of incorrect control settings
software detailed design |  | misalignment with architecture | compliance issue
software detailed design | standard | contradiction with architecture | compliance issue
software development plan |  | lack of update | data issue
software development plan |  | lack of tracking procedure | computer software issue
software development plan |  | lack of tool | computer software issue
software development plan |  | lack of standard | computer software issue
software development plan |  | lack of reference | compliance issue
software development plan |  | lack of procedure | compliance issue
software development plan |  | lack of problem resolving procedure | compliance issue
software development plan |  | lack of plan | computer software issue
software development plan |  | lack of method | improper or incorrect procedure or method
software development plan |  | lack of feedback | use of device issue
software development plan |  | lack of documenting procedure | device markings issue
software development plan |  | lack of criteria | computer software issue
software development plan |  | lack of change control | compliance issue
software development plan |  | incomplete task | use of device issue
software development plan |  | incomplete activity | computer software issue
software development plan |  | inadequate traceability | computer software issue
software development plan |  | inadequate software configuration | use of incorrect control settings
software development plan |  | inadequate information | no information
software development plan |  | fail to follow risk management process | improper or incorrect procedure or method
software development plan | software | fail to address deliverable | computer software issue
software limit | software | out of software limit | computer software issue
software maintenance, maintenance plan, maintenance plans |  | lack of tracking procedure | computer software issue
software maintenance, maintenance plan, maintenance plans |  | lack of receiving procedure | compliance issue
software maintenance, maintenance plan, maintenance plans |  | lack of procedure | compliance issue
software maintenance, maintenance plan, maintenance plans |  | lack of problem resolving procedure | compliance issue
software maintenance, maintenance plan, maintenance plans |  | lack of plan | lack of maintenance documentation or guidelines
software maintenance, maintenance plan, maintenance plans |  | lack of obsolescence procudure | lack of maintenance documentation or guidelines
software maintenance, maintenance plan, maintenance plans |  | lack of feedback | use of device issue
software maintenance, maintenance plan, maintenance plans |  | lack of evaluating procedure | use of device issue
software maintenance, maintenance plan, maintenance plans |  | lack of documenting procedure | device markings issue
software maintenance, maintenance plan, maintenance plans |  | lack of criteria | lack of maintenance documentation or guidelines
software maintenance, maintenance plan, maintenance plans |  | fail to identify gap | compliance issue
software maintenance, maintenance plan, maintenance plans |  | fail to follow risk management process | improper or incorrect procedure or method
software maintenance, maintenance plan, maintenance plans | software | fail to follow process | improper or incorrect procedure or method
software requirement |  | misalignment with the other requirement | computer software issue
software requirement |  | lack of update | data issue
software requirement |  | lack of traceability | computer software issue
software requirement |  | lack of security requirement | computer system security issue
software requirement |  | lack of risk control measure requirement | no fail-safe mechanism; computer software issue
software requirement |  | lack of regulatory requirement | compliance issue; computer software issue
software requirement |  | lack of re-evaluation | use of device issue
software requirement |  | lack of operation requirement | computer software issue
software requirement |  | lack of maintenance requirement | lack of maintenance documentation or guidelines
software requirement |  | lack of IT-network requirement | application network issue
software requirement |  | lack of interface requirement | human-device interface issue
software requirement |  | lack of installation requirement | problem with software installation
software requirement |  | lack of input and output requirement | computer software issue
software requirement |  | lack of functional requirement | computer software issue
software requirement |  | lack of data definition and database requirement | data issue
software requirement |  | lack of capability requirement | computer software issue
software requirement |  | lack of alarm or warning or operator message requirement | device alarm system issue
software requirement |  | lack of acceptance requirement | computer software issue
software requirement |  | fail to establish test criteria | incorrect or inadequate test results
software requirement |  | fail to be tested | failure of device to self-test
software requirement |  | duplicated requirement | computer software issue
software requirement |  | contradicting requirement | computer software issue
software requirement | software | ambiguity requirement | computer software issue
software safety class, software safety classification |  | lack of software safety class | compliance issue
software safety class, software safety classification |  | lack of rationale | lack of maintenance documentation or guidelines
software safety class, software safety classification | software | inappropriate software safety class | compliance issue
software upgrade, software installation |  | incomplete software upgrade | application program version or upgrade problem
software upgrade, software installation |  | fail to disclose verification of software installation | problem with software installation
software upgrade, software installation |  | lack of release test | false device output
software upgrade, software installation | software | fail to upgrade software | application program version or upgrade problem
solder joint |  | solder joint fractured | Solder Joint Failure
solder joint |  | solder joint cold | Solder Joint Failure
solder joint | mechanical | open solder joint | Solder Joint Failure
solder joint |  | improper solder joint | Solder Joint Failure
sound absorber | equipment
sound pressure, sound pressure level | usability | sound pressure out of specification | audible noise
sound pressure, sound pressure level | usability | sound pressure level out of range | audible noise | sound pressure level out of specification
sound pressure, sound pressure level | usability | inappropriate sound pressure | audible noise
SOUP configuration |  | lack of unique designator | device markings issue
SOUP configuration |  | lack of title | device markings issue
SOUP configuration | software | lack of manufacturer | device markings issue
spark gap | mechanical
speaker system |  | sound pressure out of specification | audible noise
speaker system |  | loss of audio notification | not audible alarm
speaker system | usability | frequency out of range | electronic property issue
speaker volume, speak volume, speak volumes | usability | low volume | Output below Specifications
speech processor | electronic | speech processor failure | defective component
spirometric value | ventilator | incorrect value | use of device issue; data issue
spontaneous breathing trial | ventilator | fail to conduct | failure to conduct
spring loading mechanism | system | mechanism failure | protective measure issue
square bolus flow rate | pump | flow rate out of range | inaccurate flowrate
square wave bolus |  | incorrect bolus calculation | dose calculation error due to software problem
square wave bolus |  | inappropriate delivery duration | inaccurate delivery
square wave bolus |  | inappropriate delivery amount | inaccurate delivery
square wave bolus | pump | fail to support square wave bolus | use of device issue
start time | time | lack of start time | reset issue
start time |  | incorrect start time | imprecision
stator assembly |  | winding system overheat | overheating of device or device component
stator assembly |  | winding short | defective item
stator assembly |  | winding open failed | difficult to open or close
stator assembly |  | solder joint failed | Solder Joint Failure
stator assembly | pump | lead wire solder failed | Solder Joint Failure
steel insert |  | difficult to transfer through | device handling issue
steering wire | component | steering wire problem | device misassembled during manufacturing or shipping
stent graft |  | invisible under fluoroscopy | poor quality image
step counter, step count |  | inaccurate counter number | Incorrect Software Programming Calculations
step counter, step count |  | step count out of range | use of incorrect control settings
step counter, step count | measure | fail to measure body movement | unintended movement; Missing Test Results
stepper motor | electrical
sterile barrier system |  | unacceptable particular level | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | unacceptable linting level | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | unacceptable cleanliness level | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | seal discontinuity | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | integrity compromised | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | inadequate seal strength | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | fail to prevent ingress of microorganisms | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile barrier system |  | contaminated | device contamination with biological material; delivered as unsterile product; bacterial contamination of device
sterile drape | sterility | sterile drape compromised | Failure to Sterilize; Delivered As Unsterile Product
sterile drape, sterile bagging, sterile bag, sterile bags | sterility | undesired multiple members involved | device handling issue
sterile glove | sterility | sterile glove compromised | Failure to Sterilize; Delivered As Unsterile Product
sterility assurance level | sterility | sterility compromised | Delivered As Unsterile Product; Failure to Sterilize
sterilization compatibility |  | packaging material fails to withstand sterilization process | Failure to Sterilize
sterilization cycle |  | fail to reach designated cycles | Failure to Sterilize
sterilization dose |  | insufficient dose | incorrect or inadequate test results
sterilization load
sterilizing agent |  | unqualified sterilizing agent | Failure to Sterilize
steroid plug |  | inappropriate plug size | defective component
steroid plug | lead | fall out of header cavity | detachment of device component
stimulation program | physiology | inappropriate output level | improper device output
stimulation rate | CRM | inappropriate stimulation rate | therapeutic or diagnostic output failure
stop distance | safety | stopping distance over limit | Fail-Safe Mechanism Issue
stopcock valve | component | stopcock valve failure | defective component
storage cover |  | water proof failed | fluid leak
storage cover |  | shipping test failed | failure of device to self-test
storage cover | part | insufficient storage space | inadequate storage
storage fluid |  | incorrect chemical name | chemical issue
storage position, storage pose, storage poses |  | inappropriate storage pose | device maintenance issue
storage position, storage pose, storage poses | motion | inappropriate position | malposition of device
strain relief | component | part worn out | Naturally Worn
strap material, adjustable strap, adjustable straps |  | tightening adjustment difficult | device difficult to setup or prepare
strap material, adjustable strap, adjustable straps |  | strap material torn | torn material
strap material, adjustable strap, adjustable straps | brace | position adjustment difficult | difficult to position
strength program |  | inappropriate program selected | programming issue
strength program |  | inappropriate muscle fiber rest periods | Output Issue
strength program |  | inappropriate muscle fiber explosive strength | use of device issue
strength program |  | inappropriate muscle fiber excitement | Unexpected Therapeutic Results
strength program | program | inappropriate muscle fiber contraction times | Retraction Problem
strip line | component | lack of strip line | component missing
stuck key test | test | fail to generate alarm | device alarm system issue
substrate epoxy | chemical | cracked substrate epoxy | Crack
substrate prosthesis
suction port |  | part damaged | suction issue
summative evaluation |  | misunderstandable safety information | no information
summative evaluation |  | misconduct in early stage | failure to conduct
summative evaluation |  | lack of user interface | inadequate user interface
summative evaluation |  | lack of training | Inadequate Training
summative evaluation |  | lack of summative evaluation | use of device issue
summative evaluation |  | lack of hazard related scenario | use of device issue
summative evaluation |  | lack of evaluation method | use of device issue
summative evaluation |  | inappropriate method | improper or incorrect procedure or method
summative evaluation |  | inadequate objective evidence | use of device issue
summative evaluation |  | inadequate hazard related scenario | use of device issue
summative evaluation |  | fail to validate the safe use | protective measure issue
summative evaluation | usability | fail to identify consequence of use error | use of device issue
surface acoustic wave device, SAW | equipment
surgeon chair | hardware | low weight capacity | patient-device incompatibility
surgical graft | component | surgical graft failure | defective component
surgical instrument | equipment | damage surgical instrument | device damaged prior to use
suspend mode |  | fail to suspend delivery | delivery system failure
suspend mode | pump | fail to generate periodic notification | device alarm system issue
suspend mode |  | fail to deliver at clinically insignificant rate | failure to deliver
suspend mode |  | fail to clear periodic notification | computer software issue
suspend mode |  | fail to cancel suspend mode | computer software issue
suture sleeve |  | tie down crush | detachment of device component
suture sleeve | lead | suture sleeve missized | size incorrect for patient
suture sleeve |  | suture sleeve cut | Cut in Material
suture sleeve |  | loose suture sleeve | detachment of device component
synthetic material |  | incorrect nature of chemical treatment or modification | chemical issue
synthetic material |  | incorrect generic or chemical name | chemical issue
synthetic nontextile prosthesis
synthetic textile prosthesis
syringe connection
syringe drive | component | syringe drive problem | failure to deploy
syringe marking |  | syringe marking incorrectly placed | device markings issue
syringe marking | label | inaccurate syringe marking | device markings issue
system activation | system | premature system activation | Premature Deployment
T wave, R wave, P wave | CRM | vulnerable T wave | electronic property issue
T wave, R wave, P wave | CRM | variation to detect the R wave | device sensing issue
T wave, R wave, P wave | CRM | oversensing | Over-sensing
T wave, R wave, P wave | CRM | incorrect wave amplitude | electronic property issue
T wave, R wave, P wave | CRM | incorrect triggered pulse | electronic property issue
T wave, R wave, P wave | CRM | alignment issue | failure to align
table top | component
tachy rate | CRM | tachy rate cut off low | Electromagnetic Interference (EMI)
tactile feedback | usability | no tactile feedback | use of device issue
tactile feedback |  | loss of feedback | use of device issue
tactile feedback | usability | incorrect feedback | use of device issue
tactile feedback | usability | inappropriate tactile feedback | use of device issue
tactile prompt | tactile | no tactile prompt | Inappropriate Prompts
tactile prompt | tactile | inappropriate tactile prompt | Inappropriate Prompts
tap water | material
tapered tubular vascular graft |  | usable length out of specification | physical property issue
tapered tubular vascular graft |  | pressurized internal diameter out of specification | physical property issue
tapered tubular vascular graft |  | end internal diameter out of specification | physical property issue
task analysis |  | inadequate task analysis | improper or incorrect procedure or method
task analysis | usability | fail to conduct task analysis | use of device issue
telemetry equipment | equipment
television unit | equipment
temperature compensator | instrument
temperature cycle | environment | temperature out of range | temperature issue
temperature cycle |  | insufficient time | imprecision
temperature probe | component | loose temperature probe | detachment of device component
temporary basal rate |  | lack of single basal rate | inaccurate delivery
temporary basal rate |  | lack of single basal duration | inaccurate delivery
temporary basal rate |  | fail to resume | reset issue
temporary basal rate | pump | fail to align with basal pofile rate | failure to align
tens belt |  | overuse | use of device issue
tens belt | physiology | misuse | use of device issue
tens belt |  | inappropriate usage | use of device issue
tens device | physiology | inadequate muscle training | Inadequate Training
term of service, service term, service terms, terms & conditions, term & condition | service | fail to email Terms & Conditions | device markings issue
terminal assembly | component
terminal sterilization
termination sleeve |  | termination sleeve undersized | size incorrect for patient
termination sleeve | lead | termination sleeve oversized | size incorrect for patient
test documentation |  | lack of tester | incorrect or inadequate test results
test documentation |  | lack of test tool | incorrect or inadequate test results
test documentation |  | lack of test result | incorrect or inadequate test results
test documentation |  | lack of test date | incorrect or inadequate test results
test documentation |  | lack of test configuration | incorrect or inadequate test results
test documentation |  | lack of software version | application program version or upgrade problem
test documentation | standard | fail to record anomaly | data back-up problem
test procedure |  | incorrect test procedure | incorrect or inadequate test results
test procedure | standard | inadequate test procedure | incorrect or inadequate test results
test result |  | missing test result | missing test results
test result |  | lack of test result | missing test results
test result | test | incorrect test result | incorrect or inadequate test results
test result |  | inadequate test result | incorrect or inadequate test results
test strip | component | incompatible test strip | failure of device to self-test
theory of operation, operation theory, operation theories | operation | incorrect operation theory | device handling issue
therapeutic output | clinical
thermal cut-out | safety | triggered by worst case normal operating condition | self-activation or keying
thermal insulation | safety | inadequate thermal insulation | electrical shorting
thermocouple wire | component
thoracic vein | CRM | thoracic vein punctured | use of device issue; Material Puncture
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | user uncomfortable | human-device interface issue
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | user confused | inadequate user interface
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | splint dislodgement | dislodged or dislocated
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | over bending | bent
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | missing splint | component missing
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | missing soft goods | component missing
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | low stiffness | material rigid or stiff
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | lack of ductile | material rigid or stiff
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | inadequate size | size incorrect for patient
thumb splint, finger splint, finger splints, wrist splint, wrist splints |  | difficult to remove | difficult to remove
thumb splint, finger splint, finger splints, wrist splint, wrist splints | brace | adjustment difficult | device difficult to setup or prepare
tidal volume |  | volume control failed | tidal volume fluctuations
tidal volume |  | tidal volume fluctuation | Tidal Volume Fluctuations
tidal volume |  | inappropriate measurement range | incorrect measurement; measurement system incompatibility
tidal volume |  | low volume accuracy | volume accuracy issue
tidal volume |  | fail to update tidal volume | tidal volume fluctuations
tidal volume |  | fail to set tidal volume limit | tidal volume fluctuations
tidal volume |  | fail to raise high prority alarm | device alarm system issue
tidal volume |  | fail to deliver tidal volume | tidal volume fluctuations
tidal volume |  | fail to compensate tidal volume | tidal volume fluctuations
tidal volume |  | fail to clear tidal volume | tidal volume fluctuations
tidal volume | measure | fail to calculate tidal volume | tidal volume fluctuations
tilt angle |  | tilt angle out of range | defective item
tilt angle | measure | loss of tilt angle data | loss of data
time clock | electrical | incorrect time and date | imprecision
time clock | electrical | corruption of encrypted real time clock | imprecision
time duration |  | fail to disclose warm up period | instruction for use issue
time duration | time | time out of range | imprecision
time format | usability | inadequate time format | imprecision
time verification | time | fail to verify time set | reset issue
time verification |  | fail to generate indication | no display or display failure
tip conductor coil | component
tip of the catheter, tip of catheter, tip of catheters, catheter tip, catheter tips |  | less flexible | material rigid or stiff
tip of the catheter, tip of catheter, tip of catheters, catheter tip, catheter tips |  | lack of longitudinal slits | component missing
tip of the catheter, tip of catheter, tip of catheters, catheter tip, catheter tips |  | flow blocked | no flow
tip of the catheter, tip of catheter, tip of catheters, catheter tip, catheter tips | catheter | complete blockage | blockage within device or device component
tissue force |  | permanent deformation | mechanical issue
tissue force | motion | force out of range | device handling issue
tissue load |  | permanent deformation | mechanical issue
tissue load | motion | load out of specification | mechanical issue
tissue torque |  | torque out of range | defective item
tissue torque | motion | permanent deformation | mechanical issue
titan capacitor | component | partial energy delivered | output energy incorrect
titan capacitor |  | inappropriate foil material selection | use of device issue
titan capacitor
titan capacitor
tomographic pallet | equipment | pallet cracked | Crack
tone test | test | fail to generate periodic test | failure of device to self-test
tool home, tool homing |  | Zero Position out of Range | positioning issue
tool home, tool homing | software | position repeatability out of specification | positioning issue
tool translate, tool translation | tool | tool translation failed | deployment issue
touch current | electrical | current out of range | electrical issue
touch input | usability | loss of touch input | failure to read input signal
tracking form |  | missing tracking form | device markings issue
tracking system |  | tracking system exceeds limit | optical issue
tracking system |  | tracking error | optical issue
tracking system |  | sensor damaged | break; failure to read input signal
tracking system |  | over magnetic exposure | use of device issue; environmental control or utility issue; electromagnetic compatibility issue
tracking system |  | loss of feedback | use of device issue
tracking system |  | insensitive sensing | decreased sensitivity; device sensing issue
tracking system | system | expose to over magnetic field | Radiation overexposure
tracking system |  | detect position slow | device sensing issue
tracking system |  | detect position inaccurately | device sensing issue
tracking system |  | detect orientation slow | device sensing issue; failure to align
tracking system | system | detect orientation inaccurately | device sensing issue; failure to align
traction table |  | inappropriate user | use of device issue
traditional use | user
training mode |  | fail to deactivate mode | computer software issue
training mode | usability | fail to activate mode | computer software issue | enter mode
training program | training | inappropriate training program | Inadequate Training
transducer pressure probe | measure
transducer probe | component | transducer probe overheating | overheating of device or device component
transfer set |  | transfer set leakage | fluid leak
transformer winding |  | inadequate insulation | electrical shorting
transformer winding | component | high winding temperature | overheating of device or device component
transit dose |  | over dose wasted | incorrect or inadequate test results
transmission line | electrical
transport ventilator |  | loss of function | device handling issue
transport ventilator | ventilator | fail to transport patient | device handling issue
transportation configuration |  | transportation configuration failed | Manufacturing or Shipping Issue Associated with Device
transportation configuration |  | package delayed | Packaging Issue
transportation configuration |  | loss of transportation configuration indication | no display or display failure
transportation configuration | ship | inappropriate transportation configuration | Manufacturing or Shipping Issue Associated with Device
trap zone |  | miss guards and protective measures | protective measure issue
trap zone |  | inappropriate trapping zones | defective item
trap zone |  | inappropriate safe distance | optical issue
trap zone | dimension | inappropriate gap | device misassembled during manufacturing or shipping
tube capillary | component
tube compensation | ventilator | lack of compensation | improper device output
tubular vascular graft |  | pressurized internal diameter out of specification | size incorrect for patient
tubular vascular graft |  | internal diameter out of specification | size incorrect for patient
tubular vascular graft |  | inappropriate material construction | chemical issue
tubular vascular graft |  | inappropriate direct visualization surgical technique | human-device interface issue
tubular vascular graft |  | inable to prevent blood from flowing throught implant wall | fluid leak
tubular vascular graft |  | inable to maintain adequate patency | insufficient flow or underinfusion
tubular vascular graft |  | inable to maintain adequate material and structural integrity | material integrity issue
tubular vascular graft |  | inable to anastomose to a blood vessel | size incorrect for patient
tubular vascular graft |  | implant dimension incosistency | size incorrect for patient
ultra low penetration air filter, ULPA filter |  | particle collection inefficiency | Environmental Particulates
ultra low penetration air filter, ULPA filter |  | integrity compromised | Material Disintegration
uneven floor | environment | device falling down | slippage of device or device component; component falling
unit of measure, measure unit, measurement unit |  | inappropriate measurement unit | incorrect measurement
unit of measure, measure unit, measurement unit | measure | inadequate measurement unit | incorrect measurement
universal size collar |  | inappropriate fit | fitting problem
universal size collar | brace | adjustment height out of specification | physical property issue
unsanitary condition | environment | bio-contamination | device contamination with biological material; delivered as unsterile product
uoup, interface of unknown provenance | usability | fail to evaluate UOUP | human-device interface issue
usability test |  | use device incorrectly | use of device issue
usability test |  | usability test failed | failure of device to self-test
usability test |  | lack of user profile | Instruction for Use Issue
usability test |  | lack of use condition | use of device issue
usability test |  | lack of usability test | incorrect or inadequate test results
usability test |  | lack of training | Inadequate Training
usability test |  | lack of test environment | incorrect or inadequate test results
usability test |  | lack of representative | device markings issue
usability test |  | lack of accompanying documentation | device markings issue
usability test |  | inappropriate test design | failure of device to self-test
usability test | usability | inappropriate method of collecting data | improper or incorrect procedure or method
usable length |  | length out of range | defective component
usb, removable media |  | wrong connection | misconnection
usb, removable media |  | transfer data failed | data issue
usb, removable media |  | incompatible accessory | accessory incompatible
usb, removable media |  | loss of connection | connection issue
usb, removable media | electrical | charge battery failed | Failure to Charge
use by date |  | sterility compromising | Delivered As Unsterile Product; Failure to Sterilize
use environment |  | inappropriate simulated environment setting | use of incorrect control settings
use environment |  | inappropriate lab environment setting | use of incorrect control settings
use environment | usability | inappropriate actual environment setting | use of incorrect control settings
use error, user error, user errors, operator error, operator errors |  | unqualified use | Inadequate Training
use error, user error, user errors, operator error, operator errors |  | short cut | improper or incorrect procedure or method
use error, user error, user errors, operator error, operator errors |  | routine violation | device misassembled during manufacturing or shipping
use error, user error, user errors, operator error, operator errors |  | risk from use error | use of device issue
use error, user error, user errors, operator error, operator errors |  | reversal | device misassembled during manufacturing or shipping
use error, user error, user errors, operator error, operator errors |  | operator error | device handling issue
use error, user error, user errors, operator error, operator errors |  | omission | device misassembled during manufacturing or shipping
use error, user error, user errors, operator error, operator errors |  | mistiming | imprecision
use error, user error, user errors, operator error, operator errors |  | misordering | device misassembled during manufacturing or shipping
use error, user error, user errors, operator error, operator errors |  | lack of use error consideration | use of device issue
use error, user error, user errors, operator error, operator errors |  | intrusion | material protrusion
use error, user error, user errors, operator error, operator errors |  | inadequately trained | Inadequate Training
use error, user error, user errors, operator error, operator errors |  | inadequate or improper illumination | Inadequate Lighting
use error, user error, user errors, operator error, operator errors |  | inaccessibility | human-device interface issue
use error, user error, user errors, operator error, operator errors |  | glare | use of device issue
use error, user error, user errors, operator error, operator errors |  | fatigue | use of device issue
use error, user error, user errors, operator error, operator errors |  | fail to identify use error | use of device issue
use error, user error, user errors, operator error, operator errors |  | fail to identify root cause | compliance issue
use error, user error, user errors, operator error, operator errors | usability | fail to analyze new use error | use of device issue
use error, user error, user errors, operator error, operator errors |  | application of bad rule | improper or incorrect procedure or method
use exclusion | usability | usage violated | use of device issue
use scenario |  | lack of task sequence | improper or incorrect procedure or method
use scenario |  | lack of task | use of device issue
use scenario |  | lack of high severity use scenario | use of device issue
use scenario |  | lack of hazard related use scenario | use of device issue
use scenario |  | lack of harm severity | use of device issue
use scenario |  | inappropriate use scenario selected | use of device issue
use scenario |  | inadequate hazard related use scenario | use of device issue
use scenario | usability | fail to identify hazard related use scenario | use of device issue
use specification |  | lack of use environment | use of device issue
use specification |  | lack of operating principle | improper or incorrect procedure or method
use specification |  | lack of intended user profile | Instruction for Use Issue
use specification |  | lack of intended patient population | user used incorrect product for intended use
use specification |  | lack of intended part | use of device issue
use specification |  | lack of intended medical indication | no display or display failure
use specification | usability | inadequate specification | use of device issue
user group, age group, age groups | user | inappropriate user | user used incorrect product for intended use
user interaction |  | fail to undo user interaction | application program issue
user interaction |  | fail to redo user interaction | application program issue
user interaction | usability | lack of user interaction process | human-device interface issue
user interface, User control interface, GUI |  | lack of keypad | human-device interface issue
user interface, User control interface, GUI |  | lack of display | no display or display failure
user interface, User control interface, GUI, graphical user interface |  | lack of audio output | no device output
user interface, User control interface, GUI, graphical user interface |  | inappropriate visual interface | human-device interface issue
user interface, User control interface, GUI, graphical user interface |  | inappropriate tactile display | incorrect display
user interface, User control interface, GUI, graphical user interface |  | inappropriate software interface | inadequate user interface; application program version or upgrade problem
user interface, User control interface, GUI, graphical user interface |  | inappropriate auditory interface | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | inadequate user interface | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | inadequate use safety design | protective measure issue
user interface, User control interface, GUI, graphical user interface |  | fail to identify safety | human-device interface issue
user interface, User control interface, GUI, graphical user interface |  | fail to identify primary operating function | human-device interface issue
user interface, User control interface, GUI, graphical user interface | usability | fail to identify interface characteristics | human-device interface issue
user manual |  | missing usage time | imprecision
user manual |  | missing troubleshooting info | Instruction for Use Issue
user manual |  | missing side effect notes | Instruction for Use Issue
user manual |  | missing operation info | Instruction for Use Issue
user manual |  | missing installation info | installation-related problem
user manual |  | missing contraindication | Instruction for Use Issue
user manual |  | incorrect user manual | Instruction for Use Issue
user manual | usability | inadequate user manual | Instruction for Use Issue
user profile | usability | lack of user profile | Instruction for Use Issue
user profile |  | inadequate user profile | Instruction for Use Issue
ut | regulatory | incorrect product classification | regulatory Issue
valve cam | component
valve control | software | fail to re-enable control | deployment issue
valve diaphragm | ventilator | fail to replace component | device maintenance issue
vascular patch |  | wall thickness out of specification | physical property issue
vascular patch |  | nominal width out of specification | physical property issue
vascular patch |  | nominal length out of specification | physical property issue
vascular patch |  | incompatible patch | accessory incompatible
vascular patch |  | inappropriate material construction | chemical issue
vascular patch |  | inable to prevent blood from flowing throught implant wall | fluid leak
vascular patch |  | inable to maintain adequate patency | insufficient flow or underinfusion
vascular patch |  | inable to maintain adequate material and structural integrity | material integrity issue
vascular patch |  | inable to anastomose to a blood vessel | size incorrect for patient
vascular prosthesis |  | fail to reside partially or completely within a blood vessel | use of device issue
vascular prosthesis |  | fail to form an internal bypass or shunt | use of device issue
vdc power, vdc system, vdc systems, vdc system power, vdc system powers | power | voltage out of range | electrical issue
vdc power, vdc system, vdc systems, vdc system power, vdc system powers |  | incorrect voltage value | electronic property issue
ventilation mode | ventilator | incorrect ventilation mode | use of device issue
ventilation mode |  | fail to select mode | use of device issue
ventilation mode |  | ventilation mode incomptable with circuit type | component incompatible
ventilation mode |  | inadequate ventilation mode | use of device issue
ventilator status | ventilator | incorrect ventilation mode | failure to cycle
ventricular channel | CRM | ventricular channel inhibited | communication or transmission issue
ventricular pacing | CRM | inappropriate pacing triggered | failure to pace or properly pace
ventricular pacing | CRM | disable pacing | No Pacing
ventricular rate | CRM | slow ventricular rate | capturing issue
ventricular rate |  | inappropriate rate calculation | parameter calculation error due to software problem
verification method, verification technique, verification techniques |  | inappropriate white box testing | failure of device to self-test
verification method, verification technique, verification techniques |  | inappropriate walkthroughs | protective measure issue
verification method, verification technique, verification techniques |  | inappropriate statistical testing | failure of device to self-test
verification method, verification technique, verification techniques |  | inappropriate static analysis | protective measure issue
verification method, verification technique, verification techniques |  | inappropriate inspections | protective measure issue
verification method, verification technique, verification techniques |  | inappropriate dynamic analysis | protective measure issue
verification method, verification technique, verification techniques | test | inappropriate black box testing | failure of device to self-test
verification plan |  | lack of verification technique | protective measure issue
verification plan |  | lack of verification strategy | protective measure issue
verification plan |  | lack of verification activity | protective measure issue
verification plan |  | lack of criteria | protective measure issue
verification plan |  | lack of appropriate personnel | protective measure issue
verification plan |  | inadequate verification plan | protective measure issue
verification plan |  | fail to plan verification tool | protective measure issue
verification plan |  | fail to plan verification on each function | protective measure issue
verification plan | standard | fail to plan verification at milestone | protective measure issue
verification plan |  | fail to follow plan | improper or incorrect procedure or method
vial spike, vial adapter |  | medication sprayed | fluid leak
vial spike, vial adapter |  | check vial alarm | device alarm system issue
vibrating alarm, vibration alarm, vibration alarms, vibration notification, vibration notifications | safety | loss of vibrating alarm | device alarm system issue
vibrating alarm, vibration alarm, vibration alarms, vibration notification, vibration notifications | safety | fail to configure vibrating alarm | device alarm system issue
video connection | video | loss of video connection | disconnection
video data | video | invalid encoded data | data issue
video data | video | data corrupted | data issue
video display terminal, VDT | component
video input | video | loss of video input | inadequate user interface
video output, HDMI, HDMI port | video | loss of video output | output issue
video output, HDMI, HDMI port | video | inadequate video output channel | inadequate user interface
video output, HDMI, HDMI port |  | fail to label text | device markings issue
video output, HDMI, HDMI port |  | fail to label standard symbol | device markings issue
video output, HDMI, HDMI port |  | fail to stereo mode | output issue
video output channel |  | low resolution | Image Resolution Poor
video output channel | video | loss of connection | connection issue
video output channel |  | fail to stereo mode | use of device issue
video output channel |  | fail to independent channel | communication or transmission issue
visibility angle | video | visibility angle out of specification | human-device interface issue
visibility distance | video | visibility distance out of range | human-device interface issue
visual alarm, visibly alarm, visibly alarms |  | no visual alarm | alarm not visible
visual alarm, visibly alarm, visibly alarms | safety | loss of visual alarm | alarm not visible
visual feedback | usability | no visual feedback | human-device interface issue
visual indicator |  | loss of indication | no display or display failure
visual indicator | video | confused indication | no display or display failure
visual prompt |  | unclear visual prompt | human-device interface issue
visual prompt | vision | no visual prompt | human-device interface issue
visualization computer |  | no computing | operating system becomes nonfunctional
visualization computer | video | loss of visual | human-device interface issue
visualization processor |  | no computing | operating system becomes nonfunctional
visualization processor |  | loss of visual | human-device interface issue
visualization processor |  | loss of video output | inadequate user interface
visualization processor |  | fail to simultaneously generate display outputs | no display or display failure
visualization processor |  | fail to send out software or data | application program version or upgrade problem
visualization processor |  | fail to provide video encoded data | inadequate user interface
visualization processor | component | fail to monitor power event | power conditioning issue
vital capacity | ventilator | incorrect volume calculation | parameter calculation error due to software problem
vital capacity | ventilator
voltage loss | electrical | voltage loss | electrical issue
voltage regulator, power regulator | electrical | voltage out of range | electrical issue
voltage regulator, power regulator |  | regulator failed | electrical issue
voltage regulator, power regulator |  | recommended voltage exceeded | electrical issue
volume accuracy, volume inaccuracy, volume inaccuracies | measure | low volume accuracy | volume accuracy issue
volume calculation | calculation | incorrect volume calculation | parameter calculation error due to software problem
volume control, V-limit control, volume limit control | control | volume control failed | use of device issue
volume control, V-limit control, volume limit control |  | fail to trigger volume limit control | decreased sensitivity
volume control, V-limit control, volume limit control |  | inaccurate volume limit | volume accuracy issue
volume control, V-limit control, volume limit control |  | volume out of specification | volume accuracy issue
volume control, V-limit control, volume limit control |  | fail to disable ventilation | device remains activated
volume control ventilation |  | lack of extra flow | insufficient flow or underinfusion
volume control ventilation | ventilator | inadequate flow | insufficient flow or underinfusion
volumetric carbon dioxide | ventilator | foreign material inside | Item Contaminated during Manufacturing or Shipping
warn light | safety | incorrect warning light | device alarm system issue
warning label | quality | missing warning label | device alarm system issue
wash cycle | usage | wash cycle high | device cleaning issue
waste process | environment | waste process difficult | device maintenance issue
watchdog clock, system clock, system clocks | test | clock out of tolerance | Date-related software issue
water entry pressure |  | inappropriate pressure | pressure issue
water proof, water resist, water resistant | quality | water proof failed | fluid leak
water purification system | system | system failure | device inoperable
water softener process | system | process failure | improper or incorrect procedure or method
water softener regeneration cycle | system | mistiming | imprecision
water treatment | system
wear test |  | wear test failed | failure of device to self-test
wear test |  | user uncomfortable | human-device interface issue
wear test |  | low resistancy | physical resistance
wear test | test | fail to wear | material integrity issue
wedge filter | component | wedge filter problem | Filtration Issue
patient weight, body weight |  | fail to enter weight data | use of device issue
patient weight, body weight |  | body weight value off | patient data issue
patient weight, body weight |  | patient weight missing | patient data issue
patient weight, body weight |  | null weight value | patient data issue
patient weight, body weight | pump | weight out of range | patient-device incompatibility
weld connection | mechanical | inappropriate weld connection | Solder Joint Failure
wheel lock | safety | unwanted movement | unintended movement
wheelchair tire | component
wheelchair wheel | component
wire straightener |  | out of dimension | defective component
wire straightener |  | inappropriate fit | fitting problem
wire straightener |  | difficult to transfer through | device handling issue
wireless communication |  | wireless communication out of area | wireless communication issue
wireless communication | communication | loss of wireless communication | wireless communication issue
wireless connection | communication | loss of wireless connection | wireless communication issue
wireless power | power | wireless power out of specification | Power Source Issue
wiring harness | electrical
work of breathing | ventilator | inadequate breath | improper device output
stopcock connection
adaptor connection
medication compounding adaptor connection
aseptic processing
critical surface |  | fail to identify critical surface | poor quality image
unidirectional airflow
transmissible agent
core wire, coiled wire
safety wire
tip anchor
molded tip
tip barb
shrink tube |  |  | failure to adhere or bond
pigtail catheter
conventional mode |  | inappropriate mode configuration | component incompatible; use of incorrect control settings
IUD, intrauterine contraceptive device | 3
leak compensation |  | inappropriate leak compensation setting | use of incorrect control settings
volume guarantee, V-guarantee |  | fail to adjust pressure | pressure issue
volume guarantee, V-guarantee |  | fail to disable control | device remains activated
volume guarantee, V-guarantee |  | fail to trigger volume guarantee control | decreased sensitivity
volume guarantee, V-guarantee |  | fail to deactivate | device remains activated
volume guarantee, V-guarantee |  | volume out of specification | tidal volume fluctuations
volume guarantee, V-guarantee |  | inaccurate volume | incorrect measurement
prico |  | fail to activate prico license | failure to deliver
prico |  | fail to deactivate | device remains activated
prico |  | loss of function | improper device output
prico |  | fail to maintain value in the center | improper device output; data issue
prico |  | incorrect limit setting | improper device output
inspiratory flow, I-Flow |  | inspiratory flow out of range | improper flow or infusion
inspiratory flow, I-Flow |  | fail to auto-set flow | improper flow or infusion
manual breath |  | fail to hold mandatory breath | device difficult to maintain
manual breath |  | fail to trigger breath | device inoperable
manual breath |  | impropriate pressure limit | pressure issue
manual breath |  | inappropriate inspiratory time | use of incorrect control settings
manual breath |  | long termination time | Failure to Auto Stop
pressure threshold |  | inappropriate threshold setting | pressure issue
circuit test |  | inappropriate test configuration | incorrect or inadequate test results; inappropriate or unexpected reset
circuit test |  | inappropriate flow setting | improper flow or infusion
circuit test |  | inappropriate valve setting | incorrect or inadequate test results
circuit test |  | low pressure accuracy | incorrect measurement
circuit test |  | incorrect leaking result | Incorrect or Inadequate Result
circuit test |  | lack of pressure indication | device alarm system issue
threshold test |  | inappropriate test configuration | incorrect or inadequate test results
rotary control |  | rotation failed | defective item
patient size |  | fail to specify patient size | patient data issue
patient size |  | fail to indicate patient size | patient data issue
patient size |  | fail to restore factory defaults | patient data issue
P-Insp control, Pinsp control, P-PSV control |  | inspiratory pressure out of range | pressure issue
P-High control, P-Manual control, P-Max control, Pmean control |  | pressure over upper limit | pressure issue
P-High control, P-Manual control, P-Max control, Pmean control |  | fail to update pressure | pressure issue
P-High control, P-Manual control, P-Max control, Pmean control |  | incorrect display range | incorrect display
P-High control, P-Manual control, P-Max control, Pmean control |  | pressure out of range | pressure issue
CPAP control, PSV control, FOT, HFO, forced oscillation |  | inadequate ventilation mode | failure to cycle
CPAP control, PSV control, FOT, HFO, forced oscillation |  | fail to initiate forced oscillation | failure to cycle
CPAP control, PSV control, FOT, HFO, forced oscillation |  | fail to disable forced oscillation | device remains activated
CPAP control, PSV control, FOT, HFO, forced oscillation |  | fail to export settings | failure to convert to back-up
CPAP control, PSV control, FOT, HFO, forced oscillation |  | pressure out of range | pressure issue
rate control, backup rate control |  | rate out of range | improper device output
rate control, backup rate control |  | parameter misprogrammed | use of incorrect control settings; improper device output
expiratory time, E-time, exhalation time, respiratory time |  | lack of time indication | no display or display failure
expiratory time, E-time, exhalation time, respiratory time |  | incorrect time display range | incorrect display
expiratory time, E-time, exhalation time, respiratory time |  | inadequate time measurement | incorrect measurement
expiratory time, E-time, exhalation time, respiratory time |  | low time accuracy | imprecision
expiratory time, E-time, exhalation time, respiratory time |  | expiratory time out of range | use of incorrect control settings
expiratory flow, E-Flow |  | expiratory flow out of range | use of incorrect control settings
Amp control, Ampmax |  | pressure amplitude out of range | use of incorrect control settings
O2 control, O2 flush control |  | oxygen out of specification | medical gas supply problem
backup control |  | control out of range | use of incorrect control settings
stabilization period |  | period out of range | electronic property issue
volume measurement |  | low measurement accuracy | volume accuracy issue
exhaled volume, Ve |  | volume out of specification | tidal volume fluctuations; volume accuracy issue
exhaled volume, Ve |  | incorrect display range | incorrect display
exhaled volume, Ve |  | low volume accuracy | volume accuracy issue
exhaled volume, Ve |  | incorrect volume calculation | parameter calculation error due to software problem
trigger volume |  | volume trigger failed | decreased sensitivity
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to measure | measurement system incompatibility
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to label EtCO2 text | device markings issue
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to disable measurement | device remains activated
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to raise high priority alarm | device alarm system issue
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to interface with other device | device-device incompatibility; component or accessory incompatibility
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | pressure measurement out of range | pressure issue
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | low pressure accuracy | imprecision
capnometer, capnometry device, capnometry parameter, capnometer device, CO2 parameter, EtCO2, end-tidal CO2, FiCO2, capnometer fault, capnometer connection |  | fail to setup configuration | device difficult to setup or prepare
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | inadeqaute oxygen measurement | incorrect measurement
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | inadequate oxygen display | no display or display failure
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | incorrect display range | incorrect display
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | incorrect time configuration | use of incorrect control settings; use of device issue
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | inappropriate sensitivity setting | use of incorrect control settings
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | inappropriate algorithm setting | use of incorrect control settings
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | fail to raise high priority alarm | device alarm system issue
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | fail to label SpO2 text | device markings issue
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | low oxygen accuracy | imprecision
pulse oximetry, SpO2, pulse oximeter, oximeter fault, oximeter connection |  | low measurement accuracy | imprecision; incorrect measurement
pulse rate |  | inadequate message display | no display or display failure
pulse rate |  | incorrect display range | incorrect display
pulse rate |  | low pulse rate accuracy | imprecision; incorrect measurement; electronic property issue
pulse rate |  | low measurement accuracy | imprecision; incorrect measurement
perfusion index |  | inadequate message display | no display or display failure
perfusion index |  | incorrect display range | incorrect display
trend screen, trend data |  | fail to setup timebase | use of device issue; use of incorrect control settings
trend screen, trend data |  | fail to delete data | data issue
trend screen, trend data |  | illegible data | data issue
trend screen, trend data |  | inadequate data storage | data issue
patient data screen |  | unauthorized deletion | unauthorized access to computer system
patient data screen |  | fail to enter patient data | patient data issue
tubing cassette, filter cassette |  | foreign material inside | particulates
tubing cassette, filter cassette |  | inappropriate fit | fitting problem
tubing cassette, filter cassette |  | invalid filter | component incompatible
tubing cassette, filter cassette |  | difficult to insert | Difficult to Insert
tubing cassette, filter cassette |  | difficult to remove | difficult to remove
tubing cassette, filter cassette |  | incompatible cassette | accessory incompatible
tubing cassette, filter cassette |  | cassette leakage | fluid leak
Pressure regulator |  | fail to inform regulator on problem | communication or transmission issue
Pressure regulator |  | inappropriate rated pressure | pressure issue
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | water splashed | fluid leak
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | fail to label gas name | device markings issue
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | fail to label text | device markings issue
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | fail to identify gas type | failure to sense
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | fail to disclose pressure range | instruction for use issue
gas inlet, air inlet, gas inlet port, exhaled gas inlet port |  | connector mismatched | component incompatible
electrical safety class |  | inappropriate electrical safety class | electrical issue
electrical safety class |  | lack of earth protective | electrical shorting
applied part |  | incorrect labeling | device markings issue
applied part |  | fail to solid insulation of patient | electrical shorting
applied part |  | fail to physical separate patient | electrical shorting
applied part |  | loss of electrical isolation | electrical shorting
membrane switch, membrane button |  | self-activation | self-activation or keying
membrane switch, membrane button |  | fail to mark icon | device markings issue
membrane switch, membrane button |  | fail to mark symbol | device markings issue
membrane switch, membrane button |  | lack of edge-triggered activation | difficult or delayed activation
IFU, Instruction for Use |  | confused IFU | Instruction for Use Issue
IFU, Instruction for Use |  | fail to address affirmative statement | instruction for use issue
IFU, Instruction for Use |  | fail to disclose IFU revision | instruction for use issue
IFU, Instruction for Use |  | inadequate IFU | Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
IFU, Instruction for Use |  | inadequate specification | Instruction for Use Issue
IFU, Instruction for Use |  | invalid IFU | Instruction for Use Issue
IFU, Instruction for Use |  | inadequate language | Instruction for Use Issue; inadequate user interface
IFU, Instruction for Use |  | inadequate info of pre-use checks | Instruction for Use Issue; device markings issue
IFU, Instruction for Use |  | lack of application specification | improper or incorrect procedure or method
IFU, Instruction for Use |  | lack of brief usability description | inadequate user interface
IFU, Instruction for Use |  | lack of care instructions | Instruction for Use Issue
IFU, Instruction for Use |  | lack of conditions of use | Instruction for Use Issue
IFU, Instruction for Use |  | lack of contraindications | Instruction for Use Issue
IFU, Instruction for Use |  | lack of limitations | Instruction for Use Issue
IFU, Instruction for Use |  | lack of indication | Instruction for Use Issue
IFU, Instruction for Use |  | lack of intended users | user used incorrect product for intended use
IFU, Instruction for Use |  | lack of legal manufacturer name and address | Instruction for Use Issue
IFU, Instruction for Use |  | lack of local language | Instruction for Use Issue; inadequate user interface
IFU, Instruction for Use |  | lack of physical operating principle | Instruction for Use Issue
IFU, Instruction for Use |  | lack of significant characteristics | Instruction for Use Issue
IFU, Instruction for Use |  | lack of performance characteristics | Instruction for Use Issue
IFU, Instruction for Use |  | lack of warning | Instruction for Use Issue; device alarm system issue
IFU, Instruction for Use |  | missing IFU in package | Instruction for Use Issue
safety sign |  | fail to comply with standards or regulation | nonstandard device or device component
safety sign |  | sign loss of explanation | device markings issue; instruction for use issue
safety sign |  | missing safety sign | device markings issue; instruction for use issue
Rx, Rx only, prescription use only |  | clinical setting changed | use of incorrect control settings
Rx, Rx only, prescription use only |  | inadequate labeling | device markings issue
pressure port |  | fail to label text | device markings issue
NRTL, ETL |  | lack of NRTL mark | device markings issue
NRTL, ETL |  | lack of ETL mark | device markings issue
equipotential, equipotential connection |  | fail to label equipotentiality label | device markings issue
proprietary symbol |  | fail to label proprietary symbol | device markings issue
calibration gas, inlet gas |  | inappropriate combustibility | medical gas supply problem
calibration gas, inlet gas |  | inappropriate ignition | medical gas supply problem
calibration gas, inlet gas |  | fail to compatible with Oxygen | material erosion
calibration gas, inlet gas |  | inappropriate disposal of calibration gases | instruction for use issue
ISO 80601-2-12 |  | fail to comply with ventilator safety requirement | compliance issue
ISO 80601-2-55 |  | fail to comply requirement for respiratory gas monitor | compliance issue
ISO 80601-2-61 |  | fail to comply requirement for pulse oximeter equipment | compliance issue
ISO 80601-2-87 |  | fail to comply requirement for high frequency ventilator | compliance issue
IEC 60601-1 |  | fail to comply with electrical safety requirement | compliance issue
IEC 60601-1-2 |  | fail to comply with EM safety requirement | compliance issue
IEC 60601-1-6 |  | fail to comply with usability safety requirement | compliance issue
IEC 60601-1-8 |  | fail to comply with alarm system requirement | compliance issue
IEC 62304 |  | fail to comply with software requirement | compliance issue
IEC 62366-1 |  | fail to comply with usability requirement | compliance issue
pneumatic system, pneumatic subsystem, pneumatic connection |  | lack of shut off mechanism | protective measure issue
pneumatic system, pneumatic subsystem, pneumatic connection |  | lack of turn on mechanism | device inoperable
pneumatic system, pneumatic subsystem, pneumatic connection |  | fail to control valve | improper flow or infusion
pneumatic system, pneumatic subsystem, pneumatic connection |  | fail to activate alarm | device alarm system issue
pneumatic system, pneumatic subsystem, pneumatic connection |  | fail to convert to pneumatic output | False Device Output
pneumatic system, pneumatic subsystem, pneumatic connection |  | loose connection | detachment of device component
pneumatic system, pneumatic subsystem, pneumatic connection |  | loss of suction | suction issue
pneumatic system, pneumatic subsystem, pneumatic connection |  | self-dismantling | detachment of device component
pneumatic system, pneumatic subsystem, pneumatic connection |  | fail to produce pressure | pressure issue
pneumatic system, pneumatic subsystem, pneumatic connection |  | fail to compatible with Oxygen | Incompatibility Problem
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | PCB short circuit | electrical shorting; spark
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | PCB burned | Burn of Device or Device Component; burning odor
circuit board, mainboard, main board, motherboard, printed circuit board, PC board, logic board, power board, PCBA, PCB |  | PCBA broken or damaged | break
actuation valve, HFO valve |  | fail to control valve | circuit failure
High frequency oscillation generator, HFO generator |  | frequency out of range | electronic property issue
High frequency oscillation generator, HFO generator |  | inaccurate frequency | electronic property issue
High frequency oscillation generator, HFO generator |  | hole blocked | blockage within device or device component
High frequency oscillation generator, HFO generator |  | fail to hold pressure | pressure issue
High frequency oscillation generator, HFO generator |  | inappropriate resolution | electronic property issue
mechanical horn coil |  | fail to generate pneumatic oscillation | mechanical issue
mechanical horn coil |  | fail to activate control | device inoperable
gas contact material, gas contact component |  | fail to compatible with Oxygen | material erosion
gas contact material, gas contact component |  | fail to compatible with air | material erosion
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | fail to supply range of flow | insufficient flow or underinfusion
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | fail to supply range of pressure | pressure issue
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | non-standardized connector | nonstandard device or device component
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | lack of locking mechanism | Detachment of Device or Device Component
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | fail to compatible with filter | component or accessory incompatibility
gas input system, gas input subsystem, O2 input subsystem, air input subsystem |  | fail to separate output port | protective measure issue
inlet connector, O2 inlet connector, air inlet connector |  | non-standardized connector | nonstandard device or device component
water trap |  | inadequate size | defective component
water trap |  | difficult to replace | device maintenance issue
water trap |  | fail to withstand pressure | mechanical issue; pressure issue
pneumatic connection, pneumatic material |  | incorrect chemical composition | material erosion
respiratory gas, exhalation gas |  | fail to warm gas | improper flow or infusion
respiratory gas, exhalation gas |  | inappropriate mixture of medical air and O2 | medical gas supply problem
blend ratio |  | incorrect ratio calculation | parameter calculation error due to software problem
accurate control |  | low control accuracy | imprecision
pressure waveform |  | inappropriate preset value | reset issue; pressure issue
pressure waveform |  | incorrect derived value | data issue; pressure issue
magnetic code |  | fail to read code | failure to read input signal
gas hose |  | missing color code | device markings issue
gas hose |  | inappropriate color code | device markings issue
compressed air |  | water splashed | fluid leak
heated wire |  | inadeqaute temperature resistance | temperature issue
heated wire |  | fail to compensate temperature | temperature issue
ambient valve, sub-ambient valve, passive sub-ambient valve, sub-ambient relief port, mask sub-ambient relief port |  | fail to trigger valve to reduce pressure | fail-safe design failure; protective measure issue
ambient valve, sub-ambient valve, passive sub-ambient valve, sub-ambient relief port, mask sub-ambient relief port |  | fail to trigger port to reduce pressure | fail-safe design failure; protective measure issue
ambient valve, sub-ambient valve, passive sub-ambient valve, sub-ambient relief port, mask sub-ambient relief port |  | fail to access ambient air | fail-safe design failure; protective measure issue
Safety system, Safety subsystem |  | fail to provide exit path | fail-safe design failure
Safety system, Safety subsystem |  | lack of backup valve | fail-safe design failure
Safety system, Safety subsystem |  | fail to access ambient air | fail-safe design failure
Safety system, Safety subsystem |  | fail to open valve to reduce pressure | fail-safe design failure
pressure sensor, pressure transducer, barometric pressure sensor |  | inaccurate barometric pressure measurement | incorrect measurement; pressure issue
pressure sensor, pressure transducer, barometric pressure sensor |  | inaccurate inspiratory pressure measurement | incorrect measurement; pressure issue
pressure sensor, pressure transducer, barometric pressure sensor |  | fail to identify loss of pressure gas | failure to sense
pressure sensor, pressure transducer, barometric pressure sensor |  | fail to calculate differential pressure | parameter calculation error due to software problem; pressure issue
pressure sensor, pressure transducer, barometric pressure sensor |  | fail to calculate absolute pressure | parameter calculation error due to software problem; pressure issue
pressure sensor, pressure transducer, barometric pressure sensor |  | fail to compensate flow | insufficient flow or underinfusion
pressure sensor, pressure transducer, barometric pressure sensor |  | fail to interface with other device | device-device incompatibility
pressure sensor, pressure transducer, barometric pressure sensor |  | inappropriate gas correction | incorrect measurement
pressure sensor, pressure transducer, barometric pressure sensor |  | inaccurate airway pressure measurement | incorrect measurement; pressure issue
Flow Cassette, Flow Cassette system, Flow Cassette subsystem, air Flow Cassette, oxygen Flow Cassette |  | fail to identify gas mixture | failure to sense
Flow Cassette, Flow Cassette system, Flow Cassette subsystem, air Flow Cassette, oxygen Flow Cassette |  | inappropriate gas correction | Overcorrection
Flow Cassette, Flow Cassette system, Flow Cassette subsystem, air Flow Cassette, oxygen Flow Cassette |  | fail to measure temperature | incorrect measurement
Flow Cassette, Flow Cassette system, Flow Cassette subsystem, air Flow Cassette, oxygen Flow Cassette |  | fail to measure pressure | incorrect measurement
temperature sensor |  | fail to compensate temperature | temperature issue
exhalation system, exhalation subsystem |  | fail to measure flow | insufficient flow or underinfusion; measurement system incompatibility
analog interface, digital interface |  | loss of connection | connection issue; data issue
HFO Module |  | loss of function | no device output
user interface, User control interface, GUI, graphical user interface |  | confused interface | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | design lack of intuitively implemented | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | difficult to apply device | device difficult to setup or prepare
user interface, User control interface, GUI, graphical user interface |  | fail to single GUI design | human-device interface issue
user interface, User control interface, GUI, graphical user interface |  | fail to zoom in small | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | fail to access functions | computer software issue
user interface, User control interface, GUI, graphical user interface |  | fail to access menu | computer software issue
user interface, User control interface, GUI, graphical user interface |  | inadequate message display | incorrect display
user interface, User control interface, GUI, graphical user interface |  | inappropriate interface design | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | lack of help information | no information
user interface, User control interface, GUI, graphical user interface |  | unable to freeze | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | unable to overlay | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | unable to review loop | inadequate user interface
user interface, User control interface, GUI, graphical user interface |  | unable to support color-blind | inadequate user interface
3D Model meta data, metadata |  | fail to display metadata | no display or display failure
3D Model order |  | inappropriate user | use of device issue
3D Model order |  | fail to submit request | use of device issue
3D Model order |  | fail to provide model type | use of device issue
3D Model order |  | fail to provide patient diagnosis info | use of device issue
3D Model order |  | software corrupted | programming issue; computer software issue
3D Model order |  | fail to order 3D model | use of device issue
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | inappropriate image format | data issue; image display error
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | non-dicom format image | data issue
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | incorrect dicom tag | device markings issue
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | fail to open images | image display error
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | incorrect reference line | image display error
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | incorrect image slice displayed | image display error
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | incomplete dicom header | data issue; loss of data
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | fail to display DICOM data | no display or display failure
DICOM, digital imaging and communication in medicine, 2D DICOM slice, DICOM data |  | fail to calculate slice area | incorrect software programming calculations
PACS, picture archiving and communication system
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | poor tumor size image | poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | poor tumor location image | poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | poor tumor vascularity image | poor quality image; therapeutic or diagnostic output failure
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to visualize surrounding critical structures | poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to interpret images | image resolution poor; poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | unable to render 2D views | application program issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | unable to measure point to point | application program issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | unable to render 3D views | application program issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to display patient image at OR monitor | incompatibility problem
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | unable to configure image display | application program issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | inappropriate image format | data issue; image display error
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to read images | image display error
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to manipulate visualization parameters | optical issue; use of incorrect control settings
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to visualize surrounding anatomy | poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to visualize anatomical variation | optical issue; poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to visualize aberrant anatomy | optical issue; poor quality image
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to develop detailed 3D model | optical issue; poor quality image; image display error
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to explore anatomy independently | human-device interface issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | fail to toggle image visibility | application program issue
Pre-Op patient image, pre-operative patient image, preoperative patient image, pre-operative image, pre-op image, medical imaging, advanced Imaging |  | poor medical image | poor quality image; therapeutic or diagnostic output failure
anatomy segmentation, segmentation layer, segmented anatomic region, 3D segmentation |  | fail to display anatomic infill | application program issue
data structure
Acoustic Energy, Acoustic pressure |  | sound pressure out of specification | pressure issue
Acoustic Energy, Acoustic pressure |  | annoying sound | audible noise
Acoustic Energy, Acoustic pressure |  | harmful sound | audible noise
output current, therapeutic current |  | output current out of range | therapeutic or diagnostic output failure; unexpected therapeutic results
Output Power |  | output power out of range | output energy incorrect; energy output to patient tissue incorrect
Output Voltage |  | output voltage out of range | output issue
output waveform |  | inappropriate output waveform | electronic property issue; output issue
EOS, end of service | 3 | reduced service life | shelf life exceeded; improper device output
stored pressure |  | accidental injection | fluid injection
non-ionizing radiation, electromagnetic emission, low-frequency, radiofrequency, microwave, IR, visible light, UV |  | unintended radiation exposure | Radiation overexposure
ionizing radiation, gamma, X-ray |  | unintended radiation exposure | Radiation overexposure
ionizing radiation, gamma, X-ray |  | device broken or damaged | device damaged prior to use
ionizing radiation, gamma, X-ray |  | Insufficient Visibility | human-device interface issue
ionizing radiation, gamma, X-ray |  | image dark | poor quality image
ionizing radiation, gamma, X-ray |  | unclear image | Image Resolution Poor
device state, device mode, device setting |  | unclear state | human-device interface issue
device state, device mode, device setting |  | invisible state | human-device interface issue
device state, device mode, device setting |  | unclear mode | human-device interface issue
device state, device mode, device setting |  | invisible mode | human-device interface issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | inadequate pre-op planning | improper or incorrect procedure or method
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | inadequate surgical plan | improper or incorrect procedure or method
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to create pre-op planning | application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | unable to develop surgical plan | application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | save plan to different patient | patient data issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to save surgical plan | loss of data; data back-up problem
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to refer pre-op plan intraoperatively | application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to name surgical plan | loss of data; data back-up problem
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to provide virtual clip-applying tool | human-device interface issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to use 2D images to envision a 3D model | image resolution poor; poor quality image; application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to use 2D CT images to envision a 3D model | application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to use 2D MRI images to envision a 3D model | application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to develop surgical plan independently | image resolution poor; poor quality image; application program issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | missing review of pre-operative patient images | use of device issue
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | unsafe pre-op planning | improper or incorrect procedure or method
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | inefficiency pre-op planning | improper or incorrect procedure or method
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | inefficacy pre-op planning | improper or incorrect procedure or method
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | poor medical image | poor quality image; therapeutic or diagnostic output failure
pre-op plan, pre-operative plan, surgical plan, Pre-operative 3D plan, pre-surgical plan, surgical plan, surgical procedure |  | fail to provide spatial relationship | image display error; application program issue
clinical decision |  | incorrect clinical decision | poor quality image
clinical decision |  | entering wrong place | poor quality image
clinical decision |  | moving wrong direction | poor quality image
clinical decision |  | over resection | poor quality image
clinical decision |  | less resection | poor quality image
vascular anatomy |  | poor anatomy image | optical issue; poor quality image; therapeutic or diagnostic output failure
critical structure, surrounding structure |  | fail to visualize color-coded critical structures | image display error
critical structure, surrounding structure |  | fail to visualize color-coded surrounding structures | image display error
critical structure, surrounding structure |  | fail to visualize spatial relationship | image display error
critical structure, surrounding structure |  | fail to visualize surrounding critical structures | poor quality image
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to provide virtual 3D model | application program issue; data issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to adjust structure visibility | human-device interface issue; optical issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inappropriate 3D image reconstructed | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to reconstruct 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to list 3D models | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to provide external 3D model creation | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to order 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to manage 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to rotate 3D model | application program issue; use of device issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to monitor 3D model status | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to address 3D model problems | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to download 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | load incorrect 3D model | application program issue; human-device interface issue; device handling issue; use of device issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to upload 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to manipulate 3D model | use of device issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to manage 3D model online | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to review 3D model of anatomy | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to review 3D model of pathology | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to review 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to toggle on 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to toggle off 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to select 3D model | application program issue; human-device interface issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to deselect 3D model | application program issue; human-device interface issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to use 2D images to envision a 3D model | image resolution poor; poor quality image
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to deliver 3D model | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | incorrect volume calculation | parameter calculation error due to software problem
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to scan model volume | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | unable to configure 3D model display | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to display 3D model at OR monitor | incompatibility problem
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to display volumetric sizing information | application program issue; no display or display failure
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inappropriate segnmentation alogorithm applied | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inappropriate model rendered | poor quality image
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inappropriate model orientation | image orientation incorrect
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to auto-orient model | image orientation incorrect
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inappropriate spacing of 2D slices applied | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | unclear 3D model | Image Resolution Poor; human-device interface issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | fail to manipulate visualization parameters | application program issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | inadequate 3D model data | data issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | incorrect 3D model | data issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | incorrect measurement marking | device markings issue
3D model, 3D anatomical model, anatomic model, patient model, structure, Visible Patient model container, Visible Patient |  | incomplete 3D model | image resolution poor; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | inadequate user interface | inadequate user interface
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | poor anatomy image | optical issue; poor quality image; therapeutic or diagnostic output failure
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | confused interface | inadequate user interface
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | unclear simulation | human-device interface issue; poor quality image; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | inappropriate usage environment | environmental control or utility issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | difficult to use | human-device interface issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to explore anatomy independently | human-device interface issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to apply clips to simulate resection volume | human-device interface issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to remove clips | human-device interface issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to place next clip | human-device interface issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to place virtual clip | use of device issue; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to provide real time organ volume calculation | image resolution poor; poor quality image; imprecision; application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to develop detailed 3D model | optical issue; poor quality image; image display error; application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | inappropriate color code | image resolution poor; poor quality image; use of device issue
2D image, 2D visualization, 2D image, 2D Source Image |  | image orientation incorrect | image orientation Incorrect
2D image, 2D visualization, 2D image, 2D Source Image |  | image flipped | image reversal
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to order 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to manage 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to generate 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to manage 2D images online | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to review 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to toggle on 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to toggle off 2D images | application program issue
2D image, 2D visualization, 2D image, 2D Source Image |  | poor axial image | image resolution poor; poor quality image; therapeutic or diagnostic output failure
2D image, 2D visualization, 2D image, 2D Source Image |  | poor sagittal image | image resolution poor; poor quality image; therapeutic or diagnostic output failure
2D image, 2D visualization, 2D image, 2D Source Image |  | poor coronal image | image resolution poor; poor quality image; therapeutic or diagnostic output failure
2D image, 2D visualization, 2D image, 2D Source Image |  | poor anatomical atlas image | image resolution poor; poor quality image; therapeutic or diagnostic output failure
2D image, 2D visualization, 2D image, 2D Source Image |  | fail to interpret images | application program issue
intra-operatively, intraoperatively, intraoperative |  | lack of pre-operative patient images | use of device issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to view simulation session | application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to provide simulation resection | application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to save simulation session | application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to render simulation resection | application program issue
patient simulation, patient simulation tool, Clip Apply, Clip Apply Model, virtual clip, Outcome Simulation, simulation session, simulation resection, Simulation of Resection, simulated resection |  | fail to continue the paused session | application program issue
Digital Hub |  | fail to save simulation session at digital hub | data issue
Digital Hub |  | insufficient memory size | data back-up problem
web portal |  | fail to view images through web portal | application program issue
web portal |  | fail to access through web portal | application program issue
surgeon portal |  | fail to access through web portal | application program issue
surgeon portal |  | insecure login | unauthorized access to computer system
surgeon portal |  | lack of local language | Instruction for Use Issue; inadequate user interface
surgeon portal |  | fail to order modeling service | application program issue
surgeon portal |  | fail to upload images | application program issue
surgeon portal |  | fail to initiate viewer | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to view images through web portal | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to view 3D anatomical model | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to show pre-operative images | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | unable to render 2D views | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | unable to render 3D views | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to place annotations | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to launch program | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to select image type | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to select order | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to search patient name | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to search surgeon name | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to sort patient interest list | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to sort patient list | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to sort surgeon list | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to filter patient interest list | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to compatible with radiology monitor | Incompatibility Problem
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to place measurements | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | lack of measurement tool | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | lack of video | inadequate instructions for healthcare professional
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | lack of image library | inadequate instructions for healthcare professional
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | lack of case library | inadequate instructions for healthcare professional
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | lack of backup viewer | fail-safe mechanism issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to place virtual clip | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to compatible with regular computer | Incompatibility Problem
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to verify image | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to save current view | data back-up problem; loss of data
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to replace current view | data back-up problem; loss of data
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to confirm current view | data back-up problem; loss of data
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to monitor loading images | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to provide simulation resection | application program issue
DSEP, Digital Surgery Enablement Platform |  | fail to provide external 3D model creation | application program issue
DSEP, Digital Surgery Enablement Platform |  | inadequate data storage | data issue
DSEP, Digital Surgery Enablement Platform |  | fail to manage users | application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | IEC 62304 violated | computer software issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | inappropriate medical device classification | compliance issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | inconsistent algorithm | programming issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | unresponsive click | failure to deploy; application program issue
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | inaccuracy distance measurement | incorrect measurement
Aurora viewer, digital solution viewer, DICOM viewer, image viewer, Visible Patient Planner, Visible Patient |  | fail to access across devices and locations | human-device interface issue
operating room |  | improper or incorrect use of device | use of device issue; improper or incorrect procedure or method
operating room |  | fail to follow safe practices | use of device issue; improper or incorrect procedure or method
operating room |  | incorrect patient data loaded | use of device issue; improper or incorrect procedure or method
operating room |  | ignoring orientation indicators | use of device issue; improper or incorrect procedure or method
2D image stack, image stack |  | inadequate image types | application program issue
2D image stack, image stack |  | inadequate image manipulation | human-device interface issue
2D image stack, image stack |  | unable to pan image | application program issue
2D image stack, image stack |  | unable to rotate image | application program issue
2D image stack, image stack |  | unable to zoom image | application program issue
2D image stack, image stack |  | unable to scroll image | application program issue
2D image stack, image stack |  | unable to adjust view of image | application program issue
2D image stack, image stack |  | fail to view image stack | application program issue
anatomy segmentation, segmentation layer, segmented anatomic region, 3D segmentation |  | unable to render 2D overlay on 3D segmentation | application program issue
anatomy segmentation, segmentation layer, segmented anatomic region, 3D segmentation |  | fail to display anatomic contour | application program issue
undo button |  | unresponsive click | failure to deploy; application program issue
undo button |  | fail to undo an adjustment | failure to deploy; application program issue
done button |  | fail to done an adjustment | failure to deploy; application program issue
confirm button |  | fail to confirm current view | failure to deploy; application program issue
edit button |  | fail to edit notes | human-device interface issue; application program issue
re-center button |  | unresponsive click | failure to deploy; application program issue
re-center button |  | fail to reset view | failure to reset
reset button |  | fail to reset view | failure to reset
reset button |  | fail to reset annotations | failure to reset
trash icon |  | unresponsive click | failure to deploy; application program issue
distance tool, Linear distance, annotation distance |  | fail to measure straight-line distance | application program issue; incorrect measurement
distance tool, Linear distance, annotation distance |  | fail to pick start point | human-device interface issue; application program issue
distance tool, Linear distance, annotation distance |  | fail to pick end point | human-device interface issue; application program issue
distance tool, Linear distance, annotation distance |  | fail to rename distance | device markings issue; application program issue
distance tool, Linear distance, annotation distance |  | fail to pick measurement point | human-device interface issue; application program issue
distance tool, Linear distance, annotation distance |  | incorrect distance measurement | incorrect measurement
healthy tissue, resected organ |  | fail to observe healthy tissue | human-device interface issue; application program issue
healthy tissue, resected organ |  | fail to observe resected organ | human-device interface issue; application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to pan viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to rotate 2D viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to rotate 3D viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to resize viewing panel | reset issue; failure to reset
viewing panel, view panel, 3D panel, panel view |  | unable to maximize viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to restore to default size | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to reset viewing panel | failure to reset
viewing panel, view panel, 3D panel, panel view |  | unable to hide viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to perform multiple measurements | application program issue
viewing panel, view panel, 3D panel, panel view |  | unable to select layout for viewing panel | application program issue
viewing panel, view panel, 3D panel, panel view |  | inadequate panel view | incorrect display
viewing panel, view panel, 3D panel, panel view |  | fail to display semi-transparent overlay | application program issue
viewing panel, view panel, 3D panel, panel view |  | fail to toggle image visibility | application program issue
cut plane |  | unable to advance cut plane | application program issue
cut plane |  | fail to display cut plane | no display or display failure
cut plane |  | fail to reset default settings | reset issue; failure to reset
cut plane |  | unable to rotate cut plane | application program issue
multi-planer resection, MPR |  | unable to advance MPR cut plane | application program issue
3D Model Creator |  | fail to create anatomy segmentation | application program issue; poor quality image
3D Model Creator |  | fail to create related measurement | application program issue; poor quality image
3D Model Creator |  | fail to create metadata | application program issue; poor quality image
3D Model Creator |  | fail to create 3D Model | application program issue; poor quality image
PACS, Picture Archiving and Communication System |  | fail to store medical imaging | application program issue
PACS, Picture Archiving and Communication System |  | fail to retrieve medical imaging | application program issue
PACS, Picture Archiving and Communication System |  | fail to distribute medical imaging | application program issue
PACS, Picture Archiving and Communication System |  | fail to analyze medical imaging | application program issue
PACS, Picture Archiving and Communication System |  | fail to process medical imaging | application program issue
healthcare provider |  | inadequate IFU | Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
healthcare provider |  | lack of video | Inadequate instructions for healthcare professional; Inadequate instructions for non-healthcare professional
Marketing Representative, Sales Representative |  | fail to monitor key performance | Market Issue
Marketing Representative, Sales Representative |  | fail to monitor product use | Market Issue
Marketing Representative, Sales Representative |  | fail to provide promotion strategies | Market Issue
clinical workflow |  | unclear workflow | data issue
clinical workflow |  | clinical workflow interrupted | use of device issue; disinfection or sterilization issue at user location
Hepatectomy procedure |  | Postoperative liver failure | improper or incorrect procedure or method
Hepatectomy procedure |  | incorrect procedure | improper or incorrect procedure or method
laparoscopic colorectal procedure |  | Central Nervous System Event (CNS) damaged | improper or incorrect procedure or method
laparoscopic colorectal procedure |  | incorrect procedure | improper or incorrect procedure or method
laparoscopic urologic procedure |  | Nerve Injury | improper or incorrect procedure or method
laparoscopic urologic procedure |  | Nerve damaged | improper or incorrect procedure or method
laparoscopic urologic procedure |  | Acute Bleeding | improper or incorrect procedure or method
laparoscopic urologic procedure |  | incorrect procedure | improper or incorrect procedure or method
Pulmonary Resection procedure |  | Major Vascular Injury | improper or incorrect procedure or method
Pulmonary Resection procedure |  | Bronchial Injury | improper or incorrect procedure or method
Pulmonary Resection procedure |  | Acute respiratory failure | improper or incorrect procedure or method
Pulmonary Resection procedure |  | incorrect procedure | improper or incorrect procedure or method
rigid material, semi-rigid material
clinical parameter |  | incorrect parameter | use of incorrect control settings
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | incorrect position calculation | positioning issue; computer software issue; incorrect software programming calculations
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | improper report information | application program issue
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | fail to create 2d images | application program issue
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | fail to create 3d images | application program issue
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | unable to read image | application program issue; therapeutic or diagnostic output failure
implant software, reporting software, planning software, reporting system, image review workstation software, merge |  | false output | computer software issue
health data, healthcare data, patient data, electronic medical record |  | incorrect patient information | patient data issue
health data, healthcare data, patient data, electronic medical record |  | incorrect patient record | patient data issue
hardware upgrade |  | lack of release test | false device output
pet, positron emission tomography |  | inappropriate therapy | therapeutic or diagnostic output failure
measurement data |  | measurement data corrupted | Electromagnetic Interference (EMI)
measurement data |  | inappropriate measurement data displayed | incorrect display
thumbnail tab |  | tab changed unexpectedly | application program issue
measurement mode, distance mode, interaction mode |  | fail to deactivate mode | application program issue
measurement mode, distance mode, interaction mode |  | fail to activate mode | application program issue
measurement mode, distance mode, interaction mode |  | loss of measurement | incorrect measurement
3D view |  | unable to rotate 3D view | application program issue
3D view |  | unable to render 3D view | application program issue
3D view |  | fail to display 3D view | no display or display failure
check box |  | inadequate check box | human-device interface issue
network mesh |  | fail to generate network mesh | application program issue
network mesh |  | fail to highlight color on | application program issue
network mesh |  | fail to click on the network mesh | application program issue
volume simulation |  | fail to generate volume simulation | application program issue
virtual machine |  | fail to run on virtual machine | application program issue
virtual machine |  | fail to run GPU | operating system becomes nonfunctional
virtual keyboard |  | fail to render virtual keyboard | incorrect display
virtual keyboard |  | fail to localizate virtual keyboard | incorrect display
view parameter |  | inadequate view parameters | parameter calculation error due to software
view parameter |  | inadequate 3D model parameters | parameter calculation error due to software
image system |  | fail to load image system | problem with software installation
chest wall |  | fail to create reformat | device operates differently than expected
chest wall |  | unable to move back | device operates differently than expected
calcium scoring |  | scoring overestimated | false positive result
hemo data |  | incorrect patient information | patient data issue
remote machine, service machine |  | incorrect map displayed | incorrect display
breast implant |  | silicone breast implant ruptured | material rupture
dental implant |  | fail to seat osteotomy | positioning issue
dental implant |  | unable to place implant | difficult to advance
dental implant |  | improper osteotomy position | positioning issue
dental implant |  | improper implant position | positioning issue
z-score |  | incorrect z-score value | patient data issue
orientation label |  | incorrect orientation label | device markings issue
voice clip |  | incorrect voice clip | device markings issue
offsite reading |  | reading slow | device operates differently than expected
monitor image
monitor image |  | fail to display image | no display or display failure
endoscopic image |  | halation occurred | incorrect display
call-ball |  | fail to activate alarm | use of device issue; not audible alarm

## Stop
stop words | category
capture iec | quality
operating system | software
operation position | motion
operational state | motion
operator position | human
protective measure | quality
operating room | environment
exclude | grammer | skip the rest sentence
note: | grammer | skip the rest sentence
federal communication commission | regulation
sterile adapter | equipment
endoscope sterilization tray, endoscope tray | part
touch input device | equipment
stop_phrases for below titles
introduction
purpose & scope
project name and code
general product need and description
product offering
definition, abbreviation
reference, document
responsibility

## Hazard-harm
Product_Category | Product_SubCategory | Hazard | Hazardous_Situation | Harm | Severity
CRM | Pacer | Ambient Noise Issue | Inconvenience | No harm | 1 | Stent_Graft: FDA Maude hazards + FDA hazard lists + CMS ENDO table
CRM | Pacer | Ambient Noise Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Application Program Issue | Abnormal Cardiac Pacing | Discomfort | 2
CRM | Pacer | Application Program Issue | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Application Program Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Application Program Issue | Loss of Cardiac Pacing | Prolonged Surgery | 2
CRM | Pacer | Application Program Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Issue | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Application Program Version Or Upgrade Problem | Abnormal Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Application Program Version Or Upgrade Problem | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Application Program Version Or Upgrade Problem | Abnormal Cardiac Pacing | Dizziness | 2
CRM | Pacer | Application Program Version Or Upgrade Problem | Abnormal Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Application Program Version Or Upgrade Problem | Incomplete or Misleading Diagnostic Data | Death | 5
CRM | Pacer | Application Program Version Or Upgrade Problem | Inconvenience | No harm | 1
CRM | Pacer | Application Program Version Or Upgrade Problem | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Version Or Upgrade Problem | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Artifact | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Artifact | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Artifact | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Artifact | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Battery Issue | Device Heating | Burn, Thermal | 3
CRM | Pacer | Battery Issue | Device Heating | Burn, Thermal | 2
CRM | Pacer | Battery Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Biocompatibility Issue | Biocompatibility Issue | Hypersensitivity | 3
CRM | Pacer | Biocompatibility Issue | Biocompatibility Issue | Hematoma | 2
CRM | Pacer | Biofilm Coating In Device | Surgical Complications | Seroma | 2
CRM | Pacer | Burst | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Capturing Issue | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Capturing Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Capturing Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Capturing Issue | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Circuit Failure | Abnormal Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Circuit Failure | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Communication or Transmission Issue | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Communication or Transmission Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Communication or Transmission Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Component Or Accessory Incompatibility | Inconvenience | No harm | 1
CRM | Pacer | Computer Operating System Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Computer Operating System Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Computer Operating System Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Computer Software Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Crack | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Date Related Software Issue | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Defective Component | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Defective Component | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Defective Component | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Defective Component | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Degraded | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Degraded | Unintentional Extra Cardiac Stimulation | Additional Surgical Procedure | 3
CRM | Pacer | Degraded | Unintentional Extra Cardiac Stimulation | Burn, Thermal | 2
CRM | Pacer | Delayed Charge Time | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Delayed Charge Time | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Device Abrasion From Instrument Or Another Object | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Device Alarm System Issue | Inconvenience | No harm | 1
CRM | Pacer | Device Alarm System Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Device Alarm System Issue | Loss of Patient Notification | Additional Surgical Procedure | 3
CRM | Pacer | Device Damaged Prior To Use | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Device Damaged Prior To Use | Abnormal Defibrillation Therapy | Prolonged Surgery | 1
CRM | Pacer | Device Damaged Prior To Use | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Difficult To Maintain | Inconvenience | No harm | 1
CRM | Pacer | Device Difficult To Setup Or Prepare | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Device Difficult To Setup Or Prepare | Confusion | Prolonged Surgery | 1
CRM | Pacer | Device Difficult To Setup Or Prepare | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Device Maintenance Issue | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Device Or Device Component Damaged By Another Device | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Device Or Device Component Damaged By Another Device | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Device Or Device Component Damaged By Another Device | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Or Device Component Damaged By Another Device | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Device Sensing Issue | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Sensing Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Device Sensing Issue | Inconvenience | No harm | 1
CRM | Pacer | Device Sensing Issue | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Device Sensing Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Difficult To Advance | Surgical Complications | Embolism | 3
CRM | Pacer | Difficult To Deploy | Confusion | Additional Surgical Procedure | 3
CRM | Pacer | Difficult To Insert | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Difficult To Insert | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Disinfection Or Sterilization Issue At User Location | Lack of Sterility | Infection | 3
CRM | Pacer | Dislodged or Dislocated | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Dislodged or Dislocated | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Dislodged or Dislocated | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Dislodged or Dislocated | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electrical Shorting | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electrical Shorting | Inconvenience | No harm | 1
CRM | Pacer | Electrical Shorting | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Electrical Shorting | Loss of Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Electrical Shorting | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electrical Shorting | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Electrical Shorting | Loss of Defibrillation Therapy | Additional Surgical Procedure | 4
CRM | Pacer | Electro Magnetic Interference | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Electro Magnetic Interference (EMI) | Inconvenience | No harm | 1
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Electro Magnetic Interference (EMI) | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Electro Static Discharge | Abnormal Cardiac Pacing | No harm | 1
CRM | Pacer | Electro Static Discharge | Abnormal Defibrillation Therapy | No harm | 1
CRM | Pacer | Electronic Property Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Energy Output To Patient Tissue Incorrect | Device Heating | Burn, Thermal | 2
CRM | Pacer | Energy Output To Patient Tissue Incorrect | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Entrapment Of Device Or Device Component | Surgical Complications | Hemorrhage Serious | 3
CRM | Pacer | Entrapment Of Device Or Device Component | Surgical Complications | Pneumothorax | 2
CRM | Pacer | Environmental Control Or Utility Issue | Abnormal Cardiac Pacing | No harm | 1
CRM | Pacer | Failure To Align | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Failure To Align | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Analyze Rhythm | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Failure To Analyze Rhythm | Inconvenience | No harm | 1
CRM | Pacer | Failure To Analyze Rhythm | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Auto Stop | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Failure To Capture | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Failure To Capture | Abnormal Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Failure To Capture | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Capture | Inconvenience | No harm | 1
CRM | Pacer | Failure To Capture | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Discharge | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Interrogate | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Failure To Interrogate | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Failure To Interrogate | Inaccessible Diagnostic Data | Intervention or Procedure | 2
CRM | Pacer | Failure To Interrogate | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Failure To Interrogate | Inconvenience | No harm | 1
CRM | Pacer | Failure To Pace Or Properly Pace | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Failure To Pace Or Properly Pace | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Pace Or Properly Pace | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Pace Or Properly Pace | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Failure To Pace Or Properly Pace | Inconvenience | Prolonged Surgery | 1
CRM | Pacer | Failure To Sense | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Sense | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Service | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Shock Or Properly Shock | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Failure To Spike | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Spike | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Sterilize | Lack of Sterility | Infection | 5
CRM | Pacer | Failure To Sterilize | Lack of Sterility | Infection | 3
CRM | Pacer | False Positive Result | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | False Positive Result | Inconvenience | No harm | 1
CRM | Pacer | False Reading From Device Non Compliance | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | False Reading From Device Non Compliance | Inconvenience | No harm | 1
CRM | Pacer | Fluid Leak | Extravascular Foreign Body | Additional Surgical Procedure | 4
CRM | Pacer | Fluid Leak | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Fracture | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Fracture | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | High Capture Threshold | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | High Capture Threshold | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | High Impedance | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | High Sensing Threshold | Inconvenience | No harm | 1
CRM | Pacer | Human Device Interface Issue | Inconvenience | No harm | 1
CRM | Pacer | Human Factors Issue | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Human Factors Issue | Application of Excessive Force | Dissection | 3
CRM | Pacer | Human Factors Issue | Inconvenience | No harm | 1
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Human Factors Issue | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Human Factors Issue | Surgical Complications | Cardiac Perforation | 4
CRM | Pacer | Human Factors Issue | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Image Display Error | Inconvenience | No harm | 1
CRM | Pacer | Impedance Issue | Inconvenience | No harm | 1
CRM | Pacer | Improper Alarm | Inconvenience | No harm | 1
CRM | Pacer | Improper Or Incorrect Procedure Or Method | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Improper Or Incorrect Procedure Or Method | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Improper Or Incorrect Procedure Or Method | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Improper Or Incorrect Procedure Or Method | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Inaccurate Synchronization | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Inaccurate Synchronization | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Inaccurate Synchronization | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Inaccurate Synchronization | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Inadequate Instructions For Healthcare Professional | Inappropriate Programmed Settings | Prolonged Surgery | 1
CRM | Pacer | Inadequate Service | Inconvenience | No harm | 1
CRM | Pacer | Inadequate Service | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Inadequate Service | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Inadequate Service | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Inadequate Service | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | Inadequate Service | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Inadequate Training | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Inadequate Training | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Inadequate Training | Inconvenience | No harm | 1
CRM | Pacer | Inappropriate Shock | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Interpretation Of Rhythm | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Incorrect Interpretation Of Rhythm | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Incorrect Interpretation Of Rhythm | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Incorrect Interpretation Of Rhythm | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Incorrect Measurement | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Measurement | Incomplete | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Measurement | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Measurement | Incomplete or Misleading Diagnostic Data | Intervention or Procedure | 1
CRM | Pacer | Incorrect Measurement | Inconvenience | No harm | 1
CRM | Pacer | Incorrect Measurement | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Measurement | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Incorrect Software Programming Calculations | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Software Programming Calculations | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Software Programming Calculations | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Intermittent Continuity | Inconvenience | No harm | 1
CRM | Pacer | Invalid Sensing | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Invalid Sensing | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Invalid Sensing | Inconvenience | No harm | 1
CRM | Pacer | Invalid Sensing | Loss of Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Loose Or Intermittent Connection | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Loss Of Data | Inconvenience | No harm | 1
CRM | Pacer | Low Audible Alarm | Inconvenience | No harm | 1
CRM | Pacer | Low Battery | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Low Battery | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations | Abnormal Misuse | Death | 5
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations | Device Tissue Pressure | Additional Surgical Procedure | 3
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Manufacturing Or Shipping Issue Associated With Device | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Manufacturing Or Shipping Issue Associated With Device | Device Tissue Pressure | Pocket Erosion | 3
CRM | Pacer | Material Deformation | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Material Separation | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Material Separation | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Material Separation | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Material Separation | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Misconnection | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Misconnection | Device Heating | Burn, Thermal | 2
CRM | Pacer | Misconnection | Inconvenience | No harm | 1
CRM | Pacer | Misconnection | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Misconnection | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Missing Test Results | Inconvenience | No harm | 1
CRM | Pacer | MRI Environment Compatibility Issue | Abnormal Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | MRI Environment Compatibility Issue | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue | Abnormal Cardiac Pacing | Heart Failure | 3
CRM | Pacer | MRI Environment Compatibility Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | MRI Environment Compatibility Issue | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | MRI Environment Compatibility Issue | Device Heating | Burn, Thermal | 2
CRM | Pacer | MRI Environment Compatibility Issue | Device Tissue Pressure | Injury | 2
CRM | Pacer | MRI Environment Compatibility Issue | Device Tissue Pressure | No harm | 1
CRM | Pacer | MRI Environment Compatibility Issue | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | MRI Environment Compatibility Issue | Incomplete or Misleading Diagnostic Data | Intervention or Procedure | 2
CRM | Pacer | MRI Environment Compatibility Issue | Inconvenience | No harm | 1
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | MRI Environment Compatibility Issue | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | MRI Environment Compatibility Issue | Unintentional Extra Cardiac Stimulation | Discomfort | 1
CRM | Pacer | No Pacing | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Overheating Of Device Or Device Component | Device Heating | Burn, Thermal | 3
CRM | Pacer | Overheating Of Device Or Device Component | Device Heating | Burn, Thermal | 2
CRM | Pacer | Oversensing | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Oversensing | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Oversensing | Inappropriate Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Oversensing | Inconvenience | No harm | 1
CRM | Pacer | Oversensing | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Oversensing | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Oversensing | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Pacer Found In Back Up Mode | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Pacer Found In Back Up Mode | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Pacer Found In Back Up Mode | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacer Found In Back Up Mode | Inconvenience | No harm | 1
CRM | Pacer | Pacer Found In Back Up Mode | Unintentional Extra Cardiac Stimulation | Additional Surgical Procedure | 3
CRM | Pacer | Pacer Found In Back Up Mode | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Pacing Asynchronously | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacing Inadequately | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Pacing Inadequately | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacing Inadequately | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Pacing Inadequately | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Parameter Calculation Error Due To Software Problem | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Parameter Calculation Error Due To Software Problem | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Physical Property Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Poor Quality Image | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Premature Discharge of Battery | Inappropriate Programmed Settings | Intervention or Procedure | 2
CRM | Pacer | Premature Discharge of Battery | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Premature Discharge of Battery | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Premature Discharge of Battery | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Premature Discharge of Battery | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Premature Discharge of Battery | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | Problem With Software Installation | Inconvenience | No harm | 1
CRM | Pacer | Product Quality Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Product Quality Issue | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Product Quality Issue | Device Heating | Burn, Thermal | 2
CRM | Pacer | Product Quality Issue | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Product Quality Issue | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Product Quality Issue | Loss of Patient Notification | Intervention or Procedure | 2
CRM | Pacer | Product Quality Issue | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Product Quality Issue | Unintentional Extra Cardiac Stimulation | Discomfort | 1
CRM | Pacer | Programming Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Programming Issue | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Programming Issue | Confusion | Arrhythmia | 2
CRM | Pacer | Programming Issue | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Programming Issue | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Programming Issue | Inconvenience | No harm | 1
CRM | Pacer | Programming Issue | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Radiation Overexposure | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Radio Signal Problem | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Radio Signal Problem | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Radio Signal Problem | Inappropriate Programmed Settings | Arrhythmia | 2
CRM | Pacer | Radiofrequency Interference (RFI) | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Sensing Intermittently | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Sensing Intermittently | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Shielding Failure | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Shielding Failure | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Shielding Failure | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Therapeutic Or Diagnostic Output Failure | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Unauthorized Access To Computer System | Abnormal Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unauthorized Access To Computer System | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Unauthorized Access To Computer System | Inconvenience | No harm | 1
CRM | Pacer | Unauthorized Access To Computer System | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unauthorized Access To Computer System | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Unauthorized Access To Computer System | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Undersensing | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Undersensing | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Undersensing | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Undersensing | Inconvenience | No harm | 1
CRM | Pacer | Undersensing | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Undersensing | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Unexpected Therapeutic Results | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Unstable Capture Threshold | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Unstable Capture Threshold | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unstable Capture Threshold | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Unstable Capture Threshold | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Device Issue | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Device Issue | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Use Of Device Issue | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Use Of Device Issue | Inconvenience | No harm | 1
CRM | Pacer | Use Of Device Issue | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Use Of Device Issue | Surgical Complications | Hematoma | 2
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Cardiac Pacing | Discomfort | 2
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Use Of Incorrect Control Settings | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Use Of Incorrect Control Settings | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Use Of Incorrect Control Settings | Incomplete or Misleading Diagnostic Data | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings | Inconvenience | No harm | 1
CRM | Pacer | Use Of Incorrect Control Settings | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Use Of Incorrect Control Settings | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Incorrect Control Settings | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Use Of Incorrect Control Settings | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Wireless Communication Issue | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Wireless Communication Issue | Inconvenience | No harm | 1
Therapy | Brace | Accessory Incompatible | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Accessory Incompatible | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Application Program Version or Upgrade Problem | Abnormal Treatment | Injury to Patient | 2
Therapy | Brace | Application Program Version or Upgrade Problem | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Application Security Issue | Misuse | Injury to Patient | 2
Therapy | Brace | Application Security Issue | Product Malfunction | Injury to Patient | 2
Therapy | Brace | Battery Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Battery Issue | Loss of Data | Delay of Treatment | 1
Therapy | Brace | Bent | Unintend Motion | Loss of Treatment | 2
Therapy | Brace | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
Therapy | Brace | Blockage within Device or Device Component | Inadequate Support | Discomfort | 2
Therapy | Brace | Break | Loss of Data | Delay of Treatment | 1
Therapy | Brace | Break | Loss of Mechanical Integrity | Injury to Patient | 2
Therapy | Brace | Calibration Issue | Abnormal Treatment | Injury to Patient | 2
Therapy | Brace | Compliance Issue | Loss of Compliance | No Harm | 1
Therapy | Brace | Component Incompatible | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Component Missing | Inadequate Support | Loss of Treatment | 2
Therapy | Brace | Connection Issue | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Defective Component | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Defective Component | Material Degraded | Injury to Patient | 2
Therapy | Brace | Defective Component | Unintend Motion | Injury to Patient | 2
Therapy | Brace | Detachment of Device Component | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Detachment of Device Component | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Detachment of Device or Device Component | Unintend Motion | Injury to Patient | 2
Therapy | Brace | Device Alarm System Issue | Inconvenience | No Harm | 1
Therapy | Brace | Device Alarm System Issue | Loss of Notification | Injury to Patient | 2
Therapy | Brace | Device Cleaning Issue | Abrasion to User | Skin Irriation | 2
Therapy | Brace | Device Cleaning Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Device Cleaning Issue | Loss of Structural Integrity | Discomfort | 2
Therapy | Brace | Device Cleaning Issue | Material Degraded | Injury to Patient | 2
Therapy | Brace | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Device Contamination with Biological Material | Skin Allergy or Irritation | Skin Irriation | 2
Therapy | Brace | Device Damaged Prior to Use | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Device Difficult to Setup or Prepare | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
Therapy | Brace | Device Expiration Issue | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Device Handling Issue | Loss of Material Integrity | Injury to Patient | 2
Therapy | Brace | Device Maintenance Issue | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Device Maintenance Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Device Markings Issue | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Device Markings Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
Therapy | Brace | Device Markings Issue | Inadquate IFU or Labeling | Injury to Patient | 2
Therapy | Brace | Device Markings Issue | Misuse | Delay of Treatment | 1
Therapy | Brace | Device Markings Issue | Misuse | Injury to Patient | 2
Therapy | Brace | Device Markings Issue | User Confused | Delay of Treatment | 1
Therapy | Brace | Device Markings Issue | User Confused | No Harm | 1
Therapy | Brace | Device Operates Differently than Expected | Abnormal Treatment | Injury to Patient | 2
Therapy | Brace | Device Operates Differently than Expected | Inappropriate Treatment | Injury to Patient | 2
Therapy | Brace | Device Operates Differently than Expected | Misuse | Injury to Patient | 2
Therapy | Brace | Device Operates Differently than Expected | Product Malfunction | Injury to Patient | 2
Therapy | Brace | Device Packaging Compromised | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Device Packaging Compromised | Inadequate Support | Loss of Treatment | 2
Therapy | Brace | Device Packaging Compromised | Loss of Material Integrity | Injury to Patient | 2
Therapy | Brace | Device Packaging Compromised | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Difficult to Insert | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Difficult to Position | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Difficult to Remove | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Electrical Power Problem | Loss of Data | Delay of Treatment | 1
Therapy | Brace | Environmental Control or Utility Issue | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Environmental Control or Utility Issue | Product Malfunction | Injury to Patient | 2
Therapy | Brace | Fail-Safe Mechanism Issue | Inappropriate Treatment | Injury to Patient | 2
Therapy | Brace | Failure to adhere or bond | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Fitting Problem | Improper Fit or adjustment | Discomfort | 2
Therapy | Brace | Fitting Problem | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Fitting Problem | Improper Fit or adjustment | No harm | 1
Therapy | Brace | Human Factors Issue | Abnormal Treatment | Tissue Necrosis | 3
Therapy | Brace | Human-Device Interface Issue | Abrasion to User | Bleed | 2
Therapy | Brace | Human-Device Interface Issue | Brace Applied Inappropriately | Loss of Treatment | 2
Therapy | Brace | Human-Device Interface Issue | Fail to Operate | Delay of Treatment | 1
Therapy | Brace | Human-Device Interface Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Human-Device Interface Issue | Inadequate Support | Discomfort | 2
Therapy | Brace | Human-Device Interface Issue | Inconvenience | Delay of Treatment | 1
Therapy | Brace | Human-Device Interface Issue | Inconvenience | No Harm | 1
Therapy | Brace | Human-Device Interface Issue | Product Malfunction | Injury to Patient | 2
Therapy | Brace | Image Display Error | User Confused | No Harm | 1
Therapy | Brace | Imprecision | Unintend Motion | Injury to Patient | 2
Therapy | Brace | Improper or Incorrect Procedure or Method | Inadequate IFU or Labeling | Injury to Patient | 2
Therapy | Brace | Inadequate instructions for healthcare professional | Inadequate IFU or Labeling | Delay of Treatment | 1
Therapy | Brace | Inadequate instructions for healthcare professional | Loss of Notification | Tissue Necrosis | 3
Therapy | Brace | Inadequate instructions for healthcare professional | Misuse | No Harm | 1
Therapy | Brace | Inadequate instructions for non-healthcare professional | Inadequate IFU or Labeling | Delay of Treatment | 1
Therapy | Brace | Inadequate instructions for non-healthcare professional | Inadequate IFU or Labeling | Injury to Patient | 2
Therapy | Brace | Inadequate instructions for non-healthcare professional | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Inadequate instructions for non-healthcare professional | Misuse | Burn, Thermal | 2
Therapy | Brace | Inadequate Training | Misuse | Injury to Patient | 2
Therapy | Brace | Inadequate User Interface | Misuse | Delay of Treatment | 1
Therapy | Brace | Inadequate User Interface | User Confused | No Harm | 1
Therapy | Brace | Inappropriate Prompts | User Confused | No Harm | 1
Therapy | Brace | Incorrect Display | User Confused | No Harm | 1
Therapy | Brace | Incorrect Measurement | Abnormal Treatment | Injury to Patient | 2
Therapy | Brace | Incorrect Measurement | Inappropriate Treatment | Injury to Patient | 2
Therapy | Brace | Incorrect Measurement | User Confused | No Harm | 1
Therapy | Brace | Incorrect or Inadequate Result | Loss of Mechanical Integrity | Loss of Treatment | 2
Therapy | Brace | Instruction for Use Issue | Brace Applied Inappropriately | Loss of Treatment | 2
Therapy | Brace | Instruction for Use Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
Therapy | Brace | Instruction for Use Issue | Inadequate IFU or Labeling | Injury to Patient | 2
Therapy | Brace | Instruction for Use Issue | Inadequate IFU or Labeling | No Harm | 1
Therapy | Brace | Instruction for Use Issue | Loss of Notification | Injury to Patient | 2
Therapy | Brace | Instruction for Use Issue | Loss of Notification | Swelling | 2
Therapy | Brace | Loss of Data | Inconvenience | No Harm | 1
Therapy | Brace | Low Battery | Loss of Data | Delay of Treatment | 1
Therapy | Brace | Market Issue | Inconvenience | No Harm | 1
Therapy | Brace | Material Integrity Issue | Abrasion to User | Discomfort | 2
Therapy | Brace | Material Integrity Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Material Integrity Issue | Loss of Material Integrity | Injury to Patient | 2
Therapy | Brace | Material Integrity Issue | Material Degraded | Injury to Patient | 2
Therapy | Brace | Material Integrity Issue | Material Degraded | Loss of Treatment | 2
Therapy | Brace | Material Rigid or Stiff | Brace Applied Inappropriately | Loss of Treatment | 2
Therapy | Brace | Material Rigid or Stiff | Loss of Mechanical Integrity | Loss of Treatment | 2
Therapy | Brace | Mechanical Issue | Inadequate Support | Delay of Treatment | 1
Therapy | Brace | Mechanical Issue | Loss of Mechanical Integrity | Loss of Treatment | 2
Therapy | Brace | Mechanical Issue | Loss of Mechanical Integrity | No Harm | 1
Therapy | Brace | Migration of Device or Device Component | Unintend Motion | Injury to Patient | 2
Therapy | Brace | No Display or Display Failure | Inconvenience | Delay of Treatment | 1
Therapy | Brace | No Display or Display Failure | Inconvenience | No Harm | 1
Therapy | Brace | No Display or Display Failure | User Confused | Delay of Treatment | 1
Therapy | Brace | Patient-Device Incompatibility | Improper Fit or adjustment | Discomfort | 2
Therapy | Brace | Patient-Device Incompatibility | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Patient-Device Incompatibility | Patient immobilizated | Restricted Blood Circulation | 2
Therapy | Brace | Physical Resistance | Inconvenience | No Harm | 1
Therapy | Brace | Positioning Issue | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Problem with Software Installation | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Protective Measure Issue | Misuse | Injury to Patient | 2
Therapy | Brace | regulatory Issue | Loss of Compliance | No Harm | 1
Therapy | Brace | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
Therapy | Brace | Size Incorrect for patient | User Confused | No Harm | 1
Therapy | Brace | Split | Loss of Material Integrity | Loss of Treatment | 2
Therapy | Brace | support failed | Inadequate Support | Delay of Treatment | 1
Therapy | Brace | Temperature Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Torn Material | Loss of Material Integrity | Loss of Treatment | 2
Therapy | Brace | Unable to Access | Loss of Data | Delay of Treatment | 1
Therapy | Brace | Unauthorized Access to Computer System | Misuse | Injury to Patient | 2
Therapy | Brace | Unintended Movement | Improper Fit or adjustment | Loss of Treatment | 2
Therapy | Brace | Unintended Movement | Unintend Motion | Injury to Patient | 2
Therapy | Brace | Unintended System Motion | Unintended Motion | Injury to Patient | 2
Therapy | Brace | Unstable | Loss of Structural Integrity | Injury to Patient | 2
Therapy | Brace | Unstable | Loss of Structural Integrity | Loss of Treatment | 2
Therapy | Brace | Unstable | Patient mobilizated | Loss of Treatment | 2
Therapy | Brace | Use of Device Issue | Bacteria Growth | Infection | 3
Therapy | Brace | Use of Device Issue | Brace Applied Inappropriately | Injury to Patient | 2
Therapy | Brace | Use of Device Issue | Brace Applied Inappropriately | Loss of Treatment | 2
Therapy | Brace | Use of Device Issue | Device Reused | Skin Irriation | 2
Therapy | Brace | Use of Device Issue | Fail to Operate | Loss of Treatment | 2
Therapy | Brace | Use of Device Issue | Improper Fit or adjustment | Discomfort | 2
Therapy | Brace | Use of Device Issue | Inadequate Support | Delay of Treatment | 1
Therapy | Brace | Use of Device Issue | Inadequate Support | Discomfort | 2
Therapy | Brace | Use of Device Issue | Inadequate Support | Loss of Treatment | 2
Therapy | Brace | Use of Device Issue | Inconvenience | No Harm | 1
Therapy | Brace | Use of Device Issue | Misuse | Injury to Patient | 2
Therapy | Brace | Use of Device Issue | Misuse | Loss of Treatment | 2
Therapy | Brace | Use of Device Issue | Misuse | Skin Irriation | 2
Therapy | Brace | Use of Device Issue | Skin Allergy or Irritation | Swelling | 2
Therapy | Brace | Use of Device Issue | User Confused | No Harm | 1
Therapy | Brace | Use of Incorrect Control Settings | Abnormal Treatment | Injury to Patient | 2
Therapy | Brace | Use of Incorrect Control Settings | Loss of Product Safety | Injury to Patient | 2
Therapy | Brace | User used incorrect product for intended use | Abnormal Treatment | Tissue Necrosis | 3
Therapy | Brace | Wireless Communication Issue | Loss of Data | Delay of Treatment | 1
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Surgical conversion | 4
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Abnormal blood loss | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Access site complications | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Additional EV procedure | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Cardiac arrhythmia | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Convert to femoral cut-down | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Dissection (aorta) | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Endoleak Type IA | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Endoleak Type IB | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | Prolonged procedure | 3
intravascular | Stent Graft | accessory incompatible; component or accessory incompatibility; component incompatible | Detachment of Components (Accessories) | No impact to patient | 1
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Infection (device related) | 4
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Neurologic impairment | 4
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Renal failure | 4
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Cardiac arrhythmia | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Edema | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Embolism | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Hemodynamic instability | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Prolonged procedure | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Respiratory complications | 3
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Biocompatibility Issue | Rash | 2
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Compromised Sterility | Infection (device related) | 4
intravascular | Stent Graft | Biocompatibility Issue; microbial contamination of device; Contamination During Use; Contamination of Device Ingredient or Reagent; Device Contamination with Biological Material; Bacterial Contamination of Device; Residue after Decontamination; Item Contaminated during Manufacturing or Shipping; contamination; device contaminated during manufacture or shipping; device contamination with chemical or other material; delivered as unsterile product; Corrosion; corroded; Coagulation in Device or Device Ingredient; calcified; Cross Reactivity; Improper Chemical Reaction; Particulates; Metal Shedding Debris | Cytoxic | Infection (device related) | 4
intravascular | Stent Graft | burst container or vessel; burst | Vessel or container rupture | Rupture (aorta) | 5
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Rupture (aorta) | 5
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Perforation (aorta) | 4
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Surgical conversion | 4
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Abnormal blood loss | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Access site complications | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Additional EV procedure | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Dissection (aorta) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Edema | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Embolism | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type IA | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type IB | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Hemodynamic instability | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | Prolonged procedure | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Kinked Component | No impact to patient | 1
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Surgical conversion | 4
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Additional EV procedure | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Endoleak Type IB | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Prolonged procedure | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Prolonged procedure | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Coiled; Kinked; Knotted; Looping; Material Twisted; Uncoiled; Material Distortion; Material Torqued | Surepass Wire Issues | No impact to patient | 1
intravascular | Stent Graft | component missing; misassembled | Loss/Deterioration of function | Prolonged procedure | 3
intravascular | Stent Graft | component missing; misassembled | Loss/Deterioration of function | Additional EV procedure | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Surgical conversion | 4
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Access site complications | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Embolism | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Occlusion (iliac limb) | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | Prolonged procedure | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Deflation Issue | No impact to patient | 1
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Rupture | Surgical conversion | 4
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Rupture | Additional EV procedure | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Rupture | Convert to femoral cut-down | 3
intravascular | Stent Graft | deflation problem; deflation Issue | Balloon Rupture | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak Type II | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Endoleak type II | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | Prolonged Procedure | 3
intravascular | Stent Graft | Delivery System Failure | Aspiration Issue | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Deflation Issue | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Balloon Inflation Issue | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Balloon Inflation Issue | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Inflation Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Inflation Issue | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Rupture | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Balloon Rupture | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Rupture | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Balloon Rupture | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Broken Component | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Hemodynamic instability | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Broken Component | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Rupture (aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Perforation (aorta) | 4
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Occlusion (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Component Attachment | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Hemodynamic instability | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Damaged Component | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Rupture (aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Perforation (aorta) | 4
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment Issue | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Cardiac arrhythmia | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Accessories) | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Rupture (aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Perforation (aorta) | 4
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Renal insufficiency | 4
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Hemodynamic instability | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Occlusion (aortic) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Detachment of Components (Catheter) | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Flaked | Infection (device related) | 4
intravascular | Stent Graft | Delivery System Failure | Flaked | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Flaked | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Occlusion (mesenteric) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Occlusion (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Stenosis (non-specific) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | Stenosis (renal) | 3
intravascular | Stent Graft | Delivery System Failure | Flaked | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Neurologic impairment | 4
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Cardiac arrhythmia | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Respiratory complications | 3
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | Rash | 2
intravascular | Stent Graft | Delivery System Failure | Fluid Leak | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Rupture (aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Perforation (aorta) | 4
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Abnormal blood loss | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Access site complications | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Dissection (aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Edema | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type IA | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Hemodynamic instability | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Kinked Component | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Obstruction | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Obstruction | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | Convert to femoral cut-down | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | Embolism | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Obstruction | No impact to patient | 1
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Surgical conversion | 4
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Additional EV procedure | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Endoleak Type IB | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Prolonged procedure | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Delivery System Failure | Surepass Wire Issues | No impact to patient | 1
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Renal failure | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Renal insufficiency | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Surgical conversion | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Additional EV procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Dissection (aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Embolism | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Occlusion (aortic) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | Prolonged procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Deployment Issue | No impact to patient | 1
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Rupture (aorta) | 5
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Perforation (aorta) | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Renal failure | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Renal insufficiency | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Surgical conversion | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Abnormal blood loss | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Access site complications | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Additional EV procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Convert to femoral cut-down | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Dissection (aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Embolism | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type IA | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type IB | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Hemodynamic instability | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Occlusion (aortic) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Occlusion (non-specific) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Prolonged procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Prolonged procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Respiratory complications | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (non-specific) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | Stenosis (renal) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Difficult to Deploy | No impact to patient | 1
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Rupture (aorta) | 5
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Surgical conversion | 4
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Abnormal blood loss | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Additional EV procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Dissection (aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Edema | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Embolism | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Occlusion (non-specific) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Prolonged procedure | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Stenosis (aorta) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Stenosis (non-specific) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Deployment Issue; Failure to Deploy; activation failure; premature activation; difficult or delayed activation; activation problem; self-activation or keying; Premature Deployment | Unable to Unsheath | No impact to patient | 1
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Surgical conversion | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Abnormal blood loss | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Access site complications | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Additional EV procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Cardiac arrhythmia | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Convert to femoral cut-down | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Dissection (aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Endoleak Type IA | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Endoleak Type IB | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | Prolonged procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Accessories) | No impact to patient | 1
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Rupture (aorta) | 5
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Perforation (aorta) | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Renal insufficiency | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Surgical conversion | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Abnormal blood loss | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Access site complications | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Additional EV procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Convert to femoral cut-down | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Dissection (aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Embolism | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type IA | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type IB | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Hemodynamic instability | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Occlusion (aortic) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Occlusion (non-specific) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | Prolonged procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Catheter) | No impact to patient | 1
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Rupture (aorta) | 5
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Perforation (aorta) | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Renal failure | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Renal insufficiency | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Surgical conversion | 4
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Abnormal blood loss | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Access site complications | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Additional EV procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Convert to femoral cut-down | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Dissection (aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Embolism | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type IA | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type IB | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Hemodynamic instability | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Occlusion (aortic) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Occlusion (mesenteric) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Occlusion (non-specific) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Prolonged procedure | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Stenosis (aorta) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Stenosis (mesenteric) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Stenosis (non-specific) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | detachment of device component; detachment of device or device component | Detachment of Components (Implant) | No impact to patient | 1
intravascular | Stent Graft | device difficult to setup or prepare | Patient-Device Incompatibility | Additional EV procedure | 3
intravascular | Stent Graft | device difficult to setup or prepare | Patient-Device Incompatibility | Prolonged procedure | 3
intravascular | Stent Graft | Device Markings Issue; device markings/labelling problem | User Confused | Prolonged procedure | 3
intravascular | Stent Graft | Device Markings Issue; device markings/labelling problem | User Confused | No impact to patient | 1
intravascular | Stent Graft | device operates differently than expected; Device Inoperable; defective device; device damaged by another device; unintended collision; device damaged prior to use; improper device output; unexpected therapeutic results; incorrect or inadequate test results; Incorrect Device or Component Shipped | Loss/Deterioration of function | Prolonged procedure | 3
intravascular | Stent Graft | device operates differently than expected; Device Inoperable; defective device; device damaged by another device; unintended collision; device damaged prior to use; improper device output; unexpected therapeutic results; incorrect or inadequate test results; Incorrect Device or Component Shipped | Loss/Deterioration of function | Additional EV procedure | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Surgical conversion | 4
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Abnormal blood loss | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Access site complications | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Additional EV procedure | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Convert to femoral cut-down | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | Prolonged procedure | 3
intravascular | Stent Graft | difficult or delayed separation; separation failure; Failure to Separate | Difficult to Separate (Demate) | No impact to patient | 1
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Perforation (aorta) | 4
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Surgical conversion | 4
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Abnormal blood loss | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Access site complications | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Additional EV procedure | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Convert to femoral cut-down | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Dissection (aorta) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Embolism | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type IA | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type IB | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Hemodynamic instability | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Occlusion (aortic) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Occlusion (non-specific) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | Prolonged procedure | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | No impact to patient | 1
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Difficult to Advance | No patient involvement | 1
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Perforation (aorta) | 4
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Surgical conversion | 4
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Abnormal blood loss | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Access site complications | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Additional EV procedure | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Convert to femoral cut-down | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Dissection (aorta) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Hemodynamic instability | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | Prolonged procedure | 3
intravascular | Stent Graft | Difficult to Advance; Failure to Advance; physical resistance | Unable to Advance | No impact to patient | 1
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Surgical conversion | 4
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Abnormal blood loss | 3
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Additional EV procedure | 3
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | Prolonged procedure | 3
intravascular | Stent Graft | Difficult to Fold or unfold; Failure to Unfold or Unwrap; failure to unwrap; difficult to fold unfold or collapse | Failure to Unfold (Graft) | No impact to patient | 1
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Difficult to Dispense Polymer | Surgical conversion | 4
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Difficult to Dispense Polymer | Additional EV procedure | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Difficult to Dispense Polymer | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Difficult to Dispense Polymer | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Difficult to Dispense Polymer | Endoleak Type IA | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Abnormal blood loss | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Additional EV procedure | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Convert to femoral cut-down | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Edema | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Embolism | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type IA | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type IB | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type II | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak type II | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | Prolonged procedure | 3
intravascular | Stent Graft | Dispensing Issue; Inaccurate Dispensing | Unable to Dispense Polymer | No impact to patient | 1
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Rupture (aorta) | 5
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Neurologic impairment | 4
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Abnormal blood loss | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Additional EV procedure | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Embolism | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Endoleak Type IA | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Endoleak Type IB | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Occlusion (non-specific) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Stenosis (non-specific) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | Implant Migration, Cuff | No impact to patient | 1
intravascular | Stent Graft | Electromagnetic Interference (EMI); device-device incompatibility; environmental compatibility problem | MRI Compatibility Issues | Burn, Thermal | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Blood in Polymer Line (major) | Endoleak Type IA | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Blood in Polymer Line (major) | Endoleak Type IB | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Blood in Polymer Line (minor) | Endoleak Type IA | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Blood in Polymer Line (minor) | Endoleak Type IB | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Rupture (aorta) | 5
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Perforation (aorta) | 4
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Surgical conversion | 4
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Additional EV procedure | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Dissection (aorta) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Edema | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Embolism | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak Type IA | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak Type IB | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak Type II | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Endoleak type II | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Hemodynamic instability | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | Prolonged procedure | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Endobag Compromised | No impact to patient | 1
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Rupture (aorta) | 5
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Surgical conversion | 4
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Additional EV procedure | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type IA | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type IB | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | Prolonged procedure | 3
intravascular | Stent Graft | Failure to adhere or bond; loss of or failure to bond | Loss of Seal | No impact to patient | 1
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Rupture (aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Perforation (aorta) | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Surgical conversion | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Abnormal blood loss | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Additional EV procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Dissection (aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Embolism | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | Prolonged procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment Issue | No impact to patient | 1
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Surgical conversion | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Abnormal blood loss | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Access site complications | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Additional EV procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Cardiac arrhythmia | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Dissection (aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Endoleak Type IA | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Endoleak Type IB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | Prolonged procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Accessories) | No impact to patient | 1
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Rupture (aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Perforation (aorta) | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Renal insufficiency | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Surgical conversion | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Abnormal blood loss | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Access site complications | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Additional EV procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Dissection (aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Embolism | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type IA | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type IB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Hemodynamic instability | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Occlusion (aortic) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | Prolonged procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Catheter) | No impact to patient | 1
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Rupture (aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Perforation (aorta) | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Renal failure | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Renal insufficiency | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Surgical conversion | 4
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Abnormal blood loss | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Access site complications | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Additional EV procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Dissection (aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Embolism | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type IA | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type IB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Hemodynamic instability | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Occlusion (aortic) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Occlusion (mesenteric) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Prolonged procedure | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Stenosis (aorta) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Stenosis (non-specific) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Fitting Problem; device-device incompatibility; compatibility problem | Detachment of Components (Implant) | No impact to patient | 1
intravascular | Stent Graft | flaked | Flaked | Infection (device related) | 4
intravascular | Stent Graft | flaked | Flaked | Surgical conversion | 4
intravascular | Stent Graft | flaked | Flaked | Abnormal blood loss | 3
intravascular | Stent Graft | flaked | Flaked | Access site complications | 3
intravascular | Stent Graft | flaked | Flaked | Additional EV procedure | 3
intravascular | Stent Graft | flaked | Flaked | Convert to femoral cut-down | 3
intravascular | Stent Graft | flaked | Flaked | Embolism | 3
intravascular | Stent Graft | flaked | Flaked | Occlusion (iliac limb) | 3
intravascular | Stent Graft | flaked | Flaked | Occlusion (mesenteric) | 3
intravascular | Stent Graft | flaked | Flaked | Occlusion (non-specific) | 3
intravascular | Stent Graft | flaked | Flaked | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | flaked | Flaked | Prolonged procedure | 3
intravascular | Stent Graft | flaked | Flaked | Stenosis (iliac limb) | 3
intravascular | Stent Graft | flaked | Flaked | Stenosis (mesenteric) | 3
intravascular | Stent Graft | flaked | Flaked | Stenosis (non-specific) | 3
intravascular | Stent Graft | flaked | Flaked | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | flaked | Flaked | Stenosis (renal) | 3
intravascular | Stent Graft | flaked | Flaked | No impact to patient | 1
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Balloon Inflation Issue | Surgical conversion | 4
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Balloon Inflation Issue | Additional EV procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Balloon Inflation Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Balloon Inflation Issue | Prolonged procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Abnormal blood loss | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Access site complications | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Additional EV procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Edema | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Hemodynamic instability | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | Prolonged procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Blood Leak through Device | No impact to patient | 1
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Neurologic impairment | 4
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Surgical conversion | 4
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Abnormal blood loss | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Additional EV procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Cardiac arrhythmia | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Convert to femoral cut-down | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Embolism | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Endoleak Type IA | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Endoleak Type IB | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Prolonged procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Respiratory complications | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | Rash | 2
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Fluid Leak | No impact to patient | 1
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Rupture (aorta) | 5
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Surgical conversion | 4
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Additional EV procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type IA | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type IB | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | Prolonged procedure | 3
intravascular | Stent Graft | Fluid Leak; leak; leak/splash | Loss of Seal | No impact to patient | 1
intravascular | Stent Graft | Image Reversal; Image Orientation Incorrect; Image Display Error; display or visual feedback problem; incorrect measurement; Poor Quality Image; Image Resolution Poor; Incorrect Display; Misfocusing | Inappropriate Treatment | Prolonged procedure | 3
intravascular | Stent Graft | Image Reversal; Image Orientation Incorrect; Image Display Error; display or visual feedback problem; incorrect measurement; Poor Quality Image; Image Resolution Poor; Incorrect Display; Misfocusing | Inappropriate Treatment | Additional EV procedure | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Balloon Inflation Issue | Surgical conversion | 4
intravascular | Stent Graft | inflation problem; Inflation Issue | Balloon Inflation Issue | Additional EV procedure | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Balloon Inflation Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Balloon Inflation Issue | Prolonged procedure | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Surgical conversion | 4
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Additional EV procedure | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Edema | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Embolism | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Endoleak (indeterminate) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Endoleak Type I (Other) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Endoleak Type IA | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Endoleak Type IB | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Occlusion (aortic) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Occlusion (iliac limb) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Occlusion (non-specific) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Prolonged procedure | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Stenosis (aorta) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Stenosis (iliac limb) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Stenosis (non-specific) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | inflation problem; Inflation Issue | Unable to Inflate | No impact to patient | 1
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Neurologic impairment | 4
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Surgical conversion | 4
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Access site complications | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Additional EV procedure | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Embolism | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Occlusion (aortic) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Occlusion (non-specific) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Prolonged procedure | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Stenosis (aorta) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Stenosis (non-specific) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Insufficient Flow or Underinfusion; infusion or flow problem; improper flow or infusion; restricted flow rate; no flow | Graft Patency | No impact to patient | 1
intravascular | Stent Graft | insulation degradation | Graft Patency | Neurologic impairment | 4
intravascular | Stent Graft | insulation degradation | Graft Patency | Surgical conversion | 4
intravascular | Stent Graft | insulation degradation | Graft Patency | Access site complications | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Additional EV procedure | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Embolism | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Occlusion (aortic) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Occlusion (iliac limb) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Occlusion (non-specific) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Prolonged procedure | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Stenosis (aorta) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Stenosis (iliac limb) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Stenosis (non-specific) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | insulation degradation | Graft Patency | No impact to patient | 1
intravascular | Stent Graft | material discolored | Material Degraded | Prolonged procedure | 3
intravascular | Stent Graft | material discolored | Material Degraded | No impact to patient | 1
intravascular | Stent Graft | material fragmentation; device or device fragments location unknown; calcified | Fragments within body | Infection (device related) | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Surgical conversion | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Abnormal blood loss | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Access site complications | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Additional EV procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Dissection (aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Hemodynamic instability | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | Prolonged procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Broken Component | No impact to patient | 1
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Rupture (aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Surgical conversion | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Abnormal blood loss | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Access site complications | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Additional EV procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Convert to femoral cut-down | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Dissection (aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Endoleak Type IA | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Endoleak Type IB | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Hemodynamic instability | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | Prolonged procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Compromised Graft Integrity | No impact to patient | 1
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Surgical conversion | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Abnormal blood loss | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Access site complications | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Additional EV procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Dissection (aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Hemodynamic instability | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | Prolonged procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Damaged Component | No impact to patient | 1
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Endobag Compromised | Rupture (aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Endobag Compromised | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Endobag Compromised | Perforation (aorta) | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Endobag Compromised | Surgical conversion | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Rupture (aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Perforation (aorta) | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Renal failure | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Renal insufficiency | 4
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Access site complications | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Additional EV procedure | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Dissection (aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Embolism | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type IA | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type IB | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Hemodynamic instability | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Occlusion (aortic) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (aorta) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (non-specific) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | Stenosis (renal) | 3
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | No impact to patient | 1
intravascular | Stent Graft | Material Integrity Issue; material integrity problem; stretched; product quality problem; Product Quality Issue; hole in material; break; crack; fracture; torn material; material frayed; material split cut or torn; folded; defective component; naturally worn; split; melted; Failure to Fold | Fracture (Implant) | No patient involvement | 1
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Neurologic impairment | 4
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Surgical conversion | 4
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Access site complications | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Additional EV procedure | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Embolism | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Occlusion (aortic) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Prolonged procedure | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Stenosis (aorta) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Stenosis (non-specific) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Material Invagination; Dent in Material; Pitted | Material Invagination (Graft) | No impact to patient | 1
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Rupture (aorta) | 5
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Renal failure | 4
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Renal insufficiency | 4
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Surgical conversion | 4
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Abnormal blood loss | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Access site complications | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Additional EV procedure | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Embolism | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Occlusion (aortic) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Occlusion (iliac limb) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Occlusion (mesenteric) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Occlusion (non-specific) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Prolonged procedure | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Prolonged procedure | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (aorta) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (iliac limb) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (mesenteric) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (non-specific) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | Stenosis (renal) | 3
intravascular | Stent Graft | material protrusion/extrusion; Material Puncture; material puncture/hole; extrusion; material perforation; Material Protrusion | Endobag Protrusion | No impact to patient | 1
intravascular | Stent Graft | material rupture | Balloon Rupture | Surgical conversion | 4
intravascular | Stent Graft | material rupture | Balloon Rupture | Additional EV procedure | 3
intravascular | Stent Graft | material rupture | Balloon Rupture | Convert to femoral cut-down | 3
intravascular | Stent Graft | material rupture | Balloon Rupture | Prolonged procedure | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Rupture (aorta) | 5
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Neurologic impairment | 4
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Surgical conversion | 4
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Additional EV procedure | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Embolism | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Endoleak (indeterminate) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Endoleak Type I (Other) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Endoleak Type IA | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Occlusion (iliac limb) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Stenosis (aorta) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Stenosis (iliac limb) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Stenosis (non-specific) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | material separation; separation problem; delamination; unraveled material; peeled/delaminated; Peeled; premature separation; material disintegration; Disassembly | Implant Separation | No impact to patient | 1
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Surgical conversion | 4
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Abnormal blood loss | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Access site complications | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Additional EV procedure | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Dissection (aorta) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Hemodynamic instability | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | Prolonged procedure | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Broken Component | No impact to patient | 1
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Surgical conversion | 4
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Abnormal blood loss | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Access site complications | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Additional EV procedure | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Dissection (aorta) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Hemodynamic instability | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | Prolonged procedure | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Damaged Component | No impact to patient | 1
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Rupture (aorta) | 5
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Perforation (aorta) | 4
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Renal failure | 4
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Renal insufficiency | 4
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Access site complications | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Additional EV procedure | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Dissection (aorta) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Embolism | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type IA | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type IB | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Hemodynamic instability | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Occlusion (aortic) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Occlusion (non-specific) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (aorta) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (non-specific) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | Stenosis (renal) | 3
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | No impact to patient | 1
intravascular | Stent Graft | Mechanical Issue; material deformation; deformation due to compressive stress; mechanical problem; difficult to open or close; collapse; expulsion; unintended ejection; component falling; Difficult to Insert | Fracture (Implant) | No patient involvement | 1
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Rupture (aorta) | 5
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Neurologic impairment | 4
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Abnormal blood loss | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Additional EV procedure | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Embolism | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Endoleak Type IA | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Endoleak Type IB | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Occlusion (non-specific) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Stenosis (non-specific) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Cuff | No impact to patient | 1
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Neurologic impairment | 4
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Surgical conversion | 4
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Abnormal blood loss | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Additional EV procedure | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Embolism | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Endoleak Type IB | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Occlusion (non-specific) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Prolonged procedure | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Prolonged procedure | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Limb | No impact to patient | 1
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Rupture (aorta) | 5
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Neurologic impairment | 4
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Additional EV procedure | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Embolism | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak Type IA | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak Type IB | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Occlusion (aortic) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Occlusion (non-specific) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Stenosis (non-specific) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Migration of Device or Device Component; migration or expulsion of device; migration; device dislodged or dislocated | Implant Migration, Main Body | No impact to patient | 1
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Rupture (aorta) | 5
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Perforation (aorta) | 4
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Surgical conversion | 4
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Abnormal blood loss | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Additional EV procedure | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Convert to femoral cut-down | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Dissection (aorta) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Embolism | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type IA | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type IB | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Endoleak Type IIIIB | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | Prolonged procedure | 3
intravascular | Stent Graft | misconnection; disconnection; incomplete or inadequate connection; connection problem; loose or intermittent connection; failure to disconnect | Detachment Issue | No impact to patient | 1
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Rupture (aorta) | 5
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Perforation (aorta) | 4
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Abnormal blood loss | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Additional EV procedure | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Dissection (aorta) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type II | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak type II | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Occlusion (aortic) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Prolonged procedure | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Stenosis (aorta) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Stenosis (non-specific) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | Stenosis (renal) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visibility Issue | No impact to patient | 1
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Rupture (aorta) | 5
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Perforation (aorta) | 4
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Abnormal blood loss | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Additional EV procedure | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Dissection (aorta) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type IA | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type IB | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type II | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak type II | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Occlusion (aortic) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Occlusion (non-specific) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Prolonged procedure | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Stenosis (aorta) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Stenosis (non-specific) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | Stenosis (renal) | 3
intravascular | Stent Graft | Optical Issue; Light Interference; Optical Decentration; Optical Discoloration; Optical Distortion; Optical Obstruction | Visualization Issue | No impact to patient | 1
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Compromised Sterility | Infection (device related) | 4
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Surgical conversion | 4
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Abnormal blood loss | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Access site complications | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Additional EV procedure | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Convert to femoral cut-down | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Dissection (aorta) | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Hemodynamic instability | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | Prolonged procedure | 3
intravascular | Stent Graft | Packaging Issue; Device Packaging Compromised; Incomplete or Missing Packaging; Tear, Rip or Hole in Device Packaging; Unsealed Device Packaging; packaging problem; manufacturing packaging or shipping problem; tear rip or hole in device packaging; unsealed device packaging; shipping damage or problem | Damaged Component | No impact to patient | 1
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Surgical conversion | 4
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Additional EV procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Convert to femoral cut-down | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Embolism | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Endoleak (indeterminate) | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Endoleak Type IIIIB | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | Prolonged procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Blockage within Device | No impact to patient | 1
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Surgical conversion | 4
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Additional EV procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Convert to femoral cut-down | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Embolism | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Endoleak (indeterminate) | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Endoleak Type IIIIB | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | Prolonged procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Obstruction | No impact to patient | 1
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Surgical conversion | 4
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Additional EV procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Convert to femoral cut-down | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Embolism | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Endoleak Type IA | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Endoleak Type IB | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | Prolonged procedure | 3
intravascular | Stent Graft | partial blockage; complete blockage; obstruction of flow; inability to irrigate; filling problem; flushing problem; Blockage within Device or Device Component; Blocked Connection; Device Clogged; Obstruction within Device; Occlusion within Device | Occlusion within Device (Implant) | No impact to patient | 1
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Rupture (aorta) | 5
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Infection (device related) | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Neurologic impairment | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Perforation (aorta) | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Renal failure | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Renal insufficiency | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Surgical conversion | 4
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Abnormal blood loss | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Access site complications | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Additional EV procedure | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Cardiac arrhythmia | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Convert to femoral cut-down | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Dissection (aorta) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Dissection (upper extremity) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Edema | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Embolism | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type IA | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type IB | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type II | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Hemodynamic instability | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (aortic) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (mesenteric) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (non-specific) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Occlusion (upper extremity) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Prolonged procedure | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Respiratory complications | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (aorta) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (non-specific) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (renal) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Stenosis (upper extremity) | 3
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | Rash | 2
intravascular | Stent Graft | Patient-Device Incompatibility; patient device interaction problem; human-device interface problem; inadequate user interface; Device Handling Issue | Patient-Device Incompatibility | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Renal failure | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Renal insufficiency | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Dissection (aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak Type IA | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak Type IB | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Occlusion (aortic) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Occlusion (non-specific) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | Prolonged procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Position Implant | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Abnormal blood loss | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Access site complications | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Convert to femoral cut-down | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Convert to femoral cut-down | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | Prolonged procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Dock | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Abnormal blood loss | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Access site complications | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Convert to femoral cut-down | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Convert to femoral cut-down | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | Prolonged procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Difficult to Undock | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Rupture (aorta) | 5
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Abnormal blood loss | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Endoleak Type IA | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Endoleak Type IB | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Endoleak Type II | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Endoleak type II | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | Prolonged procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Inadequate endobag-wall apposition | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Rupture (aorta) | 5
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Neurologic impairment | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Perforation (aorta) | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Renal failure | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Renal insufficiency | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Access site complications | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Convert to femoral cut-down | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Embolism | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type IA | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type IB | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type II | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Endoleak Type IIIIB | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Hemodynamic instability | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Occlusion (aortic) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Occlusion (mesenteric) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Occlusion (non-specific) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Prolonged procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (iliac limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (mesenteric) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (non-specific) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | Stenosis (renal) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Malposition of the Device | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Abnormal blood loss | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Access site complications | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Dissection (aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Dissection (upper extremity) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Edema | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Embolism | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Occlusion (non-specific) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Dock | No impact to patient | 1
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Surgical conversion | 4
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Abnormal blood loss | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Access site complications | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Additional EV procedure | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Dissection (aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Dissection (upper extremity) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Edema | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Embolism | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Occlusion (iliac limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Occlusion (non-specific) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Positioning Issue; malposition of device; activation positioning problem; difficult or delayed positioning; positioning problem; difficult to position; positioning failure; therapy delivered to incorrect body area; Dislodged or Dislocated; Failure to Align | Unable to Undock | No impact to patient | 1
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Rupture (aorta) | 5
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Perforation (aorta) | 4
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Surgical conversion | 4
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Additional EV procedure | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Dissection (aorta) | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Edema | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Embolism | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak Type IA | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak Type IB | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak Type II | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Endoleak type II | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Hemodynamic instability | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | Prolonged procedure | 3
intravascular | Stent Graft | Pressure Issue; Decrease in Pressure; Increase in Pressure; No Pressure; pressure problem | Unable to hold pressure in endobag | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Access site complications | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Convert to femoral cut-down | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak (indeterminate) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak Type I (Other) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak Type IA | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak Type IB | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Remove | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Access site complications | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Convert to femoral cut-down | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Access site complications | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Convert to femoral cut-down | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Retract | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Access site complications | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Convert to femoral cut-down | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Dissection (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Endoleak (indeterminate) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Endoleak Type I (Other) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Endoleak Type IA | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Endoleak Type IB | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Difficult to Withdraw | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Rupture (aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Dissection (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Edema | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Embolism | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Dissection (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Edema | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Embolism | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Retract | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Rupture (aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Dissection (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Edema | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Embolism | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Occlusion (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Occlusion (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Stenosis (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Stenosis (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Stenosis (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Unsheath | No impact to patient | 1
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Rupture (aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Perforation (aorta) | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Surgical conversion | 4
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Abnormal blood loss | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Access site complications | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Additional EV procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Convert to femoral cut-down | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Dissection (aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Edema | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Embolism | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Endoleak (indeterminate) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Endoleak Type I (Other) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Endoleak Type IA | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Endoleak Type IB | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Occlusion (iliac limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Occlusion (non-specific) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | Prolonged procedure | 3
intravascular | Stent Graft | retraction problem; Mechanical Jam; Mechanics Altered; Physical Resistance; Difficult to Remove | Unable to Withdraw | No impact to patient | 1
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Surgical conversion | 4
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Additional EV procedure | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak Type IA | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak Type IB | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | Prolonged procedure | 3
intravascular | Stent Graft | Size Incorrect for patient; inadequacy of device shape and/or size | Size Incorrect for Patient (Undersized) | No impact to patient | 1
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Rupture (aorta) | 5
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Perforation (aorta) | 4
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Surgical conversion | 4
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Abnormal blood loss | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Additional EV procedure | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Convert to femoral cut-down | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Dissection (aorta) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Embolism | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak (indeterminate) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type I (Other) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type IA | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type IB | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Endoleak Type IIIIB | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Occlusion (non-specific) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | Prolonged procedure | 3
intravascular | Stent Graft | sticking; physical resistance/sticking | Component Attachment | No impact to patient | 1
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Rupture (aorta) | 5
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Surgical conversion | 4
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Additional EV procedure | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Embolism | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Endoleak (indeterminate) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Endoleak Type I (Other) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Endoleak Type IA | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Occlusion (iliac limb) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Occlusion (non-specific) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Stenosis (iliac limb) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Stenosis (non-specific) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | unintended movement; device slipped; unintended system motion; Unintended Ejection | Implant movement | No impact to patient | 1
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Rupture (aorta) | 5
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Rupture (other, non-aorta) | 5
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Neurologic impairment | 4
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Perforation (aorta) | 4
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Renal failure | 4
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Renal insufficiency | 4
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Surgical conversion | 4
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Abnormal blood loss | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Access site complications | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Additional EV procedure | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Convert to femoral cut-down | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Dissection (aorta) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Dissection (iliac & distal vessels) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Embolism | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak (indeterminate) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type I (Other) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type IA | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type IB | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type IIIIA (aortic component) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type IIIIA (iliac component) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Endoleak Type IIIIB | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Hemodynamic instability | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Occlusion (aortic) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Occlusion (iliac limb) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Occlusion (non-specific) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Occlusion (pelvis lower limb) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Perforation (other, non-aorta) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Prolonged procedure | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Stenosis (aorta) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Stenosis (iliac limb) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Stenosis (non-specific) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | Stenosis (pelvis, lower limb) | 3
intravascular | Stent Graft | User used incorrect product for intended use; off-label use; device expiration issue | Incorrect Product for intended use (off-label) | No impact to patient | 1

## ProductCode_Lists
Category | Product Code | Device | Device Class | Regulation Number | Regulation Description | Life Sustaining | Implant | Client
Anesthesiology | PHZ | Abnormal Breath Sound Device | 2 | 868.1900 | Diagnostic pulmonary-function interpretation calculator. | N | N
Anesthesiology | CBL | Absorbent, Carbon-Dioxide | 1 | 868.5300 | Carbon dioxide absorbent. | N | N
Anesthesiology | BSF | Absorber, Carbon-Dioxide | 1 | 868.5310 | Carbon dioxide absorber. | N | N
Anesthesiology | MOD | Accessory To Continuous Ventilator (Respirator) | 2 | 868.5895 | Continuous ventilator. | N | N
Anesthesiology | OQU | Airway Monitoring System | 2 | 868.5730 | Tracheal tube. | N | N
Anesthesiology | OFO | Airway Suction Kit | 1 | 868.6810 | Tracheobronchial suction catheter. | N | N
Anesthesiology | CAO | Airway, Esophageal (Obturator) | 2 | 868.5650 | Esophageal obturator. | Y | N
Anesthesiology | BTQ | Airway, Nasopharyngeal | 1 | 868.5100 | Nasopharyngeal airway. | N | N
Anesthesiology | CAE | Airway, Oropharyngeal, Anesthesiology | 1 | 868.5110 | Oropharyngeal airway. | N | N
Anesthesiology | BXL | Algesimeter, Manual | 1 | 868.1030 | Manual algesimeter. | N | N
Anesthesiology | BSI | Algesimeter, Powered | 2 | 868.1040 | Powered algesimeter. | N | N
Anesthesiology | JEG | Analyzer, Gas, Argon, Gaseous-Phase | 2 | 868.1075 | Argon gas analyzer. | N | N
Anesthesiology | CCK | Analyzer, Gas, Carbon-Dioxide, Gaseous-Phase | 2 | 868.1400 | Carbon dioxide gas analyzer. | N | N
Anesthesiology | CCC | Analyzer, Gas, Carbon-Dioxide, Partial Pressure, Blood-Phase, Indwelling | 2 | 868.1150 | Indwelling blood carbon dioxide partial pressure (PCO2) analyzer. | N | N
Anesthesiology | CCJ | Analyzer, Gas, Carbon-Monoxide, Gaseous-Phase | 2 | 868.1430 | Carbon monoxide gas analyzer. | N | N
Anesthesiology | NHO | Analyzer, Gas, Desflurane, Gaseous-Phase (Anesthetic Concentration) | 2 | 868.1500 | Enflurane gas analyzer. | N | N
Anesthesiology | CBQ | Analyzer, Gas, Enflurane, Gaseous-Phase (Anesthetic Concentration) | 2 | 868.1500 | Enflurane gas analyzer. | N | N
Anesthesiology | CBS | Analyzer, Gas, Halothane, Gaseous-Phase (Anesthetic Conc.) | 2 | 868.1620 | Halothane gas analyzer. | N | N
Anesthesiology | BSE | Analyzer, Gas, Helium, Gaseous-Phase | 2 | 868.1640 | Helium gas analyzer. | N | N
Anesthesiology | NHQ | Analyzer, Gas, Isoflurane, Gaseous-Phase (Anesthetic Concentration) | 2 | 868.1500 | Enflurane gas analyzer. | N | N
Anesthesiology | JEF | Analyzer, Gas, Neon, Gaseous-Phase | 2 | 868.1670 | Neon gas analyzer. | N | N
Anesthesiology | CCI | Analyzer, Gas, Nitrogen, Gaseous-Phase | 2 | 868.1690 | Nitrogen gas analyzer. | N | N
Anesthesiology | CBR | Analyzer, Gas, Nitrous-Oxide, Gaseous Phase (Anesthetic Conc.) | 2 | 868.1700 | Nitrous oxide gas analyzer. | N | N
Anesthesiology | CCL | Analyzer, Gas, Oxygen, Gaseous-Phase | 2 | 868.1720 | Oxygen gas analyzer. | N | N
Anesthesiology | CCE | Analyzer, Gas, Oxygen, Partial Pressure, Blood-Phase, Indwelling | 2 | 868.1200 | Indwelling blood oxygen partial pressure (PO2) analyzer. | N | N
Anesthesiology | NHP | Analyzer, Gas, Sevoflurane, Gaseous-Phase (Anesthetic Concentration) | 2 | 868.1500 | Enflurane gas analyzer. | N | N
Anesthesiology | BXA | Analyzer, Gas, Water Vapor, Gaseous-Phase | 1 | 868.1975 | Water vapor analyzer. | N | N
Anesthesiology | CBZ | Analyzer, Ion, Hydrogen-Ion (Ph), Blood-Phase, Indwelling | 2 | 868.1170 | Indwelling blood hydrogen ion concentration (pH) analyzer. | N | N
Anesthesiology | MRP | Analyzer, Nitric Oxide | 2 | 868.2380 | Nitric oxide analyzer. | N | N
Anesthesiology | MRQ | Analyzer, Nitrogen Dioxide | 2 | 868.2385 | Nitrogen dioxide analyzer. | Y | N
Anesthesiology | PUG | Analyzer, Nitrogen Dioxide, Exempt | 2 | 868.2385 | Nitrogen dioxide analyzer. | Y | N
Anesthesiology | JED | Analyzer, Oxyhemoglobin Concentration, Blood-Phase, Indwelling | 3 | 868.1120 | Indwelling blood oxyhemoglobin concentration analyzer. | N | N
Anesthesiology | OFP | Anesthesia Breathing Circuit Kit (Adult & Pediatric) | 1 | 868.5240 | Anesthesia breathing circuit. | N | N
Anesthesiology | CAZ | Anesthesia Conduction Kit | 2 | 868.5140 | Anesthesia conduction kit. | N | N
Anesthesiology | OFQ | Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | CAC | Apparatus, Autotransfusion | 2 | 868.5830 | Autotransfusion apparatus. | Y | N
Anesthesiology | BWL | Apparatus, Electronanesthesia | 3 | 868.5400 | Electroanesthesia apparatus. | N | N
Anesthesiology | CBN | Apparatus, Gas-Scavenging | 2 | 868.5430 | Gas-scavenging apparatus. | N | N
Anesthesiology | MRN | Apparatus, Nitric Oxide Delivery | 2 | 868.5165 | Nitric oxide administration apparatus. | Y | N
Anesthesiology | MRO | Apparatus, Nitric Oxide, Backup Delivery | 2 | 868.5165 | Nitric oxide administration apparatus. | Y | N
Anesthesiology | CCT | Applicator (Laryngo-Tracheal), Topical Anesthesia | 2 | 868.5170 | Laryngotracheal topical anesthesia applicator. | N | N
Anesthesiology | CBT | Arterial Blood Sampling Kit | 1 | 868.1100 | Arterial blood sampling kit. | N | N
Anesthesiology | BYE | Attachment, Breathing, Positive End Expiratory Pressure | 2 | 868.5965 | Positive end expiratory pressure breathing attachment. | Y | N
Anesthesiology | CBO | Attachment, Intermittent Mandatory Ventilation (Imv) | 2 | 868.5955 | Intermittent mandatory ventilation attachment. | Y | N
Anesthesiology | BTC | Bag, Reservoir | 1 | 868.5320 | Reservoir bag. | N | N
Anesthesiology | CCO | Bed, Rocking, Breathing Assist | 2 | 868.5180 | Rocking bed. | Y | N
Anesthesiology | OFZ | Blood Specimen Collection Kit (Excludes Hiv Testing) | 1 | 868.1100 | Arterial blood sampling kit. | N | N
Anesthesiology | BYO | Bottle, Blow | 1 | 868.5220 | Blow bottle. | N | N
Anesthesiology | OGA | Brachial Plexus Anesthesia Kit | 2 | 868.5140 | Anesthesia conduction kit. | N | N
Anesthesiology | EPE | Brush, Cleaning, Tracheal Tube | 1 | 868.5795 | Tracheal tube cleaning brush. | N | N
Anesthesiology | PDT | Burn Resuscitation Decision Support Software | 2 | 868.1890 | Predictive pulmonary-function value calculator. | N | N
Anesthesiology | BRY | Cabinet, Table And Tray, Anesthesia | 1 | 868.6100 | Anesthetic cabinet, table, or tray. | N | N
Anesthesiology | NDC | Calculator, Drug Dose | 2 | 868.1890 | Predictive pulmonary-function value calculator. | N | N
Anesthesiology | BTY | Calculator, Predicted Values, Pulmonary Function | 2 | 868.1890 | Predictive pulmonary-function value calculator. | N | N
Anesthesiology | BZC | Calculator, Pulmonary Function Data | 2 | 868.1880 | Pulmonary-function data calculator. | N | N
Anesthesiology | BZM | Calculator, Pulmonary Function Interpretor (Diagnostic) | 2 | 868.1900 | Diagnostic pulmonary-function interpretation calculator. | N | N
Anesthesiology | BXX | Calibrator, Pressure, Gas | 1 | 868.2620 | Gas pressure calibrator. | N | N
Anesthesiology | BXW | Calibrator, Volume, Gas | 1 | 868.1870 | Gas volume calibrator. | N | N
Anesthesiology | CAT | Cannula, Nasal, Oxygen | 1 | 868.5340 | Nasal oxygen cannula. | N | N
Anesthesiology | OEV | Cardiopulmonary Resuscitation Aid Kit | 2 | 868.5915 | Manual emergency ventilator. | N | N
Anesthesiology | BZN | Cart, Emergency, Cardiopulmonary (Excluding Equipment) | 1 | 868.6175 | Cardiopulmonary emergency cart. | N | N
Anesthesiology | BSO | Catheter, Conduction, Anesthetic | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | BZB | Catheter, Nasal, Oxygen | 1 | 868.5350 | Nasal oxygen catheter. | N | N
Anesthesiology | MAJ | Catheter, Percutaneous, Intraspinal, Short Term | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | BSY | Catheters, Suction, Tracheobronchial | 1 | 868.6810 | Tracheobronchial suction catheter. | N | N
Anesthesiology | OGB | Caudal Anesthesia Kit | 2 | 868.5140 | Anesthesia conduction kit. | N | N
Anesthesiology | BYN | Chair, Posture, For Cardiac And Pulmonary Treatment | 1 | 868.5365 | Posture chair for cardiac or pulmonary treatment. | N | N
Anesthesiology | CBF | Chamber, Hyperbaric | 2 | 868.5470 | Hyperbaric chamber. | Y | N
Anesthesiology | LNZ | Changer, Tube, Endotracheal | 2 | 868.5730 | Tracheal tube. | N | N
Anesthesiology | CAI | Circuit, Breathing (W Connector, Adaptor, Y Piece) | 1 | 868.5240 | Anesthesia breathing circuit. | N | N
Anesthesiology | CAG | Circulator, Breathing-Circuit | 2 | 868.5250 | Breathing circuit circulator. | Y | N
Anesthesiology | BXJ | Clip, Nose | 1 | 868.6225 | Nose clip. | N | N
Anesthesiology | BTI | Compressor, Air, Portable | 2 | 868.6250 | Portable air compressor. | N | N
Anesthesiology | BZL | Computer, Oxygen-Uptake | 2 | 868.1730 | Oxygen uptake computer. | N | N
Anesthesiology | BYD | Condenser, Heat And Moisture (Artificial Nose) | 1 | 868.5375 | Heat and moisture condenser (artificial nose). | N | N
Anesthesiology | BZA | Connector, Airway (Extension) | 1 | 868.5810 | Airway connector. | N | N
Anesthesiology | NFB | Conserver, Oxygen | 2 | 868.5905 | Noncontinuous ventilator (IPPB). | N | N
Anesthesiology | OGC | Continuous Brachial Plexus Block Tray | 2 | 868.5150 | Anesthesia conduction needle. | N | N
Anesthesiology | NOU | Continuous, Ventilator, Home Use | 2 | 868.5895 | Continuous ventilator. | Y | N | Philips Resporonics
Anesthesiology | OGP | Cricothyrotomy Kit | 2 | 868.5800 | Tracheostomy tube and tube cuff. | N | N
Anesthesiology | BSK | Cuff, Tracheal Tube, Inflatable | 2 | 868.5750 | Inflatable tracheal tube cuff. | Y | N
Anesthesiology | OGD | Custom Anesthesia Tray | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | ECX | Cylinder, Compressed Gas, And Valve | 1 | 868.2700 | Pressure regulator. | N | N
Anesthesiology | PRK | Device Indicating An Exhalation Event | 2 | 868.2375 | Breathing frequency monitor. | N | N
Anesthesiology | CBH | Device, Fixation, Tracheal Tube | 1 | 868.5770 | Tracheal tube fixation device. | N | N
Anesthesiology | MZT | Device, Heimlich Maneuver Assist | 2 | 868.5115 | Device to relieve acute upper airway obstruction. | Y | N
Anesthesiology | NHJ | Device, Positive Pressure Breathing, Intermittent | 2 | 868.5905 | Noncontinuous ventilator (IPPB). | N | N
Anesthesiology | BYW | Device, Rebreathing | 1 | 868.5675 | Rebreathing device. | N | N
Anesthesiology | BYH | Drain, Tee (Water Trap) | 1 | 868.5995 | Tee drain (water trap). | N | N
Anesthesiology | BTP | Dropper, Ether | 1 | 868.5420 | Ether hook. | N | N
Anesthesiology | OGE | Epidural Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | PMB | External Airway | 2 | 868.5105 | External negative pressure airway aid. | N | N
Anesthesiology | CAH | Filter, Bacterial, Breathing-Circuit | 2 | 868.5260 | Breathing circuit bacterial filter. | N | N
Anesthesiology | BSN | Filter, Conduction, Anesthetic | 2 | 868.5130 | Anesthesia conduction filter. | N | N
Anesthesiology | BXY | Flowmeter, Calibration, Gas | 1 | 868.2350 | Gas calibration flowmeter. | N | N
Anesthesiology | CCN | Flowmeter, Nonback-Pressure Compensated, Bourdon Gauge | 1 | 868.2300 | Bourdon gauge flowmeter. | N | N
Anesthesiology | CAX | Flowmeter, Tube, Thorpe, Back-Pressure Compensated | 1 | 868.2340 | Compensated thorpe tube flowmeter. | N | N
Anesthesiology | BWB | Forceps, Tube Introduction | 1 | 868.5780 | Tube introduction forceps. | N | N
Anesthesiology | BXK | Gas, Calibration (Specified Concentration) | 1 | 868.6400 | Calibration gas. | N | N
Anesthesiology | KGK | Gas, Collecting Vessel | 1 | 868.1575 | Gas collection vessel. | N | N
Anesthesiology | BSZ | Gas-Machine, Anesthesia | 2 | 868.5160 | Gas machine for anesthesia or analgesia. | Y | N
Anesthesiology | BXH | Gauge, Gas Pressure, Cylinder/Pipeline | 1 | 868.2610 | Gas pressure gauge. | N | N
Anesthesiology | CAW | Generator, Oxygen, Portable | 2 | 868.5440 | Portable oxygen generator. | N | N | Philips Resporonics
Anesthesiology | OGF | Glossopharyngeal Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | BZE | Heater, Breathing System W/Wo Controller (Not Humidifier Or Nebulizer | 2 | 868.5270 | Breathing system heater. | N | N
Anesthesiology | QAV | High Flow/High Velocity Humidified Oxygen Delivery Device | 2 | 868.5454 | High flow humidified oxygen delivery device. | N | N
Anesthesiology | NVP | Holding Chambers, Direct Patient Interface | 2 | 868.5630 | Nebulizer. | N | N
Anesthesiology | FOG | Hood, Oxygen, Infant | 1 | 868.5700 | Nonpowered oxygen tent. | N | N
Anesthesiology | BTB | Hook, Ether | 1 | 868.5420 | Ether hook. | N | N
Anesthesiology | OGG | Humidifier Nebulizer Kit | 2 | 868.5450 | Respiratory gas humidifier. | N | N
Anesthesiology | KFZ | Humidifier, Non-Direct Patient Interface (Home-Use) | 1 | 868.5460 | Therapeutic humidifier for home use. | N | N
Anesthesiology | BTT | Humidifier, Respiratory Gas, (Direct Patient Interface) | 2 | 868.5450 | Respiratory gas humidifier. | N | N
Anesthesiology | OBN | Humidifier, Respiratory, Mask (Direct Patient Interface) | 2 | 868.5450 | Respiratory gas humidifier. | N | N
Anesthesiology | PNV | Impedance Measuring Device Utilizing Oscillation Techniques | 2 | 868.1840 | Diagnostic spirometer. | N | N
Anesthesiology | OGH | Laryngoscope Kit | 1 | 868.5540 | Rigid laryngoscope. | N | N
Anesthesiology | CAL | Laryngoscope, Non-Rigid | 1 | 868.5530 | Flexible laryngoscope. | N | N
Anesthesiology | CCW | Laryngoscope, Rigid | 1 | 868.5540 | Rigid laryngoscope. | N | N
Anesthesiology | BSJ | Mask, Gas, Anesthetic | 1 | 868.5550 | Anesthetic gas mask. | N | N
Anesthesiology | BYG | Mask, Oxygen | 1 | 868.5580 | Oxygen mask. | N | N
Anesthesiology | BYF | Mask, Oxygen, Low Concentration, Venturi | 1 | 868.5600 | Venturi mask. | N | N
Anesthesiology | KGB | Mask, Oxygen, Non-Rebreathing | 1 | 868.5570 | Nonrebreathing mask. | N | N
Anesthesiology | KHA | Mask, Scavenging | 1 | 868.5590 | Scavenging mask. | N | N
Anesthesiology | NMC | Mask, Ventilator, Non-Continuous, Reprocessed | 2 | 868.5905 | Noncontinuous ventilator (IPPB). | N | N
Anesthesiology | ONZ | Mechanical Ventilator | 2 | 868.5895 | Continuous ventilator. | N | N
Anesthesiology | OLF | Medevac Use Oxygen Generator | 2 | 868.5440 | Portable oxygen generator. | Y | N
Anesthesiology | BXR | Meter, Airway Pressure (Inspiratory Force) | 2 | 868.1780 | Inspiratory airway pressure meter. | N | N
Anesthesiology | BZH | Meter, Peak Flow, Spirometry | 2 | 868.1860 | Peak-flow meter for spirometry. | N | N
Anesthesiology | BZR | Mixer, Breathing Gases, Anesthesia Inhalation | 2 | 868.5330 | Breathing gas mixer. | Y | N
Anesthesiology | CBA | Monitor, Air Embolism, Ultrasonic | 2 | 868.2025 | Ultrasonic air embolism monitor. | N | N
Anesthesiology | CAP | Monitor, Airway Pressure (Includes Gauge And/Or Alarm) | 2 | 868.2600 | Airway pressure monitor. | N | N
Anesthesiology | FLS | Monitor, Apnea, Facility Use | 2 | 868.2377 | Apnea monitor. | Y | N
Anesthesiology | NPF | Monitor, Apnea, Home Use | 2 | 868.2377 | Apnea monitor. | Y | N
Anesthesiology | BZQ | Monitor, Breathing Frequency | 2 | 868.2375 | Breathing frequency monitor. | N | N
Anesthesiology | LKD | Monitor, Carbon-Dioxide, Cutaneous | 2 | 868.2480 | Cutaneous carbon dioxide (PcCO2) monitor. | N | N
Anesthesiology | JEZ | Monitor, Lung Water Measurement | 3 | 868.2450 | Lung water monitor. | N | N
Anesthesiology | KLK | Monitor, Oxygen, Cutaneous, For Infant Not Under Gas Anesthesia | 2 | 868.2500 | Cutaneous oxygen (PcO2) monitor. | N | N
Anesthesiology | LPP | Monitor, Oxygen, Cutaneous, For Uses Other Than For Infant Not Under Gas Anesthesia | 2 | 868.2500 | Cutaneous oxygen (PcO2) monitor. | N | N
Anesthesiology | BYP | Mouthpiece, Breathing | 1 | 868.5620 | Breathing mouthpiece. | N | N
Anesthesiology | OGI | Nasal Endotracheal Tube Holder Kit | 1 | 868.5770 | Tracheal tube fixation device. | N | N
Anesthesiology | CAF | Nebulizer (Direct Patient Interface) | 2 | 868.5630 | Nebulizer. | N | N
Anesthesiology | CCQ | Nebulizer, Medicinal, Non-Ventilatory (Atomizer) | 1 | 868.5640 | Medicinal nonventilatory nebulizer (atomizer). | N | N
Anesthesiology | NNH | Needle, Conduction, Anesthesia, Reprocessed | 2 | 868.5150 | Anesthesia conduction needle. | N | N
Anesthesiology | BSP | Needle, Conduction, Anesthetic (W/Wo Introducer) | 2 | 868.5150 | Anesthesia conduction needle. | N | N
Anesthesiology | BWC | Needle, Emergency Airway | 2 | 868.5090 | Emergency airway needle. | Y | N
Anesthesiology | MIA | Needle, Spinal, Short Term | 2 | 868.5150 | Anesthesia conduction needle. | N | N
Anesthesiology | NMR | Needle, Spinal, Short Term, Reprocessed | 2 | 868.5150 | Anesthesia conduction needle. | N | N
Anesthesiology | OGJ | Nerve Block Tray | 2 | 868.5140 | Anesthesia conduction kit. | N | N
Anesthesiology | OYI | Non-Bronchoscopic Bronchoalveolar Lavage Catheter | 1 | 868.6810 | Tracheobronchial suction catheter. | N | N
Anesthesiology | PHS | Oral Airway Assist | 1 | 868.5110 | Oropharyngeal airway. | N | N
Anesthesiology | OGK | Orotracheal Intubation Guide Kit | 2 | 868.1400 | Carbon dioxide gas analyzer. | N | N
Anesthesiology | OGL | Oxygen Administration Kit | 2 | 868.5655 | Portable liquid oxygen unit. | N | N
Anesthesiology | BYI | Percussor, Powered-Electric | 2 | 868.5665 | Powered percussor. | N | N
Anesthesiology | CCM | Plethysmograph, Pressure | 2 | 868.1750 | Pressure plethysmograph. | N | N
Anesthesiology | JEH | Plethysmograph, Volume | 2 | 868.1760 | Volume plethysmograph. | N | N
Anesthesiology | JAX | Pneumotachometer | 2 | 868.2550 | Pneumotachometer. | N | N
Anesthesiology | QBY | Positive Airway Pressure System | 2 | 868.5273 | Positive airway pressure delivery system. | N | N
Anesthesiology | PXE | Pressure Monitoring (Air/Gas) Kit | 1 | 868.2610 | Gas pressure gauge. | N | N
Anesthesiology | BRW | Protector, Dental | 1 | 868.5820 | Dental protector. | N | N
Anesthesiology | OFV | Regional Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | CAN | Regulator, Pressure, Gas Cylinder | 1 | 868.2700 | Pressure regulator. | N | N
Anesthesiology | NHK | Resuscitator, Manual, Non Self-Inflating | 2 | 868.5905 | Noncontinuous ventilator (IPPB). | Y | N
Anesthesiology | QCX | Retrograde Intubation Set | 2 | 868.5095 |  | N | N
Anesthesiology | BXQ | Rhinoanemometer (Measurement Of Nasal Decongestion) | 2 | 868.1800 | Rhinoanemometer. | N | N
Anesthesiology | BZO | Set, Tubing And Support, Ventilator (W Harness) | 1 | 868.5975 | Ventilator tubing. | N | N
Anesthesiology | NVO | Spacer, Direct Patient Interface | 2 | 868.5630 | Nebulizer. | N | N
Anesthesiology | PHY | Sparse Sample Pk Profile And Dosing Software | 2 | 868.1890 | Predictive pulmonary-function value calculator. | N | N
Anesthesiology | OFU | Spinal Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | OFT | Spinal Epidural Anesthesia Kit | 2 | 868.5120 | Anesthesia conduction catheter. | N | N
Anesthesiology | BZG | Spirometer, Diagnostic | 2 | 868.1840 | Diagnostic spirometer. | N | N
Anesthesiology | BZK | Spirometer, Monitoring (W/Wo Alarm) | 2 | 868.1850 | Monitoring spirometer. | N | N
Anesthesiology | BWF | Spirometer, Therapeutic (Incentive) | 2 | 868.5690 | Incentive spirometer. | N | N
Anesthesiology | CBG | Spreader, Cuff | 1 | 868.5760 | Cuff spreader. | N | N
Anesthesiology | BZS | Stethoscope Head | 1 | 868.1930 | Stethoscope head. | N | N
Anesthesiology | BZW | Stethoscope, Esophageal | 1 | 868.1910 | Esophageal stethoscope. | N | N
Anesthesiology | BZT | Stethoscope, Esophageal, With Electrical Conductors | 2 | 868.1920 | Esophageal stethoscope with electrical conductors. | N | N
Anesthesiology | BXM | Stimulator, Nerve, Ac-Powered | 2 | 868.2775 | Electrical peripheral nerve stimulator. | N | N
Anesthesiology | BXN | Stimulator, Nerve, Battery-Powered | 2 | 868.2775 | Electrical peripheral nerve stimulator. | N | N
Anesthesiology | KOI | Stimulator, Nerve, Peripheral, Electric | 2 | 868.2775 | Electrical peripheral nerve stimulator. | N | N
Anesthesiology | BRX | Stool, Anesthesia | 1 | 868.6700 | Anesthesia stool. | N | N
Anesthesiology | BTK | Strap, Head, Gas Mask | 1 | 868.5560 | Gas mask head strap. | N | N
Anesthesiology | BSR | Stylet, Tracheal Tube | 1 | 868.5790 | Tracheal tube stylet. | N | N
Anesthesiology | JAY | Support, Breathing Tube | 1 | 868.5280 | Breathing tube support. | N | N
Anesthesiology | CCX | Support, Patient Position | 1 | 868.6820 | Patient position support. | N | N
Anesthesiology | BYL | Tent, Oxygen | 1 | 868.5700 | Nonpowered oxygen tent. | N | N
Anesthesiology | BYK | Tent, Oxygen, Electrically Powered | 2 | 868.5710 | Electrically powered oxygen tent. | Y | N
Anesthesiology | FNC | Tent, Pediatric Aerosol | 2 | 868.5710 | Electrically powered oxygen tent. | N | N
Anesthesiology | OFS | Tracheal Suction Set | 1 | 868.6810 | Tracheobronchial suction catheter. | N | N
Anesthesiology | OFR | Tracheobronchial Suction Catheter Kit | 1 | 868.6810 | Tracheobronchial suction catheter. | N | N
Anesthesiology | OGV | Tracheostomy Cleaning Tray | 2 | 868.5800 | Tracheostomy tube and tube cuff. | N | N
Anesthesiology | OGW | Tracheostomy Kit | 2 | 868.5800 | Tracheostomy tube and tube cuff. | N | N
Anesthesiology | NXA | Tracheotomy Care Kit | 2 | 868.5800 | Tracheostomy tube and tube cuff. | N | N
Anesthesiology | BXP | Transducer, Gas Flow | 1 | 868.2885 | Gas flow transducer. | N | N
Anesthesiology | BXO | Transducer, Gas Pressure | 1 | 868.2900 | Gas pressure transducer. | N | N
Anesthesiology | BYR | Transducer, Gas Pressure, Differential | 1 | 868.2875 | Differential pressure transducer. | N | N
Anesthesiology | JOH | Tube Tracheostomy And Tube Cuff | 2 | 868.5800 | Tracheostomy tube and tube cuff. | Y | N
Anesthesiology | BTS | Tube, Bronchial (W/Wo Connector) | 2 | 868.5720 | Bronchial tube. | Y | N
Anesthesiology | BYM | Tube, Thorpe, Uncompensated | 1 | 868.2320 | Uncompensated thorpe tube flowmeter. | N | N
Anesthesiology | BTR | Tube, Tracheal (W/Wo Connector) | 2 | 868.5730 | Tracheal tube. | Y | N
Anesthesiology | NMA | Tube, Tracheal, Reprocessed | 2 | 868.5730 | Tracheal tube. | Y | N
Anesthesiology | CBI | Tube, Tracheal/Bronchial, Differential Ventilation (W/Wo Connector) | 2 | 868.5740 | Tracheal/bronchial differential ventilation tube. | Y | N
Anesthesiology | BTO | Tube, Tracheostomy (W/Wo Connector) | 2 | 868.5800 | Tracheostomy tube and tube cuff. | Y | N
Anesthesiology | BYX | Tubing, Pressure And Accessories | 1 | 868.5860 | Pressure tubing and accessories. | N | N
Anesthesiology | BYJ | Unit, Liquid-Oxygen, Portable | 2 | 868.5655 | Portable liquid oxygen unit. | N | N
Anesthesiology | CBP | Valve, Non-Rebreathing | 2 | 868.5870 | Nonrebreathing valve. | Y | N
Anesthesiology | JFE | Valve, Switching (Ploss) | 1 | 868.1965 | Switching valve (ploss). | N | N
Anesthesiology | CAD | Vaporizer, Anesthesia, Non-Heated | 2 | 868.5880 | Anesthetic vaporizer. | N | N
Anesthesiology | CBK | Ventilator, Continuous, Facility Use | 2 | 868.5895 | Continuous ventilator. | Y | N | Philips Resporonics
Anesthesiology | MNT | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use | 2 | 868.5895 | Continuous ventilator. | Y | N | Philips Resporonics
Anesthesiology | NQY | Ventilator, Continuous, Minimal Ventilatory Support, Home Use | 2 | 868.5895 | Continuous ventilator. | Y | N
Anesthesiology | MNS | Ventilator, Continuous, Non-Life-Supporting | 2 | 868.5895 | Continuous ventilator. | N | N | Philips Resporonics
Anesthesiology | BTM | Ventilator, Emergency, Manual (Resuscitator) | 2 | 868.5915 | Manual emergency ventilator. | Y | N
Anesthesiology | BTL | Ventilator, Emergency, Powered (Resuscitator) | 2 | 868.5925 | Powered emergency ventilator. | Y | N
Anesthesiology | BYT | Ventilator, External Body, Negative Pressure, Adult (Cuirass) | 2 | 868.5935 | External negative pressure ventilator. | Y | N
Anesthesiology | BZD | Ventilator, Non-Continuous (Respirator) | 2 | 868.5905 | Noncontinuous ventilator (IPPB). | N | N | Philips Resporonics
Anesthesiology | MNR | Ventilatory Effort Recorder | 2 | 868.2375 | Breathing frequency monitor. | N | N | Philips Resporonics
Anesthesiology | QEB | Ventilatory Electrical Impedance Tomograph | 2 | 868.1505 | Ventilatory electrical impedance tomograph. | N | N
Anesthesiology | CAM | Yoke Assembly, Medical Gas | 1 | 868.6885 | Medical gas yoke assembly. | N | N
cardiovascular | KRI | Accessory Equipment, Cardiopulmonary Bypass | 1 | 870.4200 | Cardiopulmonary bypass accessory equipment. | N | N
cardiovascular | NKW | Actuator, Syringe, For Injector, Reprocessed | 2 | 870.1670 | Syringe actuator for an injector. | N | N
cardiovascular | DQF | Actuator, Syringe, Injector Type | 2 | 870.1670 | Syringe actuator for an injector. | N | N
cardiovascular | DRW | Adaptor, Lead Switching, Electrocardiograph | 2 | 870.2350 | Electrocardiograph lead switching adaptor. | N | N
cardiovascular | DTL | Adaptor, Stopcock, Manifold, Fitting, Cardiopulmonary Bypass | 2 | 870.4290 | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | N | N
cardiovascular | QJB | Adjunct Monitor, Protein Calorie Malnutrition | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | PPW | Adjunctive Cardiovascular Status Indicator | 2 | 870.2200 | Adjunctive cardiovascular status indicator. | N | N
cardiovascular | QAQ | Adjunctive Predictive Cardiovascular Indicator | 2 | 870.2210 |  | N | N
cardiovascular | NAJ | Agents, Embolic, For Treatment Of Uterine Fibroids | 2 | 870.3300 | Vascular embolization device. | N | Y
cardiovascular | LIX | Aid, Cardiopulmonary Resuscitation | 2 | 870.5210 | Cardiopulmonary resuscitation (CPR) aid. | N | N
cardiovascular | DSJ | Alarm, Blood-Pressure | 2 | 870.1100 | Blood pressure alarm. | N | N
cardiovascular | DSM | Alarm, Leakage Current, Portable | 1 | 870.2640 | Portable leakage current alarm. | N | N
cardiovascular | DRR | Amplifier And Signal Conditioner, Biopotential | 2 | 870.2050 | Biopotential amplifier and signal conditioner. | N | N
cardiovascular | DRQ | Amplifier And Signal Conditioner, Transducer Signal | 2 | 870.2060 | Transducer signal amplifier and conditioner. | N | N
cardiovascular | MNW | Analyzer, Body Composition | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | PUH | Analyzer, Body Composition, Exempt | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | DTC | Analyzer, Pacemaker Generator Function | 2 | 870.3630 | Pacemaker generator function analyzer. | N | N
cardiovascular | KRE | Analyzer, Pacemaker Generator Function, Indirect | 2 | 870.3640 | Indirect pacemaker generator function analyzer. | N | N
cardiovascular | QEK | Angiographic Coronary Vascular Physiologic Simulation Software | 2 | 870.1415 | Coronary vascular physiologic simulation software device. | N | N
cardiovascular | OEQ | Angiography/Angioplasty Kit | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | OER | Angioscopic Valvulotome Kit | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | OED | Antimicrobial Blood Pressure Cuff | 2 | 870.1120 | Blood pressure cuff. | N | N
cardiovascular | PST | Aortic Annuloplasty Ring | 2 | 870.3800 | Annuloplasty ring. | Y | Y
cardiovascular | PNQ | Apical Closure Device | 2 | 870.4510 | Apical closure device. | N | Y
cardiovascular | DWM | Apparatus, Suction, Patient Care | 2 | 870.5050 | Patient care suction apparatus. | N | N
cardiovascular | QEZ | Aspiration Thrombectomy Catheter | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | MPC | Atrial Defibrillator | 3 | 870.5300 | DC-defibrillator (including paddles). | Y | Y
cardiovascular | MKJ | Automated External Defibrillators (Non-Wearable) | 3 | 870.5310 | Automated external defibrillator system. | N | N
cardiovascular | MPD | Auxiliary Power Supply (Ac Or Dc) For Low-Energy Dc-Defibrillator | 2 | 870.5300 | DC-defibrillator (including paddles). | Y | N
cardiovascular | MPE | Auxiliary Power Supply (Acor Dc) For External Transcutaneous Cardiac Pacemaker | 2 | 870.5550 | External transcutaneous cardiac pacemaker (noninvasive). | Y | N
cardiovascular | DSW | Bag, Polymeric Mesh, Pacemaker | 1 | 870.3650 | Pacemaker polymeric mesh bag. | N | Y
cardiovascular | DXR | Ballistocardiograph | 2 | 870.2320 | Ballistocardiograph. | N | N
cardiovascular | OZT | Balloon Aortic Valvuloplasty | 2 | 870.1255 | Balloon aortic valvuloplasty catheter. | N | N
cardiovascular | DXQ | Blood Pressure Cuff | 2 | 870.1120 | Blood pressure cuff. | N | N
cardiovascular | DSA | Cable, Transducer And Electrode, Patient, (Including Connector) | 2 | 870.2900 | Patient transducer and electrode cable (including connector). | N | N
cardiovascular | NCP | Cannula, Arterial, Cardiopulmonary Bypass (Cpb), Embolism Protection | 2 | 870.4210 | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | N | N
cardiovascular | DQR | Cannula, Catheter | 2 | 870.1300 | Catheter cannula. | N | N
cardiovascular | OJQ | Cardiac Allograft Gene Expression Profiling Test System | 2 | 862.1163 | Cardiac allograft gene expression profiling test system. | N | N
cardiovascular | OES | Cardiac Catheterization Kit | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | DQH | Cardiograph, Apex (Vibrocardiograph) | 2 | 870.2310 | Apex cardiograph (vibrocardiograph). | N | N
cardiovascular | OET | Cardioplegia Solution Administration Kit | 2 | 870.4240 | Cardiopulmonary bypass heat exchanger. | N | N
cardiovascular | OEU | Cardiopulmonary Bypass Catheter Kit | 2 | 870.4210 | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | N | N
cardiovascular | OEX | Cardiovascular Catheter Sheath Introducer Kit | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | OEZ | Cardiovascular Procedure Kit | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | OFA | Cardiovascular Surgical Instruments Tray | 1 | 870.4500 | Cardiovascular surgical instruments. | N | N
cardiovascular | PDU | Catheter For Crossing Total Occlusions | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | OFB | Catheter Guide Wire Kit | 2 | 870.1330 | Catheter guide wire. | N | N
cardiovascular | OFC | Catheter Introducer Kit | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | OFD | Catheter Introducer Kit | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | PJB | Catheter Remote Control System | 2 | 870.5700 | Steerable cardiac ablation catheter remote control system. | N | N
cardiovascular | NLI | Catheter, Angiography, Reprocessed | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | LIT | Catheter, Angioplasty, Peripheral, Transluminal | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | NVM | Catheter, Angioplasty, Peripheral, Transluminal, Dual-Balloon | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | NMM | Catheter, Angioplasty, Peripheral, Transluminal, Reprocessed | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | DWF | Catheter, Cannula And Tubing, Vascular, Cardiopulmonary Bypass | 2 | 870.4210 | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | N | N
cardiovascular | KRA | Catheter, Continuous Flush | 2 | 870.1210 | Continuous flush catheter. | N | N
cardiovascular | DRF | Catheter, Electrode Recording, Or Probe, Electrode Recording | 2 | 870.1220 | Electrode recording catheter or electrode recording probe. | N | N
cardiovascular | DXE | Catheter, Embolectomy | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | DYG | Catheter, Flow Directed | 2 | 870.1240 | Flow-directed catheter. | N | N
cardiovascular | MTD | Catheter, Intracardiac Mapping, High-Density Array | 2 | 870.1220 | Electrode recording catheter or electrode recording probe. | N | N
cardiovascular | NLG | Catheter, Intracardiac Mapping, High-Density, Reprocessed | 2 | 870.1220 | Electrode recording catheter or electrode recording probe. | N | N
cardiovascular | MJN | Catheter, Intravascular Occluding, Temporary | 2 | 870.4450 | Vascular clamp. | N | N
cardiovascular | DQO | Catheter, Intravascular, Diagnostic | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | OGZ | Catheter, Intravascular, Plaque Morphology Evaluation | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | NMB | Catheter, Oximeter, Fiber Optic, Reprocessed | 2 | 870.1230 | Fiberoptic oximeter catheter. | N | N
cardiovascular | DQE | Catheter, Oximeter, Fiber-Optic | 2 | 870.1230 | Fiberoptic oximeter catheter. | N | N
cardiovascular | DQY | Catheter, Percutaneous | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | PNO | Catheter, Percutaneous, Cutting/Scoring | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | QJP | Catheter, Percutaneous, Neurovasculature | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | MCW | Catheter, Peripheral, Atherectomy | 2 | 870.4875 | Intraluminal artery stripper. | N | N
cardiovascular | OBI | Catheter, Pressure Monitoring, Cardiac | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | NLH | Catheter, Recording, Electrode, Reprocessed | 2 | 870.1220 | Electrode recording catheter or electrode recording probe. | N | N
cardiovascular | DXF | Catheter, Septostomy | 2 | 870.5175 | Septostomy catheter. | N | N
cardiovascular | DRA | Catheter, Steerable | 2 | 870.1280 | Steerable catheter. | N | N
cardiovascular | NKS | Catheter, Steerable, Reprocessed | 2 | 870.1280 | Steerable catheter. | N | N
cardiovascular | NRY | Catheter, Thrombus Retriever | 2 | 870.1250 | Percutaneous catheter. | N | N | Balt
cardiovascular | OBJ | Catheter, Ultrasound, Intravascular | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | LOX | Catheters, Transluminal Coronary Angioplasty, Percutaneous | 2 | 870.5100 | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | N | N
cardiovascular | OFE | Central Venous Blood Pressure Kit | 2 | 870.1140 | Venous blood pressure manometer. | N | N
cardiovascular | OFF | Central Venous Catheter Tray | 2 | 870.1140 | Venous blood pressure manometer. | N | N
cardiovascular | QEM | Cerebral Oximeter | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | KRF | Charger, Pacemaker | 1 | 870.3670 | Pacemaker charger. | N | N
cardiovascular | DXC | Clamp, Vascular | 2 | 870.4450 | Vascular clamp. | N | N
cardiovascular | NMF | Clamp, Vascular, Reprocessed | 2 | 870.4450 | Vascular clamp. | N | N
cardiovascular | DSS | Clip, Vascular | 2 | 870.3250 | Vascular clip. | N | Y
cardiovascular | DST | Clip, Vena-Cava | 2 | 870.3260 | Vena cava clip. | N | Y
cardiovascular | DRM | Compressor, Cardiac, External | 2 | 870.5200 | External cardiac compressor. | N | N
cardiovascular | DSK | Computer, Blood-Pressure | 2 | 870.1110 | Blood pressure computer. | N | N
cardiovascular | DXG | Computer, Diagnostic, Pre-Programmed, Single-Function | 2 | 870.1435 | Single-function, preprogrammed diagnostic computer. | N | N
cardiovascular | DQK | Computer, Diagnostic, Programmable | 2 | 870.1425 | Programmable diagnostic computer. | N | N
cardiovascular | DTQ | Console, Heart-Lung Machine, Cardiopulmonary Bypass | 2 | 870.4220 | Cardiopulmonary bypass heart-lung machine console. | Y | N
cardiovascular | DWA | Control, Pump Speed, Cardiopulmonary Bypass | 2 | 870.4380 | Cardiopulmonary bypass pump speed control. | Y | N
cardiovascular | DWC | Controller, Temperature, Cardiopulmonary Bypass | 2 | 870.4250 | Cardiopulmonary bypass temperature controller. | Y | N
cardiovascular | QEX | Coronary Mechanical Thrombectomy With Aspiration | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | PJA | Coronary Vascular Physiologic Simulation Software | 2 | 870.1415 | Coronary vascular physiologic simulation software device. | N | N
cardiovascular | MJJ | Cpb Check Valve, Retrograde Flow, In-Line | 2 | 870.4400 | Cardiopulmonary bypass blood reservoir. | N | N
cardiovascular | PMK | Cpr Aid Device Without Feedback | 1 | 870.5210 | Cardiopulmonary resuscitation (CPR) aid. | N | N
cardiovascular | PMJ | Cpr Aid Feedback Device (No Software) | 2 | 870.5210 | Cardiopulmonary resuscitation (CPR) aid. | N | N
cardiovascular | QBE | Cranial Sound Monitor | 2 | 870.1875 | Stethoscope. | N | N
cardiovascular | OFG | Ct Biopsy Tray | 2 | 870.4075 | Endomyocardial biopsy device. | N | N
cardiovascular | DRK | Dc-Defibrillator, High Energy, (Including Paddles) | 3 | 870.5300 | DC-defibrillator (including paddles). | N | N
cardiovascular | LDD | Dc-Defibrillator, Low-Energy, (Including Paddles) | 2 | 870.5300 | DC-defibrillator (including paddles). | Y | N
cardiovascular | DTP | Defoamer, Cardiopulmonary Bypass | 2 | 870.4230 | Cardiopulmonary bypass defoamer. | Y | N
cardiovascular | DXM | Densitometer | 2 | 870.1450 | Densitometer. | N | N
cardiovascular | DSI | Detector And Alarm, Arrhythmia | 2 | 870.1025 | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | N | N
cardiovascular | KRL | Detector, Bubble, Cardiopulmonary Bypass | 2 | 870.4205 | Cardiopulmonary bypass bubble detector. | N | N
cardiovascular | DWZ | Device, Biopsy, Endomyocardial | 2 | 870.4075 | Endomyocardial biopsy device. | N | N
cardiovascular | DRN | Device, Counter-Pulsating, External | 2 | 870.5225 | External counter-pulsating device. | N | N
cardiovascular | MMX | Device, Percutaneous Retrieval | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | KRD | Device, Vascular, For Promoting Embolization | 2 | 870.3300 | Vascular embolization device. | N | Y
cardiovascular | OFH | Digital Angiography Tray | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | DRE | Dilator, Vessel, For Percutaneous Catheterization | 2 | 870.1310 | Vessel dilator for percutaneous catheterization. | N | N
cardiovascular | DWP | Dilator, Vessel, Surgical | 2 | 870.4475 | Surgical vessel dilator. | N | N
cardiovascular | DXJ | Display, Cathode-Ray Tube, Medical | 2 | 870.2450 | Medical cathode-ray tube display. | N | N
cardiovascular | PZS | Dual Lumen Ecmo Cannula | 2 | 870.4100 | Extracorporeal circuit and accessories for long-term respiratory/cardiopulmonary failure. | N | N
cardiovascular | DXK | Echocardiograph | 2 | 870.2330 | Echocardiograph. | N | N
cardiovascular | DPS | Electrocardiograph | 2 | 870.2340 | Electrocardiograph. | N | N
cardiovascular | QDA | Electrocardiograph Software For Over-The-Counter Use | 2 | 870.2345 |  | N | N
cardiovascular | MWJ | Electrocardiograph, Ambulatory (Without Analysis) | 2 | 870.2800 | Medical magnetic tape recorder. | N | N
cardiovascular | MLO | Electrocardiograph, Ambulatory, With Analysis Algorithm | 2 | 870.2800 | Medical magnetic tape recorder. | N | N
cardiovascular | DRX | Electrode, Electrocardiograph | 2 | 870.2360 | Electrocardiograph electrode. | N | N
cardiovascular | MLN | Electrode, Electrocardiograph, Multi-Function | 2 | 870.2360 | Electrocardiograph electrode. | N | N
cardiovascular | LDF | Electrode, Pacemaker, Temporary | 2 | 870.3680 | Cardiovascular permanent or temporary pacemaker electrode. | Y | N
cardiovascular | NHW | Electrode, Pacing And Cardioversion, Temporary, Epicardial | 2 | 870.3680 | Cardiovascular permanent or temporary pacemaker electrode. | Y | N
cardiovascular | MFY | Endothelial Cell Harvesting Kit | 2 | 870.4075 | Endomyocardial biopsy device. | N | N
cardiovascular | OTD | Endovascular Suturing System | 2 | 870.3460 | Endovascular Suturing System. | N | Y
cardiovascular | PLA | Esophageal Thermal Regulation And Gastric Suctioning Device | 2 | 870.5910 | Esophageal thermal regulation device. | N | N
cardiovascular | QJZ | Extracorporeal System For Long-Term Respiratory / Cardiopulmonary Failure | 2 | 870.4100 | Extracorporeal circuit and accessories for long-term respiratory/cardiopulmonary failure. | Y | N
cardiovascular | DTM | Filter, Blood, Cardiopulmonary Bypass, Arterial Line | 2 | 870.4260 | Cardiopulmonary bypass arterial line blood filter. | Y | N
cardiovascular | JOD | Filter, Blood, Cardiotomy Suction Line, Cardiopulmonary Bypass | 2 | 870.4270 | Cardiopulmonary bypass cardiotomy suction line blood filter. | N | N
cardiovascular | DTK | Filter, Intravascular, Cardiovascular | 2 | 870.3375 | Cardiovascular intravascular filter. | N | Y
cardiovascular | KRJ | Filter, Prebypass, Cardiopulmonary Bypass | 2 | 870.4280 | Cardiopulmonary prebypass filter. | N | N
cardiovascular | DPW | Flowmeter, Blood, Cardiovascular | 2 | 870.2100 | Cardiovascular blood flowmeter. | N | N
cardiovascular | DTX | Gas Control Unit, Cardiopulmonary Bypass | 2 | 870.4300 | Cardiopulmonary bypass gas control unit. | Y | N
cardiovascular | DXS | Gauge, Pressure, Coronary, Cardiopulmonary Bypass | 2 | 870.4310 | Cardiopulmonary bypass coronary pressure gauge. | N | N
cardiovascular | JOR | Generator, Pulsatile Flow, Cardiopulmonary Bypass | 3 | 870.4320 | Cardiopulmonary bypass pulsatile flow generator. | N | N
cardiovascular | JOQ | Generator, Pulse, Pacemaker, External Programmable (For Electrophysiological Studies Only) | 2 | 870.1750 | External programmable pacemaker pulse generator. | Y | N
cardiovascular | MZY | Graft, Vascular, Stainless Steel Tunneler | 2 | 870.3450 | Vascular graft prosthesis. | N | N
cardiovascular | MAL | Graft, Vascular, Synthetic/Biologic Composite | 2 | 870.3450 | Vascular graft prosthesis. | N | Y
cardiovascular | MOF | Guide, Wire, Catheter, Neurovasculature | 2 | 870.1330 | Catheter guide wire. | N | N
cardiovascular | NKQ | Guidewire, Catheter, Reprocessed | 2 | 870.1330 | Catheter guide wire. | N | N
cardiovascular | DTR | Heat-Exchanger, Cardiopulmonary Bypass | 2 | 870.4240 | Cardiopulmonary bypass heat exchanger. | Y | N
cardiovascular | DTJ | Holder, Heart-Valve, Prosthesis | 1 | 870.3935 | Prosthetic heart valve holder. | N | N
cardiovascular | NZE | Hyperthermia Monitor | 2 | 870.5900 | Thermal regulating system. | N | N
cardiovascular | QAF | Impedance Plethysmograph, Cranial | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | DXY | Implantable Pacemaker Pulse-Generator | 3 | 870.3610 | Implantable pacemaker pulse generator. | N | Y | St Jude Medical
cardiovascular | DXT | Injector And Syringe, Angiographic | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | NKU | Injector And Syringe, Angiographic, Balloon Inflation, Reprocessed | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | NKT | Injector And Syringe, Angiographic, Reprocessed | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | IZQ | Injector, Contrast Medium, Automatic | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | NKP | Injector, Contrast Medium, Automatic, Reprocessed | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | DXL | Injector, Indicator | 2 | 870.1660 | Indicator injector. | N | N
cardiovascular | DWS | Instruments, Surgical, Cardiovascular | 1 | 870.4500 | Cardiovascular surgical instruments. | N | N
cardiovascular | PSQ | Intracardiac Patch Or Pledget, Biologically Derived | 2 | 870.3470 | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | N | Y
cardiovascular | QFJ | Intravascular Bleed Monitor | 2 | 870.1345 |  | N | N
cardiovascular | DYB | Introducer, Catheter | 2 | 870.1340 | Catheter introducer. | N | N | Endologix(delivery)
cardiovascular | DQP | Kit, Balloon Repair, Catheter | 3 | 870.1350 | Catheter balloon repair kit. | N | N
cardiovascular | OFI | Laser Blood Flow Kit | 2 | 870.2120 | Extravascular blood flow probe. | N | N
cardiovascular | OFJ | Lead Introducer Kit | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | PNJ | Leadless Pacemaker | 3 | 870.3610 | Implantable pacemaker pulse generator. | Y | Y
cardiovascular | PFY | Low Patient Acuity Monitoring System | 2 | 870.2300 | Cardiac monitor (including cardiotachometer and rate alarm). | N | N
cardiovascular | OCR | Lung Sound Monitor | 2 | 870.1875 | Stethoscope. | N | N
cardiovascular | DTG | Magnet, Test, Pacemaker | 1 | 870.3690 | Pacemaker test magnet. | N | N
cardiovascular | KRK | Manometer, Blood-Pressure, Venous | 2 | 870.1140 | Venous blood pressure manometer. | N | N
cardiovascular | MAB | Marker, Cardiopulmonary Bypass (Vein Marker) | 2 | 870.3450 | Vascular graft prosthesis. | N | N
cardiovascular | KFJ | Materials, Repair Or Replacement, Pacemaker | 3 | 870.3710 | Pacemaker repair or replacement material. | N | Y
cardiovascular | QEY | Mechanical Thrombolysis Catheter | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | DQG | Microsphere, Trace | 3 | 870.1360 | Trace microsphere. | N | Y
cardiovascular | DTW | Monitor And/Or Control, Level Sensing, Cardiopulmonary Bypass | 2 | 870.4340 | Cardiopulmonary bypass level sensing monitor and/or control. | N | N
cardiovascular | DRY | Monitor, Blood-Gas, On-Line, Cardiopulmonary Bypass | 2 | 870.4330 | Cardiopulmonary bypass on-line blood gas monitor. | N | N
cardiovascular | DRT | Monitor, Cardiac (Incl. Cardiotachometer & Rate Alarm) | 2 | 870.2300 | Cardiac monitor (including cardiotachometer and rate alarm). | N | N
cardiovascular | OBH | Monitor, Extracellular Fluid, Lymphedema, Extremity | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | DRI | Monitor, Line Isolation | 1 | 870.2620 | Line isolation monitor. | N | N
cardiovascular | MWI | Monitor, Physiological, Patient (Without Arrhythmia Detection Or Alarms) | 2 | 870.2300 | Cardiac monitor (including cardiotachometer and rate alarm). | N | N
cardiovascular | MHX | Monitor, Physiological, Patient(With Arrhythmia Detection Or Alarms) | 2 | 870.1025 | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | N | N
cardiovascular | MLC | Monitor, St Segment | 2 | 870.2340 | Electrocardiograph. | N | N
cardiovascular | MLD | Monitor, St Segment With Alarm | 2 | 870.1025 | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | N | N
cardiovascular | PLB | Multivariate Vital Signs Index | 2 | 870.2300 | Cardiac monitor (including cardiotachometer and rate alarm). | N | N
cardiovascular | PFA | Non-Invasive Central Venous Manometer | 2 | 870.1140 | Venous blood pressure manometer. | N | N
cardiovascular | DQT | Occluder, Catheter Tip | 2 | 870.1370 | Catheter tip occluder. | N | N
cardiovascular | ORD | Optical Coherence Tomography, Intravascular Catheter | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | DRZ | Oscillometer | 2 | 870.2675 | Oscillometer. | N | N
cardiovascular | NSA | Over-The-Counter Automated External Defibrillator | 3 | 870.5310 | Automated external defibrillator system. | Y | N
cardiovascular | DQA | Oximeter | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | DPZ | Oximeter, Ear | 2 | 870.2710 | Ear oximeter. | N | N
cardiovascular | OCH | Oximeter, Infrared, Sporting, Aviation | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | NLF | Oximeter, Reprocessed | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | MUD | Oximeter, Tissue Saturation | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | NMD | Oximeter, Tissue Saturation, Reprocessed | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | PGJ | Oximeter, Wellness | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | DTZ | Oxygenator, Cardiopulmonary Bypass | 2 | 870.4350 | Cardiopulmonary bypass oxygenator. | Y | N
cardiovascular | BYS | Oxygenator, Long Term Support Greater Than 6 Hours | 2 | 870.4100 | Extracorporeal circuit and accessories for long-term respiratory/cardiopulmonary failure. | Y | N
cardiovascular | DSZ | Pacemaker Battery | 3 | 870.3610 | Implantable pacemaker pulse generator. | N | Y
cardiovascular | DTD | Pacemaker Lead Adaptor | 2 | 870.3620 | Pacemaker lead adaptor. | N | Y
cardiovascular | DRO | Pacemaker, Cardiac, External Transcutaneous (Non-Invasive) | 2 | 870.5550 | External transcutaneous cardiac pacemaker (noninvasive). | Y | N
cardiovascular | QAB | Pacing System Analyzer | 2 | 870.3605 | Pacing system analyzer. | Y | N
cardiovascular | DXZ | Patch, Pledget And Intracardiac, Petp, Ptfe, Polypropylene | 2 | 870.3470 | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | N | Y
cardiovascular | QLJ | Patient Monitor With Independent Remote Display And Operation | 2 | 870.1025 | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | N | N
cardiovascular | PCK | Pediatric Ventricular Assist Device | HDE | 870.3545 | Ventricular bypass (assist) device. | Y | Y
cardiovascular | OFK | Percutaneous Atrial Catheter Kit | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | PQK | Percutaneous Catheter For Creation Of An Arteriovenous Fistula For Hemodialysis Access | 2 | 870.1252 |  | N | N
cardiovascular | PPN | Percutaneous Catheter, Ultrasound | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | OFL | Percutaneous Sheath Introducer Kit | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | OMH | Pericardial Patch To Facilitate Revision Surgeries | 2 | 870.3470 | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Y | Y
cardiovascular | QEW | Peripheral Mechanical Thrombectomy With Aspiration | 2 | 870.5150 | Embolectomy catheter. | N | N
cardiovascular | DTB | Permanent Pacemaker Electrode | 3 | 870.3680 | Cardiovascular permanent or temporary pacemaker electrode. | N | Y
cardiovascular | DQB | Phlebograph, Impedance | 2 | 870.2750 | Impedance phlebograph. | N | N
cardiovascular | OFM | Phlebotomy Blood Collection Kit | 2 | 870.2750 | Impedance phlebograph. | N | N
cardiovascular | DQC | Phonocardiograph | 1 | 870.2390 | Phonocardiograph. | N | N
cardiovascular | QDB | Photoplethysmograph Analysis Software For Over-The-Counter Use | 2 | 870.2790 |  | N | N
cardiovascular | OAC | Plethysmograph, Air Displacement For Body Composition Analysis | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | DSB | Plethysmograph, Impedance | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | JOM | Plethysmograph, Photoelectric, Pneumatic Or Hydraulic | 2 | 870.2780 | Hydraulic, pneumatic, or photoelectric plethysmographs. | N | N
cardiovascular | DPT | Probe, Blood-Flow, Extravascular | 2 | 870.2120 | Extravascular blood flow probe. | N | N
cardiovascular | LXN | Probe, Test, Heart-Valve | 1 | 870.4500 | Cardiovascular surgical instruments. | N | N
cardiovascular | KRB | Probe, Thermodilution | 2 | 870.1915 | Thermodilution probe. | N | N
cardiovascular | KRG | Programmer, Pacemaker | 3 | 870.3700 | Pacemaker programmers. | N | N
cardiovascular | DSY | Prosthesis, Vascular Graft, Of 6mm And Greater Diameter | 2 | 870.3450 | Vascular graft prosthesis. | N | Y
cardiovascular | DYF | Prosthesis, Vascular Graft, Of Less Then 6mm Diameter | 2 | 870.3450 | Vascular graft prosthesis. | N | Y
cardiovascular | OMZ | Pulmonary (Pulmonic) Valvuloplasty Catheters/Percutaneous Valvuloplasty Catheter | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | PAL | Pulmonic Replacement Heart Valve | HDE | 870.3925 | Replacement heart valve. | Y | Y
cardiovascular | OVJ | Pulse Generator, External Pacemaker, Dual Chamber | 2 | 870.3600 | External pacemaker pulse generator. | Y | N
cardiovascular | QLS | Pulse Oximeter For Low-Acuity Settings | 2 | 870.2700 | Oximeter. | N | N
cardiovascular | DTE | Pulse-Generator, Pacemaker, External | 3 | 870.3600 | External pacemaker pulse generator. | N | N
cardiovascular | KFM | Pump, Blood, Cardiopulmonary Bypass, Non-Roller Type | 2 | 870.4360 | Nonroller-type blood pump. | N | N
cardiovascular | DWB | Pump, Blood, Cardiopulmonary Bypass, Roller Type | 2 | 870.4370 | Roller-type cardiopulmonary bypass blood pump. | Y | N
cardiovascular | MXD | Recorder, Event, Implantable Cardiac, (With Arrhythmia Detection) | 2 | 870.1025 | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | N | Y
cardiovascular | MXC | Recorder, Event, Implantable Cardiac, (Without Arrhythmia Detection) | 2 | 870.2800 | Medical magnetic tape recorder. | N | Y
cardiovascular | DSH | Recorder, Magnetic Tape, Medical | 2 | 870.2800 | Medical magnetic tape recorder. | N | N
cardiovascular | DSF | Recorder, Paper Chart | 1 | 870.2810 | Paper chart recorder. | N | N
cardiovascular | DYE | Replacement Heart-Valve | 3 | 870.3925 | Replacement heart valve. | N | Y
cardiovascular | NPP | Reprocessed Blood Pressure Cuff | 2 | 870.1120 | Blood pressure cuff. | N | N
cardiovascular | PNE | Reprocessed Catheter Introducer | 2 | 870.1340 | Catheter introducer. | N | N
cardiovascular | OWQ | Reprocessed Intravascular Ultrasound Catheter | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | DTN | Reservoir, Blood, Cardiopulmonary Bypass | 2 | 870.4400 | Cardiopulmonary bypass blood reservoir. | N | N
cardiovascular | PTN | Reservoir, Blood, Cardiopulmonary Bypass, Exempt | 2 | 870.4400 | Cardiopulmonary bypass blood reservoir. | N | N
cardiovascular | QJH | Reverse Central Venous Recanalization System | 2 | 870.1342 |  | N | N
cardiovascular | KRH | Ring, Annuloplasty | 2 | 870.3800 | Annuloplasty ring. | N | Y
cardiovascular | MOP | Rotator, Prosthetic Heart Valve | 1 | 870.3935 | Prosthetic heart valve holder. | N | N
cardiovascular | DTY | Sensor, Blood-Gas, In-Line, Cardiopulmonary Bypass | 2 | 870.4410 | Cardiopulmonary bypass in-line blood gas sensor. | N | N
cardiovascular | NQH | Sensor, Pressure, Aneurysm, Implantable | 2 | 870.2855 | Implantable Intra-aneurysm Pressure Measurement System. | N | Y
cardiovascular | PNS | Short-Term Intravascular Filter Catheter | 2 | 870.3375 | Cardiovascular intravascular filter. | N | Y
cardiovascular | OEY | Single Lead Over-The-Counter Electrocardiograph | 2 | 870.2340 | Electrocardiograph. | N | N
cardiovascular | DTI | Sizer, Heart-Valve, Prosthesis | 1 | 870.3945 | Prosthetic heart valve sizer. | N | N
cardiovascular | PPS | Sleeve, Head And Neck, Compressible | 2 | 870.5800 | Compressible limb sleeve. | N | N
cardiovascular | JOW | Sleeve, Limb, Compressible | 2 | 870.5800 | Compressible limb sleeve. | N | N
cardiovascular | MWS | Stabilizer, Heart | 1 | 870.4500 | Cardiovascular surgical instruments. | N | N
cardiovascular | NQG | Stabilizer, Heart, Non-Compression, Reprocessed | 1 | 870.4500 | Cardiovascular surgical instruments. | N | N
cardiovascular | DQD | Stethoscope, Electronic | 2 | 870.1875 | Stethoscope. | N | N
cardiovascular | LDE | Stethoscope, Manual | 1 | 870.1875 | Stethoscope. | N | N
cardiovascular | DSR | Stimulator, Carotid Sinus Nerve | 3 | 870.3850 | Carotid sinus nerve stimulator. | N | Y
cardiovascular | DWX | Stripper, Artery, Intraluminal | 2 | 870.4875 | Intraluminal artery stripper. | N | N
cardiovascular | DWQ | Stripper, Vein, External | 2 | 870.4885 | External vein stripper. | N | N
cardiovascular | NLJ | Stripper, Vein, External, Reprocessed | 2 | 870.4885 | External vein stripper. | N | N
cardiovascular | DRB | Stylet, Catheter | 2 | 870.1380 | Catheter stylet. | N | N
cardiovascular | DTS | Sucker, Cardiotomy Return, Cardiopulmonary Bypass | 2 | 870.4420 | Cardiopulmonary bypass cardiotomy return sucker. | N | N
cardiovascular | DWD | Suction Control, Intracardiac, Cardiopulmonary Bypass | 2 | 870.4430 | Cardiopulmonary bypass intracardiac suction control. | N | N
cardiovascular | MAV | Syringe, Balloon Inflation | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | PTM | Syringe, Balloon Inflation, Exempt | 2 | 870.1650 | Angiographic injector and syringe. | N | N
cardiovascular | DSP | System, Balloon, Intra-Aortic And Control | 2 | 870.3535 | Intra-aortic balloon and control system. | N | N
cardiovascular | NKO | System, Balloon, Intra-Aortic And Control, Reprocessed | 3 | 870.3535 | Intra-aortic balloon and control system. | N | N
cardiovascular | NKR | System, Catheter Control, Reprocessed | 2 | 870.1290 | Steerable catheter control system. | N | N
cardiovascular | DXX | System, Catheter Control, Steerable | 2 | 870.1290 | Steerable catheter control system. | N | N
cardiovascular | NDQ | System, Catheter Or Guidewire, Steerable (Magnetic) | 2 | 870.1290 | Steerable catheter control system. | N | N
cardiovascular | NCX | System, Hypothermia, Intravenous, Cooling | 2 | 870.5900 | Thermal regulating system. | N | N
cardiovascular | DXN | System, Measurement, Blood-Pressure, Non-Invasive | 2 | 870.1130 | Noninvasive blood pressure measurement system. | N | N
cardiovascular | MSX | System, Network And Communication, Physiological Monitors | 2 | 870.2300 | Cardiac monitor (including cardiotachometer and rate alarm). | Y | N
cardiovascular | DXW | System, Phonocatheter, Intracavitary | 2 | 870.1270 | Intracavitary phonocatheter system. | N | N
cardiovascular | DRJ | System, Signal Isolation | 1 | 870.2600 | Signal isolation system. | N | N
cardiovascular | DWJ | System, Thermal Regulating | 2 | 870.5900 | Thermal regulating system. | N | N
cardiovascular | MIH | System, Endovascular Graft, Aortic Aneurysm Treatment | 3 |  |  | N | Y | Endologix (stent_graft)
cardiovascular | PBL | Temporary Cardiac Support Blood Pump | 3 | 870.4360 | Nonroller-type blood pump. | Y | N
cardiovascular | NTE | Temporary Carotid Catheter For Embolic Capture | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | PUM | Temporary Catheter, Embolic Protection, Transcatheter Intracardiac Procedures | 2 | 870.1251 | Temporary catheter for embolic protection during transcatheter intracardiac procedures. | N | N
cardiovascular | NFA | Temporary Coronary Saphenous Vein Bypass Graft For Embolic Protection | 2 | 870.1250 | Percutaneous catheter. | N | N
cardiovascular | PYX | Temporary Non-Roller Type Right Heart Support Blood Pump | 3 | 870.4360 | Nonroller-type blood pump. | Y | N
cardiovascular | DRL | Tester, Defibrillator | 2 | 870.5325 | Defibrillator tester. | N | N
cardiovascular | KRC | Tester, Electrode, Surface, Electrocardiographic | 2 | 870.2370 | Electrocardiograph surface electrode tester. | N | N
cardiovascular | DTA | Tester, Pacemaker Electrode Function | 2 | 870.3720 | Pacemaker electrode function tester. | N | N
cardiovascular | DTF | Tools, Pacemaker Service | 1 | 870.3730 | Pacemaker service tools. | N | N
cardiovascular | DRP | Tourniquet, Automatic Rotating | 2 | 870.5925 | Automatic rotating tourniquet. | N | N
cardiovascular | JON | Transducer, Apex Cardiographic | 2 | 870.2840 | Apex cardiographic transducer. | N | N
cardiovascular | DRS | Transducer, Blood-Pressure, Extravascular | 2 | 870.2850 | Extravascular blood pressure transducer. | N | N
cardiovascular | JOO | Transducer, Heart Sound | 2 | 870.2860 | Heart sound transducer. | N | N
cardiovascular | DXO | Transducer, Pressure, Catheter Tip | 2 | 870.2870 | Catheter tip pressure transducer. | N | N
cardiovascular | JOP | Transducer, Ultrasonic | 2 | 870.2880 | Ultrasonic transducer. | N | N
cardiovascular | DXP | Transducer, Vessel Occlusion | 2 | 870.2890 | Vessel occlusion transducer. | N | N
cardiovascular | DXH | Transmitters And Receivers, Electrocardiograph, Telephone | 2 | 870.2920 | Telephone electrocardiograph transmitter and receiver. | N | N
cardiovascular | DRG | Transmitters And Receivers, Physiological Signal, Radiofrequency | 2 | 870.2910 | Radiofrequency physiological signal transmitter and receiver. | N | N
cardiovascular | DRC | Trocar | 2 | 870.1390 | Trocar. | N | N
cardiovascular | NMK | Trocar, Reprocessed | 2 | 870.1390 | Trocar. | N | N
cardiovascular | LHX | Trousers, Anti-Shock | 2 | 870.5800 | Compressible limb sleeve. | N | N
cardiovascular | DWE | Tubing, Pump, Cardiopulmonary Bypass | 2 | 870.4390 | Cardiopulmonary bypass pump tubing. | N | N
cardiovascular | OMV | Ultrasonic Body Composition Analyzer | 2 | 870.2770 | Impedance plethysmograph. | N | N
cardiovascular | NUI | Ultrasound, Infusion, System | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | MNJ | Valve, Pressure Relief, Cardiopulmonary Bypass | 2 | 870.4400 | Cardiopulmonary bypass blood reservoir. | N | N
cardiovascular | MGZ | Valvulotome | 2 | 870.4885 | External vein stripper. | N | N
cardiovascular | DYC | Vectorcardiograph | 2 | 870.2400 | Vectorcardiograph. | N | N
cardiovascular | OKR | Ventricular Bypass (Assist) Device | 3 | 870.3545 | Ventricular bypass (assist) device. | N | N
cardiovascular | OMR | Vessel Guard Or Cover | 2 | 870.3470 | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | N | Y
cardiovascular | MFX | Vessel Guard Or Cover, To Facilitate Revision Surgeries | 2 | 870.3470 | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | N | N
cardiovascular | OFN | Winged Intravenous Catheterization Kit | 2 | 870.1200 | Diagnostic intravascular catheter. | N | N
cardiovascular | DQX | Wire, Guide, Catheter | 2 | 870.1330 | Catheter guide wire. | N | N
cardiovascular | PTL | Wire, Guide, Catheter, Exempt | 2 | 870.1330 | Catheter guide wire. | N | N
cardiovascular | DQI | Withdrawal/Infusion Pump | 2 | 870.1800 | Withdrawal-infusion pump. | N | N
cardiovascular | NIK | SENSOLOG/DIALOG/REGENCY FAMILY OF PACEMAKERS | 3 |  | ICD | N | Y | St Jude Medical
cardiovascular | LOZ | Artificial Heart | 3 |  | TEMPORARY CARDIO WEST TOTAL ARTIFICIAL HEART (TAH-T) | Y | Y | Syncardia
Dental | KHR | Absorber, Saliva, Paper | 1 | 872.6050 | Saliva absorber. | N | N
Dental | NHA | Abutment, Implant, Dental, Endosseous | 2 | 872.3630 | Endosseous dental implant abutment. | N | Y
Dental | NDP | Accessories, Implant, Dental, Endosseous | 1 | 872.3980 | Endosseous dental implant accessories. | N | N
Dental | EIF | Accessories, Retractor, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EBZ | Activator, Ultraviolet, For Polymerization | 2 | 872.6070 | Ultraviolet activator for polymerization. | N | N
Dental | PZY | Additively Manufactured, Preformed, Resin Denture Tooth | 2 | 872.3590 | Preformed plastic denture tooth. | N | N
Dental | DYH | Adhesive, Bracket And Tooth Conditioner, Resin | 2 | 872.3750 | Bracket adhesive resin and tooth conditioner. | N | N
Dental | KOM | Adhesive, Denture, Acacia And Karaya With Sodium Borate | 1 | 872.3400 | Karaya and sodium borate with or without acacia denture adhesive. | N | N
Dental | MMU | Adhesive, Denture, Acacia And Karaya With Sodium Borate > 12% By Weight | 3 | 872.3400 | Karaya and sodium borate with or without acacia denture adhesive. | N | N
Dental | KOL | Adhesive, Denture, Carboxymethylcellulose Sodium (32%) And Ethylene-Oxide Homopolymer | 1 | 872.3410 | Ethylene oxide homopolymer and/or carboxymethylcellulose sodium denture adhesive. | N | N
Dental | KOQ | Adhesive, Denture, Carboxymethylcellulose Sodium (40-100&) | 1 | 872.3410 | Ethylene oxide homopolymer and/or carboxymethylcellulose sodium denture adhesive. | N | N
Dental | KXW | Adhesive, Denture, Carboxymethylcellulose Sodium (49%) And Ethylene-Oxide Homopolymer | 1 | 872.3410 | Ethylene oxide homopolymer and/or carboxymethylcellulose sodium denture adhesive. | N | N
Dental | KOS | Adhesive, Denture, Carboxymethylcellulose Sodium And Cationic Polyacrylamide Polymer | 3 | 872.3420 | Carboxymethylcellulose sodium and cationic polyacrylamide polymer denture adhesive. | N | N
Dental | KOP | Adhesive, Denture, Karaya | 1 | 872.3450 | Ethylene oxide homopolymer and/or karaya denture adhesive. | N | N
Dental | KXX | Adhesive, Denture, Karaya And Ethylene-Oxide Homopolymer | 1 | 872.3450 | Ethylene oxide homopolymer and/or karaya denture adhesive. | N | N
Dental | KOR | Adhesive, Denture, Karaya With Sodium Borate | 3 | 872.3400 | Karaya and sodium borate with or without acacia denture adhesive. | N | N
Dental | KON | Adhesive, Denture, Polyacrylamide Polymer (Modified Cationic) | 3 | 872.3480 | Polyacrylamide polymer (modified cationic) denture adhesive. | N | N
Dental | KOO | Adhesive, Denture, Polyvinyl Methylether Maleic Acid Calcium-Sodium Double Salt | 1 | 872.3490 | Carboxymethylcellulose sodium and/or polyvinylmethylether maleic acid calcium-sodium double salt denture adhesive. | N | N
Dental | EJR | Agent, Polishing, Abrasive, Oral Cavity | 1 | 872.6030 | Oral cavity abrasive polishing agent. | N | N
Dental | KLE | Agent, Tooth Bonding, Resin | 2 | 872.3200 | Resin tooth bonding agent. | N | N
Dental | KOJ | Airbrush | 2 | 872.6080 | Airbrush. | N | N
Dental | EHA | Aligner, Beam, X-Ray | 1 | 872.1820 | Dental x-ray exposure alignment device. | N | N
Dental | ECQ | Aligner, Bracket, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NXC | Aligner, Sequential | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | EJJ | Alloy, Amalgam | 2 | 872.3070 | Dental amalgam, mercury, and amalgam alloy. | N | N
Dental | EJT | Alloy, Gold-Based Noble Metal | 2 | 872.3060 | Noble metal alloy. | N | N
Dental | EJH | Alloy, Metal, Base | 2 | 872.3710 | Base metal alloy. | N | N
Dental | EJS | Alloy, Other Noble Metal | 2 | 872.3060 | Noble metal alloy. | N | N
Dental | EFD | Amalgamator, Dental, Ac-Powered | 1 | 872.3100 | Dental amalgamator. | N | N
Dental | EJX | Anchor, Preformed | 1 | 872.3130 | Preformed anchor. | N | N
Dental | EIT | Applicator, Rapid Wax, Dental | 2 | 872.3060 | Noble metal alloy. | N | N
Dental | KXR | Applicator, Resin | 1 | 872.3140 | Resin applicator. | N | N
Dental | EJP | Articulators | 1 | 872.3150 | Articulator. | N | N
Dental | EGG | Attachment, Precision, All | 1 | 872.3165 | Precision attachment. | N | N
Dental | ECI | Band, Elastic, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | DYO | Band, Material, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | ECM | Band, Preformed, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EHO | Bar, Preformed | 1 | 872.3165 | Precision attachment. | N | N
Dental | NPL | Barrier, Animal Source, Intraoral | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | NPK | Barrier, Synthetic, Intraoral | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | NQA | Biologic Material, Dental | 3 | 872.3930 | Bone grafting material. | N | Y
Dental | NRQ | Blade-Form Endosseous Dental Implant | 2 | 872.3640 | Endosseous dental implant. | N | Y
Dental | NPM | Bone Grafting Material, Animal Source | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | NPZ | Bone Grafting Material, Dental, With Biologic Component | 3 | 872.3930 | Bone grafting material. | N | Y
Dental | LPK | Bone Grafting Material, For Dental Bone Repair | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | NUN | Bone Grafting Material, Human Source | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | LYC | Bone Grafting Material, Synthetic | 2 | 872.3930 | Bone grafting material. | N | Y
Dental | NJM | Bracket, Ceramic, Orthodontic | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | EJF | Bracket, Metal, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | NQS | Bracket, Metal, Orthodontic, Reprocessed | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | DYW | Bracket, Plastic, Orthodontic | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | NLC | Bracket, Plastic, Orthodontic, Reprocessed | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | EKW | Broach, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EJL | Bur, Dental | 1 | 872.3240 | Dental bur. | N | N
Dental | NME | Bur, Diamond Coated, Reprocessed | 1 | 872.3240 | Dental bur. | N | N
Dental | EKJ | Burnisher, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DZS | Capsule, Dental, Amalgam | 1 | 872.3110 | Dental amalgam capsule. | N | N
Dental | KOT | Carboxymethylcellulose Sodium Or Polyvinyl Methylether Maleic Acid Calcium-Sodium | 1 | 872.3490 | Carboxymethylcellulose sodium and/or polyvinylmethylether maleic acid calcium-sodium double salt denture adhesive. | N | N
Dental | NTK | Caries Detector, Laser Light, Transmission | 2 | 872.1745 | Laser fluorescence caries detection device. | N | N
Dental | EKI | Carrier, Amalgam, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKH | Carver, Dental Amalgam, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EIK | Carver, Wax, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMA | Cement, Dental | 2 | 872.3275 | Dental cement. | N | N
Dental | NEA | Cement, Ear, Nose And Throat | 2 | 872.3275 | Dental cement. | N | Y
Dental | EAG | Cephalometer | 2 | 872.1830 | Cephalometer. | N | N
Dental | KLC | Chair, Dental, With Operative Unit | 1 | 872.6250 | Dental chair and accessories. | N | N
Dental | NRU | Chair, Dental, Without Operative Unit | 1 | 872.6250 | Dental chair and accessories. | N | N
Dental | EML | Chisel, Bone, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMM | Chisel, Osteotome, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EEF | Clamp, Rubber Dam | 1 | 872.6300 | Rubber dam and accessories. | N | N
Dental | ECN | Clamp, Wire, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EHP | Clasp, Preformed | 1 | 872.3285 | Preformed clasp. | N | N
Dental | EJW | Clasp, Wire | 1 | 872.3285 | Preformed clasp. | N | N
Dental | JER | Cleaner, Denture, Mechanical | 1 | 872.3530 | Mechanical denture cleaner. | N | N
Dental | EFT | Cleanser, Denture, Over The Counter | 1 | 872.3520 | OTC denture cleanser. | N | N
Dental | NUX | Cleanser, Denture, Prescription | 1 | 872.3520 | OTC denture cleanser. | N | N
Dental | QCJ | Closed Loop Auto Titration Device For Oral Appliances | 2 | 872.5571 | Auto titration device for oral appliances. | N | N
Dental | EBE | Coating, Denture Hydrophilic, Resin | 2 | 872.3300 | Hydrophilic resin coating for dentures. | N | N
Dental | EBD | Coating, Filling Material, Resin | 2 | 872.3310 | Coating material for resin fillings. | N | N
Dental | EHB | Collimator, X-Ray | 1 | 872.1840 | Dental x-ray position indicating device. | N | N
Dental | QGO | Composite Resin Warmer | 1 | 872.6100 | Anesthetic warmer. | N | N
Dental | EKG | Condenser, Amalgam And Foil, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EAH | Cone, Radiographic, Lead-Lined | 1 | 872.1850 | Lead-lined position indicator. | N | N
Dental | EBW | Controller, Foot, Handpiece And Cord | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | EFN | Cotton, Roll | 1 | 872.6050 | Saliva absorber. | N | N
Dental | POW | Crown And Bridge, Long-Term Temporary, Resin | 2 | 872.3770 | Temporary crown and bridge resin. | N | N
Dental | EBG | Crown And Bridge, Temporary, Resin | 2 | 872.3770 | Temporary crown and bridge resin. | N | N
Dental | ELZ | Crown, Preformed | 1 | 872.3330 | Preformed crown. | N | N
Dental | EHK | Cup, Prophylaxis | 1 | 872.6290 | Prophylaxis cup. | N | N
Dental | EKT | Curette, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKE | Curette, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMS | Curette, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMK | Curette, Surgical, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EHS | Cushion, Denture, Over The Counter | 2 | 872.3540 | OTC denture cushion or pad. | N | N
Dental | NKJ | Cushion, Pad, Denture, Wax Impregnated Cotton, Over The Counter | 1 | 872.3540 | OTC denture cushion or pad. | N | N
Dental | ELO | Cusp, Gold And Stainless Steel | 1 | 872.3350 | Gold or stainless steel cusp. | N | N
Dental | EHQ | Cusp, Preformed | 1 | 872.3360 | Preformed cusp. | N | N
Dental | EIE | Dam, Rubber | 1 | 872.6300 | Rubber dam and accessories. | N | N
Dental | PNP | Dental Abutment Design Software For Dental Laboratory | 2 | 872.3630 | Endosseous dental implant abutment. | N | Y
Dental | OIV | Dental Amalgam | 2 | 872.3070 | Dental amalgam, mercury, and amalgam alloy. | N | N
Dental | MZW | Dental Cement W/Out Zinc-Oxide Eugenol As An Ulcer Covering For Pain Relief | 2 | 872.3275 | Dental cement. | N | N
Dental | NXZ | Dental Hygiene Kit | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | OFY | Dental Implant Surgical Tray | 1 | 872.3980 | Endosseous dental implant accessories. | N | N
Dental | NYC | Dental Prophylaxis Kit | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | PLV | Dental Stereotaxic Instrument | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | EKO | Denture Preformed (Partially Prefabricated Denture) | 2 | 872.3600 | Partially fabricated denture kit. | N | N
Dental | EBO | Denture Repair Kit | 2 | 872.3570 | OTC denture repair kit. | N | N
Dental | ELM | Denture, Plastic, Teeth | 2 | 872.3590 | Preformed plastic denture tooth. | N | N
Dental | EAQ | Detector, Ultraviolet | 2 | 872.6350 | Ultraviolet detector. | N | N
Dental | LRK | Device, Anti-Snoring | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | LFC | Device, Caries Detection | 2 | 872.1740 | Caries detection device. | N | N
Dental | NFP | Device, Dental Sonography, For Diagnosis Of Tmj / Mpd Disorders | 2 | 872.2050 | Dental sonography device. | N | N
Dental | NFQ | Device, Dental Sonography, For Monitoring Jaw Sounds | 1 | 872.2050 | Dental sonography device. | N | N
Dental | MVH | Device, Detection, Sulfide | 2 | 872.1870 | Sulfide detection device. | N | N
Dental | LQZ | Device, Jaw Repositioning | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | NFR | Device, Jaw Tracking, For Diagnosis Of Tmj / Mpd Disorders | 2 | 872.2060 | Jaw tracking device. | N | N
Dental | NFS | Device, Jaw Tracking, For Monitoring Jaw Positions | 1 | 872.2060 | Jaw tracking device. | N | N
Dental | NXV | Diagnostic Light, Soft Tissue Detector | 2 | 872.6350 | Ultraviolet detector. | N | N
Dental | PHR | Diammine Silver Fluoride Dental Hypersensitivity Varnish | 2 | 872.3260 | Cavity varnish. | N | N
Dental | EHJ | Disk, Abrasive | 1 | 872.6010 | Abrasive device and accessories. | N | N
Dental | EHE | Dispenser, Mercury And/Or Alloy | 1 | 872.3080 | Mercury and alloy dispenser. | N | N
Dental | DZI | Drill, Bone, Powered | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | DZA | Drill, Dental, Intraoral | 1 | 872.4130 | Intraoral dental drill. | N | N
Dental | ECT | Driver, Band, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DZJ | Driver, Wire, And Bone Drill, Manual | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | EMJ | Elevator, Surgical, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NYD | Endodontic Kit | 2 | 872.3820 | Root canal filling resin. | N | N
Dental | MAU | Eraser, Dental Stain | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | EHZ | Evacuator, Oral Cavity | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | EKC | Excavator, Dental, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | OHP | Expiratory Resistance Valve, Intranasal, For Obstructive Sleep Apnea | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | EKB | Explorer, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | PME | External Cleaning Solution | 2 | 872.3260 | Cavity varnish. | N | N
Dental | MQN | External Mandibular Fixator And/Or Distractor | 2 | 872.4760 | Bone plate. | N | Y
Dental | KCR | Facebow | 1 | 872.3220 | Facebow. | N | N
Dental | EMI | File, Bone, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKA | File, Margin Finishing, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMR | File, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKS | File, Pulp Canal, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NYE | Fixture Mount Kit | 1 | 872.3980 | Endosseous dental implant accessories. | N | N
Dental | JES | Floss, Dental | 1 | 872.6390 | Dental floss. | N | N
Dental | EFK | Forceps, Articulation Paper | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EFL | Forceps, Dressing, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMH | Forceps, Rongeur, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EJG | Forceps, Rubber Dam Clamp | 1 | 872.6300 | Rubber dam and accessories. | N | N
Dental | EMG | Forceps, Tooth Extractor, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EJE | Frame, Rubber Dam | 1 | 872.6300 | Rubber dam and accessories. | N | N
Dental | EIL | Gauge, Depth, Instrument, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EAS | Gel, Electrode, For Pulp Tester | 1 | 872.1730 | Electrode gel for pulp testers. | N | N
Dental | OFX | General Purpose Dental Tray | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | NYF | Gingival Retraction Kit | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | MPI | Glenoid Fossa Prosthesis | 3 | 872.3950 | Glenoid fossa prosthesis. | N | Y
Dental | EEJ | Guard, Disk | 1 | 872.6010 | Abrasive device and accessories. | N | N
Dental | EKM | Gutta-Percha | 1 | 872.3850 | Gutta percha. | N | N
Dental | EJB | Handle, Instrument, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EFB | Handpiece, Air-Powered, Dental | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | NYL | Handpiece, Air-Powered, Root Canal Irrigation | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | EFA | Handpiece, Belt And/Or Gear Driven, Dental | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | EGS | Handpiece, Contra- And Right-Angle Attachment, Dental | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | EKX | Handpiece, Direct Drive, Ac-Powered | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | KMW | Handpiece, Rotary Bone Cutting | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | EKY | Handpiece, Water-Powered | 1 | 872.4200 | Dental handpiece and accessories. | N | N
Dental | DZB | Headgear, Extraoral, Orthodontic | 2 | 872.5500 | Extraoral orthodontic headgear. | N | N
Dental | EEG | Heat Source For Bleaching Teeth | 1 | 872.6475 | Heat source for bleaching teeth. | N | N
Dental | EMD | Hemostat, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMQ | Hoe, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EGZ | Holder, Film, X-Ray | 1 | 872.1905 | Dental x-ray film holder. | N | N
Dental | OAT | Implant, Endosseous, Orthodontic | 2 | 872.3640 | Endosseous dental implant. | N | Y
Dental | DZE | Implant, Endosseous, Root-Form | 2 | 872.3640 | Endosseous dental implant. | N | Y
Dental | ELE | Implant, Subperiosteal | 2 | 872.3645 | Subperiosteal implant material. | N | Y
Dental | MDL | Implant, Transmandibular | 2 | 872.4760 | Bone plate. | N | Y
Dental | QJJ | Indirect Bonding Tray | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EGQ | Injector, Jet, Gas-Powered | 2 | 872.4465 | Gas-powered jet injector. | N | N
Dental | EGM | Injector, Jet, Mechanical-Powered | 2 | 872.4475 | Spring-powered jet injector. | N | N
Dental | KZO | Ink, Arch Tracing | 1 | 872.3150 | Articulator. | N | N
Dental | EKF | Instrument, Contouring, Matrix, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKD | Instrument, Cutting, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DZP | Instrument, Diamond, Dental | 1 | 872.4535 | Dental diamond instrument. | N | N
Dental | NLD | Instrument, Diamond, Dental, Reprocessed | 1 | 872.4535 | Dental diamond instrument. | N | N
Dental | EIY | Instrument, Filling, Plastic, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ELA | Instrument, Hand, Calculus Removal | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ECP | Instrument, Ligature Tucking, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DZN | Instruments, Dental Hand | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | MPJ | Interarticular Disc Prosthesis (Interpositional Implant) | 3 | 872.3970 | Interarticular disc prosthesis (interpositional implant). | N | Y
Dental | QJK | Intraoral Camera | 2 | 872.3661 | Optical Impression Systems for CAD/CAM. | N | N
Dental | OZR | Intraoral Pressure Gradient Device | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | PFL | Intraoral, Dental Foam Protector | 1 | 872.6890 | Intraoral dental wax. | N | N
Dental | LZD | Joint, Temporomandibular, Implant | 3 | 872.3940 | Total temporomandibular joint prosthesis. | N | Y
Dental | EJZ | Knife, Margin Finishing, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EMO | Knife, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NYG | Lapping Tool Kit | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NBL | Laser, Fluorescence Caries Detection | 2 | 872.1745 | Laser fluorescence caries detection device. | N | N
Dental | EAY | Light, Fiber Optic, Dental | 1 | 872.4620 | Fiber optic dental light. | N | N
Dental | EAZ | Light, Operating, Dental | 1 | 872.4630 | Dental operating light. | N | N
Dental | EBA | Light, Surgical Headlight | 1 | 872.4630 | Dental operating light. | N | N
Dental | EJK | Liner, Cavity, Calcium Hydroxide | 2 | 872.3250 | Calcium hydroxide cavity liner. | N | N
Dental | DYX | Lock, Wire, And Ligature, Intraoral | 2 | 872.4600 | Intraoral ligature and wire lock. | N | Y
Dental | DYT | Maintainer, Space Preformed, Orthodontic | 1 | 872.5525 | Preformed tooth positioner. | N | N
Dental | MPL | Mandibular Condyle Prosthesis | 3 | 872.3960 | Mandibular condyle prosthesis. | N | Y
Dental | EMP | Marker, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ELW | Material, Impression | 2 | 872.3660 | Impression material. | N | N
Dental | EBH | Material, Impression Tray, Resin | 1 | 872.3670 | Resin impression tray material. | N | N
Dental | EBF | Material, Tooth Shade, Resin | 2 | 872.3690 | Tooth shade resin material. | N | N
Dental | NFE | Materials, Polytetrafluoroethylene Vitreous Carbon, For Maxillofacial Alveolar Ridge Augmentation | 2 | 872.3680 | Polytetrafluoroethylene (PTFE) vitreous carbon materials. | N | N
Dental | JEO | Measurer, Gingival Fluid | 1 | 872.1500 | Gingival fluid measurer. | N | N
Dental | ELY | Mercury | 2 | 872.3070 | Dental amalgam, mercury, and amalgam alloy. | N | N
Dental | EAX | Mirror, Mouth | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DYN | Mouthpiece, Saliva Ejector | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | DZM | Needle, Dental | 1 | 872.4730 | Dental injecting needle. | N | N
Dental | NMW | Needle, Dental, Reprocessed | 1 | 872.4730 | Dental injecting needle. | N | N
Dental | OGT | Oral Irrigation Kit | 1 | 872.6510 | Oral irrigation unit. | N | N
Dental | PWL | Oral Surgery Tray | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | PLH | Orthodontic Led Accessory | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | PNN | Orthodontic Software | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | OYH | Orthodontic Vibratory Accessory | 2 | 872.5470 | Orthodontic plastic bracket. | N | N
Dental | EHR | Pad, Denture, Over The Counter | 2 | 872.3540 | OTC denture cushion or pad. | N | N
Dental | KCS | Pantograph | 1 | 872.3730 | Pantograph. | N | N
Dental | EFH | Paper, Articulation | 1 | 872.6140 | Articulation paper. | N | N
Dental | EGI | Parallelometer | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NSB | Patient Personal Hygiene Kit | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | JET | Pick, Massaging | 1 | 872.6650 | Massaging pick or tip for oral hygiene. | N | N
Dental | MYB | Pillow, Cervical (For Mild Sleep Apnea) | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | EBL | Pin, Retentive And Splinting, And Accessory Instruments | 1 | 872.3740 | Retentive and splinting pin. | N | N
Dental | QJC | Plaque Disclosing Agent | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | NYH | Plaque Disclosing Kit | 2 | 872.1740 | Caries detection device. | N | N
Dental | EEA | Plate, Base, Shellac | 1 | 872.6200 | Base plate shellac. | N | N
Dental | JEY | Plate, Bone | 2 | 872.4760 | Bone plate. | N | Y
Dental | JEX | Plier, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EJY | Pliers, Operative | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKR | Plugger, Root Canal, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EHL | Point, Abrasive | 1 | 872.6010 | Abrasive device and accessories. | N | N
Dental | EKN | Point, Paper, Endodontic | 1 | 872.3830 | Endodontic paper point. | N | N
Dental | EKL | Point, Silver, Endodontic | 1 | 872.3840 | Endodontic silver point. | N | N
Dental | KXY | Polyvinyl Methylether Maleic Anhydride &/Or Acid Copolymer & Carboxymethylce | 3 | 872.3500 | Polyvinylmethylether maleic anhydride (PVM-MA), acid copolymer, and carboxymethylcellulose sodium (NACMC) denture adhesive. | N | N
Dental | KMY | Positioner, Tooth, Preformed | 1 | 872.5525 | Preformed tooth positioner. | N | N
Dental | ELR | Post, Root Canal | 1 | 872.3810 | Root canal post. | N | N
Dental | EIH | Powder, Porcelain | 2 | 872.6660 | Porcelain powder for clinical use. | N | N
Dental | QIA | Powered Tongue Scraper | 1 | 872.6865 | Powered toothbrush. | N | N
Dental | ORY | Prelude Tongue Suspension System | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | Y
Dental | EKQ | Preparer, Root Canal Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EIX | Probe, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | PIP | Prophy Powder, Airbrush Accessory | 2 | 872.6080 | Airbrush. | N | N
Dental | NEI | Prosthesis, Condyle, Mandibular, Temporary | 2 | 872.4770 | Temporary mandibular condyle reconstruction plate. | N | Y
Dental | EFX | Protector, Silicate | 1 | 872.6670 | Silicate protector. | N | N
Dental | EME | Punch, Biopsy, Surgical | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ECS | Pusher, Band, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EKP | Reamer, Pulp Canal, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EBP | Reliner, Denture, Over The Counter | 2 | 872.3560 | OTC denture reliner. | N | N
Dental | EIS | Remover, Crown | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EBI | Resin, Denture, Relining, Repairing, Rebasing | 2 | 872.3760 | Denture relining, repairing, or rebasing resin. | N | N
Dental | KIF | Resin, Root Canal Filling | 2 | 872.3820 | Root canal filling resin. | N | N
Dental | MMT | Resin, Root Canal Filling Containing Chloroform | 3 | 872.3820 | Root canal filling resin. | N | N
Dental | OFW | Restorative Instrument And Component Tray | 2 | 872.3690 | Tooth shade resin material. | N | N
Dental | JEP | Retainer, Matrix | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | DYJ | Retainer, Screw Expansion, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EIG | Retractor, All Types | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | KKO | Ring, Teething, Fluid-Filled | 2 | 872.5550 | Teething ring. | N | N
Dental | MEF | Ring, Teething, Non-Fluid Filled | 1 | 872.5550 | Teething ring. | N | N
Dental | NTO | Rinse, Oral, Antibacterial (By Physical Means) | 2 | 872.5580 | Oral rinse to reduce the adhesion of dental plaque. | N | N
Dental | DZH | Saw, Bone, Ac-Powered | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | EMN | Scaler, Periodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ELB | Scaler, Rotary | 2 | 872.4840 | Rotary scaler. | N | N
Dental | ELC | Scaler, Ultrasonic | 2 | 872.4850 | Ultrasonic scaler. | N | N
Dental | KZN | Scanner, Color | 2 | 872.3661 | Optical Impression Systems for CAD/CAM. | N | N
Dental | EIR | Scissors, Collar And Crown | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EGN | Scissors, Surgical Tissue, Dental | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | LCN | Scraper, Tongue | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | DZL | Screw, Fixation, Intraosseous | 2 | 872.4880 | Intraosseous fixation screw or wire. | N | Y
Dental | EBC | Sealant, Pit And Fissure, And Conditioner | 2 | 872.3765 | Pit and fissure sealant and conditioner. | N | N
Dental | ECR | Setter, Band, Orthodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | PLC | Sleep Appliances With Patient Monitoring | 2 | 872.5570 | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | N | N
Dental | KZP | Solution, Cement Dissolving | 2 | 872.3750 | Bracket adhesive resin and tooth conditioner. | N | N
Dental | ELS | Splint, Endodontic Stabilizing | 2 | 872.3890 | Endodontic stabilizing splint. | N | N
Dental | EKK | Spreader, Pulp Canal Filling Material, Endodontic | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | ECO | Spring, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | ECG | Sterilizer, Boiling Water | 1 | 872.6710 | Boiling water sterilizer. | N | N
Dental | KOK | Sterilizer, Endodontic Dry Heat | 3 | 872.6730 | Endodontic dry heat sterilizer. | N | N
Dental | ECC | Sterilizer, Glass Bead | 3 | 872.6730 | Endodontic dry heat sterilizer. | N | N
Dental | LTF | Stimulator, Salivary System | 2 | 872.5560 | Electrical salivary stimulatory system. | N | N
Dental | EHM | Strip, Polishing Agent | 1 | 872.6010 | Abrasive device and accessories. | N | N
Dental | EJI | Syringe, Cartridge | 2 | 872.6770 | Cartridge syringe. | N | N
Dental | EIB | Syringe, Irrigating (Dental) | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EIC | Syringe, Periodontic, Endodontic, Irrigating | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | EID | Syringe, Restorative And Impression Material | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | MXF | System, Dental, Hydrokinetic, Caries Removal & Cavity Preparation | 2 | 872.4120 | Bone cutting instrument and accessories. | N | N
Dental | NOF | System, Optical Impression, Computer Assisted Design And Manufacturing (Cad/Cam) Of Dental Restorations | 2 | 872.3661 | Optical Impression Systems for CAD/CAM. | N | N
Dental | MUH | System, X-Ray, Extraoral Source, Digital | 2 | 872.1800 | Extraoral source x-ray system. | N | N
Dental | ELK | Teeth, Artificial, Backing And Facing | 1 | 872.3910 | Backing and facing for an artificial tooth. | N | N
Dental | ELJ | Teeth, Artificial, Posterior With Metal Insert | 1 | 872.3900 | Posterior artificial tooth with a metal insert. | N | N
Dental | ELL | Teeth, Porcelain | 2 | 872.3920 | Porcelain tooth. | N | N
Dental | ELN | Teeth, Preformed Gold Denture | 1 | 872.3580 | Preformed gold denture tooth. | N | N
Dental | EAT | Tester, Pulp | 2 | 872.1720 | Pulp tester. | N | N
Dental | JEW | Tip, Rubber, Oral Hygiene | 1 | 872.6650 | Massaging pick or tip for oral hygiene. | N | N
Dental | MMD | Toothbrush, Ionic, Battery-Powered | 1 | 872.6865 | Powered toothbrush. | N | N
Dental | EFW | Toothbrush, Manual | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | JEQ | Toothbrush, Powered | 1 | 872.6865 | Powered toothbrush. | N | N
Dental | KMT | Tray, Fluoride, Disposable | 1 | 872.6870 | Disposable fluoride tray. | N | N
Dental | EHY | Tray, Impression, Preformed | 1 | 872.6880 | Preformed impression tray. | N | N
Dental | KCQ | Tube Impression And Matrix | 1 | 872.6570 | Impression tube. | N | N
Dental | DZD | Tube, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EKZ | Unit, Electrosurgical, And Accessories, Dental | 2 | 872.4920 | Dental electrosurgical unit and accessories. | N | N
Dental | EIA | Unit, Operative Dental | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | NRD | Unit, Operative Dental, Accessories | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | EFS | Unit, Oral Irrigation | 1 | 872.6510 | Oral irrigation unit. | N | N
Dental | EBR | Unit, Suction Operatory | 1 | 872.6640 | Dental operative unit and accessories. | N | N
Dental | ECB | Unit, Syringe, Air And/Or Water | 1 | 872.4565 | Dental hand instrument. | N | N
Dental | NOB | Unit, Ultraviolet Sanitation/Sterilization (For Toothbrushes), Non-Sterile | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | MCF | Unit, Ultraviolet Sanitation/Sterilization (For Toothbrushes), Sterile | 1 | 872.6855 | Manual toothbrush. | N | N
Dental | EHD | Unit, X-Ray, Extraoral With Timer | 2 | 872.1800 | Extraoral source x-ray system. | N | N
Dental | EAP | Unit, X-Ray, Intraoral | 2 | 872.1810 | Intraoral source x-ray system. | N | N
Dental | LBH | Varnish, Cavity | 2 | 872.3260 | Cavity varnish. | N | N
Dental | EFC | Warmer, Anesthetic Tube | 1 | 872.6100 | Anesthetic warmer. | N | N
Dental | EGD | Wax, Dental, Intraoral | 1 | 872.6890 | Intraoral dental wax. | N | N
Dental | EJQ | Wheel, Polishing Agent | 1 | 872.6010 | Abrasive device and accessories. | N | N
Dental | DZK | Wire, Fixation, Intraosseous | 2 | 872.4880 | Intraosseous fixation screw or wire. | N | Y
Dental | DZC | Wire, Orthodontic | 1 | 872.5410 | Orthodontic appliance and accessories. | N | N
Dental | EMB | Zinc Oxide Eugenol | 1 | 872.3275 | Dental cement. | N | N
Ear noise & throat | PFO | Active Implantable Bone Conduction Hearing System | 1 | 874.3340 | Active implantable bone conduction hearing system. | N | Y
Ear noise & throat | ETR | Adaptor, Short Increment Sensitivity Index (Sisi) | 1 | 874.1070 | Short increment sensitivity index (SISI) adapter. | N | N
Ear noise & throat | KBH | Adenotome | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | LYN | Apparatus, Visual Reinforcement Audiometric | 2 | 874.1050 | Audiometer. | N | N
Ear noise & throat | KCJ | Applicator, Ent | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | LRD | Applicator, Ent Drug | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | EWO | Audiometer | 2 | 874.1050 | Audiometer. | N | N
Ear noise & throat | EMX | Balloon, Epistaxis | 1 | 874.4100 | Epistaxis balloon. | N | N
Ear noise & throat | QGK | Balloon, Nasal Airway | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KCC | Bistoury, Tracheal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNR | Block, Cutting, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXS | Block, Cutting, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KCL | Blower, Powder, Ent | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | MXP | Bone Particle Collector | 1 | 874.4800 | Bone particle collector. | N | N
Ear noise & throat | KCD | Bougie, Esophageal, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBI | Bougie, Eustachian | 1 | 874.4175 | Nasopharyngeal catheter. | N | N
Ear noise & throat | EOQ | Bronchoscope (Flexible Or Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | KTI | Bronchoscope Accessory | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | BTG | Brush, Biopsy, Bronchoscope (Non-Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | EQJ | Bur, Ear, Nose And Throat | 1 | 874.4140 | Ear, nose, and throat bur. | N | N
Ear noise & throat | NLZ | Bur, Ent, Diamond Coated, Single Use, Reprocessed | 1 | 874.4140 | Ear, nose, and throat bur. | N | N
Ear noise & throat | ETW | Calibrator, Hearing Aid / Earphone And Analysis Systems | 2 | 874.3310 | Hearing aid calibrator and analysis system. | N | N
Ear noise & throat | ENW | Catheter, Nasopharyngeal | 1 | 874.4175 | Nasopharyngeal catheter. | N | N
Ear noise & throat | EWC | Chamber, Acoustic (For Audiometric Testing) | 1 | 874.1060 | Acoustic chamber for audiometric testing. | N | N
Ear noise & throat | JYE | Chisel, Middle-Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYF | Clamp, Ossicle Holding | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JEI | Claw, Foreign Body, Bronchoscope (Non-Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | NNN | Crimper, Wire, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXT | Crimper, Wire, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KBJ | Curette, Adenoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | BST | Curette, Biopsy, Bronchoscope (Non-Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | JEL | Curette, Biopsy, Bronchoscope (Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | JYG | Curette, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAO | Curette, Ethmoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAP | Curette, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBK | Curette, Salpingeal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ETT | Cushion, Earphone (For Audiometric Testing) | 1 | 874.1100 | Earphone cushion for audiometric testing. | N | N
Ear noise & throat | KBL | Depressor, Metal Tongue, Surgical | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | EWT | Device, Antichoke, Suction | 3 | 874.5350 | Suction antichoke device. | N | N
Ear noise & throat | EWW | Device, Antichoke, Tongs | 3 | 874.5370 | Tongs antichoke device. | N | N
Ear noise & throat | KTH | Device, Anti-Stammering | 1 | 874.5840 | Antistammering device. | N | N
Ear noise & throat | LZI | Device, Assistive Listening | 2 | 874.3320 | Group hearing aid or group auditory trainer. | N | N
Ear noise & throat | MCK | Device, Voice Amplification | 2 | 874.3730 | Laryngeal prosthesis (Taub design). | N | N
Ear noise & throat | NNP | Die, Wire Bending, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXW | Die, Wire Bending, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KCF | Dilator, Esophageal, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | LWF | Dilator, Nasal | 1 | 874.3900 | Nasal dilator. | N | N
Ear noise & throat | KCG | Dilator, Tracheal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBM | Dissector, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ERL | Drill, Surgical, Ent (Electric Or Pneumatic) Including Handpiece | 2 | 874.4250 | Ear, nose, and throat electric or pneumatic surgical drill. | N | N
Ear noise & throat | KCM | Dropper, Ent | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | KCN | Ear Wick | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | OGR | Ear, Nose, And Throat Surgical Tray | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | MIB | Elastomer, Silicone Block | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | KLX | Electroglottograph | 2 | 874.1325 | Electroglottograph. | N | N
Ear noise & throat | KAD | Elevator, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | EOX | Esophagoscope (Flexible Or Rigid) | 2 | 874.4710 | Esophagoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | PNZ | Eustachian Tube Balloon Dilation Device | 2 | 874.4180 | Eustachian tube balloon dilation system. | N | N
Ear noise & throat | JYH | Excavator, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | PKA | External Upper Esophageal Sphincter (Ues) Compression Device | 2 | 874.5900 | External upper esophageal sphincter compression device. | N | N
Ear noise & throat | LRB | Face Plate Hearing Aid | 1 | 874.3300 | Hearing Aid. | N | N
Ear noise & throat | BWH | Forceps, Biopsy, Bronchoscope (Non-Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | NLE | Forceps, Biopsy, Bronchoscope (Non-Rigid), Reprocessed | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | JEK | Forceps, Biopsy, Bronchoscope (Rigid) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | KAE | Forceps, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNQ | Forceps, Wire Closure, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXX | Forceps, Wire Closure, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KBN | Gag, Mouth | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYI | Gauge, Mastoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYJ | Gauge, Measuring | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ETS | Generator, Electronic Noise (For Audiometric Testing) | 2 | 874.1120 | Electronic noise generator for audiometric testing. | N | N
Ear noise & throat | KAQ | Gouge, Nasal, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBO | Guillotine, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNJ | Gustometer, Non-Sterile | 1 | 874.1500 | Gustometer. | N | N
Ear noise & throat | ETM | Gustometer, Sterile | 1 | 874.1500 | Gustometer. | N | N
Ear noise & throat | ESD | Hearing Aid, Air Conduction | 1 | 874.3300 | Hearing Aid. | N | N
Ear noise & throat | OSM | Hearing Aid, Air Conduction With Wireless Technology | 2 | 874.3305 | Wireless air-conduction hearing aid. | N | N
Ear noise & throat | NIX | Hearing Aid, Air Conduction, Transcutaneous System | 2 | 874.3950 | Transcutaneous air conduction hearing aid system. | N | Y
Ear noise & throat | LXB | Hearing Aid, Bone Conduction | 2 | 874.3300 | Hearing Aid. | N | N
Ear noise & throat | MAH | Hearing Aid, Bone Conduction, Implanted | 2 | 874.3300 | Hearing Aid. | N | Y
Ear noise & throat | EPF | Hearing Aid, Group And Auditory Trainer | 2 | 874.3320 | Group hearing aid or group auditory trainer. | N | N
Ear noise & throat | KHL | Hearing Aid, Master | 2 | 874.3330 | Master hearing aid. | N | N
Ear noise & throat | JYL | Hook, Microsurgical Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBP | Hook, Tonsil Suturing | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KCH | Hook, Tracheal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | MIC | Implant, Muscle, Pectoralis | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | JYM | Inserter, Myringotomy Tube | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYN | Inserter, Sacculotomy Tack | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | LRC | Instrument, Ent Manual Surgical | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N | Acclarent
Ear noise & throat | NNS | Instrument, Prosthesis Modification For Ossicular Replacement Surgery, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | LRE | Instrument, Prosthesis Modification For Ossicular Replacement Surgery, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KMA | Irrigator, Powered Nasal | 1 | 874.5550 | Powered nasal irrigator. | N | N
Ear noise & throat | KAR | Irrigator, Sinus | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNO | Jig, Piston Cutting, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXY | Jig, Piston Cutting, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | LDG | Kit, Earmold, Impression | 1 | 874.3300 | Hearing Aid. | N | N
Ear noise & throat | NRK | Kit, Test, Olfactory | 2 | 874.1600 | Olfactory test device. | N | N
Ear noise & throat | JYO | Knife, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KTG | Knife, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZY | Knife, Laryngeal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYP | Knife, Myringotomy | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAS | Knife, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBQ | Knife, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | EQN | Laryngoscope, Nasopharyngoscope | 2 | 874.4760 | Nasopharyngoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | EQL | Laryngostroboscope | 1 | 874.4750 | Laryngostroboscope. | N | N
Ear noise & throat | ESE | Larynx, Artificial (Battery-Powered) | 1 | 874.3375 | Battery-powered artificial larynx. | N | N
Ear noise & throat | EWG | Laser, Ent Microsurgical Carbon-Dioxide | 2 | 874.4500 | Ear, nose, and throat microsurgical carbon dioxide laser. | N | N
Ear noise & throat | LXR | Laser, Microsurgical Argon, For Use In Otology | 2 | 874.4490 | Argon laser for otology, rhinology, and laryngology. | N | N
Ear noise & throat | LMS | Laser, Microsurgical Argon, For Uses Other Than Otology, Including Laryngology & General Use In Otolaryngology | 2 | 874.4490 | Argon laser for otology, rhinology, and laryngology. | N | N
Ear noise & throat | LLO | Laser, Neodymium:Yag, Pulmonary Surgery | 2 | 874.4500 | Ear, nose, and throat microsurgical carbon dioxide laser. | N | N
Ear noise & throat | JYQ | Loop, Wire | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KLW | Masker, Tinnitus | 2 | 874.3400 | Tinnitus masker. | N | N
Ear noise & throat | EWY | Mediastinoscope, Surgical | 2 | 874.4720 | Mediastinoscope and accessories. | N | N
Ear noise & throat | NLY | Microdebrider, Ent, High Speed, Single Use, Reprocessed | 1 | 874.4140 | Ear, nose, and throat bur. | N | N
Ear noise & throat | KAH | Microrule, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAI | Mirror, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAJ | Mobilizer, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ETC | Mold, Middle-Ear | 2 | 874.3430 | Middle ear mold. | N | Y
Ear noise & throat | OGS | Myringotomy Procedure Kit | 2 | 874.3880 | Tympanostomy tube. | N | N
Ear noise & throat | KCO | Nasal Spray, Ent Delivery | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | EOB | Nasopharyngoscope (Flexible Or Rigid) | 2 | 874.4760 | Nasopharyngoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | KBR | Needle, Tonsil Suturing | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | PDQ | Neurosurgical Nerve Locator | 2 | 874.1820 | Surgical nerve stimulator/locator. | N | N
Ear noise & throat | JYR | Nipper, Malleus | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ERA | Otoscope | 1 | 874.4770 | Otoscope. | N | N
Ear noise & throat | KAT | Perforator, Antrum | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYS | Perforator, Ear-Lobe | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYT | Pick, Microsurgical Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | OGU | Pierced Ear/Entry Cleaning Tray | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NHB | Polymer, Ear, Nose And Throat, Synthetic, Absorbable | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | N
Ear noise & throat | JOF | Polymer, Ent Synthetic, Porous Polyethylene | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | ESH | Polymer, Ent Synthetic-Pife, Silicon Elastomer, Polyethylene, Polyurethane | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | KHJ | Polymer, Ent Synthetic-Polyamide (Mesh Or Foil Material) | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | LBM | Porous Polyethylene Ossicular Replacement | 2 | 874.3450 | Partial ossicular replacement prosthesis. | N | Y
Ear noise & throat | JYW | Press, Vein | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAK | Probe, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JAZ | Prosthesis, Facial, Mandibular Implant | 2 | 874.3695 | Mandibular implant facial prosthesis. | N | Y
Ear noise & throat | EWL | Prosthesis, Laryngeal (Taub) | 2 | 874.3730 | Laryngeal prosthesis (Taub design). | N | N
Ear noise & throat | FWN | Prosthesis, Larynx (Stents And Keels) | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | N
Ear noise & throat | ETB | Prosthesis, Partial Ossicular Replacement | 2 | 874.3450 | Partial ossicular replacement prosthesis. | N | Y
Ear noise & throat | JPW | Pump, Nebulizer, Electrically Powered | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | EPN | Pump, Nebulizer, Manual | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | KBS | Punch, Adenoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAW | Punch, Antrum | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYX | Punch, Attic | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KTF | Punch, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAX | Punch, Ethmoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNT | Punch, Gelfoam, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JXZ | Punch, Gelfoam, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KAY | Punch, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBT | Punch, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYY | Rasp, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAZ | Rasp, Frontal-Sinus | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBA | Rasp, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | LBP | Replacement, Ossicular (Stapes) Using Absorbable Gelatin Material | 2 | 874.3450 | Partial ossicular replacement prosthesis. | N | Y
Ear noise & throat | ETA | Replacement, Ossicular Prosthesis, Total | 2 | 874.3495 | Total ossicular replacement prosthesis. | N | Y
Ear noise & throat | LBN | Replacement, Total Ossicular, Prosthesis, Porous, Polyethylene | 2 | 874.3495 | Total ossicular replacement prosthesis. | N | Y
Ear noise & throat | KAL | Retractor, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JYZ | Rod, Measuring Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZA | Rongeur, Mastoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBB | Rongeur, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZZ | Saw, Laryngeal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBC | Saw, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | EWQ | Saw, Surgical, Ent (Electric Or Pneumatic) | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZB | Scissors, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBD | Scissors, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NNU | Scissors, Wire Cutting, Ent, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JYA | Scissors, Wire Cutting, Ent, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | KBW | Screw, Oral | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBX | Screw, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZC | Searcher, Mastoid | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | QDD | Self-Fitting Air-Conduction Hearing Aid | 2 | 874.3325 |  | N | N
Ear noise & throat | EWA | Set, Audiometer Calibration | 1 | 874.1080 | Audiometer calibration set. | N | N
Ear noise & throat | KBY | Set, Filliform, Eustachian | 1 | 874.4175 | Nasopharyngeal catheter. | N | N
Ear noise & throat | KAA | Set, Laryngeal Injection | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | JZD | Snare, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBE | Snare, Nasal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBZ | Snare, Tonsil | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | EQH | Source, Carrier, Fiberoptic Light | 1 | 874.4350 | Ear, nose, and throat fiberoptic light source and carrier. | N | N
Ear noise & throat | LYA | Splint, Intranasal Septal | 1 | 874.4780 | Intranasal splint. | N | N
Ear noise & throat | QJL | Splint, Intranasal Septal, Cber Led | 1 | 874.4780 | Intranasal splint. | N | N
Ear noise & throat | EPP | Splint, Nasal | 1 | 874.5800 | External nasal splint. | N | N
Ear noise & throat | JZE | Spoon, Ear | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KHH | Stimulator, Caloric-Air | 1 | 874.1800 | Air or water caloric stimulator. | N | N
Ear noise & throat | ETP | Stimulator, Caloric-Water | 1 | 874.1800 | Air or water caloric stimulator. | N | N
Ear noise & throat | ETN | Stimulator, Nerve | 2 | 874.1820 | Surgical nerve stimulator/locator. | N | N
Ear noise & throat | KCP | Syringe, Ent | 1 | 874.5220 | Ear, nose, and throat drug administration device. | N | N
Ear noise & throat | MIX | System, Vocal Cord Medialization | 2 | 874.3620 | Ear, nose, and throat synthetic polymer material. | N | Y
Ear noise & throat | ESX | Tack, Sacculotomy (Cody Tack) | 2 | 874.3760 | Sacculotomy tack (Cody tack) | N | Y
Ear noise & throat | ENZ | Telescope, Laryngeal-Bronchial | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | ETY | Tester, Auditory Impedance | 2 | 874.1090 | Auditory impedance tester. | N | N
Ear noise & throat | PTO | Tester, Auditory Impedance, Exempt | 2 | 874.1090 | Auditory impedance tester. | N | N
Ear noise & throat | KCA | Tonsillectome | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | LJW | Tracheotome | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | MQW | Transilluminator, Fiber Optic | 1 | 874.4350 | Ear, nose, and throat fiberoptic light source and carrier. | N | N
Ear noise & throat | MMO | Tray, Surgical, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBF | Trephine, Sinus | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KTE | Trocar, Ent | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAB | Trocar, Laryngeal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KBG | Trocar, Sinus | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KCI | Trocar, Tracheal | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | NLB | Trocar, Tracheal, Laryngeal, Sinus, Reprocessed | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KTR | Tube, Aspirating, Bronchoscope, Rigid | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | JZF | Tube, Ear Suction | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | KAC | Tube, Laryngectomy | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ESZ | Tube, Shunt, Endolymphatic | 2 | 874.3820 | Endolymphatic shunt. | N | Y
Ear noise & throat | KLZ | Tube, Shunt, Endolymphatic With Valve | 2 | 874.3850 | Endolymphatic shunt tube with valve. | N | Y
Ear noise & throat | KCB | Tube, Tonsil Suction | 1 | 874.4420 | Ear, nose, and throat manual surgical instrument. | N | N
Ear noise & throat | ETK | Tube, Toynbee Diagnostic | 1 | 874.1925 | Toynbee diagnostic tube. | N | N
Ear noise & throat | ETD | Tube, Tympanostomy | 2 | 874.3880 | Tympanostomy tube. | N | Y
Ear noise & throat | KQL | Tube, Tympanostomy With Semi-Permeable Membrane | 2 | 874.3930 | Tympanostomy tube with semipermeable membrane. | N | Y
Ear noise & throat | LBL | Tube, Tympanostomy, Porous Polyethylene | 2 | 874.3880 | Tympanostomy tube. | N | Y
Ear noise & throat | JEJ | Tubing, Instrumentation, Bronchoscope (Brush Sheath A/O Aspirating) | 2 | 874.4680 | Bronchoscope (flexible or rigid) and accessories. | N | N
Ear noise & throat | PLK | Tympanic Membrane Direct Contact Hearing Aid | 2 | 874.3315 | Tympanic membrane contact hearing aid. | N | Y
Ear noise & throat | NAS | Tympanometer | 2 | 874.1090 | Auditory impedance tester. | N | N
Ear noise & throat | PTP | Tympanometer, Exempt | 2 | 874.1090 | Auditory impedance tester. | N | N
Ear noise & throat | ETF | Unit, Examining/Treatment, Ent | 1 | 874.5300 | Ear, nose, and throat examination and treatment unit. | N | N
Ear noise & throat | NNV | Vise, Ossicular Finger, Non-Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Ear noise & throat | JYB | Vise, Ossicular Finger, Sterile | 1 | 874.3540 | Prosthesis modification instrument for ossicular replacement surgery. | N | N
Gastroenterology/urology | KNZ | Accessories, A-V Shunt | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | KOC | Accessories, Blood Circuit, Hemodialysis | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | KNY | Accessories, Catheter, G-U | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | MNL | Accessories, Cleaning Brushes, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FEB | Accessories, Cleaning, For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NZA | Accessories, Germicide, Cleaning, For Endoscopes | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FEM | Accessories, Photographic, For Endoscope (Exclude Light Sources) | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KNR | Adapter, A-V Shunt Or Fistula | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | FHC | Adaptor To The Cord, For Transurethral Surgical Instrument | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FFY | Adaptor, Bulbs, Miscellaneous, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FKM | Adaptor, Fistula | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FKN | Adaptor, Shunt | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | EYI | Adaptor, Ureteral Catheter | 1 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FJP | Adaptor, Y | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | NOY | Agents, Embolic, For Treatment Of Benign Prostatic Hyperplasia | 2 | 876.5550 | Prostatic artery embolization device. | N | Y
Gastroenterology/urology | KPN | Alarm, Conditioned Response Enuresis | 2 | 876.2040 | Enuresis alarm. | N | N
Gastroenterology/urology | FIW | Alarm, Pillow Pressure | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | LYK | Angioscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FER | Anoscope And Accessories | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCT | Anti Fog Solution And Accessories, Endoscopy | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PRO | Anti-Regurgitation Blood Tubing Kit | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FLD | Apparatus, Hemoperfusion, Sorbent | 2 | 876.5870 | Sorbent hemoperfusion system. | N | N
Gastroenterology/urology | FDP | Apparatus, Pneumoperitoneum, Automatic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | EZS | Appliance, Colostomy, Disposable | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | OOW | Applicator For Rectal Suppository | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FEJ | Attachment, Binocular, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDZ | Attachment, Eyepiece, For Insertion Of Prescription Lens | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FEA | Attachment, Teaching, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODX | Autonomous Extracorporeal Blood Leak Detector/Alarm | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | EXF | Bag, Bile Collecting | 2 | 876.5010 | Biliary catheter and accessories. | N | N
Gastroenterology/urology | FON | Bag, Drainage, With Adhesive, Ostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | FCB | Bag, Hemostatic | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | GDS | Bag, Stomal | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | EXH | Bag, Urinary, Ileostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | EXG | Bag, Urine Collecting, Ureterostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | NNW | Bag, Urine Collection, Leg, For External Use, Non-Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | FAQ | Bag, Urine Collection, Leg, For External Use, Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | FOC | Bag, Urine Collection, Newborn | 2 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | PRP | Balloon Gastrostomy Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | NGU | Balloon, Liver Tamponade | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FCD | Barium Enema Kit | 1 | 876.5210 | Enema kit. | N | N
Gastroenterology/urology | PHT | Basin For Endoscope Reprocessing | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PTT | Biliary Catheter For Irrigation And Contrast Injection, Exempt | 2 | 876.5010 | Biliary catheter and accessories. | N | N
Gastroenterology/urology | GCA | Biliary Catheter For Stone Removal That May Also Allow For Irrigation And Contrast Injection | 2 | 876.5010 | Biliary catheter and accessories. | N | N
Gastroenterology/urology | PNB | Biliary Stent System For Benign Strictures | 2 | 876.5011 | Metallic biliary stent system for benign strictures. | N | Y
Gastroenterology/urology | QGE | Biliary Stent, Drain, And Dilator Accessories | 1 | 876.5012 | Biliary stent, drain, and dilator accessories. | N | N
Gastroenterology/urology | FCG | Biopsy Needle | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | POZ | Biopsy Needle Kit | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | PPA | Bladder Irrigation Kit | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FAT | Bougie, Esophageal, And Gastrointestinal, Gastro-Urology | 2 | 876.5365 | Esophageal dilator. | N | N
Gastroenterology/urology | FAX | Bougie, Urological | 1 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | FCP | Box, Battery, Pocket | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCO | Box, Battery, Rechargeable | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCY | Bulb, Inflation, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FKD | Canister, Coil | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FBM | Cannula And Trocar, Suprapubic, Non-Disposable | 1 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | FIQ | Cannula, A-V Shunt | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | FGS | Carrier, Sponge, Endoscopic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KDH | Catheter (Gastric, Colonic, Etc.), Irrigation And Aspiration | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FEZ | Catheter And Tube, Suprapubic | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | Y
Gastroenterology/urology | OHR | Catheter Care Tray | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | NML | Catheter, Biliary, Reprocessed | 2 | 876.5010 | Biliary catheter and accessories. | N | N
Gastroenterology/urology | EZC | Catheter, Coude | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FGH | Catheter, Double Lumen Female Urethrographic | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | LFK | Catheter, Femoral | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | MSD | Catheter, Hemodialysis, Implanted | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | NYU | Catheter, Hemodialysis, Implanted, Coated | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | MPB | Catheter, Hemodialysis, Non-Implanted | 2 | 876.5540 | Blood access device and accessories. | Y | N
Gastroenterology/urology | NQJ | Catheter, Hemodialysis, Non-Implanted, Ultrafiltration, For Peripheral Use | 2 | 876.5540 | Blood access device and accessories. | Y | N
Gastroenterology/urology | NIF | Catheter, Hemodialysis, Triple Lumen, Implanted | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | NIE | Catheter, Hemodialysis, Triple Lumen, Non-Implanted | 2 | 876.5540 | Blood access device and accessories. | Y | N
Gastroenterology/urology | FEW | Catheter, Malecot | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | FKO | Catheter, Peritoneal Dialysis, Single Use | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | PTV | Catheter, Peritoneal Dialysis, Single Use, Exempt | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | FJS | Catheter, Peritoneal, Long-Term Indwelling | 2 | 876.5630 | Peritoneal dialysis system and accessories. | N | Y
Gastroenterology/urology | PTU | Catheter, Peritoneal, Long-Term Indwelling, Exempt | 2 | 876.5630 | Peritoneal dialysis system and accessories. | N | Y
Gastroenterology/urology | GBT | Catheter, Rectal | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | KPH | Catheter, Rectal For Continent Ileostomy | 1 | 876.5030 | Continent ileostomy catheter. | N | N
Gastroenterology/urology | EZK | Catheter, Retention Type | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | EZL | Catheter, Retention Type, Balloon | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FGD | Catheter, Retention, Barium Enema With Bag | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | EZD | Catheter, Straight | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | LFJ | Catheter, Subclavian | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | NQV | Catheter, Suction, Tracheobronchial, Reprocessed | 1 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | KOB | Catheter, Suprapubic (And Accessories) | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | Y
Gastroenterology/urology | EYC | Catheter, Upper Urinary Tract | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FGF | Catheter, Ureteral Disposable (X-Ray) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | EYB | Catheter, Ureteral, Gastro-Urology | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | GBL | Catheter, Ureteral, General & Plastic Surgery | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | GBM | Catheter, Urethral | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FGI | Catheter, Urethrographic, Male | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | KOD | Catheter, Urological | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | MJC | Catheter, Urological (Antimicrobial) And Accessories | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | EZR | Cement, Stomal Appliance, Ostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | FIA | Chair, Dialysis, Unpowered, Without Scales | 1 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | PSX | Chlorine Meter | 2 | 876.5665 | Water purification system for hemodialysis. | Y | N
Gastroenterology/urology | FBN | Choledochoscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FKC | Clamp, Cannula | 1 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | NOO | Clamp, Dialysis Arm | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FGW | Clamp, Electrical | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FKK | Clamp, Line | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FFN | Clamp, Non-Electrical | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FHA | Clamp, Penile | 1 | 876.5160 | Urological clamp for males. | N | N
Gastroenterology/urology | FIG | Clamp, Tubing, Blood, Automatic | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | EXB | Collector, Ostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | KNX | Collector, Urine, (And Accessories) For Indwelling Catheter | 2 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | FFH | Collector, Urine, Pediatric, For Indwelling Catheter | 2 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | NZU | Collector, Urine, Powered, Non Indwelling Catheter | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | PGD | Colon Capsule Imaging System | 2 | 876.1330 | Colon capsule endoscopy system. | N | N
Gastroenterology/urology | KPL | Colonic Irrigation System | 2 | 876.5220 | Colonic irrigation system. | N | N
Gastroenterology/urology | NUB | Colonic Irrigation System, General Well Being | 3 | 876.5220 | Colonic irrigation system. | N | N
Gastroenterology/urology | FDF | Colonoscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FTJ | Colonoscope, General & Plastic Surgery | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OWN | Confocal Optical Imaging | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FKB | Connector, Blood Tubing, Infusion T | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJQ | Connector, Shunt | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FKY | Connector, Tubing, Dialysate | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | EYK | Connector, Ureteral Catheter | 1 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | KDK | Container, Transport, Kidney | 2 | 876.5880 | Isolated kidney perfusion and transport system and accessories. | N | N
Gastroenterology/urology | FBJ | Cord, Electric For Transurethral Surgical Instrument | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FFZ | Cord, Electric, For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FFF | Cover, Biopsy Forceps | 1 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FHH | Crusher, Spur, Colostomy | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | QAJ | Cutaneous Electrode Stimulator For Urinary Incontinence | 2 | 876.5330 |  | N | N
Gastroenterology/urology | EXQ | Cystometer, Electrical Recording | 2 | 876.1620 | Urodynamics measurement system. | N | N
Gastroenterology/urology | FAP | Cystometric Gas (Carbon-Dioxide) On Hydraulic Device | 2 | 876.1620 | Urodynamics measurement system. | N | N
Gastroenterology/urology | FAJ | Cystoscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FBO | Cystourethroscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FJZ | Declotting Tray, Kit (Including Contents) | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FHZ | Desiccator, Transurethral | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FJF | Detector, Air Bubble | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | KQR | Detector, Air Or Foam | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJC | Detector, Blood Level | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJB | Detector, Dialysate Level | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJD | Detector, Leak, Blood | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FEN | Device, Cystometric, Hydraulic | 2 | 876.1620 | Urodynamics measurement system. | N | N
Gastroenterology/urology | MFJ | Device, Endoscopic Suturing | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | LKY | Device, External Penile Rigidity | 2 | 876.5020 | External penile rigidity devices. | N | N
Gastroenterology/urology | FHW | Device, Impotence, Mechanical/Hydraulic | 3 | 876.3350 | Penile inflatable implant. | N | Y
Gastroenterology/urology | EZY | Device, Incontinence, Mechanical/Hydraulic | 3 | 876.5280 | Implanted mechanical/hydraulic urinary continence device. | N | Y
Gastroenterology/urology | NNX | Device, Incontinence, Urosheath Type, Non-Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | EXJ | Device, Incontinence, Urosheath Type, Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | FFR | Device, Locking, For Intestinal Clamp | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FDY | Device, Measuring, For Panendoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NOA | Device, Paste-On For Incontinence, Non-Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | EXI | Device, Paste-On For Incontinence, Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | NEQ | Device, Telemedicine, Robotic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FFG | Device, Urine Flow Rate Measuring, Non-Electrical, Disposable | 2 | 876.1800 | Urine flow or volume measuring system. | N | N
Gastroenterology/urology | KPO | Dialysate Concentrate For Hemodialysis (Liquid Or Powder) | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | PEV | Dialysis Administration Kit | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | LIF | Dialyzer Reprocessing System | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FJI | Dialyzer, Capillary, Hollow Fiber | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJH | Dialyzer, Disposable | 2 | 876.5830 | Hemodialyzer with disposable insert (Kiil type). | Y | N
Gastroenterology/urology | KDI | Dialyzer, High Permeability With Or Without Sealed Dialysate System | 2 | 876.5860 | High permeability hemodialysis system. | Y | N
Gastroenterology/urology | FJG | Dialyzer, Parallel Flow | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FHS | Dialyzer, Single Coil | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FJJ | Dialyzer, Twin Coil | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | EZN | Dilator, Catheter, Ureteral | 2 | 876.5470 | Ureteral dilator. | N | N
Gastroenterology/urology | KNQ | Dilator, Esophageal | 2 | 876.5365 | Esophageal dilator. | N | N
Gastroenterology/urology | EZM | Dilator, Esophageal (Metal Olive) Gastro-Urology | 2 | 876.5365 | Esophageal dilator. | N | N
Gastroenterology/urology | FFP | Dilator, Rectal | 1 | 876.5450 | Rectal dilator. | N | N
Gastroenterology/urology | KOE | Dilator, Urethral | 2 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | FAH | Dilator, Urethral, Mechanical | 2 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | FKA | Dilator, Vessel | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | NII | Disinfectant, Dialysate Delivery System | 2 | 876.5860 | High permeability hemodialysis system. | N | N
Gastroenterology/urology | NIH | Disinfectant, Subsystem, Water Purification | 2 | 876.5665 | Water purification system for hemodialysis. | N | N
Gastroenterology/urology | FFL | Dislodger, Stone, Basket, Ureteral, Metal | 2 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | NQU | Dislodger, Stone, Basket, Ureteral, Metal, Reprocessed | 2 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | LQR | Dislodger, Stone, Biliary | 2 | 876.5010 | Biliary catheter and accessories. | N | N
Gastroenterology/urology | FGO | Dislodger, Stone, Flexible | 2 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | NQT | Dislodger, Stone, Flexible, Ureteral, Reprocessed | 2 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | ONJ | Dissolvable Gel For Preventing Ureteral Stone Migration | 2 | 876.4680 | Ureteral stone dislodger. | N | N
Gastroenterology/urology | FDT | Duodenoscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PJH | Dynamic Rectal Control System | 2 | 876.5930 | Rectal control system. | N | N
Gastroenterology/urology | FAS | Electrode, Electrosurgical, Active, Urological | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | NLW | Electrode, Electrosurgical, Active, Urological, Reprocessed | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FEH | Electrode, Flexible Suction Coagulator | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | NLV | Electrode, Flexible Suction Coagulator, Reprocessed | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FFT | Electrode, Ph, Stomach | 1 | 876.1400 | Stomach pH electrode. | N | N
Gastroenterology/urology | NWI | Electrosurgical Electrode Kit | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | MPA | Endoilluminator | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NKZ | Endoilluminator, Reprocessed | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODC | Endoscope Channel Accessory | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCV | Endoscope Holder | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NWU | Endoscope Introducer Kit | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PUP | Endoscope Maintenance System | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PEA | Endoscope, Accessories, Image Post-Processing For Color Enhancement | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NWB | Endoscope, Accessories, Narrow Band Spectrum | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCP | Endoscope, Ac-Powered And Accessories | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCS | Endoscope, Battery-Powered And Accessories | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GDB | Endoscope, Fiber Optic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCQ | Endoscope, Flexible | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCO | Endoscope, Mirror | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KYH | Endoscope, Ophthalmic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCN | Endoscope, Prism | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCM | Endoscope, Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FED | Endoscopic Access Overtube, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | MNK | Endoscopic Bite Block | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODA | Endoscopic Central Control Unit | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NWT | Endoscopic Cholangiogram Kit | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODB | Endoscopic Contamination Prevention Sheath | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDX | Endoscopic Cytology Brush | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCZ | Endoscopic Grasping/Cutting Instrument, Non-Powered | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PTS | Endoscopic Grasping/Cutting Instrument, Non-Powered, Exempt | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCY | Endoscopic Guidewire, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FBK | Endoscopic Injection Needle, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCX | Endoscopic Irrigation/Suction System | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PCV | Endoscopic Leak Tester | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCC | Endoscopic Magnetic Retriever | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PTQ | Endoscopic Magnetic Retriever, Exempt | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODD | Endoscopic Retrograde Cholangiopancreatography (Ercp) Cannula | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCU | Endoscopic Storage Cover | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ODE | Endoscopic Suture/Plication System, Gastroesophageal Reflux Disease (Gerd) | 2 | 876.1500 | Endoscope and accessories. | N | Y
Gastroenterology/urology | OCW | Endoscopic Tissue Approximation Device | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | QIJ | Endoscopic Transhepatic Venous Access Needle | 2 | 876.1050 |  | N | N
Gastroenterology/urology | ODG | Endoscopic Ultrasound System, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FET | Endoscopic Video Imaging System/Component, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OCS | Endoscopic Video Imaging System/Component, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCE | Enema Kit | 1 | 876.5210 | Enema kit. | N | N
Gastroenterology/urology | PRM | Enteral Administration Kit (Adult & Pediatric) | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PIO | Enteral Specific Transition Connectors | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PNR | Enteral Syringes With Enteral Specific Connectors | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FDA | Enteroscope And Accessories | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PLQ | Enzyme Packed Cartridge | 2 | 876.5985 | Enzyme packed cartridge. | N | N
Gastroenterology/urology | EXZ | Equipment, Photographic, For Physiologic Function Monitor | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PID | Esophageal Dilator Balloon With Or Without Electrode Sensors | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRQ | Esophageal Dilator Kit | 2 | 876.5365 | Esophageal dilator. | N | N
Gastroenterology/urology | PIE | Esophageal Dilator With Balloon And Electrode Sensors | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | QGG | Esophageal Sheath With Distal Balloon Anchor | 2 | 876.1510 |  | N | N
Gastroenterology/urology | QIS | Esophageal, Mucosal, Electrical Characterization | 2 | 876.1450 |  | N | N
Gastroenterology/urology | GCL | Esophagoscope, General & Plastic Surgery | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDW | Esophagoscope, Rigid, Gastro-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FFD | Evacuator, Bladder, Manually Operated | 1 | 876.4370 | Gastroenterology-urology evacuator. | N | N
Gastroenterology/urology | KQT | Evacuator, Gastro-Urology | 2 | 876.4370 | Gastroenterology-urology evacuator. | N | N
Gastroenterology/urology | MNG | External Urethral Occluder, Urinary Incontinence-Control, Female | 1 | 876.5160 | Urological clamp for males. | N | N
Gastroenterology/urology | PRR | Feeding Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FBW | Filiform And Filiform Follower | 1 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | FKJ | Filter, Blood, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FIS | Flowmeter, Dialysate | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | PPB | Foley Catheter Kit (Excludes Hiv Testing) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | KGE | Forceps, Biopsy, Electric | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | QEC | Forceps, Biopsy, Electric Surgical Hemostasis Within Tracheobronchial Tree | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | NLU | Forceps, Biopsy, Electric, Reprocessed | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FCL | Forceps, Biopsy, Non-Electric | 1 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | NON | Forceps, Biopsy, Non-Electric, Reprocessed | 1 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FJR | Forceps, Disconnect | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | PKI | Fully Subcutaneous Implanted Hemodialysis Catheter | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | EYQ | Garment, Protective, For Incontinence | 1 | 876.5920 | Protective garment for incontinence. | N | N
Gastroenterology/urology | PRS | Gastric Colonic Catheter Irrigation Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRT | Gastric Irrigation And Aspiration Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRU | Gastric Lavage Kit (Adult Or Pediatric Use) | 2 | 876.4370 | Gastroenterology-urology evacuator. | N | N
Gastroenterology/urology | QFZ | Gastroenterology-Urology Accessories To A Biopsy Instrument | 1 | 876.1080 | Gastroenterology-urology accessories to a biopsy instrument. | N | N
Gastroenterology/urology | NYV | Gastrointestinal Motility System, Capsule | 2 | 876.1725 | Gastrointestinal motility monitoring system. | N | N
Gastroenterology/urology | PRW | Gastrointestinal Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PIF | Gastrointestinal Tubes With Enteral Specific Connectors | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PLI | Gastrointestional Tube Holder | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FDS | Gastroscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCK | Gastroscope, General & Plastic Surgery | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PRX | Gastrostomy Guide Wire Placement Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRY | Gastrostomy Jejunostomy Feeding Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRZ | Gastrostomy Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PRV | Gastro-Urology Biopsy Kit | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FJM | Guard, Shunt | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FKF | Gun, Tie, Dialysis | 1 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | PEH | Hemodialysis Catheter Luer End Cap | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | ONW | Hemodialysis System For Home Use | 2 | 876.5860 | High permeability hemodialysis system. | N | N
Gastroenterology/urology | PPI | Hemodialysis Tray (Not A Start/Stop Kit) | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | MSF | Hemodialyzer, Re-Use, High Flux | 2 | 876.5860 | High permeability hemodialysis system. | Y | N
Gastroenterology/urology | MSE | Hemodialyzer, Re-Use, Low Flux | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | PKL | Hemostatic Metal Clip For The Gi Tract | 2 | 876.4400 | Hemorrhoidal ligator. | Y | Y
Gastroenterology/urology | PLP | High Intensity Ultrasound System For Prostate Tissue Ablation | 2 | 876.4340 | High intensity ultrasound system for prostate tissue ablation. | N | N
Gastroenterology/urology | FHQ | Holder, Needle, Gastroenterologic | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | EYJ | Holder, Ureteral Catheter | 1 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FHB | Hook, Gastro-Urology | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FFS | Image, Illumination, Fiberoptic, For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PEW | Implantable Transprostatic Tissue Retractor System | 2 | 876.5530 | Implantable transprostatic tissue retractor system. | N | Y
Gastroenterology/urology | QGD | Implanted Mechanical/Hydraulic Urinary Continence Device Surgical Accessories | 1 | 876.5290 | Implanted mechanical/hydraulic urinary continence device surgical accessories. | N | N
Gastroenterology/urology | QFQ | Ingested, Transient, Space Occupying Device For Weight Management And/Or Weight Loss | 2 | 876.5982 |  | N | N
Gastroenterology/urology | FKL | Insert, Pump, Blood | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | KNW | Instrument, Biopsy | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | QHC | Instrument, Biopsy, Core, Soft Tissue, Female Reproductive Organs | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FCF | Instrument, Biopsy, Mechanical, Gastrointestinal | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | NLS | Instrument, Biopsy, Reprocessed | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FCK | Instrument, Biopsy, Suction | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FEX | Instrument, Catheter, Punch | 1 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | FEI | Instrument, Special Lens, For Endoscope | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCX | Insufflator, Automatic Carbon-Dioxide For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | LCG | Intestinal Splinting Tubes | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FJT | Intracatheter, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | EXD | Irrigator, Ostomy | 2 | 876.5895 | Ostomy irrigator. | N | N
Gastroenterology/urology | FHY | Jelly, Contact, For Transurethral Surgical Instrument | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FHX | Jelly, Lubricating, For Transurethral Surgical Instrument | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PIN | Kidney Perfusion Kit | 2 | 876.5880 | Isolated kidney perfusion and transport system and accessories. | N | N
Gastroenterology/urology | NWQ | Kit, Catheter, External, Male (Excludes Hiv Testing) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | NWR | Kit, Catheter, Foley (Excludes Hiv Testing) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | NWO | Kit, Catheter, Urinary (Exludes Hiv Testing) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | NWS | Kit, Choledochoscope, Flexible And Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FGA | Kit, Nephroscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NFK | Kit, Repair, Catheter, Hemodialysis | 2 | 876.5540 | Blood access device and accessories. | Y | N
Gastroenterology/urology | FTI | Lamp, Endoscope, Incandescent | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCJ | Laparoscope, General & Plastic Surgery | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NLM | Laparoscope, General & Plastic Surgery, Reprocessed | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | QJN | Laparoscopic Accessories, Esophageal Sizing | 2 | 876.5360 |  | N | N
Gastroenterology/urology | PSD | Laparoscopic Gastrostomy Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PSA | Laparoscopic Jejunostomy Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | OTJ | Laparoscopic Single Port Access Device | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDE | Laparoscopy Kit | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | GCI | Laryngoscope, Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NTN | Led Light Source | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | MND | Ligator, Esophageal | 2 | 876.4400 | Hemorrhoidal ligator. | N | N
Gastroenterology/urology | FHN | Ligator, Hemorrhoidal | 2 | 876.4400 | Hemorrhoidal ligator. | N | N
Gastroenterology/urology | GCT | Light Source, Endoscope, Xenon Arc | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCW | Light Source, Fiberoptic, Routine | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCQ | Light Source, Incandescent, Diagnostic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCR | Light Source, Photographic, Fiberoptic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FCS | Light, Catheter, Fiberoptic, Glass, Ureteral | 2 | 876.4020 | Fiberoptic light ureteral catheter. | N | N
Gastroenterology/urology | LQC | Lithotriptor, Biliary Mechanical | 2 | 876.4500 | Mechanical lithotriptor. | N | N
Gastroenterology/urology | FFK | Lithotriptor, Electro-Hydraulic | 2 | 876.4480 | Electrohydraulic lithotriptor. | N | N
Gastroenterology/urology | LNS | Lithotriptor, Extracorporeal Shock-Wave, Urological | 2 | 876.5990 | Extracorporeal shock wave lithotripter. | N | N
Gastroenterology/urology | FEO | Lithotriptor, Ultrasonic | 2 | 876.4480 | Electrohydraulic lithotriptor. | N | N
Gastroenterology/urology | QKZ | Magnetically Maneuvered Capsule Endoscopy System | 2 | 876.1310 |  | N | N
Gastroenterology/urology | PPC | Male External Catheterization Kit (Excludes Hiv Testing) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FJA | Manometer, Water | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | NBG | Marker, Colon | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NDS | Media, Culture, Ex Vivo, Tissue And Cell | 2 | 876.5885 | Tissue culture media for human ex vivo tissue and cell culture processing applications. | N | N
Gastroenterology/urology | GCH | Mediastinoscope, Diagnostic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FLB | Meter, Conductivity, Induction, Remote Type | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FIZ | Meter, Conductivity, Non-Remote | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | ODF | Mini Endoscope, Gastroenterology-Urology | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | QGY | Mobile / Tablet Software Application To Control Settings Of Surgical And Endoscopic Camera | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KLA | Monitor, Esophageal Motility, And Tube | 2 | 876.1725 | Gastrointestinal motility monitoring system. | N | N
Gastroenterology/urology | FLA | Monitor, Temperature, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | PPO | Multi-Lumen Hemodialysis Catheterization Kit | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | PSB | Nasogastric Feed Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PSC | Nasogastric-Gastrojejunal Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | NMX | Needle And Needle Set, Gastro-Urology, Reprocessed | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | PSE | Needle Catheter Jejunostomy Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | NMY | Needle, Endoscopic, Reprocessed | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FIE | Needle, Fistula | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | NNE | Needle, Fistula, Reprocessed | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FHP | Needle, Pneumoperitoneum, Simple | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | NLX | Needle, Pneumoperitoneum, Spring Loaded, Reprocessed | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PPE | Nephroscope Set | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PSF | Non-Balloon Replacement Gastrostomy System | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | QHH | Non-Implanted Nerve Stimulator For Pain Associated With Irritable Bowel Syndrome (Ibs) | 2 | 876.5340 |  | N | N
Gastroenterology/urology | FEC | Obturator, For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | ONY | Oral Removable Retainer For Weight Management | 2 | 876.5981 | Oral removable palatal space occupying device for weight management and/or weight loss. | N | N
Gastroenterology/urology | OLH | Orchidometer | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | PQE | Ostomy Care Kit | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | EZT | Pacemaker, Bladder | 3 | 876.5270 | Implanted electrical urinary continence device. | N | Y
Gastroenterology/urology | PCU | Pancreatic Stent, Covered, Metallic, Removable | 2 | 876.5015 | Pancreatic drainage stent and delivery system. | N | Y
Gastroenterology/urology | FAK | Panendoscope (Gastroduodenoscope) | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FAL | Panendoscope (Urethroscope) | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PTK | Pediatric Esophageal Atresia Anastomosis Device | 3 | 876.5980 | Gastrointestinal tube and accessories. | Y | N
Gastroenterology/urology | QIR | Pediatric Hemodialysis System | 2 | 876.5861 |  | Y | N
Gastroenterology/urology | QGA | Penile Implant Surgical Accessories | 1 | 876.3500 | Penile implant surgical accessories. | N | N
Gastroenterology/urology | PSG | Percutaneous Endoscopic Gastrostomy Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PSH | Percutaneous Replacement Gastrostomy Tube Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PSI | Perineal Irrigation Kit | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | KPP | Peritoneal Dialysate Filter | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | PPJ | Peritoneal Dialysis Administration Kit | 2 | 876.5630 | Peritoneal dialysis system and accessories. | N | N
Gastroenterology/urology | PPK | Peritoneal Dialysis Tubing Kit | 2 | 876.5630 | Peritoneal dialysis system and accessories. | N | N
Gastroenterology/urology | PNG | Peritoneal, Drainage Catheter For Refractory Ascites, Long-Term Indwelling | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | Y
Gastroenterology/urology | GCG | Peritoneoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | LRI | Ph Meter For Dialysis Hydrogen Ion Concentration | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | EXW | Pill, Radio | 2 | 876.1725 | Gastrointestinal motility monitoring system. | N | N
Gastroenterology/urology | FDB | Plate, Patient | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FJY | Plier, Crimp | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FJO | Plier, Tube | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FHO | Pneumoperitoneum Needle | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | EZQ | Pouch, Colostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | QEF | Pressure Ulcer Management Tool | 2 | 876.2100 |  | N | N
Gastroenterology/urology | FGM | Probe And Director, Gastro-Urology | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | EXX | Probe, Rectal, Non-Powered | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | GCF | Proctoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OQT | Prostate Lesion, Documentation, System | 2 | 876.2050 | Prostate lesion documentation system. | N | N
Gastroenterology/urology | FAE | Prosthesis, Penile | 2 | 876.3630 | Penile rigidity implant. | N | Y
Gastroenterology/urology | JCW | Prosthesis, Penis, Inflatable | 3 | 876.3350 | Penile inflatable implant. | N | Y
Gastroenterology/urology | FTQ | Prosthesis, Penis, Rigid Rod | 2 | 876.3630 | Penile rigidity implant. | N | Y
Gastroenterology/urology | FTO | Prosthesis, Testicle | 3 | 876.3750 | Testicular prosthesis. | N | Y
Gastroenterology/urology | FAF | Prosthesis, Testicular | 3 | 876.3750 | Testicular prosthesis. | N | Y
Gastroenterology/urology | EXE | Protector, Ostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | FIB | Protector, Transducer, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FEQ | Pump, Air, Non-Manual, For Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FIR | Pump, Blood, Extra-Luminal | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FIH | Pump, Infusion Or Syringe, Extra-Luminal | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FCI | Punch, Biopsy | 2 | 876.1075 | Gastroenterology-urology biopsy instrument. | N | N
Gastroenterology/urology | FES | Recorder, External, Pressure, Amplifier & Transducer | 2 | 876.1725 | Gastrointestinal motility monitoring system. | N | N
Gastroenterology/urology | PSJ | Rectal Dilator Kit | 1 | 876.5450 | Rectal dilator. | N | N
Gastroenterology/urology | PBP | Rectal Insert | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | PSK | Replacement Gastrostomy Feeding Catheter Tray | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FJL | Resectoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDC | Resectoscope, Working Element | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FDG | Retractor, Fiberoptic | 1 | 876.4530 | Gastroenterology-urology fiberoptic retractor. | N | N
Gastroenterology/urology | FGN | Retractor, Non-Self-Retaining | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FFO | Retractor, Self-Retaining | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | EYS | Ribdam | 1 | 876.4560 | Ribdam. | N | N
Gastroenterology/urology | FJX | Ring, Crimp | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FJW | Ring, Joint | 1 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FHI | Ring, Laparotomy | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | EZP | Rod, Colostomy | 2 | 876.4270 | Colostomy rod. | N | N
Gastroenterology/urology | FBI | Rongeur, Cystoscopic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KDO | Rongeur, Cystoscopic, Hot | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KGD | Scissors For Cystoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PTR | Scissors For Cytoscope, Exempt | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FHL | Scoop | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | EXA | Selector, Size, Ostomy | 1 | 876.5900 | Ostomy pouch and accessories. | N | N
Gastroenterology/urology | KDJ | Set, Administration, For Peritoneal Dialysis, Disposable | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | LBW | Set, Dialysis, Single Needle (Co-Axial Flow) | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | FIF | Set, Dialysis, Single Needle With Uni-Directional Pump | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | NNG | Set, Dialysis, Single Needle With Uni-Directional Pump, Reprocessed | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FKI | Set, Dialyzer Holder | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FHT | Set, Gavage, Infant, Sterile | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | LEY | Set, Oral Administration | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | KDL | Set, Perfusion, Kidney, Disposable | 2 | 876.5880 | Isolated kidney perfusion and transport system and accessories. | N | N
Gastroenterology/urology | FJK | Set, Tubing, Blood, With And Without Anti-Regurgitation Valve | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | NNY | Sheath, Corrugated Rubber, For Non-Indwelling Catheter, Non-Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | EYT | Sheath, Corrugated Rubber, For Nonindwelling Catheter, Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | KPM | Shunt, Peritoneal | 2 | 876.5955 | Peritoneo-venous shunt. | N | Y
Gastroenterology/urology | FAM | Sigmoidoscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FAN | Sigmoidoscope, Rigid, Electrical | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KDM | Sigmoidoscope, Rigid, Non-Electrical | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PPL | Single Needle (Co-Axial Flow) Dialysis Kit | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | PPH | Single Needle With Uni-Directional Pump Kit | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FDI | Snare, Flexible | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | NLT | Snare, Flexible, Reprocessed | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FGX | Snare, Non-Electrical | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FDJ | Snare, Rigid Self-Opening | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FKH | Solution-Test Standard-Conductivity, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FBS | Sound, Metal, Interconnected | 1 | 876.4590 | Interlocking urethral sound. | N | N
Gastroenterology/urology | FBX | Sound, Urethral, Metal Or Plastic | 1 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | PGU | Spatial Imaging For Display Of Endoscope Position | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FFQ | Speculum, Rectal | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FDR | Sphyncteroscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FHK | Spreader, Bladder Neck | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FJN | Stabilizer, Shunt | 2 | 876.5540 | Blood access device and accessories. | N | N
Gastroenterology/urology | FLC | Station, Dialysis Control, Negative Pressure Type | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FAD | Stent, Ureteral | 2 | 876.4620 | Ureteral stent. | N | Y
Gastroenterology/urology | FGE | Stents, Drains And Dilators For The Biliary Ducts | 2 | 876.5010 | Biliary catheter and accessories. | Y | Y
Gastroenterology/urology | EZW | Stimulator, Electrical, Implantable, For Incontinence | 3 | 876.5270 | Implanted electrical urinary continence device. | N | Y
Gastroenterology/urology | KPI | Stimulator, Electrical, Non-Implantable, For Incontinence | 2 | 876.5320 | Nonimplanted electrical continence device. | N | N
Gastroenterology/urology | NAM | Stimulator, Peripheral Nerve, Non-Implanted, For Pelvic Floor Dysfunction | 2 | 876.5310 | Nonimplanted, peripheral electrical continence device. | N | N
Gastroenterology/urology | EYD | Stirrups | 1 | 876.4890 | Urological table and accessories. | N | N
Gastroenterology/urology | FFW | String And Tubes, Gastrointestinal, To Locate Internal Bleeding | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | MNV | Strip, Dialysate Ph Indicator | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | NTZ | Strip, Dialysate, Ph, Bicarbonate, Glucose, Acid, Indicator | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | MSZ | Strip, Test, Reagent, Residuals For Dialysate | 2 | 876.5665 | Water purification system for hemodialysis. | Y | N
Gastroenterology/urology | MSY | Strip, Test, Reagent, Residuals For Dialysate, Disinfectant | 2 | 876.5665 | Water purification system for hemodialysis. | Y | N
Gastroenterology/urology | EZB | Stylet For Catheter, Gastro-Urology | 1 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | EYA | Stylet, Ureteral | 1 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | PLL | Submucosal Injection Agent | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | FKR | Subsystem, Proportioning | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FIP | Subsystem, Water Purification | 2 | 876.5665 | Water purification system for hemodialysis. | N | N
Gastroenterology/urology | EXN | Support, Hernia | 1 | 876.5970 | Hernia support. | N | N
Gastroenterology/urology | QGC | Suprapubic Catheter Accessories | 1 | 876.5100 | Suprapubic catheter accessories. | N | N
Gastroenterology/urology | NWP | Suprapubic Cystostomy Catheter Kit | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | KOA | Surgical Instruments, G-U, Manual (And Accessories) | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | KXM | System Accessories, Extracorporeal | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | MSB | System And Accessories, Isolated Heart, Transport And Preservation | 2 | 876.5880 | Isolated kidney perfusion and transport system and accessories. | N | N
Gastroenterology/urology | FFI | System, Alarm, Electrosurgical | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | LLB | System, Blood, Extracorporeal And Accessories | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FKQ | System, Dialysate Delivery, Central Multiple Patient | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FIK | System, Dialysate Delivery, Recirculating | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FIJ | System, Dialysate Delivery, Recirculating, Single Pass | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FII | System, Dialysate Delivery, Sealed | 2 | 876.5860 | High permeability hemodialysis system. | Y | N
Gastroenterology/urology | KPF | System, Dialysate Delivery, Semi-Automatic, Peritoneal | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | FIL | System, Dialysate Delivery, Single Pass | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FKP | System, Dialysate Delivery, Single Patient | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | FKT | System, Dialysate Delivery, Sorbent Regenerated | 2 | 876.5600 | Sorbent regenerated dialysate delivery system for hemodialysis. | Y | N
Gastroenterology/urology | MYE | System, Electrogastrography (Egg) | 2 | 876.1735 | Electrogastrography system. | N | N
Gastroenterology/urology | FHF | System, Evacuator, Fluid | 2 | 876.4370 | Gastroenterology-urology evacuator. | N | N
Gastroenterology/urology | FFX | System, Gastrointestinal Motility (Electrical) | 2 | 876.1725 | Gastrointestinal motility monitoring system. | N | N
Gastroenterology/urology | MQS | System, Hemodialysis, Access Recirculation Monitoring | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | MON | System, Hemodialysis, Remote Accessories | 1 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | NSI | System, Imaging, Esophageal, Wireless, Capsule | 2 | 876.1300 | Ingestible telemetric gastrointestinal capsule imaging system. | N | N
Gastroenterology/urology | NEZ | System, Imaging, Gastrointestinal, Wireless, Capsule | 2 | 876.1300 | Ingestible telemetric gastrointestinal capsule imaging system. | N | N
Gastroenterology/urology | LJH | System, Irrigation, Urological | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | KDN | System, Perfusion, Kidney | 2 | 876.5880 | Isolated kidney perfusion and transport system and accessories. | N | N
Gastroenterology/urology | FKX | System, Peritoneal, Automatic Delivery | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | PTW | System, Peritoneal, Automatic Delivery, Exempt | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | NAY | System, Surgical, Computer Controlled Instrument | 2 | 876.1500 | Endoscope and accessories. | N | N | Verb surgical
Gastroenterology/urology | NNZ | System, Urine Drainage, Closed, For Non-Indwelling Catheter, Non-Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | EYZ | System, Urine Drainage, Closed, For Nonindwelling Catheter, Sterile | 1 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | FFE | System, Water Jet Catheter, Renal | 2 | 876.4650 | Water jet renal stone dislodger system. | N | N
Gastroenterology/urology | NHV | System, Water Purification, General Medical Use | 2 | 876.5665 | Water purification system for hemodialysis. | N | N
Gastroenterology/urology | EYH | Table, Cysto, Non-Electrical | 1 | 876.4890 | Urological table and accessories. | N | N
Gastroenterology/urology | MMZ | Table, Cystometric, Electric | 2 | 876.4890 | Urological table and accessories. | N | N
Gastroenterology/urology | KQS | Table, Cystometric, Non-Electric And Accessories | 1 | 876.4890 | Urological table and accessories. | N | N
Gastroenterology/urology | FIN | Tank, Holding, Dialysis | 2 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FBP | Telescope, Rigid, Endoscopic | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | QKA | Temporarily-Placed Urethral Opening System For Symptoms Of Benign Prostatic Hyperplasia | 2 | 876.5510 |  | N | N
Gastroenterology/urology | FKE | Tie, Dialysis | 1 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FKW | Tip, Vessel | 2 | 876.5540 | Blood access device and accessories. | N | Y
Gastroenterology/urology | GCW | Transformer, Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OEJ | Transurethral Electrosurgical Unit, Benign Prostatic Hyperplasia | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FCH | Tray, Biopsy, Without Biopsy Instruments | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | FCM | Tray, Catheterization, Sterile Urethral, With Or Without Catheter (Kit) | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | EYN | Tray, Irrigation, Sterile | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FKG | Tray, Start/Stop (Including Contents), Dialysis | 1 | 876.5820 | Hemodialysis system and accessories. | N | N
Gastroenterology/urology | FGK | Tripsor, Stone, Bladder | 2 | 876.4500 | Mechanical lithotriptor. | N | N
Gastroenterology/urology | FBQ | Trocar, Gastro-Urology | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | EXM | Truss, Umbilical | 1 | 876.5970 | Hernia support. | N | N
Gastroenterology/urology | FEG | Tube, Double Lumen For Intestinal Decompression And/Or Intubation | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FFA | Tube, Drainage, Suprapubic | 2 | 876.5090 | Suprapubic urological catheter and accessories. | N | N
Gastroenterology/urology | FPD | Tube, Feeding | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | KGC | Tube, Gastro-Enterostomy | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FRQ | Tube, Levine | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | BSS | Tube, Nasogastric | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FEF | Tube, Single Lumen, With Mercury Wt Balloon For Intestinal Intubation And / Or Decompression | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FCZ | Tube, Smoke Removal, Endoscopic | 1 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | KNT | Tubes, Gastrointestinal (And Accessories) | 2 | 876.5980 | Gastrointestinal tube and accessories. | N | N
Gastroenterology/urology | FID | Tubing, Dialysate | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | KQQ | Tubing, Dialysate (And Connector) | 2 | 876.5820 | Hemodialysis system and accessories. | Y | N
Gastroenterology/urology | ODN | Ultrafiltration-Controlled Nocturnal Dialysate Delivery System | 2 | 876.5860 | High permeability hemodialysis system. | Y | N
Gastroenterology/urology | FAR | Unit, Electrosurgical | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | KNS | Unit, Electrosurgical, Endoscopic (With Or Without Accessories) | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | NLR | Unit, Electrosurgical, Endoscopic (With Or Without Accessories), Reprocessed | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | PPD | Universal Drainage Tray | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | QGB | Ureteral Stent Accessories | 1 | 876.4630 | Ureteral stent accessories. | N | N
Gastroenterology/urology | FGB | Ureteroscope And Accessories, Flexible/Rigid | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | PIH | Urethral Insert With Pump For Bladder Drainage | 2 | 876.5140 | Urethral insert with pump for bladder drainage. | N | Y
Gastroenterology/urology | FBR | Urethrometer | 1 | 876.5520 | Urethral dilator. | N | N
Gastroenterology/urology | FGC | Urethroscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | EZO | Urethrotome | 2 | 876.4770 | Urethrotome. | N | N
Gastroenterology/urology | PPG | Urinary Drainage Collection Kit | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | FCN | Urinary Drainage Collection Kit, For Indwelling Catheter | 2 | 876.5250 | Urine collector and accessories. | N | N
Gastroenterology/urology | PPF | Urinary Irrigation Kit | 2 | 876.5130 | Urological catheter and accessories. | N | N
Gastroenterology/urology | EXS | Urinometer, Electrical | 2 | 876.1800 | Urine flow or volume measuring system. | N | N
Gastroenterology/urology | EXR | Urinometer, Mechanical | 2 | 876.1800 | Urine flow or volume measuring system. | N | N
Gastroenterology/urology | EXY | Uroflowmeter | 2 | 876.1800 | Urine flow or volume measuring system. | N | N
Gastroenterology/urology | PVY | Vasovasostomy Set | 1 | 876.4730 | Manual gastroenterology-urology surgical instrument and accessories. | N | N
Gastroenterology/urology | PFH | Venous Window Needle Guide | 2 | 876.5540 | Blood access device and accessories. | Y | Y
Gastroenterology/urology | PIA | Vibrator For Climax Control Of Premature Ejaculation | 2 | 876.5025 | Vibrator for climax control of premature ejaculation. | N | N
Gastroenterology/urology | MLW | Warmer, Peritoneal Dialysate | 2 | 876.5630 | Peritoneal dialysis system and accessories. | Y | N
Gastroenterology/urology | NVE | Washer, Cleaner, Automated, Endoscope | 2 | 876.1500 | Endoscope and accessories. | N | N
Gastroenterology/urology | OEK | Water-Induced Thermotherapy System, Benign Prostatic Hyperplasia | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
Gastroenterology/urology | FDL | Wristlet, Patient Return | 2 | 876.4300 | Endoscopic electrosurgical unit and accessories. | N | N
General & plastic surgery | PMX | Absorbable Collagen Hemostatic Agent With Thrombin | 3 | 878.4490 | Absorbable hemostatic agent and dressing. | N | Y
General & plastic surgery | OMT | Absorbable Lung Biopsy Plug | 2 | 878.4755 | Absorbable lung biopsy plug. | N | Y
General & plastic surgery | KGZ | Accessories, Catheter | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | LYU | Accessory, Surgical Apparel | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | GCE | Adaptor, Catheter | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | LMF | Agent, Absorbable Hemostatic, Collagen Based | 3 | 878.4490 | Absorbable hemostatic agent and dressing. | N | Y
General & plastic surgery | LMG | Agent, Absorbable Hemostatic, Non-Collagen Based | 3 | 878.4490 | Absorbable hemostatic agent and dressing. | N | Y
General & plastic surgery | ORC | Air Filter Portable Apparatus | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | QGI | Air-Handling Apparatus Accessory | 1 | 878.5080 | Air-handling apparatus accessory. | N | N
General & plastic surgery | NEH | Anchor, Fascial | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | PLY | Antimicrobial Drapes | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | FZG | Apparatus, Air Handling, Bench | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | FZI | Apparatus, Air Handling, Enclosure | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | FZH | Apparatus, Air Handling, Room | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | HAM | Apparatus, Electrosurgical | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | FYD | Apparatus, Exhaust, Surgical | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | GCY | Apparatus, Suction, Single Patient Use, Portable, Nonpowered | 1 | 878.4680 | Nonpowered, single patient, portable suction apparatus. | N | N
General & plastic surgery | JCX | Apparatus, Suction, Ward Use, Portable, Ac-Powered | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | FHM | Apparatus, Suturing, Stomach And Intestinal | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FYI | Appliance, Facial Fracture, External | 1 | 878.3250 | External facial fracture fixation appliance. | N | N
General & plastic surgery | NVJ | Applicator, Transurethral, Radio Frequency, For Stress Urinary Incontinence In Women | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | HBT | Applier, Hemostatic Clip | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEF | Applier, Staple, Surgical, | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDO | Applier, Surgical, Clip | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OJT | Aspiration Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | QJF | Autofluorescence Imaging Adjunct Tool For Wounds | 2 | 878.4550 |  | N | N
General & plastic surgery | KGY | Bag, Intestine | 1 | 878.4100 | Organ bag. | N | N
General & plastic surgery | PXN | Basic Set-Up Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MOQ | Battery, Replacement, Rechargeable | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | KOZ | Beads, Hydrophilic, For Wound Exudate Absorption | 1 | 878.4018 | Hydrophilic wound dressing. | N | N
General & plastic surgery | NNA | Biopsy Device, Percutaneous, Reprocessed | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFG | Bit, Surgical | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GFA | Blade, Saw, General & Plastic Surgery, Surgical | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | DWH | Blade, Saw, Surgical, Cardiovascular | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GES | Blade, Scalpel | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | NOD | Board, Arm (With Cover), Non-Sterile | 1 | 878.3910 | Noninflatable extremity splint. | N | N
General & plastic surgery | BTX | Board, Arm (With Cover), Sterile | 1 | 878.3910 | Noninflatable extremity splint. | N | N
General & plastic surgery | OKG | Body Fluid Clean Up Kit | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | OKF | Body Fluid Disposal Kit | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | OKH | Body Fluids Barrier Kit | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | PXO | Bone Marrow Biopsy Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LEJ | Booth, Sun Tan | 2 | 878.4635 | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | N | N
General & plastic surgery | PXP | Breast Biopsy / Localization Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEE | Brush, Biopsy, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GED | Brush, Dermabrasion, Manual | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFE | Brush, Dermabrasion, Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GEC | Brush, Scrub, Operating-Room | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFF | Bur, Surgical, General & Plastic Surgery | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | PXQ | Burn Dressing Tray | 1 | 878.4022 | Hydrogel wound dressing and burn dressing. | N | N
General & plastic surgery | OJJ | Burn Kit Without Drug | 1 | 878.4022 | Hydrogel wound dressing and burn dressing. | N | N
General & plastic surgery | KGL | Cabinet, Phototherapy (Puva) | 2 | 878.4630 | Ultraviolet lamp for dermatologic disorders. | N | N
General & plastic surgery | FTS | Camera, Cine, Endoscopic, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWL | Camera, Cine, Endoscopic, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWK | Camera, Cine, Microsurgical, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWJ | Camera, Cine, Microsurgical, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWI | Camera, Cine, Surgical, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWH | Camera, Cine, Surgical, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FXM | Camera, Still, Endoscopic | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FTH | Camera, Still, Microsurgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FTT | Camera, Still, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | KQM | Camera, Surgical And Accessories | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWG | Camera, Television, Endoscopic, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWF | Camera, Television, Endoscopic, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWE | Camera, Television, Microsurgical, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWD | Camera, Television, Microsurgical, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWC | Camera, Television, Surgical, With Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FWB | Camera, Television, Surgical, Without Audio | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | JYC | Cannula, Ear | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FGY | Cannula, Injection | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | KAM | Cannula, Sinus | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEA | Cannula, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KCE | Cannulae, Bronchial | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FYF | Cap, Surgical | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | PQN | Carbon Dioxide Gas Controlled Tissue Expander | 2 | 878.3510 |  | N | Y
General & plastic surgery | GEJ | Carrier, Ligature | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GBA | Catheter, Balloon Type | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBZ | Catheter, Cholangiography | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBQ | Catheter, Continuous Irrigation | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | OAJ | Catheter, Drainage, Intraoral/Extraoral | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBY | Catheter, Eustachian, General & Plastic Surgery | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | JCY | Catheter, Infusion | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBX | Catheter, Irrigation | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBP | Catheter, Multiple Lumen | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBO | Catheter, Nephrostomy, General & Plastic Surgery | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBN | Catheter, Pediatric, General & Plastic Surgery | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBW | Catheter, Peritoneal | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | GBS | Catheter, Ventricular, General & Plastic Surgery | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | PXR | Cement And Bone Removal Kit | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | LZN | Cement Obturator | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | GBB | Chair, Surgical, Ac-Powered | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | FZK | Chair, Surgical, Non-Electrical | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | KPJ | Chamber, Oxygen, Topical, Extremity | 2 | 878.5650 | Topical oxygen chamber for extremities. | N | N
General & plastic surgery | KDG | Chisel (Osteotome) | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | JYD | Chisel, Mastoid | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KAN | Chisel, Nasal | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FZO | Chisel, Surgical, Manual | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PXS | Cholecystectomy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HXD | Clamp | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDJ | Clamp, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MCH | Clip, Hemostatic | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | FZP | Clip, Implantable | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | NCA | Clip, Implantable, For Coronary Artery Bypass Graft (Cabg) | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | NMJ | Clip, Implantable, Reprocessed | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | FZQ | Clip, Removable (Skin) | 1 | 878.4320 | Removable skin clip. | N | N
General & plastic surgery | NJC | Clip, Vas Deferens | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | MGO | Closure, Wound, Adhesive | 1 | 878.4018 | Hydrophilic wound dressing. | N | N
General & plastic surgery | PIJ | Collagen Surgical Mesh Containing Drugs | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PBD | Composite Cultured Skin | HDE | 878.4024 |  | N | Y
General & plastic surgery | GCD | Connector, Catheter | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | FZR | Contractor, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MMP | Cover, Barrier, Protective | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | FXP | Cover, Shoe, Operating-Room | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | KGQ | Cream, Gloving, Surgeon'S | 1 | 878.4470 | Surgeon's gloving cream. | N | N
General & plastic surgery | ORE | Cryogenic, Skin Tag Removal, Otc | 2 | 878.4350 | Cryosurgical unit and accessories. | N | N
General & plastic surgery | HTF | Curette | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FZS | Curette, Surgical, General Use | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OMD | Cutaneous Tissue Adhesive With Mesh | 2 | 878.4011 | Tissue adhesive with adjunct wound closure device for topical approximation of skin. | N | N
General & plastic surgery | FZT | Cutter, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MAC | Decontamination Kit | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | KOY | Degreaser, Skin, Surgical | 1 | 878.4730 | Surgical skin degreaser or adhesive tape solvent. | N | N
General & plastic surgery | PEM | Dental Barriers And Sleeves | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | OOK | Dermal Cooling Pack/Vacuum/Massager | 2 | 878.4340 | Contact cooling system for aesthetic use. | N | N
General & plastic surgery | GFD | Dermatome | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | MVR | Device, Anastomotic, Microvascular | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | MJG | Device, Percutaneous, Biopsy | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GCC | Dilator, Catheter | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | PPP | Disposable Joint Aspiration/Injection Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDI | Dissector, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OEI | Drainage Catheter With Antibiotic | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | KGT | Drape, Adhesive, Aerosol | 1 | 878.4380 | Drape adhesive. | N | N
General & plastic surgery | HMW | Drape, Microscope, Ophthalmic | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | HMT | Drape, Patient, Ophthalmic | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | EYX | Drape, Pure Latex Sheet, With Self-Retaining Finger Cot | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | KKX | Drape, Surgical | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | ERY | Drape, Surgical, Ent | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | PUI | Drape, Surgical, Exempt | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | EYY | Drape, Urological, Disposable | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | FYE | Dress, Surgical | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | NAE | Dressing, Wound, Hydrogel Without Drug And/Or Biologic | 1 | 878.4022 | Hydrogel wound dressing and burn dressing. | N | N
General & plastic surgery | NAC | Dressing, Wound, Hydrophilic | 1 | 878.4018 | Hydrophilic wound dressing. | N | N
General & plastic surgery | NAD | Dressing, Wound, Occlusive | 1 | 878.4020 | Occlusive wound dressing. | N | N
General & plastic surgery | NWF | Drill, Battery Powered, Subungual Hematoma | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GFC | Driver, Surgical, Pin | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | MDA | Elastomer, Silicone, For Scar Management | 1 | 878.4025 | Silicone sheeting. | N | N
General & plastic surgery | JOS | Electrode, Electrosurgical | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | JOT | Electrode, Gel, Electrosurgical | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | ONQ | Electrosurgical Coagulation For Aesthetic | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | DWG | Electrosurgical Device | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | ODR | Electrosurgical Patient Return Electrode | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | MUK | Electrosurgical Radiofrequency System, Stress Urinary Incontinence, Female, Transvaginal Or Laparoscopic, Pelvic Tissue | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | PDG | Electrosurgical Vessel And/Or Tissue Sealer. With Built-In Generator. | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | GEI | Electrosurgical, Cutting & Coagulation & Accessories | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | NUJ | Electrosurgical, Cutting & Coagulation Accessories, Laparoscopic & Endoscopic, Reprocessed | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | HTE | Elevator | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEG | Elevator, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OKI | Emergency Response Safety Kit | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | QHF | Energy Based Device For Treatment Of Tattoos | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | KCW | Epilator, High Frequency, Needle-Type | 1 | 878.5350 | Needle-type epilator. | N | N
General & plastic surgery | KCX | Epilator, High Frequency, Tweezer-Type | 1 | 878.5360 | Tweezer-type epilator. | N | N
General & plastic surgery | FZW | Expander, Surgical, Skin Graft | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PZL | Extracorporeal Shock Wave Device For Treatment Of Diabetic Foot Ulcers | 2 | 878.4685 | Extracorporeal shock wave device for treatment of chronic wounds. | N | N
General & plastic surgery | NJZ | Extractor, Vein | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | ODU | Facial Implant | 2 | 878.3500 | Polytetrafluoroethylene with carbon fibers composite implant material. | N | Y
General & plastic surgery | OLI | Fat Reducing Low Level Laser | 2 | 878.5400 | Low level laser system for aesthetic use | N | N
General & plastic surgery | GEO | File, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FXT | Film, Camera, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | OHO | First Aid Kit Without Drug | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | PXU | Fluid Drainage Tray | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | PZP | Fluid Jet Removal System | 2 | 876.4350 | Fluid jet system for prostate tissue removal. | N | N
General & plastic surgery | OHV | Focused Ultrasound For Tissue Heat Or Mechanical Cellular Disruption | 2 | 878.4590 | Focused ultrasound stimulator system for aesthetic use. | N | N
General & plastic surgery | HTD | Forceps | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEN | Forceps, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HYA | Forceps, Wire Holding | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FXS | Frame, Camera, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | NXF | Gas, Laser Generating | 1 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | NAB | Gauze / Sponge,Nonresorbable For External Use | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | GDY | Gauze/Sponge, Internal, X-Ray Detectable | 1 | 878.4450 | Nonabsorbable gauze for internal use. | N | N
General & plastic surgery | PZQ | General Laparoscopic Power Morcellation Containment System | 2 | 878.4825 |  | N | N
General & plastic surgery | PXV | General Purpose Instrument Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LRO | General Surgery Tray | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | QGX | General Use Pneumatic Tourniquet | 1 | 878.5910 | Pneumatic tourniquet. | N | N
General & plastic surgery | MUL | Generator, Electrosurgical, Coagulation, Cancer | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | GDH | Gouge, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FYC | Gown, Isolation, Surgical | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | FYB | Gown, Patient | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | FYA | Gown, Surgical | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | LXZ | Guard, Instrument | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LZM | Guard, Skin Graft | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | LXI | Guide, Drill, Ligament | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GDF | Guide, Needle, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FZX | Guide, Surgical, Instrument | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FZY | Hammer, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDZ | Handle, Scalpel | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFB | Head, Surgical, Hammer | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | FXZ | Helmet, Surgical | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | HRQ | Hemostat | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | QAU | Hemostatic Device For Endoscopic Gastrointestinal Use | 2 | 878.4456 | Hemostatic device for intraluminal gastrointestinal use. | N | Y
General & plastic surgery | PXW | Herniorraphy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FXR | Holder, Camera, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | JYK | Holder, Ear Speculum | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | KAG | Holder, Speculum, Ent | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | FXY | Hood, Surgical | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | KIK | Hook, Bone | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDG | Hook, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HBI | Illuminator, Fiberoptic, Surgical Field | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FTF | Illuminator, Non-Remote | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FTG | Illuminator, Remote | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | LZK | Implant, Malar | 2 | 878.3550 | Chin prosthesis. | N | Y
General & plastic surgery | MNF | Implant, Temporal | 2 | 878.3550 | Chin prosthesis. | N | Y
General & plastic surgery | PXX | Incision Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OUB | Instrument For Treatment Of Hyperhidrosis | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | HTZ | Instrument, Cutting, Orthopedic | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HCF | Instrument, Ligature Passing And Knot Tying | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MDM | Instrument, Manual, Surgical, General Use | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KDC | Instrument, Surgical, Disposable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HWE | Instrument, Surgical, Orthopedic, Ac-Powered Motor And Accessory/Attachment | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | KIJ | Instrument, Surgical, Orthopedic, Dc-Powered Motor And Accessory/Attachment | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | HSZ | Instrument, Surgical, Orthopedic, Pneumatic Powered & Accessory/Attachment | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | PDW | Internal Tissue Marker | 2 | 878.4670 | Internal tissue marker. | N | N
General & plastic surgery | PXY | Irrigation Kit (Wounds) | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | OVR | Kit, First Aid, Talking | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | KDD | Kit, Surgical Instrument, Disposable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDN | Knife, Amputation | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | EMF | Knife, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LYV | Label Or Tag, Sterile | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PKD | Laceration Tray | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | HJE | Lamp, Fluorescein, Ac-Powered | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FQP | Lamp, Operating-Room | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FTD | Lamp, Surgical | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | GBC | Lamp, Surgical, Incandescent | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | MXG | Lamp, Uvc, (For Treating Skin And Wounds) | 2 | 878.4630 | Ultraviolet lamp for dermatologic disorders. | N | N
General & plastic surgery | FMK | Lancet, Blood | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PXZ | Laparascopic Sphincterotomy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PKE | Laparoscopic Cholecystectomy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | QCY | Laser Absorbing Particles | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | ORK | Laser Assisted Lipolysis | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PKT | Laser For Disruption Of Adipocyte Cells For Aesthetic Use | 2 | 878.5400 | Low level laser system for aesthetic use | N | N
General & plastic surgery | LNK | Laser For Gastro-Urology Use | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | OEL | Laser, Benign Prostatic Hyperplasia | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | OYW | Laser, Cellulite Appearance | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | NVK | Laser, Dental, Soft Tissue | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PDZ | Lasers For Temporary Increase Of Clear Nail In Patients With Onychomycosis | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PZX | Left Atrial Appendage Clip, Implantable | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | FXQ | Lens, Camera, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | PSN | Light Based Imaging | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | OHS | Light Based Over The Counter Wrinkle Reduction | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | OHT | Light Based Over-The-Counter Hair Removal | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | OKJ | Light Based Treatment For Cold Sores Herpes Simplex Virus-1 | 2 | 878.4860 | Light based energy source device for topical application. | N | N
General & plastic surgery | FTA | Light, Surgical, Accessories | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSZ | Light, Surgical, Carrier | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSY | Light, Surgical, Ceiling Mounted | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSX | Light, Surgical, Connector | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSW | Light, Surgical, Endoscopic | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FST | Light, Surgical, Fiberoptic | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSS | Light, Surgical, Floor Standing | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FSQ | Light, Surgical, Instrument | 2 | 878.4580 | Surgical lamp. | N | N
General & plastic surgery | FTC | Light, Ultraviolet, Dermatological | 2 | 878.4630 | Ultraviolet lamp for dermatologic disorders. | N | N
General & plastic surgery | QKL | Lipoaspirate Washing System For Aesthetic Body Contouring | 2 | 878.5040 | Suction lipoplasty system. | N | N
General & plastic surgery | PYA | Liver Biopsy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FSP | Loupe, Diagnostic/Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OAB | Low Energy Direct Current Thermal Ablation System | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | PNL | Magnetic Surgical System | 2 | 878.4815 | Magnetic surgical instrument system. | N | N
General & plastic surgery | HXL | Mallet | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFJ | Mallet, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KPK | Marker, Ostia, Aorto-Saphenous Vein | 2 | 878.4650 |  | N | Y
General & plastic surgery | NEU | Marker, Radiographic, Implantable | 2 | 878.4300 | Implantable clip. | N | Y
General & plastic surgery | FZZ | Marker, Skin | 1 | 878.4660 | Skin marker. | N | N
General & plastic surgery | FXX | Mask, Surgical | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | NUV | Massager, Vacuum, Light Induced Heating | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PBX | Massager, Vacuum, Radio Frequency Induced Heat | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | NOK | Material, External Aesthetic Restoration, No Adhesive | 1 | 878.3800 | External aesthetic restoration prosthesis. | N | N
General & plastic surgery | GBI | Material, External Aesthetic Restoration, Used With Adhesive | 1 | 878.3800 | External aesthetic restoration prosthesis. | N | N
General & plastic surgery | KKY | Material, Polytetrafluoroethylene Vitreous Carbon, For Maxillofacial Reconstruction | 2 | 878.3500 | Polytetrafluoroethylene with carbon fibers composite implant material. | N | Y
General & plastic surgery | PIR | Mercy Tape 2d And 3d Models | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FTM | Mesh, Surgical | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWT | Mesh, Surgical, Absorbable, Abdominal Hernia | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXM | Mesh, Surgical, Absorbable, Fistula | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXI | Mesh, Surgical, Absorbable, Large Abdominal Wall Defects | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXL | Mesh, Surgical, Absorbable, Organ Support | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWW | Mesh, Surgical, Absorbable, Orthopaedics, Reinforcement Of Tendon | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXF | Mesh, Surgical, Absorbable, Plastic And Reconstructive Surgery | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXC | Mesh, Surgical, Absorbable, Staple Line Reinforcement | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWZ | Mesh, Surgical, Absorbable, Thoracic, Chest Wall Reconstruction | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | JDJ | Mesh, Surgical, Acetabular, Hip, Prosthesis | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWV | Mesh, Surgical, Collagen, Diaphragmatic Hernia | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXN | Mesh, Surgical, Collagen, Fistula | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXK | Mesh, Surgical, Collagen, Large Abdominal Wall Defects | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWY | Mesh, Surgical, Collagen, Orthopaedics, Reinforcement Of Tendon | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXH | Mesh, Surgical, Collagen, Plastic And Reconstructive Surgery | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWS | Mesh, Surgical, Collagen, Plastics, Facial Implants | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXE | Mesh, Surgical, Collagen, Staple Line Reinforcement | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXB | Mesh, Surgical, Collagen, Thoracic, Chest Wall Reconstruction | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | ORQ | Mesh, Surgical, Deployer | 2 | 878.3300 | Surgical mesh. | N | N
General & plastic surgery | OQL | Mesh, Surgical, Deployment Balloon | 2 | 878.3300 | Surgical mesh. | N | N
General & plastic surgery | OTM | Mesh, Surgical, For Stress Urinary Incontinence, Male | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | EZX | Mesh, Surgical, Metal | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWU | Mesh, Surgical, Non-Absorbable, Diaphragmatic Hernia | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWR | Mesh, Surgical, Non-Absorbable, Facial Implants For Plastic Surgery | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXJ | Mesh, Surgical, Non-Absorbable, Large Abdominal Wall Defects | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OWX | Mesh, Surgical, Non-Absorbable, Orthopaedics, Reinforcement Of Tendon | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXG | Mesh, Surgical, Non-Absorbable, Plastic And Reconstructive Surgery | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXD | Mesh, Surgical, Non-Absorbable, Staple Line Reinforcement | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OXA | Mesh, Surgical, Non-Absorbable, Thoracic, Chest Wall Reconstruction | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PAJ | Mesh, Surgical, Non-Synthetic, Urogynecologic, For Apical Vaginal And Uterine Prolapse, Transabdominally Placed | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PAG | Mesh, Surgical, Non-Synthetic, Urogynecologic, For Stress Urinary Incontinence, Retropubic Or Transobturator | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | FTL | Mesh, Surgical, Polymeric | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OTO | Mesh, Surgical, Synthetic, Urogynecologic, For Apical Vaginal And Uterine Prolapse, Transabdominally Placed | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PAH | Mesh, Surgical, Synthetic, Urogynecologic, For Stress Urinary Incontinence, Female, Mini-Sling | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | OTN | Mesh, Surgical, Synthetic, Urogynecologic, For Stress Urinary Incontinence, Retropubic Or Transobturator | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | HRM | Microscope, Operating & Accessories, Ac-Powered, Ophthalmic | 1 | 878.4700 | Surgical microscope and accessories. | N | N
General & plastic surgery | HRB | Microscope, Operating, Non-Electric, Ophthalmic | 1 | 878.4700 | Surgical microscope and accessories. | N | N
General & plastic surgery | EPT | Microscope, Surgical | 1 | 878.4700 | Surgical microscope and accessories. | N | N
General & plastic surgery | FSO | Microscope, Surgical, General & Plastic Surgery | 1 | 878.4700 | Surgical microscope and accessories. | N | N
General & plastic surgery | PYB | Minor Surgical Procedures Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FTX | Mirror, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GEY | Motor, Surgical Instrument, Ac-Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GET | Motor, Surgical Instrument, Pneumatic Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | ONT | N95 Respirator With Antimicrobial/Antiviral Agent | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | GAA | Needle, Aspiration And Injection, Disposable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDM | Needle, Aspiration And Injection, Reusable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | NNC | Needle, Aspiration And Needle, Reprocessed | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | DWO | Needle, Biopsy, Cardiovascular | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | NLK | Needle, Biopsy, Cardiovascular, Reprocessed | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GCB | Needle, Catheter | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | NMT | Needle, Catheter, Reprocessed | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | FHR | Needle, Gastro-Urology | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | NMU | Needle, Gastro-Urology, Reprocessed | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAB | Needle, Suturing, Disposable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GDL | Needle, Suturing, Reusable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MIJ | Needle, Tumor Localization | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | QFC | Negative Pressure Wound Therapy Device For Reduction Of Wound Complications | 2 | 878.4783 |  | N | N
General & plastic surgery | OKO | Negative Pressure Wound Therapy Non-Powered Suction Apparatus | 2 | 878.4683 | Non-Powered suction apparatus device intended for negative pressure wound therapy. | N | N
General & plastic surgery | OMP | Negative Pressure Wound Therapy Powered Suction Pump | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | ONO | Neurosurgical Laser With Mr Thermography | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PAW | Nonabsorbable Expanded Polytetrafluoroethylene Surgical Suture For Chordae Tendinae Repair Or Replacement | 2 | 878.5035 | Nonabsorbable expanded polytetrafluoroethylene surgical suture. | Y | Y
General & plastic surgery | PGZ | Non-Absorbable, Expandable, Hemostatic Sponge For Temporary Internal Use | 2 | 878.4452 | Nonabsorbable expandable hemostatic sponge for temporary internal use. | N | N
General & plastic surgery | OEA | Non-Surgical Isolation Gown | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | FWZ | Operating Room Accessories Table Tray | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | OJH | Orthopedic Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HWM | Osteotome | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GFI | Osteotome, Manual | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OLP | Over-The-Counter Powered Light Based Laser For Acne | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | PAY | Over-The-Counter Radiofrequency Coagulation Device For Wrinkle Reduction | 2 | 878.4420 | Electrosurgical device for over-the-counter aesthetic use. | N | N
General & plastic surgery | HMP | Pad, Eye | 1 | 878.4440 | Eye pad. | N | N
General & plastic surgery | FNW | Pad, Kelly | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | PYC | Paracentesis Tray | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | QDF | Parathyroid Autofluorescence Detection Device | 2 | 878.4550 |  | N | N
General & plastic surgery | QDG | Parathyroid Autofluorescence Imaging Device | 2 | 878.4550 |  | N | N
General & plastic surgery | OXZ | Pediatric/Child Facemask | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | PYD | Pelviscopy Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OXT | Percutaneous Surgical Set With Attachments | 2 | 878.4805 |  | N | N
General & plastic surgery | FTN | Plastic Surgery And Accessories Kit | 1 | 878.3925 | Plastic surgery kit and accessories. | N | N
General & plastic surgery | HTC | Pliers, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | ESF | Polymer, Composite Synthetic Ptfe With Carbon-Fiber, Ent | 2 | 878.3500 | Polytetrafluoroethylene with carbon fibers composite implant material. | N | Y
General & plastic surgery | OPA | Powder-Free Non-Natural Rubber Latex Surgeon'S Gloves | 1 | 878.4460 | Non-powdered surgeon's glove. | N | N
General & plastic surgery | GEX | Powered Laser Surgical Instrument | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | ONG | Powered Laser Surgical Instrument With Microbeam\Fractional Output | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | ONE | Powered Light Based Non-Laser Surgical Instrument | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | ONF | Powered Light Based Non-Laser Surgical Instrument With Thermal Effect | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | QAI | Powered Microneedle Device | 2 | 878.4430 | Microneedling device for aesthetic use. | N | N
General & plastic surgery | OUP | Powered Surgical Instrument For Improvement In The Appearance Of Cellulite | 2 | 878.4790 | Powered surgical instrument for improvement in the appearance of cellulite. | N | N
General & plastic surgery | GBJ | Prosthesis, Adhesive, External | 1 | 878.3750 | External prosthesis adhesive. | N | N
General & plastic surgery | NOJ | Prosthesis, Breast, External, No Adhesive | 1 | 878.3800 | External aesthetic restoration prosthesis. | N | N
General & plastic surgery | KCZ | Prosthesis, Breast, External, Used With Adhesive | 1 | 878.3800 | External aesthetic restoration prosthesis. | N | N
General & plastic surgery | FWM | Prosthesis, Breast, Inflatable, Internal, Saline | 3 | 878.3530 | Silicone inflatable breast prosthesis. | N | Y
General & plastic surgery | FTR | Prosthesis, Breast, Noninflatable, Internal, Silicone Gel-Filled | 3 | 878.3540 | Silicone gel-filled breast prosthesis. | N | Y
General & plastic surgery | FWP | Prosthesis, Chin, Internal | 2 | 878.3550 | Chin prosthesis. | N | Y
General & plastic surgery | FZD | Prosthesis, Ear, Internal | 2 | 878.3590 | Ear prosthesis. | N | Y
General & plastic surgery | ESW | Prosthesis, Esophageal | 2 | 878.3610 | Esophageal prosthesis. | N | Y
General & plastic surgery | JCQ | Prosthesis, Esophagus | 2 | 878.3610 | Esophageal prosthesis. | N | Y
General & plastic surgery | JDK | Prosthesis, Hip, Cement Restrictor | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | FZE | Prosthesis, Nose, Internal | 2 | 878.3680 | Nose prosthesis. | N | Y
General & plastic surgery | KDA | Prosthesis, Ptfe/Carbon-Fiber | 2 | 878.3500 | Polytetrafluoroethylene with carbon fibers composite implant material. | N | Y
General & plastic surgery | JCT | Prosthesis, Tracheal, Expandable | 2 | 878.3720 | Tracheal prosthesis. | Y | Y
General & plastic surgery | NYT | Prosthesis, Tracheal, Expandable, Polymeric | 2 | 878.3720 | Tracheal prosthesis. | N | Y
General & plastic surgery | NWA | Prosthesis, Tracheal, Preformed/Molded | 2 | 878.3720 | Tracheal prosthesis. | N | Y
General & plastic surgery | BTA | Pump, Portable, Aspiration (Manual Or Powered) | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | LRY | Punch, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HTR | Rasp | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAC | Rasp, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LWK | Razor, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MSH | Respirator, Surgical | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | GCZ | Retainer, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KGS | Retention Device, Suture | 1 | 878.4930 | Suture retention device. | N | N
General & plastic surgery | GAD | Retractor | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KGW | Ring (Wound Protector), Drape Retention, Internal | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | HSO | Saw | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | DWI | Saw, Electrically Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | GDR | Saw, Manual And Accessories | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KFK | Saw, Pneumatically Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | HAB | Saw, Powered, And Accessories | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | OLC | Scaffold, Partial Medial Meniscal Defects Extending Into The Red/White Zone, Resorbable Bovine Collagen | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PMC | Scalp Cooling System | 2 | 878.4360 | Scalp cooling system to reduce the likelihood of chemotherapy-induced alopecia. | N | N
General & plastic surgery | GDX | Scalpel, One-Piece | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LRW | Scissors, General, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LXK | Scrapper, Skin Specimen | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | LRZ | Screwdriver, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | NZP | Sealant, Microbial | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | OLL | Septal Stapler/Absorbable Staples | 2 | 878.4750 | Implantable staple. | N | Y
General & plastic surgery | OJW | Shave Prep Kit | 2 | 878.4370 | Surgical drape and drape accessories. | N | N
General & plastic surgery | BWP | Shoe And Shoe Cover, Conductive | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | FXW | Shoes, Operating-Room | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | PFP | Silicone Hydrogel For Scar Management | 1 | 878.4025 | Silicone sheeting. | N | N
General & plastic surgery | OJV | Single Use Instrument Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OJU | Skin Prep Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | OUH | Skin Resurfacing Rf Applicator | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | GAE | Snare, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | KOX | Solvent, Adhesive Tape | 1 | 878.4730 | Surgical skin degreaser or adhesive tape solvent. | N | N
General & plastic surgery | HXR | Spatula, Orthopedic | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAF | Spatula, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FXG | Specula Accessories | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | EPY | Speculum, Ent | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | FXF | Speculum, Illuminated | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | FXE | Speculum, Non-Illuminated | 1 | 878.1800 | Speculum and accessories. | N | N
General & plastic surgery | FZF | Splint, Extremity, Inflatable, External | 1 | 878.3900 | Inflatable extremity splint. | N | N
General & plastic surgery | NOC | Splint, Extremity, Non-Inflatable, External, Non-Sterile | 1 | 878.3910 | Noninflatable extremity splint. | N | N
General & plastic surgery | FYH | Splint, Extremity, Noninflatable, External, Sterile | 1 | 878.3910 | Noninflatable extremity splint. | N | N
General & plastic surgery | PYU | Sponge, Eyelid Cleaning, Powered | 1 | 878.4820 | Surgical instrument motors and accessories/attachments. | N | N
General & plastic surgery | PYE | Staple Removal Kit | 1 | 878.4760 | Removable skin staple. | N | N
General & plastic surgery | GDW | Staple, Implantable | 2 | 878.4750 | Implantable staple. | N | Y
General & plastic surgery | NLL | Staple, Implantable, Reprocessed | 2 | 878.4750 | Implantable staple. | N | N
General & plastic surgery | GDT | Staple, Removable (Skin) | 1 | 878.4760 | Removable skin staple. | N | N
General & plastic surgery | GAG | Stapler, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MQR | Stent, Colonic, Metallic, Expandable | 2 | 878.3610 | Esophageal prosthesis. | N | Y
General & plastic surgery | MUM | Stent, Metallic, Expandable, Duodenal | 2 | 878.3610 | Esophageal prosthesis. | N | Y
General & plastic surgery | PYF | Sterile Puncture Tray | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FZM | Stool, Operating-Room | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | GAJ | Stripper, Vein, Disposable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAI | Stripper, Vein, Reusable | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAH | Stylet, Surgical, General & Plastic Surgery | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PYG | Suction Catheter Kit | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | FXO | Suit, Surgical | 1 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | RAB | Sunlamp Products (Pre-Standard) | 2 | 878.4635 | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | N | N
General & plastic surgery | REF | Suntan Bed | 2 | 878.4635 | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | N | N
General & plastic surgery | REG | Suntan Lamp | 2 | 878.4635 | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | N | N
General & plastic surgery | EPW | Support, Head, Surgical, Ent | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | KGO | Surgeon'S Gloves | 1 | 878.4460 | Non-powdered surgeon's glove. | N | N
General & plastic surgery | OCL | Surgical Device, For Cutting, Coagulation, And/Or Ablation Of Tissue, Including Cardiac Tissue | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | OOD | Surgical Film | 2 | 878.3300 | Surgical mesh. | Y | Y
General & plastic surgery | OUK | Surgical Mask With Antimicrobial/Antiviral Agent | 2 | 878.4040 | Surgical apparel. | N | N
General & plastic surgery | PYH | Surgical Retractor Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PQM | Surgical Smoke Precipitator | 2 | 878.5050 | Surgical smoke precipitator. | N | N
General & plastic surgery | LWG | Surgical Table Cushion | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | OVN | Suture Kit | 2 | 878.4830 | Absorbable surgical gut suture. | N | N
General & plastic surgery | MCZ | Suture Removal Kit | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | GAK | Suture, Absorbable | 2 | 878.4830 | Absorbable surgical gut suture. | N | Y
General & plastic surgery | GAL | Suture, Absorbable, Natural | 2 | 878.4830 | Absorbable surgical gut suture. | N | Y
General & plastic surgery | GAN | Suture, Absorbable, Synthetic | 2 | 878.4830 | Absorbable surgical gut suture. | N | Y
General & plastic surgery | GAM | Suture, Absorbable, Synthetic, Polyglycolic Acid | 2 | 878.4493 | Absorbable poly(glycolide/l-lactide) surgical suture. | N | Y
General & plastic surgery | NVH | Suture, Knot, Mechanical | 2 | 878.5000 | Nonabsorbable poly(ethylene terephthalate) surgical suture. | N | Y
General & plastic surgery | GAO | Suture, Nonabsorbable | 2 | 878.4495 | Stainless steel suture. | N | Y
General & plastic surgery | NJU | Suture, Nonabsorbable, Nitinol | 2 | 878.4495 | Stainless steel suture. | N | Y
General & plastic surgery | GAP | Suture, Nonabsorbable, Silk | 2 | 878.5030 | Natural nonabsorbable silk surgical suture. | N | Y
General & plastic surgery | GAQ | Suture, Nonabsorbable, Steel, Monofilament And Multifilament, Sterile | 2 | 878.4495 | Stainless steel suture. | N | Y
General & plastic surgery | PTX | Suture, Nonabsorbable, Steel, Monofilament And Multifilament, Sterile, Exempt | 2 | 878.4495 | Stainless steel suture. | N | Y
General & plastic surgery | GAR | Suture, Nonabsorbable, Synthetic, Polyamide | 2 | 878.5020 | Nonabsorbable polyamide surgical suture. | N | Y
General & plastic surgery | GAS | Suture, Nonabsorbable, Synthetic, Polyester | 2 | 878.5000 | Nonabsorbable poly(ethylene terephthalate) surgical suture. | N | Y
General & plastic surgery | GAT | Suture, Nonabsorbable, Synthetic, Polyethylene | 2 | 878.5000 | Nonabsorbable poly(ethylene terephthalate) surgical suture. | N | Y
General & plastic surgery | GAW | Suture, Nonabsorbable, Synthetic, Polypropylene | 2 | 878.5010 | Nonabsorbable polypropylene surgical suture. | N | Y
General & plastic surgery | NWJ | Suture, Recombinant Technology | 2 | 878.4494 | Absorbable poly(hydroxybutyrate) surgical suture produced by recombinant DNA technology. | N | Y
General & plastic surgery | NEW | Suture, Surgical, Absorbable, Polydioxanone | 2 | 878.4840 | Absorbable polydioxanone surgical suture. | N | Y
General & plastic surgery | NBY | Suture, Surgical, Nonabsorbable, Expanded, Polytetraflouroethylene | 2 | 878.5035 | Nonabsorbable expanded polytetrafluoroethylene surgical suture. | N | Y
General & plastic surgery | MXW | Suture, Surgical, Nonabsorbable, Poly (Vinylidene Fluoride) | 2 | 878.5010 | Nonabsorbable polypropylene surgical suture. | N | Y
General & plastic surgery | NEY | System, Ablation, Microwave And Accessories | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | NTB | System, Ablation, Ultrasound And Accessories | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | NEJ | System, Cryosurgical, Liquid Nitrogen, For Gastroenterology | 2 | 878.4350 | Cryosurgical unit and accessories. | N | N
General & plastic surgery | FAZ | System, Cryosurgical, Liquid Nitrogen, For Urology | 2 | 878.4350 | Cryosurgical unit and accessories. | N | N
General & plastic surgery | MWY | System, Microwave, Hair Removal | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | MKY | System, Skin Closure | 1 | 878.4320 | Removable skin clip. | N | N
General & plastic surgery | MUU | System, Suction, Lipoplasty | 2 | 878.5040 | Suction lipoplasty system. | N | N
General & plastic surgery | BWN | Table And Attachments, Operating-Room | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | LGX | Table, Examination, Medical, Powered | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | FQO | Table, Operating-Room, Ac-Powered | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | GDC | Table, Operating-Room, Electrical | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | FSE | Table, Operating-Room, Manual | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | FWX | Table, Operating-Room, Mechanical | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | FWY | Table, Operating-Room, Non-Electrical | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | FWW | Table, Operating-Room, Pneumatic | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | JEA | Table, Surgical With Orthopedic Accessories, Ac-Powered | 1 | 878.4960 | Operating tables and accessories and operating chairs and accessories. | N | N
General & plastic surgery | JEB | Table, Surgical With Orthopedic Accessories, Manual | 1 | 878.4950 | Manual operating table and accessories and manual operating chair and accessories. | N | N
General & plastic surgery | REH | Tabletop Sunlamp System | 2 | 878.4635 | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | N | N
General & plastic surgery | FXN | Tape, Camera, Surgical | 1 | 878.4160 | Surgical camera and accessories. | N | N
General & plastic surgery | FTY | Tape, Measuring, Rulers And Calipers | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | HXT | Tape, Orthopedic | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | MCA | Tape, Surgical, Internal | 2 | 878.3300 | Surgical mesh. | N | Y
General & plastic surgery | PBY | Temporary Tissue Marker | 2 | 878.4300 | Implantable clip. | N | N
General & plastic surgery | POD | Temporary, Internal Use Hemostatic | 2 | 878.4454 | Non-absorbable, hemostatic gauze for temporary internal use. | Y | N
General & plastic surgery | PYI | Thoracic Catheter Insertion Tray | 1 | 878.4200 | Introduction/drainage catheter and accessories. | N | N
General & plastic surgery | PJK | Tissue Adhesive For Internal Use | 3 | 878.4010 | Tissue adhesive. | N | Y
General & plastic surgery | MPN | Tissue Adhesive For The Topical Approximation Of Skin | 2 | 878.4010 | Tissue adhesive. | N | N
General & plastic surgery | OJS | Tourniquet Kit | 1 | 878.5900 | Nonpneumatic tourniquet. | N | N
General & plastic surgery | EYR | Tourniquet, Gastro-Urology | 1 | 878.5900 | Nonpneumatic tourniquet. | N | N
General & plastic surgery | GAX | Tourniquet, Nonpneumatic | 1 | 878.5900 | Nonpneumatic tourniquet. | N | N
General & plastic surgery | KCY | Tourniquet, Pneumatic | 1 | 878.5910 | Pneumatic tourniquet. | N | N
General & plastic surgery | PKO | Transparent Patch For Use In Treatment Of Tattoos | 2 | 878.4810 | Laser surgical instrument for use in general and plastic surgery and in dermatology. | N | N
General & plastic surgery | LRP | Tray, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FSM | Tray, Surgical, Instrument | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | FSH | Tray, Surgical, Needle | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | PYJ | Ulcer Wound Tray | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | NRJ | Ultrasound, Skin Permeation | 2 | 878.4410 | Low energy ultrasound wound cleaner. | N | N
General & plastic surgery | FLJ | Unit, Airflow, Laminar | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | GEH | Unit, Cryosurgical, Accessories | 2 | 878.4350 | Cryosurgical unit and accessories. | N | N
General & plastic surgery | BWA | Unit, Electrosurgical And Coagulation, With Accessories | 2 | 878.4400 | Electrosurgical cutting and coagulation device and accessories. | N | N
General & plastic surgery | FRT | Unit, Isolation, Surgical | 2 | 878.5070 | Air-handling apparatus for a surgical operating room. | N | N
General & plastic surgery | OJR | Vacuum Powered Body Fluid Collection Kit | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | PYK | Venous Ulcer Kit | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | LSA | Vise, Bench, Surgical | 1 | 878.4800 | Manual surgical instrument for general use. | N | N
General & plastic surgery | QCR | Wound Autofluorescence Imaging Device | 2 | 878.4165 | Wound autofluorescence imaging device. | N | N
General & plastic surgery | NRB | Wound Cleaner, Ultrasound | 2 | 878.4410 | Low energy ultrasound wound cleaner. | N | N
General & plastic surgery | OTK | Wound Drain Catheter System | 2 | 878.4780 | Powered suction pump. | N | N
General & plastic surgery | PYL | Wound Drainage Kit | 1 | 878.4014 | Nonresorbable gauze/sponge for external use. | N | N
General & plastic surgery | NYS | Wound Dressing With Poly(Diallyl Dimethyl Ammonium Chloride)(Pdadmac) | 2 | 878.4015 | Wound dressing with poly (diallyl dimethyl ammonium chloride) (pDADMAC) additive. | N | N
General & plastic surgery | PQI | Wound Retraction And Protection System | 2 | 878.4371 | Irrigating wound retractor device. | N | N
General hospital | QKM | A Chemical Vapor Sterilization Multivariable Chemical Indicator | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | PEQ | Accessories To Examination Light | 1 | 880.6320 | AC-powered medical examination light. | N | N
General hospital | MRZ | Accessories, Pump, Infusion | 2 | 880.5725 | Infusion pump. | N | N
General hospital | ODI | Administration Set Docking Station | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | PUK | Administration Set Securement Device | 1 | 880.5210 | Intravascular catheter securement device. | N | N
General hospital | PWH | Administrations Sets With Neuraxial Connectors | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | PWM | Admission Kit (Patient Utensil) | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | OSZ | Antimicrobial Keyboard | 2 | 880.6600 | Ultraviolet (UV) radiation chamber disinfection device. | N | N
General hospital | OPI | Antimicrobial Medical Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | GCX | Apparatus, Suction, Operating-Room, Wall Vacuum Powered | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | KXF | Applicator, Absorbent Tipped, Non-Sterile | 1 | 880.6025 | Absorbent tipped applicator. | N | N
General hospital | KXG | Applicator, Absorbent Tipped, Sterile | 1 | 880.6025 | Absorbent tipped applicator. | N | N
General hospital | PWN | Baby Care Kit | 2 | 880.5700 | Neonatal phototherapy unit. | N | N
General hospital | KYR | Bag, Ice | 1 | 880.6050 | Ice bag. | N | N
General hospital | FQM | Bandage, Elastic | 1 | 880.5075 | Elastic bandage. | N | N
General hospital | KMF | Bandage, Liquid | 1 | 880.5090 | Liquid bandage. | N | N
General hospital | NEC | Bandage, Liquid, Skin Protectant | 1 | 880.5090 | Liquid bandage. | N | N
General hospital | OSI | Bariatric Bed | 2 | 880.5100 | AC-powered adjustable hospital bed. | N | N
General hospital | FNY | Basin, Emesis | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | NZG | Bassinet, Hospital | 2 | 880.5140 | Pediatric medical crib. | N | N
General hospital | FNL | Bed, Ac-Powered Adjustable Hospital | 2 | 880.5100 | AC-powered adjustable hospital bed. | N | N
General hospital | FNK | Bed, Hydraulic, Adjustable Hospital | 1 | 880.5110 | Hydraulic adjustable hospital bed. | N | N
General hospital | FNJ | Bed, Manual | 1 | 880.5120 | Manual adjustable hospital bed. | N | N
General hospital | FMS | Bed, Pediatric Open Hospital | 2 | 880.5140 | Pediatric medical crib. | N | N
General hospital | LLI | Bed, Therapeutic, Ac-Powered, Adjustable Home-Use | 2 | 880.5100 | AC-powered adjustable hospital bed. | N | N
General hospital | KME | Bedding, Disposable, Medical | 1 | 880.6060 | Medical disposable bedding. | N | N
General hospital | FOB | Bedpan | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | FSD | Binder, Abdominal | 1 | 880.5160 | Therapeutic medical binder. | N | N
General hospital | HEF | Binder, Breast | 1 | 880.5160 | Therapeutic medical binder. | N | N
General hospital | KMO | Binder, Elastic | 1 | 880.5160 | Therapeutic medical binder. | N | N
General hospital | MDR | Binder, Medical, Therapeutic | 1 | 880.5160 | Therapeutic medical binder. | N | N
General hospital | FQK | Binder, Perineal | 1 | 880.5160 | Therapeutic medical binder. | N | N
General hospital | OWP | Biological Sterilization Process Indicator With Recombinant-Dna Plasmid | 2 | 880.2805 |  | N | N
General hospital | PDH | Blanket, Neonatal Phototherapy | 2 | 880.5700 | Neonatal phototherapy unit. | N | N
General hospital | PWO | Blood Administration Kit | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | PWP | Blood Borne Pathogen Response Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | POQ | Blood Transfusion Kit | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | FPS | Board, Bed | 1 | 880.6070 | Bed board. | N | N
General hospital | FOA | Board, Cardiopulmonary | 1 | 880.6080 | Cardiopulmonary resuscitation board. | N | N
General hospital | CBC | Bottle, Collection And Trap, Breathing System (Uncalibrated) | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | CBD | Bottle, Collection, Breathing System (Calibrated) | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | KDQ | Bottle, Collection, Vacuum | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | FPF | Bottle, Hot/Cold Water | 1 | 880.6085 | Hot/cold water bottle. | N | N
General hospital | PWQ | Buret Administration Intravenous Kit | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | FLI | Cabinet, Ethylene-Oxide Gas Aerator | 2 | 880.6100 | Ethylene oxide gas aerator cabinet. | N | N
General hospital | PZW | Catheter Access Cover, Tamper-Resistant | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | N
General hospital | JOL | Catheter And Tip, Suction | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | PWR | Catheter Repair Kit | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | N
General hospital | LJS | Catheter, Intravascular, Therapeutic, Long-Term Greater Than 30 Days | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | Y
General hospital | FOZ | Catheter, Intravascular, Therapeutic, Short-Term Less Than 30 Days | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | FOS | Catheter, Umbilical Artery | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | PEZ | Central Venous Catheter Dressing Change Kit | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | KMN | Chair And Table, Medical | 1 | 880.6140 | Medical chair and table. | N | N
General hospital | FML | Chair, Blood Donor | 1 | 880.6140 | Medical chair and table. | N | N
General hospital | FRK | Chair, Examination And Treatment | 1 | 880.6140 | Medical chair and table. | N | N
General hospital | FRJ | Chair, Geriatric | 1 | 880.6140 | Medical chair and table. | N | N
General hospital | LGM | Chamber, Patient Isolation | 2 | 880.5450 | Patient care reverse isolation chamber. | N | N
General hospital | LGN | Chamber, Patient Transport Isolation | 2 | 880.5450 | Patient care reverse isolation chamber. | N | N
General hospital | FRR | Chamber, Reverse Isolation, Patient Care | 2 | 880.5450 | Patient care reverse isolation chamber. | N | N
General hospital | MJF | Check Valve, Retrograde Flow (In-Line) | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | PWS | Chemotherapy Administration Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PWT | Chemotherapy Spill Clean-Up Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PAD | Chest Drainage Kit | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | FRF | Cleaner, Air, Medical Recirculating | 2 | 880.5045 | Medical recirculating air cleaner. | N | N
General hospital | FLG | Cleaner, Ultrasonic, Medical Instrument | 1 | 880.6150 | Ultrasonic cleaner for medical instruments. | N | N
General hospital | MDZ | Cleaners, Medical Devices | 2 | 880.6992 | Medical washer-disinfector. | N | N
General hospital | PWU | Clean-Up Kit | 1 | 880.6890 | General purpose disinfectants. | N | N
General hospital | ONB | Closed Antineoplastic And Hazardous Drug Reconstitution And Transfer System | 2 | 880.5440 | Intravascular administration set. | N | N | ICU Medical | chemolock, chemoclave(spiros) for oncology
General hospital | KPE | Container, I.V. | 2 | 880.5025 | I.V. container. | N | N
General hospital | KYW | Container, Liquid Medication, Graduated | 1 | 880.6430 | Liquid medication dispenser. | N | N
General hospital | MMK | Container, Sharps | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | LDR | Controller, Infusion, Intravascular, Electronic | 2 | 880.5725 | Infusion pump. | N | N
General hospital | LWH | Counter, Sponge, Surgical | 1 | 880.2740 | Surgical sponge scale. | N | N
General hospital | KIA | Cover, Cast | 1 | 880.6185 | Cast cover. | N | N
General hospital | FMW | Cover, Mattress (Medical Purposes) | 1 | 880.6190 | Mattress cover for medical purposes. | N | N
General hospital | NUP | Cream, Nasal, Topical, Mechanical Allergen Particle Barrier | 2 | 880.5045 | Medical recirculating air cleaner. | N | N
General hospital | PWV | Delivery Room Apparel Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | FMA | Depressor, Tongue, Non-Surgical | 1 | 880.6230 | Tongue depressor. | N | N
General hospital | LJL | Detectors And Removers, Lice, (Including Combs) | 1 | 880.5960 | Lice removal kit. | N | N
General hospital | MKB | Device, Germicidal, Ultraviolet | 2 | 880.6500 | Medical ultraviolet air purifier. | N | N
General hospital | KMK | Device, Intravascular Catheter Securement | 1 | 880.5210 | Intravascular catheter securement device. | N | N
General hospital | KZF | Device, Medical Examination, Ac Powered | 1 | 880.6320 | AC-powered medical examination light. | N | N
General hospital | FOD | Device, Occlusion, Umbilical | 1 | 880.5950 | Umbilical occlusion device. | N | N
General hospital | LDS | Device, Pasteurization, Hot Water | 2 | 880.6991 | Medical washer. | N | N
General hospital | FRZ | Device, Patient Transfer, Powered | 2 | 880.6775 | Powered patient transfer device. | N | N
General hospital | FMR | Device, Transfer, Patient, Manual | 1 | 880.6785 | Manual patient transfer device. | N | N
General hospital | KZA | Device, Vein Location, Liquid Crystal | 1 | 880.6970 | Liquid crystal vein locator. | N | N
General hospital | LBJ | Device, Vein Stabilization | 1 | 880.6980 | Vein stabilizer. | N | N
General hospital | LRJ | Disinfectant, Medical Devices | 1 | 880.6890 | General purpose disinfectants. | N | N
General hospital | MEC | Disinfector, Medical Devices | 2 | 880.6992 | Medical washer-disinfector. | N | N
General hospital | KYX | Dispenser, Liquid Medication | 1 | 880.6430 | Liquid medication dispenser. | N | N
General hospital | OXQ | Dressing Change Tray | 1 | 880.5075 | Elastic bandage. | N | N
General hospital | MHW | Dressing, Compression | 1 | 880.5075 | Elastic bandage. | N | N
General hospital | OGQ | Ear Irrigation Kit | 1 | 880.6960 | Irrigating syringe. | N | N
General hospital | PQX | Epinephrine Syringe | 2 | 880.5860 | Piston syringe. | N | N
General hospital | PJP | Fall Prevention Alarm/Sensor Attached Only | 1 | 880.2400 | Bed-patient monitor. | N | N
General hospital | PJO | Fall Prevention Alarm/Sensor Combination Attached Or Unattached | 1 | 880.2400 | Bed-patient monitor. | N | N
General hospital | QDO | Fentanyl And Other Opioid Protection Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PWW | Fever Monitoring Kit | 2 | 880.2910 | Clinical electronic thermometer. | N | N
General hospital | FRL | Fiber, Medical, Absorbent | 1 | 880.5300 | Medical absorbent fiber. | N | N
General hospital | MZM | Fibrin Sealant Preparation Device | 2 | 880.5860 | Piston syringe. | N | N
General hospital | FPB | Filter, Infusion Line | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | LZB | Finger Cot | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | FME | Gown, Examination | 1 | 880.6265 | Examination gown. | N | N
General hospital | PEF | Heparin Flush In 0.45% Sodium Chloride | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | NZW | Heparin, Vascular Access Flush | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | FRP | Holder, Infant Position | 1 | 880.5680 | Pediatric position holder. | N | N
General hospital | LRS | I.V. Start Kit | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | PBZ | Image Processing Device For Estimation Of External Blood Loss | 2 | 880.2750 | Image processing device for estimation of external blood loss. | N | N
General hospital | PKP | Immunoglobulin G (Igg) Infusion System | 2 | 880.5725 | Infusion pump. | N | N
General hospital | NRV | Implantable Radio Frequency Transponder System | 2 | 880.6300 | Implantable radiofrequency transponder system for patient identification and health information. | N | Y
General hospital | OKC | Implanted Subcutaneous Securement Catheter | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | Y
General hospital | FMZ | Incubator, Neonatal | 2 | 880.5400 | Neonatal incubator. | N | N
General hospital | FPL | Incubator, Neonatal Transport | 2 | 880.5410 | Neonatal transport incubator. | N | N
General hospital | FRC | Indicator, Biological Sterilization Process | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | MRB | Indicator, Biological, Liquid Chemical Sterilization Process | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | MTC | Indicator, Chemical, Enzymatic, Sterilization Process | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | JOJ | Indicator, Physical/Chemical Sterilization Process | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | OCI | Indicator, Physical/Chemical, Storage Temperature | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | LRT | Indicator, Sterilization | 2 | 880.2800 | Sterilization process indicator. | N | N
General hospital | OUW | Infant Positioner, Rx, Use In Highly Monitored Setting | 1 | 880.5680 | Pediatric position holder. | N | N
General hospital | QJY | Infusion Pump, Drug Specific, Pharmacy-Filled | 2 | 880.5725 | Infusion pump. | N | N
General hospital | PHC | Infusion Safety Management Software | 2 | 880.5725 | Infusion pump. | N | N
General hospital | KZD | Infusor, Pressure, For I.V. Bags | 1 | 880.5420 | Pressure infusor for an I.V. bag. | N | N
General hospital | OZW | Ingestible Event Marker | 2 | 880.6305 | Ingestible event marker. | N | N
General hospital | KZE | Injector, Fluid, Non-Electrically Powered | 2 | 880.5430 | Nonelectrically powered fluid injector. | N | N
General hospital | NSC | Injector, Pen | 2 | 880.5860 | Piston syringe. | N | N
General hospital | PWX | Insect Sting Emergency Kit | 1 | 880.5740 | Suction snakebite kit. | N | N
General hospital | KYS | Insoles, Medical | 1 | 880.6280 | Medical insole. | N | N
General hospital | OKL | Intravascular Administration Set, Automated Air Removal System | 2 | 880.5445 | Intravascular administration set, automated air removal system. | N | N
General hospital | OJA | Intravenous Extension Tubing Set | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | KZH | Introducer, Syringe Needle | 2 | 880.6920 | Syringe needle introducer. | N | N
General hospital | PQH | Iodinated Contrast Media Transfer Tubing Set | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | QFB | Isocapnic Ventilation Device | 2 | 868.5480 |  | N | N
General hospital | KYQ | Kit, Chemical Snake-Bite | 3 | 880.5760 | Chemical cold pack snakebite kit. | N | N
General hospital | LYY | Latex Patient Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | FQH | Lavage, Jet | 2 | 880.5475 | Jet lavage. | N | N
General hospital | LLK | Legging, Compression, Non-Inflatable | 2 | 880.5780 | Medical support stocking. | N | N
General hospital | PUN | Lice Removal Kit Containing A Solution | 1 | 880.5960 | Lice removal kit. | N | N
General hospital | FNG | Lift, Patient, Ac-Powered | 2 | 880.5500 | AC-powered patient lift. | N | N
General hospital | FSA | Lift, Patient, Non-Ac-Powered | 1 | 880.5510 | Non-AC-powered patient lift. | N | N
General hospital | KYT | Light, Examination, Medical, Battery Powered | 1 | 880.6350 | Battery-powered medical examination light. | N | N
General hospital | OVY | Liquid Chemical Processing System | 2 | 880.6887 |  | N | N
General hospital | BWJ | Locator, Acupuncture Point | 2 | 880.5580 | Acupuncture needle. | N | N
General hospital | POT | Lockbox | 2 | 880.5725 | Infusion pump. | N | N
General hospital | KMJ | Lubricant, Patient | 1 | 880.6375 | Patient lubricant. | N | N
General hospital | MMS | Lubricant, Vaginal, Patient | 1 | 880.6375 | Patient lubricant. | N | N
General hospital | PWY | Lumbar Puncture Tray (Adult & Pediatric) | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | FMJ | Manometer, Spinal-Fluid | 2 | 880.2500 | Spinal fluid manometer. | N | N
General hospital | FNM | Mattress, Air Flotation, Alternating Pressure | 2 | 880.5550 | Alternating pressure air flotation mattress. | N | N
General hospital | IKY | Mattress, Flotation Therapy, Non-Powered | 1 | 880.5150 | Nonpowered flotation therapy mattress. | N | N
General hospital | FOH | Mattress, Water, Temperature Regulated | 1 | 880.5560 | Temperature regulated water mattress. | N | N
General hospital | OUG | Medical Device Data System | 1 | 880.6310 | Medical device data system. | N | N
General hospital | OPJ | Medical Gloves With Chemotherapy Labeling Claims - Test For Use With Chemotherapy Drugs | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | NZH | Medication Management System, Remote | 2 | 880.6315 | Remote Medication Management System. | N | N
General hospital | CAK | Microfilter, Blood Transfusion | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | QDM | Midazolam Syringe | 2 | 880.5860 | Piston syringe. | N | N
General hospital | PND | Midline Catheter | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | PWZ | Mid-Stream Collection Kit | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | KMI | Monitor, Bed Patient | 1 | 880.2400 | Bed-patient monitor. | N | N
General hospital | FLN | Monitor, Electric For Gravity Flow Infusion Systems | 2 | 880.2420 | Electronic monitor for gravity flow infusion systems. | N | N
General hospital | CAR | Monitor, Spinal-Fluid Pressure, Electrically Powered | 2 | 880.2460 | Electrically powered spinal fluid pressure monitor. | N | N
General hospital | ORW | N95 Respirator With Antimicrobial/Antiviral Agent For Use By The General Public In Public Health Medical Emergencies | 2 | 880.6260 | Filtering facepiece respirator for use by the general public in public health medical emergencies. | N | N
General hospital | MQX | Needle, Acupuncture, Single Use | 2 | 880.5580 | Acupuncture needle. | N | N
General hospital | FMI | Needle, Hypodermic, Single Lumen | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | NKK | Needle, Hypodermic, Single Lumen, Reprocessed | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | MHC | Needle, Interosseous | 2 | 880.5570 | Hypodermic single lumen needle. | N | Y
General hospital | PYR | Neuraxial Administration Set - Intrathecal Delivery | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | FNN | Nipple, Lambs Feeding | 1 | 880.5640 | Lamb feeding nipple. | N | N
General hospital | PTI | Non-Coring (Huber) Needle | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | PVZ | Non-Stainless Steel Needle | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | NXH | Nursing Pad, Cohesive Gel | 1 | 880.5630 | Nipple shield. | N | N
General hospital | NTC | Nursing Pad, Hydrogel | 1 | 880.5630 | Nipple shield. | N | N
General hospital | NTX | Nursing Pad, Polyurethane Pad With Starch Copolymer, Glycerin, And Surfactant (F-68) | 1 | 880.5630 | Nipple shield. | N | N
General hospital | QLF | On-Body Injector | 2 | 880.5860 | Piston syringe. | N | N
General hospital | QLY | Ophthalmic Syringe | 2 | 880.5860 | Piston syringe. | N | N
General hospital | PXA | Oral Administration Set | 1 | 880.6430 | Liquid medication dispenser. | N | N
General hospital | ONK | Oral Lubricant | 1 | 880.6375 | Patient lubricant. | N | N
General hospital | FOK | Pad, Neonatal Eye | 1 | 880.5270 | Neonatal eye pad. | N | N
General hospital | POR | Parenteral Administration Kit | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | OYS | Patient Bed With Canopy/Restraints | 1 | 880.6760 | Protective restraint. | N | N
General hospital | FMC | Patient Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | LZC | Patient Examination Glove, Specialty | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PRN | Pediatric Position Holder | 1 | 880.5680 | Pediatric position holder. | N | N
General hospital | PXB | Percutaneous Introducer Kit | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | OMF | Percutaneous, Implanted, Long-Term Intravascular Catheter Accessory For Catheter Position | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | Y
General hospital | OWL | Peripheral Catheter Insertion Kit | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | PMS | Peripheral Intravenous (Piv) Infiltration Monitor | 2 | 880.5725 | Infusion pump. | N | N
General hospital | PXC | Personal Protection Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | OHY | Pill Crusher/Cutter | 1 | 880.6430 | Liquid medication dispenser. | N | N
General hospital | QBL | Piston Syringe Lever | 2 | 880.5860 | Piston syringe. | N | N
General hospital | QEH | Piston Syringe With Neuraxial Connector - Epidural And Peripheral Delivery | 2 | 880.5860 | Piston syringe. | N | N
General hospital | LZA | Polymer Patient Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | LJT | Port & Catheter, Implanted, Subcutaneous, Intravascular | 2 | 880.5965 | Subcutaneous, implanted, intravascular infusion port and catheter. | N | Y
General hospital | OKE | Port Introducer Kit | 2 | 880.5965 | Subcutaneous, implanted, intravascular infusion port and catheter. | N | N
General hospital | OBK | Port, Protector/Cushion | 2 | 880.5965 | Subcutaneous, implanted, intravascular infusion port and catheter. | N | N
General hospital | OIG | Powder-Free Guayle Rubber Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | OPC | Powder-Free Polychloroprene Patient Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PXD | Prep Kit | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | FMP | Protector, Skin Pressure | 1 | 880.6450 | Skin pressure protectors. | N | N
General hospital | FRN | Pump, Infusion | 2 | 880.5725 | Infusion pump. | N | N | BD carefusion Alaris
General hospital | LZF | Pump, Infusion, Analytical Sampling | 2 | 880.5725 | Infusion pump. | N | N
General hospital | MEB | Pump, Infusion, Elastomeric | 2 | 880.5725 | Infusion pump. | N | N
General hospital | LZH | Pump, Infusion, Enteral | 2 | 880.5725 | Infusion pump. | N | N
General hospital | MHD | Pump, Infusion, Gallstone Dissolution | 2 | 880.5725 | Infusion pump. | N | N
General hospital | LZG | Pump, Infusion, Insulin | 2 | 880.5725 | Infusion pump. | N | N | Medtronic Minimed
General hospital | OPP | Pump, Infusion, Insulin Bolus | 2 | 880.5725 | Infusion pump. | N | N
General hospital | MRH | Pump, Infusion, Ophthalmic | 2 | 880.5725 | Infusion pump. | N | N
General hospital | MEA | Pump, Infusion, Pca | 2 | 880.5725 | Infusion pump. | N | N
General hospital | FRA | Purifier, Air, Ultraviolet, Medical | 2 | 880.6500 | Medical ultraviolet air purifier. | N | N
General hospital | KMG | Purifier, Water, Ultraviolet, Medical | 2 | 880.6710 | Medical ultraviolet water purifier. | N | N
General hospital | ONM | Radiant Infant Warmer Cabinet | 2 | 880.5130 | Infant radiant warmer. | N | N
General hospital | OPH | Radiation Attenuating Medical Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | KDP | Regulator, Vacuum | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | NZJ | Respirator, N95, For Use By The General Public In Public Health Medical Emergencies | 2 | 880.6260 | Filtering facepiece respirator for use by the general public in public health medical emergencies. | N | N
General hospital | BRT | Restraint, Patient, Conductive | 1 | 880.6760 | Protective restraint. | N | N
General hospital | FMQ | Restraint, Protective | 1 | 880.6760 | Protective restraint. | N | N
General hospital | PYQ | Rfid Chip For Dental Appliance | 2 | 880.6300 | Implantable radiofrequency transponder system for patient identification and health information. | N | N
General hospital | FNS | Ring Cutter | 1 | 880.6200 | Ring cutter. | N | N
General hospital | NGT | Saline, Vascular Access Flush | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | FRW | Scale, Patient | 1 | 880.2720 | Patient scale. | N | N
General hospital | MRL | Scale, Sponge, Surgical, Electrically-Powered | 1 | 880.2740 | Surgical sponge scale. | N | N
General hospital | FRI | Scale, Stand-On, Patient | 1 | 880.2700 | Stand-on patient scale. | N | N
General hospital | FQA | Scale, Surgical Sponge | 1 | 880.2740 | Surgical sponge scale. | N | N
General hospital | JOK | Scissors, Medical, Disposable | 1 | 880.6820 | Medical disposable scissors. | N | N
General hospital | FPA | Set, Administration, Intravascular | 2 | 880.5440 | Intravascular administration set. | N | N | ICU Medical | chemolock, Diana, spinning spiros, tego
General hospital | BRZ | Set, Blood Transfusion | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | LHI | Set, I.V. Fluid Transfer | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | MTV | Sharps Needle Destruction Device | 2 | 880.6210 | Sharps needle destruction device. | N | N
General hospital | FPY | Sheet, Burn | 1 | 880.5180 | Burn sheet. | N | N
General hospital | HFS | Shield, Nipple | 1 | 880.5630 | Nipple shield. | N | N
General hospital | PXF | Sitz Bath Kit | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | KYP | Snake Bite Suction Kit | 1 | 880.5740 | Suction snakebite kit. | N | N
General hospital | LFE | Solution, Cold Sterilizing | 2 | 880.6885 | Liquid chemical sterilants/high level disinfectants. | N | N
General hospital | PXG | Spill Kit | 1 | 880.6890 | General purpose disinfectants. | N | N
General hospital | FOX | Stand, Infusion | 1 | 880.6990 | Infusion stand. | N | N
General hospital | MED | Sterilant, Medical Devices | 2 | 880.6885 | Liquid chemical sterilants/high level disinfectants. | N | N
General hospital | KCT | Sterilization Wrap Containers, Trays, Cassettes & Other Accessories | 2 | 880.6850 | Sterilization wrap. | N | N
General hospital | PEC | Sterilizer Automated Loading System | 2 | 880.6880 | Steam sterilizer. | N | N
General hospital | MLR | Sterilizer, Chemical | 2 | 880.6860 | Ethylene oxide gas sterilizer. | N | N
General hospital | KMH | Sterilizer, Dry Heat | 2 | 880.6870 | Dry-heat sterilizer. | N | N
General hospital | FLF | Sterilizer, Ethylene-Oxide Gas | 2 | 880.6860 | Ethylene oxide gas sterilizer. | N | N
General hospital | FLE | Sterilizer, Steam | 2 | 880.6880 | Steam sterilizer. | N | N
General hospital | NRW | Stimulator, Intramuscular, Automatic | 2 | 880.5580 | Acupuncture needle. | N | N
General hospital | FQL | Stocking, Medical Support (For General Medical Purposes) | 1 | 880.5780 | Medical support stocking. | N | N
General hospital | DWL | Stocking, Medical Support (To Prevent Pooling Of Blood In Legs) | 2 | 880.5780 | Medical support stocking. | N | N
General hospital | FMG | Stopcock, I.V. Set | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | FPP | Stretcher, Hand-Carried | 1 | 880.6900 | Hand-carried stretcher. | N | N
General hospital | NZD | Stretcher, Patient Restraint | 1 | 880.6900 | Hand-carried stretcher. | N | N
General hospital | FPO | Stretcher, Wheeled | 2 | 880.6910 | Wheeled stretcher. | N | N
General hospital | FPX | Strip, Adhesive, Closure, Skin | 1 | 880.5240 | Medical adhesive tape and adhesive bandage. | N | N
General hospital | KPD | Strip, Temperature, Forehead, Liquid Crystal | 2 | 880.2200 | Liquid crystal forehead temperature strip. | N | N
General hospital | PTD | Subcutaneous Implanted Apheresis Port | 2 | 880.5965 | Subcutaneous, implanted, intravascular infusion port and catheter. | N | Y
General hospital | FQJ | Support, Scrotal, Therapeutic | 1 | 880.5820 | Therapeutic scrotal support. | N | N
General hospital | MEG | Syringe, Antistick | 2 | 880.5860 | Piston syringe. | N | N
General hospital | KYZ | Syringe, Irrigating (Non Dental) | 1 | 880.6960 | Irrigating syringe. | N | N
General hospital | FMF | Syringe, Piston | 2 | 880.5860 | Piston syringe. | N | N
General hospital | NKN | Syringe, Piston, Reprocessed | 2 | 880.5860 | Piston syringe. | N | N
General hospital | NEP | System/Device, Pharmacy Compounding | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | KGX | Tape And Bandage, Adhesive | 1 | 880.5240 | Medical adhesive tape and adhesive bandage. | N | N
General hospital | PYO | Tape And Bandage, Adhesive, Adjustable Closing Mechanism, Otc Use | 1 | 880.5240 | Medical adhesive tape and adhesive bandage. | N | N
General hospital | PXH | Thermometer Kit | 2 | 880.2910 | Clinical electronic thermometer. | N | N
General hospital | FQZ | Thermometer, Clinical Color Change | 1 | 880.2900 | Clinical color change thermometer. | N | N
General hospital | FLK | Thermometer, Clinical Mercury | 2 | 880.2920 | Clinical mercury thermometer. | N | N
General hospital | FLL | Thermometer, Electronic, Clinical | 2 | 880.2910 | Clinical electronic thermometer. | N | N
General hospital | OZK | Thermometer, Exhaled Breath | 2 | 880.2910 | Clinical electronic thermometer. | N | N
General hospital | PXI | Thoracentesis Tray | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | LHB | Timer, Apgar | 1 | 880.2930 | Apgar timer. | N | N
General hospital | MZR | Topical Approximation System | 1 | 880.5240 | Medical adhesive tape and adhesive bandage. | N | N
General hospital | BYZ | Trap, Sterile Specimen | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | BYY | Tube, Aspirating, Flexible, Connecting | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | FPK | Tubing, Fluid Delivery | 2 | 880.5440 | Intravascular administration set. | N | N
General hospital | GAZ | Tubing, Noninvasive | 2 | 880.6740 | Vacuum-powered body fluid suction apparatus. | N | N
General hospital | ODY | Tunneled Catheter Remover | 2 | 880.5970 | Percutaneous, implanted, long-term intravascular catheter. | N | N
General hospital | PJJ | Two Or More Sterilant Sterilizer | 2 | 880.6860 | Ethylene oxide gas sterilizer. | N | N
General hospital | PXJ | Umbilical Catheter Insertion Tray | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | LBI | Unit, Neonatal Phototherapy | 2 | 880.5700 | Neonatal phototherapy unit. | N | N
General hospital | FNP | Urinal | 1 | 880.6730 | Body waste receptacle. | N | N
General hospital | PUR | Vacuum Syringe | 2 | 880.5860 | Piston syringe. | N | N
General hospital | PXK | Vascular Access Port Kit | 2 | 880.5965 | Subcutaneous, implanted, intravascular infusion port and catheter. | N | N
General hospital | PXL | Venipuncture Kit | 2 | 880.5570 | Hypodermic single lumen needle. | N | N
General hospital | LYZ | Vinyl Patient Examination Glove | 1 | 880.6250 | Non-powdered patient examination glove. | N | N
General hospital | PXM | Vomitus Clean-Up Kit | 1 | 880.6890 | General purpose disinfectants. | N | N
General hospital | FMT | Warmer, Infant Radiant | 2 | 880.5130 | Infant radiant warmer. | N | N
General hospital | LHF | Warmer, Microwave, Infusion Fluid | 2 | 880.5725 | Infusion pump. | N | N
General hospital | LGZ | Warmer, Thermal, Infusion Fluid | 2 | 880.5725 | Infusion pump. | N | N
General hospital | FLH | Washer Of Body Waste Receptacles | 1 | 880.6800 | Washers for body waste receptacles. | N | N
General hospital | NZX | Water, Vascular Access Flush | 2 | 880.5200 | Intravascular catheter. | N | N
General hospital | MCY | Wound Dressing Kit | 2 | 880.5075 | Elastic bandage. | N | N
General hospital | FRG | Wrap, Sterilization | 2 | 880.6850 | Sterilization wrap. | N | N
Immunology | PCG | 21-Hydroxylase Antibody (21-Ohab) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | DCF | Albumin, Antigen, Antiserum, Control | 2 | 866.5040 | Albumin immunological test system. | N | N
Immunology | DDZ | Albumin, Fitc, Antigen, Antiserum, Control | 2 | 866.5040 | Albumin immunological test system. | N | N
Immunology | DFJ | Albumin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5040 | Albumin immunological test system. | N | N
Immunology | DBC | Alpha 2, 2n-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5425 | Alpha-2-glycoproteins immunological test system. | N | N
Immunology | MGA | Alpha-1 Microglobulin, Antigen, Antiserum, Control | 1 | 866.5400 | Alpha-globulin immunological test system. | N | N
Immunology | LKL | Alpha-1-Acid-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5420 | Alpha-1-glycoproteins immunological test system. | N | N
Immunology | DFF | Alpha-1-Antichymotrypsin, Antigen, Antiserum, Control | 2 | 866.5080 | Alpha-1-antichymotrypsin immunological test system. | N | N
Immunology | OBZ | Alpha-1-Antitrypsin Kit, Qualitative Phenotype | 2 | 866.5130 | Alpha-1-antitrypsin immunological test system. | N | N
Immunology | DEM | Alpha-1-Antitrypsin, Antigen, Antiserum, Control | 2 | 866.5130 | Alpha-1-antitrypsin immunological test system. | N | N
Immunology | DEI | Alpha-1-Antitrypsin, Fitc, Antigen, Antiserum, Control | 2 | 866.5130 | Alpha-1-antitrypsin immunological test system. | N | N
Immunology | DFB | Alpha-1-Antitrypsin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5130 | Alpha-1-antitrypsin immunological test system. | N | N
Immunology | DEX | Alpha-1-B-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5420 | Alpha-1-glycoproteins immunological test system. | N | N
Immunology | DER | Alpha-1-Lipoprotein, Antigen, Antiserum, Control | 2 | 866.5580 | Alpha-1-lipoprotein immunological test system. | N | N
Immunology | DEN | Alpha-1-T-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5420 | Alpha-1-glycoproteins immunological test system. | N | N
Immunology | DAW | Alpha-2-Ap-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5425 | Alpha-2-glycoproteins immunological test system. | N | N
Immunology | DEJ | Alpha-2-Glycoproteins, Antigen, Antiserum, Control | 1 | 866.5425 | Alpha-2-glycoproteins immunological test system. | N | N
Immunology | DEF | Alpha-2-Hs-Glycoprotein, Antigen, Antiserum, Control | 1 | 866.5425 | Alpha-2-glycoproteins immunological test system. | N | N
Immunology | DEB | Alpha-2-Macroglobulin, Antigen, Antiserum, Control | 2 | 866.5620 | Alpha-2-macroglobulin immunological test system. | N | N
Immunology | DDY | Alpha-2-Macroglobulin, Fitc, Antigen, Antiserum, Control | 2 | 866.5620 | Alpha-2-macroglobulin immunological test system. | N | N
Immunology | DDT | Alpha-2-Macroglobulin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5620 | Alpha-2-macroglobulin immunological test system. | N | N
Immunology | DCO | Alpha-Globulin, Antigen, Antiserum, Control | 1 | 866.5400 | Alpha-globulin immunological test system. | N | N
Immunology | NJF | Antibodies, Anti-Actin, Anti-Smooth Muscle | 2 | 866.5120 | Antismooth muscle antibody immunological test system. | N | N
Immunology | NHX | Antibodies, Anti-Cyclic Citrullinated Peptide (Ccp) | 2 | 866.5775 | Rheumatoid factor immunological test system. | N | N
Immunology | MST | Antibodies, Gliadin | 2 | 866.5750 | Radioallergosorbent (RAST) immunological test system. | N | N
Immunology | OKM | Antibodies, Outer-Membrane Proteins | 2 | 866.5785 | Anti-Saccharomyces cerevisiae (S. cerevisiae) antibody (ASCA) test systems. | N | N
Immunology | NBT | Antibodies, Saccharomyces Cerevisiae (S.Cerevisiae) | 2 | 866.5785 | Anti-Saccharomyces cerevisiae (S. cerevisiae) antibody (ASCA) test systems. | N | N
Immunology | LRM | Anti-Dna Antibody (Enzyme-Labeled), Antigen, Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | LSW | Anti-Dna Antibody, Antigen And Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | KTL | Anti-Dna Indirect Immunofluorescent Solid Phase | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | NPH | Antigen, Cartilage Oligomeric Matrix Protein (Comp) | 2 | 866.5775 | Rheumatoid factor immunological test system. | N | N
Immunology | OSK | Anti-Glutamate Receptor (Type Nmda) Ifa | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | DBM | Antimitochondrial Antibody, Indirect Immunofluorescent, Antigen, Control | 2 | 866.5090 | Antimitochondrial antibody immunological test system. | N | N
Immunology | OQZ | Anti-Mutated Citrullinated Vimentin Igg | 2 | 866.5775 | Rheumatoid factor immunological test system. | N | N
Immunology | LJM | Antinuclear Antibody (Enzyme-Labeled), Antigen, Controls | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | LKJ | Antinuclear Antibody, Antigen, Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | DHN | Antinuclear Antibody, Indirect Immunofluorescent, Antigen, Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | DBJ | Antiparietal Antibody, Indirect Immunofluorescent, Antigen, Control | 2 | 866.5110 | Antiparietal antibody immunological test system. | N | N
Immunology | MQA | Anti-Ribosomal P Antibodies | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | LKO | Anti-Rnp Antibody, Antigen And Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | LKP | Anti-Sm Antibody, Antigen And Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | DBE | Antismooth Muscle Antibody, Indirect Immunofluorescent, Antigen, Control | 2 | 866.5120 | Antismooth muscle antibody immunological test system. | N | N
Immunology | OBE | Anti-Ss-A 52 Autoantibodies | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | PET | Anti-Ss-A Ro60 Autoantibodies | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | PNI | Aquaporin-4 Autoantibody | 2 | 866.5665 | Aquaporin-4 autoantibody immunological test system. | N | N
Immunology | NHS | Assay, Genotype, Hiv Drug Resistance, In Vitro | 2 | 866.3950 | In vitro human immunodeficiency virus (HIV) drug resistance genotype assay. | N | N
Immunology | NST | Autoantibodies, Acetylcholine Receptor, Acetylcholine Blocking And Non-Blocking | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NYO | Autoantibodies, Anti-Ribonucleic Acid Polymerase (Rnap) Iii Antibody | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | NIY | Autoantibodies, Anti-Soluble Liver Antigen (Sla), Autoimmune Hepatitis | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | MVM | Autoantibodies, Endomysial(Tissue Transglutaminase) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NWG | Autoantibodies, Glutamic Acid Decarboxylase (Gad) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NBS | Autoantibodies, Lkm-1 (Liver/Kidney Microsome, Type 1) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NUM | Autoantibodies, Nuclear Body Protein, Sp100 | 2 | 866.5090 | Antimitochondrial antibody immunological test system. | N | N
Immunology | NRI | Autoantibodies, Nuclear Pore Glycoprotein Gp210 | 2 | 866.5090 | Antimitochondrial antibody immunological test system. | N | N
Immunology | OEG | Autoantibodies, Skin (Bullous Pemphigoid 180 And Bullous Pemphigoid 230 | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NBO | Autoantibodies, Skin (Desmoglein 1 And Desmoglein 3) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | OPN | Auto-Antibodies; Phosphatidylserine, Prothrombin, Phosphatidylserine/Prothrombin Complex | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | PIV | Automated Indirect Immunofluorescence Microscope And Software-Assisted System For Clinical Use | 2 | 866.4750 | Automated indirect immunofluorescence microscope and software-assisted system. | N | N
Immunology | PKB | Autosomal Recessive Carrier Screening Gene Mutation Detection System | 2 | 866.5940 | Autosomal recessive carrier screening gene mutation detection system. | N | N
Immunology | CZQ | Bence-Jones Protein, Antigen, Antiserum, Control | 2 | 866.5150 | Bence-Jones proteins immunological test system. | N | N
Immunology | DDN | Beta-2-Glycoprotein I, Antigen, Antiserum, Control | 1 | 866.5430 | Beta-2-glycoprotein I immunological test system. | N | N
Immunology | DDK | Beta-2-Glycoprotein Iii, Antigen, Antiserum, Control | 1 | 866.5440 | Beta-2-glycoprotein III immunological test system. | N | N
Immunology | DCJ | Beta-Globulin, Antigen, Antiserum, Control | 1 | 866.5160 | Beta-globulin immunological test system. | N | N
Immunology | MVE | Biosensor, Immunoassay, Myoglobin | 2 | 866.5680 | Myoglobin immunological test system. | N | N
Immunology | QAT | Brain Trauma Assessment Test | 2 | 866.5830 | Brain trauma assessment test. | N | N
Immunology | DGM | Breast Milk, Antigen, Antiserum, Control | 1 | 866.5170 | Breast milk immunological test system. | N | N
Immunology | DGN | Breast Milk, Fitc, Antigen, Antiserum, Control | 1 | 866.5170 | Breast milk immunological test system. | N | N
Immunology | DGI | Breast Milk, Rhodamine, Antigen, Antiserum, Control | 1 | 866.5170 | Breast milk immunological test system. | N | N
Immunology | KTM | C3-Indirect Immunofluorescent Solid Phase | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | NXO | Calprotectin, Fecal | 2 | 866.5180 | Fecal calprotectin immunological test system. | N | N
Immunology | NVA | Cancer Monitoring Test System, Soluble Mesothelin-Related Peptides, Mesothelioma | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | DDH | Carbonic Anhydrase B, Antigen, Antiserum, Control | 1 | 866.5200 | Carbonic anhydrase B and C immunological test system. | N | N
Immunology | DDE | Carbonic Anhydrase C, Antigen, Antiserum, Control | 1 | 866.5200 | Carbonic anhydrase B and C immunological test system. | N | N
Immunology | NQD | Cardiac C-Reactive Protein, Antigen, Antiserum, And Control | 2 | 866.5270 | C-reactive protein immunological test system. | N | N
Immunology | DDB | Ceruloplasmin, Antigen, Antiserum, Control | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | DCY | Ceruloplasmin, Fitc, Antigen, Antiserum, Control | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | DCT | Ceruloplasmin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | OVQ | Chronic Lymphocytic Leukemia Fish Probe Kit | 2 | 866.6040 | Gene expression profiling test system for breast cancer prognosis. | N | N
Immunology | NYI | Classifier, Prognostic, Recurrence Risk Assessment, Rna Gene Expression, Breast Cancer | 2 | 866.6040 | Gene expression profiling test system for breast cancer prognosis. | N | N
Immunology | DGA | Cohn Fraction Ii, Antigen, Antiserum, Control | 1 | 866.5220 | Cohn fraction II immunological test system. | N | N
Immunology | DGJ | Colostrum, Antigen, Antiserum, Control | 1 | 866.5230 | Colostrum immunological test system. | N | N
Immunology | NRX | Complement 4d (C4d) And Complement Receptor 1 (Cr1) | 2 | 866.5820 | Systemic lupus erythematosus immunological test system. | N | N
Immunology | DBA | Complement C1 Inhibitor (Inactivator), Antigen, Antiserum, Control | 2 | 866.5250 | Complement C[bdi2] inhibitor (inactivator) immunological test system. | N | N
Immunology | DAK | Complement C1q, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | DAI | Complement C1r, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | CZY | Complement C1s, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | CZW | Complement C3, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | KTP | Complement C3b Activator Immunoassay Reagents | 2 | 866.5260 | Complement C3b inactivator immunological test system. | N | N
Immunology | DAC | Complement C3b Activator, Antigen, Antiserum, Control | 2 | 866.5260 | Complement C3b inactivator immunological test system. | N | N
Immunology | DBI | Complement C4, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | DAY | Complement C5, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | DAG | Complement C8, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | DAE | Complement C9, Antigen, Antiserum, Control | 2 | 866.5240 | Complement components immunological test system. | N | N
Immunology | DCK | C-Reactive Protein, Antigen, Antiserum, And Control | 2 | 866.5270 | C-reactive protein immunological test system. | N | N
Immunology | OVK | Cytokeratin Fragments 21-1 Eia Kit | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | LYE | Cytokeratins | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DHF | D/Km-1, Antigen, Antiserum, Control | 1 | 866.5065 | Human allotypic marker immunological test system. | N | N
Immunology | OAU | Des-Gamma-Carboxy-Prothrombin (Dcp), Risk Assessment, Hepatocellular Carcinoma | 2 | 866.6030 | AFP-L3% immunological test system. | N | N
Immunology | LOQ | Device, General Purpose, Hematology | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | MVJ | Devices, Measure, Antibodies To Glomerular Basement Membrane (Gbm) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | OYU | Dna-Probe Kit, Human Chromosome | 2 | 866.4700 | Automated fluorescence in situ hybridization (FISH) enumeration systems. | N | N
Immunology | OXP | Dna-Probe Kit, Human Chromosome X And Y, Bmt Engraftment | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | OWK | Early Growth Response 1 (Egr) Fish Probe Kit | 2 | 866.6040 | Gene expression profiling test system for breast cancer prognosis. | N | N
Immunology | JZS | Electrophoresis Instrumentation | 1 | 866.4500 | Immunoelectrophoresis equipment. | N | N
Immunology | MLE | Enzyme Immunoassay, Antiparietal Cell Antibody, Antigen, Control | 2 | 866.5110 | Antiparietal antibody immunological test system. | N | N
Immunology | JZX | Equipment, Rocket Immunoelectrophoresis | 1 | 866.4830 | Rocket immunoelectrophoresis equipment. | N | N
Immunology | LLL | Extractable Antinuclear Antibody, Antigen And Control | 2 | 866.5100 | Antinuclear antibody immunological test system. | N | N
Immunology | DCE | Fab, Antigen, Antiserum, Control | 1 | 866.5520 | Immunoglobulin G (Fab fragment specific) immunological test system. | N | N
Immunology | DCB | Fab, Fitc, Antigen, Antiserum, Control | 1 | 866.5520 | Immunoglobulin G (Fab fragment specific) immunological test system. | N | N
Immunology | DBY | Fab, Rhodamine, Antigen, Antiserum, Control | 1 | 866.5520 | Immunoglobulin G (Fab fragment specific) immunological test system. | N | N
Immunology | JZH | Factor B, Antigen, Antiserum, Control | 2 | 866.5320 | Properdin factor B immunological test system. | N | N
Immunology | DBT | Factor Xiii A, S, Antigen, Antiserum, Control | 1 | 866.5330 | Factor XIII, A, S, immunological test system. | N | N
Immunology | DBN | Fc, Antigen, Antiserum, Control | 1 | 866.5530 | Immunoglobulin G (Fc fragment specific) immunological test system. | N | N
Immunology | DBK | Fc, Fitc, Antigen, Antiserum, Control | 1 | 866.5530 | Immunoglobulin G (Fc fragment specific) immunological test system. | N | N
Immunology | DBH | Fc, Rhodamine, Antigen, Antiserum, Control | 1 | 866.5530 | Immunoglobulin G (Fc fragment specific) immunological test system. | N | N
Immunology | DBF | Ferritin, Antigen, Antiserum, Control | 2 | 866.5340 | Ferritin immunological test system. | N | N
Immunology | DAN | Fibrinopeptide A, Antigen, Antiserum, Control | 2 | 866.5350 | Fibrinopeptide A immunological test system. | N | N
Immunology | JZT | Fluorometer | 1 | 866.4520 | Immunofluorometer equipment. | N | N
Immunology | DAL | Fraction Iv-5, Antigen, Antiserum, Control | 1 | 866.5360 | Cohn fraction IV immunological test system. | N | N
Immunology | KHT | Fraction V, Antigen, Antiserum, Control | 1 | 866.5370 | Cohn fraction V immunological test system. | N | N
Immunology | DAJ | Free Secretory Component, Antigen, Antiserum, Control | 2 | 866.5380 | Free secretory component immunological test system. | N | N
Immunology | DAH | Gamma Globulin, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DAF | Gamma Globulin, Fitc, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | PTA | Genetic Variant Detection And Health Risk Assessment System | 2 | 866.5950 | Genetic health risk assessment system. | N | N
Immunology | DAD | Haptoglobin, Antigen, Antiserum, Control | 2 | 866.5460 | Haptoglobin immunological test system. | N | N
Immunology | DAB | Haptoglobin, Fitc, Antigen, Antiserum, Control | 2 | 866.5460 | Haptoglobin immunological test system. | N | N
Immunology | CZZ | Haptoglobin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5460 | Haptoglobin immunological test system. | N | N
Immunology | DAM | Hemoglobin, Chain Specific, Antigen, Antiserum, Control | 2 | 866.5470 | Hemoglobin immunological test system. | N | N
Immunology | CZX | Hemopexin, Antigen, Antiserum, Control | 2 | 866.5490 | Hemopexin immunological test system. | N | N
Immunology | CZT | Hemopexin, Fitc, Antigen, Antiserum, Control | 2 | 866.5490 | Hemopexin immunological test system. | N | N
Immunology | CZR | Hemopexin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5490 | Hemopexin immunological test system. | N | N
Immunology | PFF | High Throughput Dna Sequence Analyzer | 2 | 862.2265 | High throughput genomic sequence analyzer for clinical use. | N | N
Immunology | CZP | Iga, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CZM | Iga, Ferritin, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CZN | Iga, Fitc, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CZL | Iga, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CZK | Iga, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | OPY | Igal Heavy & Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CZJ | Igd, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGG | Igd, Fitc, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGH | Igd, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGE | Igd, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGC | Ige, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DFM | Ige, Ferritin, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGP | Ige, Fitc, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGO | Ige, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGL | Ige, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DFK | Igg (Fab Fragment Specific), Antigen, Antiserum, Control | 1 | 866.5520 | Immunoglobulin G (Fab fragment specific) immunological test system. | N | N
Immunology | DAS | Igg (Fc Fragment Specific), Antigen, Antiserum, Control | 1 | 866.5530 | Immunoglobulin G (Fc fragment specific) immunological test system. | N | N
Immunology | DAQ | Igg (Fd Fragment Specific), Antigen, Antiserum, Control | 1 | 866.5540 | Immunoglobulin G (Fd fragment specific) immunological test system. | N | N
Immunology | DFZ | Igg (Gamma Chain Specific), Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DEW | Igg, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGD | Igg, Ferritin, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGK | Igg, Fitc, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DAA | Igg, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DFO | Igg, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DAO | Igm (Mu Chain Specific), Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DFT | Igm, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DFL | Igm, Ferritin, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DFS | Igm, Fitc, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEY | Igm, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEZ | Igm, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | JKM | Immunochemical, Bence-Jones Protein | 2 | 866.5150 | Bence-Jones proteins immunological test system. | N | N
Immunology | CHN | Immunochemical, Ceruloplasmin | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | JNL | Immunochemical, Thyroglobulin Autoantibody | 2 | 866.5870 | Thyroid autoantibody immunological test system. | N | N
Immunology | JNM | Immunochemical, Transferrin | 2 | 866.5880 | Transferrin immunological test system. | N | N
Immunology | CFF | Immunoelectrophoretic, Immunoglobulins, (G, A, M) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | OPX | Immunoglobulin A Kappa Heavy & Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | PCN | Immunoglobulin G Kappa Heavy And Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | PCO | Immunoglobulin G Lambda Heavy And Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | PDE | Immunoglobulin M Kappa Heavy And Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | PDF | Immunoglobulin M Lambda Heavy And Light Chain Combined | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | JFR | Indirect Copper Assay, Ceruloplasmin | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | OYV | Inherited Nucleotide Repeat Disorder Dna Test | 2 | 866.5970 |  | N | N
Immunology | OCN | Insulin Autoantibody Kit | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | CZO | Inter-Alpha Trypsin Inhibitor, Antigen, Antiserum, Control | 1 | 866.5890 | Inter-alpha trypsin inhibitor immunological test system. | N | N
Immunology | DEC | Inter-Alpha Trypsin Inhibitor, Fitc, Antigen, Antiserum, Control | 1 | 866.5890 | Inter-alpha trypsin inhibitor immunological test system. | N | N
Immunology | DFH | Kappa, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEO | Kappa, Fitc, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DFD | Kappa, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEK | Kappa, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | MUT | Kit, Test For Nuclear Matrix (Numa) Protein (For Monitoring And Management Of Colorectal Cancer) | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | LOJ | Kit, Test,Alpha-Fetoprotein For Testicular Cancer | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | MVS | Kit, Typing, Hla-Dqb | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | DET | Lactic Dehydrogenase, Antigen, Antiserum, Control | 1 | 866.5560 | Lactic dehydrogenase immunological test system. | N | N
Immunology | DEG | Lactoferrin, Antigen, Antiserum, Control | 1 | 866.5570 | Lactoferrin immunological test system. | N | N
Immunology | DEH | Lambda, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DES | Lambda, Fitc, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEP | Lambda, Peroxidase, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DFG | Lambda, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5550 | Immunoglobulin (light chain specific) immunological test system. | N | N
Immunology | DEL | Lipoprotein X, Antigen, Antiserum, Control | 1 | 866.5590 | Lipoprotein X immunological test system. | N | N
Immunology | DFC | Lipoprotein, Low-Density, Antigen, Antiserum, Control | 2 | 866.5600 | Low-density lipoprotein immunological test system. | N | N
Immunology | PZJ | Lynch Syndrome Test System | 2 | 864.1866 | Lynch syndrome test systems. | N | N
Immunology | CGM | Method, Immunodiffusion, Immunoglobulins (G, A, M) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CFN | Method, Nephelometric, Immunoglobulins (G, A, M) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DBL | Multiple Autoantibodies, Indirect Immunofluorescent, Antigen, Control | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NTV | Myeloperoxidase, Immunoassay, System, Test | 2 | 866.5600 | Low-density lipoprotein immunological test system. | N | N
Immunology | DDR | Myoglobin, Antigen, Antiserum, Control | 2 | 866.5680 | Myoglobin immunological test system. | N | N
Immunology | DEA | Myoglobin, Fitc, Antigen, Antiserum, Control | 2 | 866.5680 | Myoglobin immunological test system. | N | N
Immunology | DDO | Myoglobin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5680 | Myoglobin immunological test system. | N | N
Immunology | JZW | Nephelometer | 1 | 866.4540 | Immunonephelometer equipment. | N | N
Immunology | DGX | Ng1m(A), Antigen, Antiserum, Control | 1 | 866.5065 | Human allotypic marker immunological test system. | N | N
Immunology | DHI | Ng3m(Bo), Antigen, Antiserum, Control | 1 | 866.5065 | Human allotypic marker immunological test system. | N | N
Immunology | DHQ | Ng3m(G), Antigen, Antiserum, Control | 1 | 866.5065 | Human allotypic marker immunological test system. | N | N
Immunology | DHY | Ng4m(A), Antigen, Antiserum, Control | 1 | 866.5065 | Human allotypic marker immunological test system. | N | N
Immunology | DDX | Plasminogen, Antigen, Antiserum, Control | 1 | 866.5715 | Plasminogen immunological test system. | N | N
Immunology | JZQ | Plates And Equipment, Radial Immunodiffusion | 1 | 866.4800 | Radial immunodiffusion plate. | N | N
Immunology | JZP | Plates, Ouchterlony Agar | 1 | 866.4600 | Ouchterlony agar plate. | N | N
Immunology | JFQ | P-Phenyl-Enediamine/Edta (Spectrophotometric), Ceruloplasmin | 2 | 866.5210 | Ceruloplasmin immunological test system. | N | N
Immunology | JZJ | Prealbumin, Antigen, Antiserum, Control | 1 | 866.5060 | Prealbumin immunological test system. | N | N
Immunology | DDS | Prealbumin, Fitc, Antigen, Antiserum, Control | 1 | 866.5060 | Prealbumin immunological test system. | N | N
Immunology | LTJ | Prostate-Specific Antigen (Psa) For Management Of Prostate Cancers | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | OWM | Prostate-Specific Antigen (Psa) For Prognostic, Recurrence Risk Assessment Of Prostate Cancers | 2 | 866.6040 | Gene expression profiling test system for breast cancer prognosis. | N | N
Immunology | DHL | Protein, Complement, Antigen, Antiserum, Control | 1 | 866.4100 | Complement reagent. | N | N
Immunology | DDF | Prothrombin, Antigen, Antiserum, Control | 1 | 866.5735 | Prothrombin immunological test system. | N | N
Immunology | JMG | Radioimmunoassay (Two-Site Solid Phase), Ferritin | 2 | 866.5340 | Ferritin immunological test system. | N | N
Immunology | JHR | Radioimmunoassay, Immunoglobulins (D, E) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | CFQ | Radioimmunoassay, Immunoglobulins (G, A, M) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | KTQ | Reagent, Complement | 1 | 866.4100 | Complement reagent. | N | N
Immunology | KTK | Reagent, Immunoassay, Carbonic Anhydrase B And C | 1 | 866.5200 | Carbonic anhydrase B and C immunological test system. | N | N
Immunology | KTO | Reagent, Immunoassay, Igg | 1 | 866.5530 | Immunoglobulin G (Fc fragment specific) immunological test system. | N | N
Immunology | CZS | Retinol-Binding Protein, Antigen, Antiserum, Control | 1 | 866.5765 | Retinol-binding protein immunological test system. | N | N
Immunology | KTS | Second Antibody (Species Specific Anti-Animal Gamma Globulin) | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | DGB | Seminal Fluid, Antigen, Antiserum, Control | 1 | 866.5800 | Seminal fluid (sperm) immunological test system. | N | N
Immunology | PZH | Serpina1 Variant Detection System | 2 | 866.5130 | Alpha-1-antitrypsin immunological test system. | N | N
Immunology | PJI | Severe Combined Immunodeficiency Disorder (Scid) Newborn Screening Test System | 2 | 866.5930 | Newborn screening test for severe combined immunodeficiency disorder (SCID). | N | N
Immunology | DFQ | Sperm, Antigen, Antiserum, Control | 1 | 866.5800 | Seminal fluid (sperm) immunological test system. | N | N
Immunology | DFX | Sperm, Fitc, Antigen, Antiserum, Control | 1 | 866.5800 | Seminal fluid (sperm) immunological test system. | N | N
Immunology | MLH | Strip, Hama Igg, Elisa, In Vitro Test System | 2 | 866.5510 | Immunoglobulins A, G, M, D, and E immunological test system. | N | N
Immunology | JZR | Support Gels | 1 | 866.4900 | Support gel. | N | N
Immunology | NTH | System, Automated Scanning Microscope And Image Analysis For Fluorescence In Situ Hybridization (Fish) Assays | 2 | 866.4700 | Automated fluorescence in situ hybridization (FISH) enumeration systems. | N | N
Immunology | NTW | System, Blood Collection, Rna Stabilization, Rna Purification, Rt-Pcr Molecular Diagnostic Test | 2 | 866.4070 | RNA Preanalytical Systems. | N | N
Immunology | PFR | System, Cystic Fibrosis Transmembrane Conductance Regulator Gene, Mutations & Variants Panel Sequencing Detection | 2 | 866.5900 | Cystic fibrosis transmembrane conductance regulator (CFTR) gene mutation detection system. | N | N
Immunology | PFS | System, Cystic Fibrosis Transmembrane Conductance Regulator Gene, Variant Gene Sequence Detection | 2 | 866.5900 | Cystic fibrosis transmembrane conductance regulator (CFTR) gene mutation detection system. | N | N
Immunology | NUA | System, Cystic Fibrosis Transmembrane Conductance Regulator, Gene Mutation Detection | 2 | 866.5900 | Cystic fibrosis transmembrane conductance regulator (CFTR) gene mutation detection system. | N | N
Immunology | NQI | System, Immunomagnetic, Circulating Cancer Cell, Enumeration | 2 | 866.6020 | Immunomagnetic circulating cancer cell selection and enumeration system. | N | N
Immunology | PFX | System, Microarray-Based, Genome-Wide, Postnatal Chromosomal Abnormality Detection | 2 | 866.5920 |  | N | N
Immunology | MID | System, Test, Anticardiolipin Immunological | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | JZG | System, Test, Beta-2-Microglobulin Immunological | 2 | 866.5630 | Beta-2-microglobulin immunological test system. | N | N
Immunology | NIG | System, Test, Carbohydrate Antigen (Ca19-9), For Monitoring And Management Of Pancreatic Cancer | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | DHX | System, Test, Carcinoembryonic Antigen | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | DCN | System, Test, C-Reactive Protein | 2 | 866.5270 | C-reactive protein immunological test system. | N | N
Immunology | DCH | System, Test, C-Reactive Protein, Rhodamine | 2 | 866.5270 | C-reactive protein immunological test system. | N | N
Immunology | NTY | System, Test, Fibrin/Fibrinogen Degradation Products For Monitoring Of Colorectal Cancer | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | NCW | System, Test, Her-2/Neu, Monitoring | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | DGW | System, Test, Hypersensitivity Pneumonitis | 2 | 866.5500 | Hypersensitivity pneumonitis immunological test system. | N | N
Immunology | MOI | System, Test, Immunological, Antigen, Tumor | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | KTN | System, Test, Infectious Mononucleosis | 2 | 866.5640 | Infectious mononucleosis immunological test system. | N | N
Immunology | DHB | System, Test, Radioallergosorbent (Rast) Immunological | 2 | 866.5750 | Radioallergosorbent (RAST) immunological test system. | N | N
Immunology | PUW | System, Test, Radioallergosorbent (Rast) Immunological, Exempt Allergens | 2 | 866.5750 | Radioallergosorbent (RAST) immunological test system. | N | N
Immunology | DHR | System, Test, Rheumatoid Factor | 2 | 866.5775 | Rheumatoid factor immunological test system. | N | N
Immunology | DHC | System, Test, Systemic Lupus Erythematosus | 2 | 866.5820 | Systemic lupus erythematosus immunological test system. | N | N
Immunology | MSW | System, Test, Thyroglobulin | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | JZO | System, Test, Thyroid Autoantibody | 2 | 866.5870 | Thyroid autoantibody immunological test system. | N | N
Immunology | MMW | System, Test, Tumor Marker, Monitoring, Bladder | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | MSV | System,Test,Antibodies,B2 - Glycoprotein I (B2 - Gpi) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | MOB | Test System, Antineutrophil Cytoplasmic Antibodies (Anca) | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | NSF | Test, Alpha Fetoprotein L3 Subfraction (Afp-L3%), For Hepatocellular Carcinoma Risk Assessment | 2 | 866.6030 | AFP-L3% immunological test system. | N | N
Immunology | OIU | Test, Epithelial Ovarian Tumor Associated Antigen (He4) | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | LTK | Test, Epithelial Ovarian Tumor-Associated Antigen (Ca125) | 2 | 866.6010 | Tumor-associated antigen immunological test system. | N | N
Immunology | NOE | Test, System, Immunoassay, Lipoprotein-Associated Phospholipase A2 | 2 | 866.5600 | Low-density lipoprotein immunological test system. | N | N
Immunology | DDC | Thyroglobulin, Antigen, Antiserum, Control | 2 | 866.5870 | Thyroid autoantibody immunological test system. | N | N
Immunology | DDJ | Thyroglobulin, Fitc, Antigen, Antiserum, Control | 2 | 866.5870 | Thyroid autoantibody immunological test system. | N | N
Immunology | DDL | Thyroglobulin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5870 | Thyroid autoantibody immunological test system. | N | N
Immunology | OZF | Tissue Rna Preservative For Collection, Storage, And Transportation | 2 | 866.4070 | RNA Preanalytical Systems. | N | N
Immunology | DFI | Total Spinal-Fluid, Antigen, Antiserum, Control | 1 | 866.5860 | Total spinal fluid immunological test system. | N | N
Immunology | DDG | Transferrin, Antigen, Antiserum, Control | 2 | 866.5880 | Transferrin immunological test system. | N | N
Immunology | DDI | Transferrin, Fitc, Antigen, Antiserum, Control | 2 | 866.5880 | Transferrin immunological test system. | N | N
Immunology | DDD | Transferrin, Rhodamine, Antigen, Antiserum, Control | 2 | 866.5880 | Transferrin immunological test system. | N | N
Immunology | OUY | Trichomonas Vaginalis Nucleic Acid Amplification Test System | 2 | 866.3860 | Trichomonas vaginalis nucleic acid assay. | N | N
Immunology | OYL | Tryptase Assay System | 2 | 866.5760 | Tryptase test system. | N | N
Immunology | OIF | Tyrosine Phosphatase (Ia-2) Autoantibody Assay | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | PAF | Voltage Gated Calcium Channel (Vgcc) Antibody Assay | 2 | 866.5660 | Multiple autoantibodies immunological test system. | N | N
Immunology | DGQ | Whole Blood Plasma, Antigen, Antiserum, Control | 1 | 866.5700 | Whole human plasma or serum immunological test system. | N | N
Immunology | DGR | Whole Human Serum, Antigen, Antiserum, Control | 1 | 866.5700 | Whole human plasma or serum immunological test system. | N | N
Immunology | PHF | Zinc Transporter 8 Auto-Antibody | 2 | 866.5670 | Zinc transporter 8 autoantibody immunological test system. | N | N
Microbiology | OQW | 2009 H1n1 Influenza Virus (Swine Origin), Nucleic Acid Or Antigen, Detection And Identification | 2 | 866.3332 | Reagents for detection of specific novel influenza A viruses. | N | N
Microbiology | JTM | Anaerobic Box Glove | 1 | 866.2120 | Anaerobic chamber. | N | N
Microbiology | LJN | Antibody Igm, If, Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | LKQ | Antibody Igm,If, Cytomegalovirus Virus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | GTR | Antideoxyribonuclease, Streptococcus Spp. | 1 | 866.3720 | Streptococcus spp. exoenzyme reagents. | N | N
Microbiology | GPF | Antigen, Agglutinating, Echinococcus Spp. | 1 | 866.3200 | Echinococcus spp. serological reagents. | N | N
Microbiology | GOT | Antigen, B. Parapertussis | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GOX | Antigen, B. Pertussis | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GPI | Antigen, Bentonite Flocculation, Trichinella Spiralis | 1 | 866.3850 | Trichinella spiralis serological reagents. | N | N
Microbiology | LSH | Antigen, Blastomyces Dermatitidis, Other | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | MCB | Antigen, C. Difficile | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | GQH | Antigen, Cf (Including Cf Control), Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | GNQ | Antigen, Cf (Including Cf Control), Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | GQN | Antigen, Cf (Including Cf Control), Herpesvirus Hominis 1,2 | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | GRC | Antigen, Cf (Including Cf Control), Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GON | Antigen, Cf (Including Cf Control), Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | GQG | Antigen, Cf (Including Cf Controls), Respiratory Syncytial Virus | 1 | 866.3480 | Respiratory syncytial virus serological reagents. | N | N
Microbiology | GMI | Antigen, Cf And / Or Id, Coccidioides Immitis | 2 | 866.3135 | Coccidioides immitis serological reagents. | N | N
Microbiology | GRJ | Antigen, Cf, (Including Cf Control), Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | GQW | Antigen, Cf, (Including Cf Control), Varicella-Zoster | 2 | 866.3900 | Varicella-zoster virus serological reagents. | N | N
Microbiology | JWT | Antigen, Cf, Aspergillus Spp. | 1 | 866.3040 | Aspergillus spp. serological reagents. | N | N
Microbiology | JWW | Antigen, Cf, B. Dermatitidis | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | GQK | Antigen, Cf, Lymphocytic Choriomeningitis Virus | 1 | 866.3360 | Lymphocytic choriomeningitis virus serological reagents. | N | N
Microbiology | GPW | Antigen, Cf, Psittacosis (Chlamydia Group) | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | GPS | Antigen, Cf, Q Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GPQ | Antigen, Cf, Spotted Fever Group | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GNF | Antigen, Cf, T. Cruzi | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | GPO | Antigen, Cf, Typhus Fever Group | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | MCD | Antigen, Ebv, Capsid | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | MDU | Antigen, Elisa, Cryptococcus | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | LIR | Antigen, Enzyme Linked Immunoabsorbent Assay, Neisseria Gonorrhoeae | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | GSF | Antigen, Erysipelothrix Rhusiopathiae | 1 | 866.3250 | Erysipelothrix rhusiopathiae serological reagents. | N | N
Microbiology | GNH | Antigen, Fluorescent Antibody Test, Schistosoma Mansoni | 1 | 866.3600 | Schistosoma spp. serological reagents. | N | N
Microbiology | NOM | Antigen, Galactomannan, Aspergillus Spp. | 1 | 866.3040 | Aspergillus spp. serological reagents. | N | N
Microbiology | GQY | Antigen, Ha (Including Ha Control), Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GOL | Antigen, Ha (Including Ha Control), Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | GRH | Antigen, Ha (Including Ha Control), Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | LHK | Antigen, Id, Candida Albicans | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | KHW | Antigen, Id, Ha, Cep, Entamoeba Histolytica & Rel. Spp. | 2 | 866.3220 | Entamoeba histolytica serological reagents. | N | N
Microbiology | LJO | Antigen, Iha, Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | GND | Antigen, Iha, T. Cruzi | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | GPD | Antigen, Indirect Fluorescent Antibody Test, Echinococcus Granulosus | 1 | 866.3200 | Echinococcus spp. serological reagents. | N | N
Microbiology | NTM | Antigen, Inflammatory Response Marker, Sepsis | 2 | 866.3210 | Endotoxin assay. | N | N
Microbiology | NQZ | Antigen, Invasive Fungal Pathogens | 2 | 866.3050 | Beta-glucan serological assays. | N | N
Microbiology | GMG | Antigen, Latex Agglutination, Coccidioides Immitis | 2 | 866.3135 | Coccidioides immitis serological reagents. | N | N
Microbiology | GMO | Antigen, Latex Agglutination, Entamoeba Histolytica & Rel. Spp. | 2 | 866.3220 | Entamoeba histolytica serological reagents. | N | N
Microbiology | GNE | Antigen, Latex Agglutination, T. Cruzi | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | GPG | Antigen, Latex Agglutination, Trichinella Spiralis | 1 | 866.3850 | Trichinella spiralis serological reagents. | N | N
Microbiology | JWK | Antigen, Positive Control, Cryptococcus Neoformans | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | LSJ | Antigen, Rubella, Other | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | JWL | Antigen, Treponema Pallidum For Fta-Abs Test | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | GSO | Antigens (Febrile), Agglutination, Brucella Spp. | 2 | 866.3085 | Brucella spp. serological reagents. | N | N
Microbiology | GRL | Antigens, All Groups, Salmonella Spp. | 2 | 866.3550 | Salmonella spp. serological reagents. | N | N
Microbiology | LIA | Antigens, All Groups, Shigella Spp. | 2 | 866.3660 | Shigella spp. serological reagents. | N | N
Microbiology | GTY | Antigens, All Groups, Streptococcus Spp. | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | GMZ | Antigens, All Types, Escherichia Coli | 1 | 866.3255 | Escherichia coli serological reagents. | N | N
Microbiology | GRY | Antigens, All, Leptospira Spp. | 2 | 866.3350 | Leptospira spp. serological reagents. | N | N
Microbiology | GOD | Antigens, Cf (Including Cf Control), Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GNG | Antigens, Cf (Including Cf Control), Coxsackievirus A 1-24, B 1-6 | 1 | 866.3145 | Coxsackievirus serological reagents. | N | N
Microbiology | GNL | Antigens, Cf (Including Cf Control), Echovirus 1-34 | 1 | 866.3205 | Echovirus serological reagents. | N | N
Microbiology | GQD | Antigens, Cf (Including Cf Control), Equine Encephalitis Virus, Eee, Wee | 1 | 866.3240 | Equine encephalomyelitis virus serological reagents. | N | N
Microbiology | GNX | Antigens, Cf (Including Cf Control), Influenza Virus A, B, C | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | GQS | Antigens, Cf (Including Cf Control), Parainfluenza Virus 1-4 | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | GOH | Antigens, Cf (Including Cf Control), Poliovirus 1-3 | 1 | 866.3405 | Poliovirus serological reagents. | N | N
Microbiology | GQB | Antigens, Cf (Including Cf Control), Reovirus 1-3 | 1 | 866.3470 | Reovirus serological reagents. | N | N
Microbiology | GSB | Antigens, Cf, All, Mycoplasma Spp. | 1 | 866.3375 | Mycoplasma spp. serological reagents. | N | N
Microbiology | GMN | Antigens, Cf, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | GSZ | Antigens, Febrile (Weil-Felix), All Groups | 1 | 866.3410 | Proteus spp. (Weil-Felix) serological reagents. | N | N
Microbiology | GNC | Antigens, Febrile, Slide And Tube, All Groups, Salmonella Spp. | 2 | 866.3550 | Salmonella spp. serological reagents. | N | N
Microbiology | GOB | Antigens, Ha (Including Ha Control), Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GNT | Antigens, Ha (Including Ha Control), Influenza Virus A, B, C | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | GQR | Antigens, Ha (Including Ha Control), Parainfluenza Virus 1-4 | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | GQA | Antigens, Ha (Including Ha Control), Reovirus 1-3 | 1 | 866.3470 | Reovirus serological reagents. | N | N
Microbiology | GNJ | Antigens, Ha, Echovirus 1-34 | 1 | 866.3205 | Echovirus serological reagents. | N | N
Microbiology | GMT | Antigens, Ha, Treponema Pallidum | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | GMJ | Antigens, Histoplasma Capsulatum, All | 2 | 866.3320 | Histoplasma capsulatum serological reagents. | N | N
Microbiology | GLZ | Antigens, If, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | GMM | Antigens, Iha, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | LKC | Antigens, Indirect Hemagglutination (Iha) Herpes Simplex Virus | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | GMQ | Antigens, Nontreponemal, All | 2 | 866.3820 | Treponema pallidum nontreponemal test reagents. | N | N
Microbiology | GSI | Antigens, Slide And Tube, All Types, Listeria Monocytogenes | 1 | 866.3355 | Listeria spp. serological reagents. | N | N
Microbiology | GSL | Antigens, Slide And Tube, Francisella Tularensis | 2 | 866.3280 | Francisella tularensis serological reagents. | N | N
Microbiology | GMS | Anti-Human Globulin, Fta-Abs Test | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | OYP | Anti-Jcv Antibody Detection Assay | 2 | 866.3336 | John Cunningham Virus serological reagents. | N | N
Microbiology | PGV | Anti-Phospholipase A2 Receptor | 2 | 866.5780 |  | Y | N
Microbiology | GSX | Antisera, Acinetobacter Calcoaceticus, All Varieties | 1 | 866.3010 | Acinetobacter calcoaceticus serological reagents. | N | N
Microbiology | GOW | Antisera, Agglutinating, B. Parapertussis | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GOY | Antisera, Agglutinating, B. Pertussis, All | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GRW | Antisera, All Fluorescent, Leptospira Spp. | 2 | 866.3350 | Leptospira spp. serological reagents. | N | N
Microbiology | GTJ | Antisera, All Groups, N. Meningitidis | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | GRM | Antisera, All Groups, Salmonella Spp. | 2 | 866.3550 | Salmonella spp. serological reagents. | N | N
Microbiology | GTZ | Antisera, All Groups, Streptococcus Spp. | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | GRX | Antisera, All Leptospira Spp. | 2 | 866.3350 | Leptospira spp. serological reagents. | N | N
Microbiology | GSA | Antisera, All Mycoplasma Spp. | 1 | 866.3375 | Mycoplasma spp. serological reagents. | N | N
Microbiology | GNA | Antisera, All Types, Escherichia Coli | 1 | 866.3255 | Escherichia coli serological reagents. | N | N
Microbiology | GRP | Antisera, All Types, H. Influenza | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | GTC | Antisera, All Types, Klebsiella Spp. | 1 | 866.3340 | Klebsiella spp. serological reagents. | N | N
Microbiology | GSH | Antisera, All Types, Listeria Monocytogenes | 1 | 866.3355 | Listeria spp. serological reagents. | N | N
Microbiology | GNB | Antisera, All Types, Shigella Spp. | 2 | 866.3660 | Shigella spp. serological reagents. | N | N
Microbiology | GWC | Antisera, All Types, Streptococcus Pneumoniae | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | GTA | Antisera, All, Serratia Marcesans | 1 | 866.3630 | Serratia spp. serological reagents. | N | N
Microbiology | GTE | Antisera, Arizona Spp., All | 1 | 866.3035 | Arizona spp. serological reagents. | N | N
Microbiology | GTF | Antisera, Bethesda - Ballerup Polyvalent, Citrobacter Spp. | 1 | 866.3125 | Citrobacter spp. serological reagents. | N | N
Microbiology | KLH | Antisera, C. Acnes | 1 | 866.3140 | Corynebacterium spp. serological reagents. | N | N
Microbiology | GOP | Antisera, C. Acnes (553, 605) | 1 | 866.3140 | Corynebacterium spp. serological reagents. | N | N
Microbiology | GOA | Antisera, Cf, Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GNO | Antisera, Cf, Coxsackievirus A 1-24, B 1-6 | 1 | 866.3145 | Coxsackievirus serological reagents. | N | N
Microbiology | GNK | Antisera, Cf, Echovirus 1-34 | 1 | 866.3205 | Echovirus serological reagents. | N | N
Microbiology | GQC | Antisera, Cf, Equine Encephalitis Virus, Eee, Wee | 1 | 866.3240 | Equine encephalomyelitis virus serological reagents. | N | N
Microbiology | GQO | Antisera, Cf, Herpesvirus Hominis 1,2 | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | GNW | Antisera, Cf, Influenza Virus A, B, C | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | GQT | Antisera, Cf, Parainfluenza Virus 1-4 | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | GOG | Antisera, Cf, Poliovirus 1-3 | 1 | 866.3405 | Poliovirus serological reagents. | N | N
Microbiology | GPZ | Antisera, Cf, Reovirus 1-3 | 1 | 866.3470 | Reovirus serological reagents. | N | N
Microbiology | GOM | Antisera, Cf, Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | LIN | Antisera, Conjugated Fluorescent, Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | GMP | Antisera, Control For Nontreponemal Tests | 2 | 866.3820 | Treponema pallidum nontreponemal test reagents. | N | N
Microbiology | GMX | Antisera, Fluorescent Antibody For Fta-Abs Test | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | GNY | Antisera, Fluorescent, Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GSY | Antisera, Fluorescent, All Globulins, Proteus Spp. | 1 | 866.3410 | Proteus spp. (Weil-Felix) serological reagents. | N | N
Microbiology | GOO | Antisera, Fluorescent, All Globulins, Salmonella Spp. | 2 | 866.3550 | Salmonella spp. serological reagents. | N | N
Microbiology | GTD | Antisera, Fluorescent, All Globulins, Shigella Spp. | 2 | 866.3660 | Shigella spp. serological reagents. | N | N
Microbiology | GTI | Antisera, Fluorescent, All Groups, N. Meningitidis | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | GTX | Antisera, Fluorescent, All Groups, Streptococcus Spp. | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | GMY | Antisera, Fluorescent, All Types, Escherichia Coli | 1 | 866.3255 | Escherichia coli serological reagents. | N | N
Microbiology | GRO | Antisera, Fluorescent, All Types, Hemophilus Spp. | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | GTB | Antisera, Fluorescent, All Types, Klebsiella Spp. | 1 | 866.3340 | Klebsiella spp. serological reagents. | N | N
Microbiology | GSG | Antisera, Fluorescent, All Types, Listeria Monocytogenes | 1 | 866.3355 | Listeria spp. serological reagents. | N | N
Microbiology | GTN | Antisera, Fluorescent, All Types, Staphylococcus Spp. | 1 | 866.3700 | Staphylococcus aureus serological reagents. | N | N
Microbiology | GWB | Antisera, Fluorescent, All Types, Streptococcus Pneumoniae | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | GRZ | Antisera, Fluorescent, All, Mycoplasma Spp. | 1 | 866.3375 | Mycoplasma spp. serological reagents. | N | N
Microbiology | JRW | Antisera, Fluorescent, B. Parapertussis | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GOZ | Antisera, Fluorescent, B. Pertussis | 1 | 866.3065 | Bordetella spp. serological reagents. | N | N
Microbiology | GSM | Antisera, Fluorescent, Brucella Spp. | 2 | 866.3085 | Brucella spp. serological reagents. | N | N
Microbiology | LKI | Antisera, Fluorescent, Chlamydia Spp. | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | GNM | Antisera, Fluorescent, Coxsackievirus A 1-24, B 1-6 | 1 | 866.3145 | Coxsackievirus serological reagents. | N | N
Microbiology | GME | Antisera, Fluorescent, Cryptococcus Neoformans | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | GRK | Antisera, Fluorescent, Echovirus 1-34 | 1 | 866.3205 | Echovirus serological reagents. | N | N
Microbiology | GQL | Antisera, Fluorescent, Herpesvirus Hominis 1,2 | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | GML | Antisera, Fluorescent, Histoplasma Capsulatum | 2 | 866.3320 | Histoplasma capsulatum serological reagents. | N | N
Microbiology | OMG | Antisera, Fluorescent, Human Metapneumovirus | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | GOE | Antisera, Fluorescent, Poliovirus 1-3 | 1 | 866.3405 | Poliovirus serological reagents. | N | N
Microbiology | GSS | Antisera, Fluorescent, Pseudomonas Aeruginosa | 2 | 866.3415 | Pseudomonas spp. serological reagents. | N | N
Microbiology | GSR | Antisera, Fluorescent, Pseudomonas Pseudomallei | 2 | 866.3415 | Pseudomonas spp. serological reagents. | N | N
Microbiology | GMA | Antisera, Fluorescent, Sporothrix Schenekii | 1 | 866.3680 | Sporothrix schenckii serological reagents. | N | N
Microbiology | GOK | Antisera, Hai (Including Hai Control), Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | GOC | Antisera, Hai, Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GNS | Antisera, Hai, Influenza Virus A, B, C | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | GQQ | Antisera, Hai, Parainfluenza Virus 1-4 | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | GPY | Antisera, Hai, Reovirus 1-3 | 1 | 866.3470 | Reovirus serological reagents. | N | N
Microbiology | LJK | Antisera, If, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | LKH | Antisera, Immunoperoxidase, Chlamydia Spp. | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | GMD | Antisera, Latex Agglutination, Cryptococcus Neoformans | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | GNZ | Antisera, Neutralization, Adenovirus 1-33 | 1 | 866.3020 | Adenovirus serological reagents. | N | N
Microbiology | GQE | Antisera, Neutralization, All Types, Rhinovirus | 1 | 866.3490 | Rhinovirus serological reagents. | N | N
Microbiology | GNN | Antisera, Neutralization, Coxsackievirus A 1-24, B 1-6 | 1 | 866.3145 | Coxsackievirus serological reagents. | N | N
Microbiology | GNI | Antisera, Neutralization, Echovirus 1-34 | 1 | 866.3205 | Echovirus serological reagents. | N | N
Microbiology | GQM | Antisera, Neutralization, Herpesvirus Hominis | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | GNR | Antisera, Neutralization, Influenza Virus A, B, C | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | GQP | Antisera, Neutralization, Parainfluenza Virus 1-4 | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | GOF | Antisera, Neutralization, Poliovirus 1-3 | 1 | 866.3405 | Poliovirus serological reagents. | N | N
Microbiology | GOJ | Antisera, Neutralization, Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | GPX | Antisera, Neutralizion, Reovirus 1-3 | 1 | 866.3470 | Reovirus serological reagents. | N | N
Microbiology | GPE | Antisera, Positive Control, Echinococcus Spp. | 1 | 866.3200 | Echinococcus spp. serological reagents. | N | N
Microbiology | GPH | Antiserum, Bentonite Flocculation, Trichinella Spiralis | 1 | 866.3850 | Trichinella spiralis serological reagents. | N | N
Microbiology | LSI | Antiserum, Blastomyces Dermatitidis, Other | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | GQI | Antiserum, Cf, Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | GNP | Antiserum, Cf, Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | GQJ | Antiserum, Cf, Lymphocytic Choriomeningitis Virus | 1 | 866.3360 | Lymphocytic choriomeningitis virus serological reagents. | N | N
Microbiology | GRB | Antiserum, Cf, Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GPT | Antiserum, Cf, Psittacosis (Chlamydia Group) | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | GPR | Antiserum, Cf, Q Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GRF | Antiserum, Cf, Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | GQX | Antiserum, Cf, Varicella-Zoster | 2 | 866.3900 | Varicella-zoster virus serological reagents. | N | N
Microbiology | LIC | Antiserum, Coagglutination (Direct) Neisseria Gonorrhoeae | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | GSE | Antiserum, Erysipelothrix Rhusiopathiae | 1 | 866.3250 | Erysipelothrix rhusiopathiae serological reagents. | N | N
Microbiology | GSW | Antiserum, Flavobacterium Meningosepticum, All Groups | 1 | 866.3270 | Flavobacterium spp. serological reagents. | N | N
Microbiology | GTH | Antiserum, Fluorescent (Direct Test), All Groups, N. Gonorrhoeae | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | GOS | Antiserum, Fluorescent, C. Diphtheriae | 1 | 866.3140 | Corynebacterium spp. serological reagents. | N | N
Microbiology | GSP | Antiserum, Fluorescent, Campylobacter Fetus | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | LJP | Antiserum, Fluorescent, Chlamydia Trachomatis | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | JRY | Antiserum, Fluorescent, Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | GSD | Antiserum, Fluorescent, Erysipelothrix Rhusiopathiae | 1 | 866.3250 | Erysipelothrix rhusiopathiae serological reagents. | N | N
Microbiology | GSJ | Antiserum, Fluorescent, Francisella Tularensis | 2 | 866.3280 | Francisella tularensis serological reagents. | N | N
Microbiology | GRA | Antiserum, Fluorescent, Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GRT | Antiserum, Fluorescent, Mycobacterium Tuberculosis | 1 | 866.3370 | Mycobacterium tuberculosis immunofluorescent reagents. | N | N
Microbiology | GPJ | Antiserum, Fluorescent, Q Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GOI | Antiserum, Fluorescent, Rabies Virus | 2 | 866.3460 | Rabiesvirus immunofluorescent reagents. | N | N
Microbiology | GRE | Antiserum, Fluorescent, Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | GSK | Antiserum, Francisella Tularensis | 2 | 866.3280 | Francisella tularensis serological reagents. | N | N
Microbiology | GRD | Antiserum, Hai, Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GRG | Antiserum, Hai, Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | GPM | Antiserum, Murine Typhus Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GQZ | Antiserum, Neutralization, Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | GQF | Antiserum, Neutralization, Respiratory Syncytial Virus | 1 | 866.3480 | Respiratory syncytial virus serological reagents. | N | N
Microbiology | GRI | Antiserum, Neutralization, Rubeola | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | GSN | Antiserum, Positive And Negative Febrile Antigen Control Serum | 2 | 866.3085 | Brucella spp. serological reagents. | N | N
Microbiology | KFG | Antiserum, Positive Control, Aspergillus Spp. | 1 | 866.3040 | Aspergillus spp. serological reagents. | N | N
Microbiology | KFH | Antiserum, Positive Control, Blastomyces Dermatitidis | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | GMH | Antiserum, Positive Control, Coccidioides Immitis | 2 | 866.3135 | Coccidioides immitis serological reagents. | N | N
Microbiology | GMK | Antiserum, Positive Control, Histoplasma Capsulatum | 2 | 866.3320 | Histoplasma capsulatum serological reagents. | N | N
Microbiology | GST | Antiserum, Pseudomonas Pseudomallei | 2 | 866.3415 | Pseudomonas spp. serological reagents. | N | N
Microbiology | GPK | Antiserum, Rickettsial Pox | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GPP | Antiserum, Rocky Mountain Spotted Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GPN | Antiserum, Typhus Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | GSQ | Antiserum, Vibrio Cholerae, All Varieties | 2 | 866.3930 | Vibrio cholerae serological reagents. | N | N
Microbiology | GTO | Anti-Streptokinase | 1 | 866.3720 | Streptococcus spp. exoenzyme reagents. | N | N
Microbiology | GTQ | Antistreptolysin - Titer/Streptolysin O Reagent | 1 | 866.3720 | Streptococcus spp. exoenzyme reagents. | N | N
Microbiology | QFA | Assay For Detection Of Proteases In Chronic Wounds | 2 | 866.3231 |  | N | N
Microbiology | PTF | Assay To Measure Pct To Aid In The Risk Assessment Of Critically Ill Patients On Their First Day Of Icu Admission | 2 | 866.3215 | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | N | N
Microbiology | NIA | Assay, Direct, Nucleic Acid Amplification, Influenza Virus | 1 | 866.3330 | Influenza virus serological reagents. | N | N
Microbiology | NHZ | Assay, Direct, Nucleic Acid Amplification, Parainfluenza Virus | 1 | 866.3400 | Parainfluenza virus serological reagents. | N | N
Microbiology | OVF | Assay, Direct, Nucleic Acid Amplification, Q Fever | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | NHY | Assay, Direct, Nucleic Acid Amplification, Respiratory Syncytial Virus | 1 | 866.3480 | Respiratory syncytial virus serological reagents. | N | N
Microbiology | NGS | Assay, Endotoxin Activity, Chemiluminescent | 2 | 866.3210 | Endotoxin assay. | N | N
Microbiology | OAI | Assay, Enterovirus Nucleic Acid | 2 | 866.3225 | Enterovirus nucleic acid assay. | N | N
Microbiology | NHT | Assay, Nucleic Acid Amplification, Bacillus Anthracis | 2 | 866.3045 | In vitro diagnostic device for <I>Bacillus</I> spp. detection. | N | N
Microbiology | NDZ | Assay, Nucleic Acid Amplification, Growth Identification, Mycobacterium Tuberculosis | 1 | 866.3370 | Mycobacterium tuberculosis immunofluorescent reagents. | N | N
Microbiology | PMN | Assayed External Control Material For Microbiology Nucleic Acid Amplification (Nat) Assays | 2 | 866.3920 | Assayed quality control material for clinical microbiology assays. | N | N
Microbiology | QCH | Assayed Quality Control Material For Clinical Microbiology Assays | 2 | 866.3920 | Assayed quality control material for clinical microbiology assays. | N | N
Microbiology | KZB | Automated Colony Counter | 1 | 866.2170 | Automated colony counter. | N | N
Microbiology | NVQ | Bacteriophage And Controls, B. Anthracis Lysis | 2 | 866.3045 | In vitro diagnostic device for <I>Bacillus</I> spp. detection. | N | N
Microbiology | JTQ | Bath, Incubators/Water, All | 1 | 866.2540 | Microbiological incubator. | N | N
Microbiology | OZZ | Bordetella Pertussis Dna Assay System | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | OMN | C. Difficile Nucleic Acid Amplification Test Assay | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | OZN | C. Difficile Toxin Gene Amplification Assay | 2 | 866.3130 | Clostridium difficile toxin gene amplification assay. | N | N
Microbiology | MGM | C. Trachomatis (Chlamydia Group) | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | LQP | Campylobacter Spp. | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | PII | Candida Species Nucleic Acid Detection System | 2 | 866.3960 | Nucleic acid-based device for the amplification, detection, and identification of microbial pathogens directly from whole blood specimens. | N | N
Microbiology | LSG | Candida Species, Antibody Detection | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | LRF | Candida Spp., Direct Antigen, Id | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | OZY | Chlamydophila Pneumoniae Dna Assay System | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | QAA | Coccidioides Spp. Nucleic Acid Detection System For Respiratory Specimens | 2 | 866.3376 |  | N | N
Microbiology | MHJ | Cryptosporidium Spp. | 2 | 866.3220 | Entamoeba histolytica serological reagents. | N | N
Microbiology | JSL | Culture Media, Anaerobic Transport | 1 | 866.2390 | Transport culture medium. | N | N
Microbiology | JSA | Culture Media, Antibiotic Assay | 1 | 866.2350 | Microbiological assay culture medium. | N | N
Microbiology | MJE | Culture Media, Antifungal, Susceptibility Test | 2 | 866.1700 | Culture medium for antimicrobial susceptibility tests. | N | N
Microbiology | LKA | Culture Media, Antimicrobial Susceptibility Test | 2 | 866.1700 | Culture medium for antimicrobial susceptibility tests. | N | N
Microbiology | JSO | Culture Media, Antimicrobial Susceptibility Test, Excluding Mueller Hinton Agar | 2 | 866.1700 | Culture medium for antimicrobial susceptibility tests. | N | N
Microbiology | JTZ | Culture Media, Antimicrobial Susceptibility Test, Mueller Hinton Agar/Broth | 2 | 866.1700 | Culture medium for antimicrobial susceptibility tests. | N | N
Microbiology | MJD | Culture Media, Antimycobacteria, Susceptibility Test | 2 | 866.1700 | Culture medium for antimicrobial susceptibility tests. | N | N
Microbiology | KZI | Culture Media, Enriched | 1 | 866.2330 | Enriched culture medium. | N | N
Microbiology | JTY | Culture Media, For Isolation Of Pathogenic Neisseria | 2 | 866.2410 | Culture medium for pathogenic Neisseria spp. | N | N
Microbiology | JSC | Culture Media, General Nutrient Broth | 1 | 866.2350 | Microbiological assay culture medium. | N | N
Microbiology | JSE | Culture Media, Multiple Biochemical Test | 1 | 866.2320 | Differential culture medium. | N | N
Microbiology | JSM | Culture Media, Non-Propagating Transport | 1 | 866.2390 | Transport culture medium. | N | N
Microbiology | JSH | Culture Media, Non-Selective And Differential | 1 | 866.2320 | Differential culture medium. | N | N
Microbiology | JSG | Culture Media, Non-Selective And Non-Differential | 1 | 866.2300 | Multipurpose culture medium. | N | N
Microbiology | JSN | Culture Media, Propagating Transport | 1 | 866.2390 | Transport culture medium. | N | N
Microbiology | JSI | Culture Media, Selective And Differential | 1 | 866.2360 | Selective culture medium. | N | N
Microbiology | JSJ | Culture Media, Selective And Non-Differential | 1 | 866.2360 | Selective culture medium. | N | N
Microbiology | JSD | Culture Media, Selective Broth | 1 | 866.2360 | Selective culture medium. | N | N
Microbiology | JSF | Culture Media, Single Biochemical Test | 1 | 866.2320 | Differential culture medium. | N | N
Microbiology | JSB | Culture Media, Vitamin Assay | 1 | 866.2350 | Microbiological assay culture medium. | N | N
Microbiology | QIV | Culture-Based Short-Term Incubation Antimicrobial Resistance Assay | 2 | 866.1655 |  | N | N
Microbiology | OZB | Dengue Nucleic Acid Amplification Assay (Naat) | 2 | 866.3946 | Dengue virus nucleic acid amplification test reagents. | N | N
Microbiology | OSU | Dengue Serological Reagents | 2 | 866.3945 | Dengue virus serological reagents. | N | N
Microbiology | QCU | Dengue Virus Antigen Assay | 2 | 866.3945 | Dengue virus serological reagents. | N | N
Microbiology | PMT | Device To Detect And Measure Procalcitonin (Pct) In Human Clinical Specimens | 2 | 866.3215 | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | N | N
Microbiology | QID | Device To Detect Antigens Of Biothreat Microbial Agents In Human Clinical Specimens | 2 | 866.4002 |  | Y | N
Microbiology | OJY | Device To Detect Or Measure Nucleic Acid From Viruses Associated With Head And Neck Cancers | 2 | 866.3236 |  | N | N
Microbiology | LJF | Device, Antimicrobial Drug Removal | 1 | 866.2560 | Microbial growth monitor. | N | N
Microbiology | KZJ | Device, Gas Generating | 1 | 866.2580 | Gas-generating device. | N | N
Microbiology | LIB | Device, General Purpose, Microbiology, Diagnostic | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JTB | Device, Media Dispensing/Stacking | 1 | 866.2440 | Automated medium dispensing and stacking device. | N | N
Microbiology | JTC | Device, Microtiter Diluting/Dispensing | 1 | 866.2500 | Microtiter diluting and dispensing device. | N | N
Microbiology | LGA | Device, Oxidase Test For Gonorrhea | 3 | 866.2420 | Oxidase screening test for gonorrhea. | N | N
Microbiology | LKS | Device, Parasite Concentration | 1 | 866.2900 | Microbiological specimen collection and transport device. | N | N
Microbiology | LIO | Device, Specimen Collection | 1 | 866.2900 | Microbiological specimen collection and transport device. | N | N
Microbiology | PSZ | Devices Detecting Influenza A, B, And C Virus Antigens | 2 | 866.3328 | Influenza virus antigen detection test system. | N | N
Microbiology | LLA | Direct Agglutination Test, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | QBX | Direct Blood Bacterial Nucleic Acid Detection System | 2 | 866.3960 | Nucleic acid-based device for the amplification, detection, and identification of microbial pathogens directly from whole blood specimens. | N | N
Microbiology | LTX | Discs, Elution | 2 | 866.1620 | Antimicrobial susceptibility test disc. | N | N
Microbiology | JTO | Discs, Strips And Reagents, Microorganism Differentiation | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | MJM | Dna Probe, Gardnerella Vaginalis | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | MKZ | Dna Probe, Nucleic Acid Amplification, Chlamydia | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | MJK | Dna Probe, Trichomonas Vaginalis | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | MLA | Dna Probe, Yeast | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | MDC | Dna-Probe - Blastomyces Dermatitidis | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | MCG | Dna-Probe, Agent, Listeria | 1 | 866.3355 | Listeria spp. serological reagents. | N | N
Microbiology | MCC | Dna-Probe, Haemophilus Spp. | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | MBT | Dna-Probe, Reagent, Histoplasma Capsulatum | 2 | 866.3320 | Histoplasma capsulatum serological reagents. | N | N
Microbiology | MDF | Dna-Probe, Reagents, Coccidioides Immitis | 2 | 866.3135 | Coccidioides immitis serological reagents. | N | N
Microbiology | MDE | Dna-Probe, Reagents, Cryptococcal | 2 | 866.3165 | Cryptococcus neoformans serological reagents. | N | N
Microbiology | MDK | Dna-Probe, Reagents, Streptococcal | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | MCS | Dna-Probe, Staphylococcus Aureus | 1 | 866.3700 | Staphylococcus aureus serological reagents. | N | N
Microbiology | MCT | Dna-Probe, Strep Pneumoniae | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | LQO | Dna-Reagents, Campylobacter Spp. | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | LSK | Dna-Reagents, Chlamydia | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | LSO | Dna-Reagents, Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | LSF | Dna-Reagents, Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | LQH | Dna-Reagents, Legionella | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | LQF | Dna-Reagents, Mycobacterium Spp. | 1 | 866.3370 | Mycobacterium tuberculosis immunofluorescent reagents. | N | N
Microbiology | LQG | Dna-Reagents, Mycoplasma Spp. | 1 | 866.3375 | Mycoplasma spp. serological reagents. | N | N
Microbiology | LSL | Dna-Reagents, Neisseria | 2 | 866.3390 | Neisseria spp. direct serological test reagents. | N | N
Microbiology | LSM | Dna-Reagents, Salmonella Spp. | 2 | 866.3550 | Salmonella spp. serological reagents. | N | N
Microbiology | LSN | Dna-Reagents, Shigella Spp. | 2 | 866.3660 | Shigella spp. serological reagents. | N | N
Microbiology | MJL | Eia, Blastomyces Dermatitidis | 2 | 866.3060 | Blastomyces dermatitidis serological reagents. | N | N
Microbiology | NOP | Elisa, Antibody, West Nile Virus | 2 | 866.3940 | West Nile virus serological reagents. | N | N
Microbiology | MDT | Elisa, Trichinella Spiralis | 1 | 866.3850 | Trichinella spiralis serological reagents. | N | N
Microbiology | LJC | Enzyme Linked Immunoabsorbent Assay, (Chlamydiae Group) | 1 | 866.3120 | Chlamydia serological reagents. | N | N
Microbiology | NRL | Enzyme Linked Immunoabsorbent Assay, Antibody, B. Anthracis | 2 | 866.3045 | In vitro diagnostic device for <I>Bacillus</I> spp. detection. | N | N
Microbiology | LFZ | Enzyme Linked Immunoabsorbent Assay, Cytomegalovirus | 2 | 866.3175 | Cytomegalovirus serological reagents. | N | N
Microbiology | LGC | Enzyme Linked Immunoabsorbent Assay, Herpes Simplex Virus, Non-Specific | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | MIZ | Enzyme Linked Immunoabsorbent Assay, Histoplasma Capsulatum | 2 | 866.3320 | Histoplasma capsulatum serological reagents. | N | N
Microbiology | LJY | Enzyme Linked Immunoabsorbent Assay, Mumps Virus | 1 | 866.3380 | Mumps virus serological reagents. | N | N
Microbiology | LJZ | Enzyme Linked Immunoabsorbent Assay, Mycoplasma Spp. | 1 | 866.3375 | Mycoplasma spp. serological reagents. | N | N
Microbiology | LIQ | Enzyme Linked Immunoabsorbent Assay, Rotavirus | 1 | 866.3405 | Poliovirus serological reagents. | N | N
Microbiology | LFX | Enzyme Linked Immunoabsorbent Assay, Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | LJB | Enzyme Linked Immunoabsorbent Assay, Rubeola Igg | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | PCL | Enzyme Linked Immunoabsorbent Assay, Rubeola Igm | 1 | 866.3520 | Rubeola (measles) virus serological reagents. | N | N
Microbiology | LGD | Enzyme Linked Immunoabsorbent Assay, Toxoplasma Gondii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | LFY | Enzyme Linked Immunoabsorbent Assay, Varicella-Zoster | 2 | 866.3900 | Varicella-zoster virus serological reagents. | N | N
Microbiology | LIP | Enzyme Linked Immunoabsorption Assay, Treponema Pallidum | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | MIY | Enzyme Linked Immunosorbent Assay, Coccidioides Immitis | 2 | 866.3135 | Coccidioides immitis serological reagents. | N | N
Microbiology | MXJ | Enzyme Linked Immunosorbent Assay, Herpes Simplex Virus, Hsv-1 | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | MYF | Enzyme Linked Immunosorbent Assay, Herpes Simplex Virus, Hsv-2 | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | MIU | Enzyme Linked Immunosorbent Assay, T. Cruzi | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | LSE | Epstein-Barr Virus, Other | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | LHW | Erythrocyte Suspension, Multi Species, Serological Reagent And Equipment | 1 | 866.3720 | Streptococcus spp. exoenzyme reagents. | N | N
Microbiology | GTP | Exoenzymes, Multiple, Streptococcal | 1 | 866.3720 | Streptococcus spp. exoenzyme reagents. | N | N
Microbiology | PYN | Fish (Fluorescence In Situ Hybridization) Kit, Nucleic Acid, Plasmodium Sp. | 2 | 866.3367 |  | N | N
Microbiology | OAH | Fish (Fluorescent In Situ Hybridization) Kit, Protein Nucleic Acid, Enterococcus Faecalis | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | NXX | Fish (Fluorescent In Situ Hybridization) Kit, Protein Nucleic Acid, Rna, Staphylococcus Aureus | 1 | 866.3700 | Staphylococcus aureus serological reagents. | N | N
Microbiology | NZS | Fish (Fluorescent In Situ Hybridization) Kit, Protein Nucleic Acid, Rna, Yeast | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | PEO | Fungal Organisms, Nucleic Acid-Based Assay | 2 | 866.3365 | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | N | N
Microbiology | NWZ | Gas Chromatography, Bacillus Anthracis Membrane Fatty Acids | 2 | 866.3045 | In vitro diagnostic device for <I>Bacillus</I> spp. detection. | N | N
Microbiology | PCI | Gastrointestinal Bacterial Panel Multiplex Nucleic Acid-Based Assay System | 2 | 866.3990 | Gastrointestinal microorganism multiplex nucleic acid-based assay. | N | N
Microbiology | PCH | Gastrointestinal Pathogen Panel Multiplex Nucleic Acid-Based Assay System | 2 | 866.3990 | Gastrointestinal microorganism multiplex nucleic acid-based assay. | N | N
Microbiology | PQZ | Gbs Culture Media, Selective And Differential | 1 | 866.2360 | Selective culture medium. | N | N
Microbiology | MHI | Giardia Spp. | 2 | 866.3220 | Entamoeba histolytica serological reagents. | N | N
Microbiology | LGB | Gonococcal Antibody Tests | 3 | 866.3290 | Gonococcal antibody test (GAT). | N | N
Microbiology | LQM | Gram Negative Identification Panel | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | LQL | Gram Positive Identification Panel | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | PEN | Gram-Negative Bacteria And Associated Resistance Markers | 2 | 866.3365 | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | N | N
Microbiology | PAM | Gram-Positive Bacteria And Their Resistance Markers | 2 | 866.3365 | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | N | N
Microbiology | OYZ | Group A Streptococcus Nucleic Acid Amplification Assay System | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | PGX | Groups A, C And G Beta-Hemolytic Streptococcus Nucleic Acid Amplification System | 2 | 866.2680 | Streptococcus spp. nucleic acid-based assay. | N | N
Microbiology | LYR | Helicobacter Pylori | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | LOL | Hepatitis A Test (Antibody And Igm Antibody) | 2 | 866.3310 | Hepatitis A virus (HAV) serological assays. | N | N
Microbiology | OQO | Herpes Simplex Virus Nucleic Acid Amplification Assay | 2 | 866.3305 | Herpes simplex virus serological assays. | N | N
Microbiology | PGI | Herpes Virus (Vzv, Hsv1, Hsv2), Dna Detection Assay For Cutaneous And Mucocutaneous Lesion Samples | 2 | 866.3309 | Herpes virus nucleic acid-based cutaneous and mucocutaneous lesion panel. | N | N
Microbiology | PGH | Hsv-1 And Hsv-2 Cns Nucleic-Acid Based Panel | 2 | 866.3307 |  | Y | N
Microbiology | OEM | Human Metapneumovirus (Hmpv) Rna Assay System | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | MIV | Immunofluorescent Assay, T. Cruzi | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | OWF | Immunohistochemical Assay, Helicobacter Pylori | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | GWD | Indirect Fluorescent Antibody Test, Entamoeba Histolytica & Rel Sp | 2 | 866.3220 | Entamoeba histolytica serological reagents. | N | N
Microbiology | OZE | Influenza A And Influenza B Multiplex Nucleic Acid Assay | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | OEP | Influenza A Virus Subtype Differentiation Nucleic Acid Assay | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | LRG | Instrument For Auto Reader & Interpretation Of Overnight Suscept. Systems | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | LRH | Instrument For Auto Reader Of Overnight Microorganism Identification System | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | OEH | Joint Biological Agent Identification And Diagnostic System (Jbaids) Tularemia Detection Kit | 2 | 866.3280 | Francisella tularensis serological reagents. | N | N
Microbiology | JSP | Kit, Anaerobic Identification | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JST | Kit, Fastidious Organisms | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSR | Kit, Identification, Dermatophyte | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSS | Kit, Identification, Enterobacteriaceae | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSW | Kit, Identification, Glucose Nonfermenter | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSY | Kit, Identification, Mycobacteria | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSX | Kit, Identification, Neisseria Gonorrhoeae | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JSZ | Kit, Identification, Pseudomonas | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JXB | Kit, Identification, Yeast | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | NPO | Kit, Immunochromatographic, Bacillus Anthracis Differential Antibody | 2 | 866.3045 | In vitro diagnostic device for <I>Bacillus</I> spp. detection. | N | N
Microbiology | JTR | Kit, Quality Control For Culture Media | 1 | 866.2480 | Quality control kit for culture media. | N | N
Microbiology | JWX | Kit, Screening, Staphylococcus Aureus | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JWZ | Kit, Screening, Trichomonas | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JXA | Kit, Screening, Urine | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JXC | Kit, Screening, Yeast | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | QGN | Lateral Flow Immunochromatography Assay For Host Infection Biomarkers | 2 | 866.3230 |  | N | N
Microbiology | LQN | Latex Agglutination Assay, Rubella | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | MJH | Legionella, Spp., Elisa | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | PIT | Leishmania Spp. Antigen Detection Assay | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | GMB | Light, Wood'S, Fluorescence | 1 | 866.2600 | Wood's fluorescent lamp. | N | N
Microbiology | QBH | Lower Respiratory Microbial Nucleic Acid Detection System | 2 | 866.3985 | Device to detect and identify microorganisms and associated resistance marker nucleic acids directly in respiratory specimens. | N | N
Microbiology | JWY | Manual Antimicrobial Susceptibility Test Systems | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | KZC | Manual Colony Counter | 1 | 866.2180 | Manual colony counter. | N | N
Microbiology | QBN | Mass Spectrometry, Maldi Tof, Microorganism Identification, Cultured Isolates | 2 | 866.3378 |  | N | N
Microbiology | JRZ | Media, Culture, Amino Acid Assay | 1 | 866.2350 | Microbiological assay culture medium. | N | N
Microbiology | PLO | Meningitis/Encephalitis Pathogen Multiplex Nucleic Acid Detection System | 2 | 866.3970 | Device to detect and identify microbial pathogen nucleic acids in cerebrospinal fluid. | N | N
Microbiology | OUS | Methicillin Resistant Staphylococcus Aureus/Methicillin Susceptible Staphylococcus Aureus Blood Culture Test Bt | 1 | 866.2050 | Staphylococcal typing bacteriophage. | N | N
Microbiology | PPU | Microbial Colony Image Assessment System | 2 | 866.2190 | Automated image assessment system for microbial colonies on solid culture media. | N | N
Microbiology | QBD | Microbial Nucleic Acid Storage And Stabilization Device | 2 | 866.2950 |  | N | N
Microbiology | JTA | Monitor, Microbial Growth | 1 | 866.2560 | Microbial growth monitor. | N | N
Microbiology | QFS | Monocyte Distribution Width Quantitative Measurement For Sepsis Risk Assessment | 2 | 866.3215 | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | N | N
Microbiology | OMI | Multiplex Flow Immunoassay, T. Gondii, Rubella And Cmv | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | PUQ | Multiplex Flow Immunoassay, T. Gondii, Rubella, Cmv Igm | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | OPL | Multiplex Immunoassay For Measles Virus, Mumps Virus, Rubella And Varicella Zoster Virus | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | OPM | Multiplex Immunoassay For T. Gondii, Rubella, Cytomegalovirus And Herpes Simplex Virus 1 And 2 | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | OZX | Mycoplasma Pneumoniae Dna Assay System | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | LSC | Neisseria Controls | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | OTG | Non-Sars Coronavirus Multiplex Nucleic Acid Assay | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | PBK | Non-Variola Orthopoxvirus Real-Time Pcr Primer And Probe Set | 2 | 866.3315 |  | N | N
Microbiology | OUC | Norovirus Serological Reagent | 2 | 866.3395 | Norovirus serological reagents. | N | N
Microbiology | OMS | Novel Influenza A Virus, A/H5 Ns1 Protein | 2 | 866.3332 | Reagents for detection of specific novel influenza A viruses. | N | N
Microbiology | NJR | Nucleic Acid Amplification Assay System, Group B Streptococcus, Direct Specimen Test | 1 | 866.3740 | Streptococcus spp. serological reagents. | N | N
Microbiology | NXD | Nucleic Acid Amplification, Novel Influenza A Virus, A/H5 (Asian Lineage) Rna | 2 | 866.3332 | Reagents for detection of specific novel influenza A viruses. | N | N
Microbiology | QEP | Nucleic Acid Detection System For Non-Viral Microorganism(S) Causing Sexually Transmitted Infections | 2 | 866.3393 |  | N | N
Microbiology | OUZ | Nucleic Amplification Assays For The Detection Of Leishmania Nucleic Acids | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | QHO | Nucleic-Acid Tests Used For Infectious Disease Screening Of Blood And Plasma Donations | Unclassified | 610.40 | Test requirements. | N | N
Microbiology | LTT | Panels, Test, Susceptibility, Antimicrobial | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | OOU | Parainfluenza Multiplex Nucleic Acid Assay | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | GTL | Phages, Staphylococcal Typing, All Types | 1 | 866.2050 | Staphylococcal typing bacteriophage. | N | N
Microbiology | PTJ | Phenotypic Test Kit, Non-Susceptible/Elevated Mic Organisms, Cultured Isolates | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | JTL | Plasma, Coagulase, Human, Horse And Rabbit | 1 | 866.2160 | Coagulase plasma. | N | N
Microbiology | OAX | Plasmodium Spp. Detection Reagents | 2 | 866.3402 | Plasmodium species antigen detection assays. | N | N
Microbiology | LYF | Pneumocystis Carinii | 2 | 866.3780 | Toxoplasma gondii serological reagents. | N | N
Microbiology | PRH | Positive Blood Culture Identification And Ast Kit | 2 | 866.1650 |  | N | N
Microbiology | PRI | Procalcitonin Assay | 2 | 866.3215 | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | N | N
Microbiology | QDZ | Qualitative Cytomegalovirus Nucleic Acid-Based Detection Device For Congenital Cytomegalovirus Infection | 2 | 866.3181 |  | N | N
Microbiology | NZB | Quality Control Material, Genetics, Dna | 2 | 866.5910 | Quality control material for cystic fibrosis nucleic acid assays. | N | N
Microbiology | QIG | Quality Control Material, Genetics, Rna | 2 | 866.5910 | Quality control material for cystic fibrosis nucleic acid assays. | N | N
Microbiology | LJG | Quality Control Slides | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | KZK | Reader, Zone, Automated | 1 | 866.2850 | Automated zone reader. | N | N
Microbiology | LSR | Reagent, Borrelia Serological Reagent | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | LOO | Reagent, Leishmanii Serological | 1 | 866.3870 | Trypanosoma spp. serological reagents. | N | N
Microbiology | LSQ | Reagent, Rickettsia Serological | 1 | 866.3500 | Rickettsia serological reagents. | N | N
Microbiology | LIE | Reagent/Device, Inoculum Calibration | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | PIQ | Reagents For Detection Of Norovirus Nucleic Acid | 2 | 866.3990 | Gastrointestinal microorganism multiplex nucleic acid-based assay. | N | N
Microbiology | LHL | Reagents, Antibody, Legionella, Direct & Indirect Fluorescent | 2 | 866.3300 | Haemophilus spp. serological reagents. | N | N
Microbiology | LLH | Reagents, Clostridium Difficile Toxin | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | MDJ | Reagents, Cysticercosis | 1 | 866.3200 | Echinococcus spp. serological reagents. | N | N
Microbiology | QDP | Respiratory Panel | 2 | 866.3985 | Device to detect and identify microorganisms and associated resistance marker nucleic acids directly in respiratory specimens. | N | N
Microbiology | MCE | Respiratory Syncytial Virus - Elisa | 1 | 866.3480 | Respiratory syncytial virus serological reagents. | N | N
Microbiology | LKT | Respiratory Syncytial Virus, Antigen, Antibody, Ifa | 1 | 866.3480 | Respiratory syncytial virus serological reagents. | N | N
Microbiology | OCC | Respiratory Virus Panel Nucleic Acid Assay System | 2 | 866.3980 | Respiratory viral panel multiplex nucleic acid assay. | N | N
Microbiology | PVQ | Rickettsia Spp. Nucleic Acid Based Detection Assay | 2 | 866.3316 |  | N | N
Microbiology | PRE | Rt-Qpcr Assay For Mrna Transcript Immune Biomarkers | 2 | 866.3215 | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | N | N
Microbiology | LSD | Rubella, Other Assays | 2 | 866.3510 | Rubella virus serological reagents. | N | N
Microbiology | GMR | Sera, Reactive And Non-Specific Control, Fta-Abs Test | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | QHM | Serology Tests Used For Infectious Disease Screening Of Blood And Plasma Donations | Unclassified | 610.40 | Test requirements. | N | N
Microbiology | OID | Sex Crimes / Sexual Assault / Suspect Evidence Collection Kit (Excludes Hiv Testing) | 1 | 866.2900 | Microbiological specimen collection and transport device. | N | N
Microbiology | LOP | Solution, Antimicrobial | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | LHJ | Staphylococcus Aureus Protein A Insoluble | 1 | 866.3700 | Staphylococcus aureus serological reagents. | N | N
Microbiology | LHT | Staphylococcus Aureus Somatic Antigens | 1 | 866.3700 | Staphylococcus aureus serological reagents. | N | N
Microbiology | KFI | Strip, Virulence, Corynebacterium Diphtheriae | 1 | 866.3140 | Corynebacterium spp. serological reagents. | N | N
Microbiology | JSK | Supplement, Culture Media | 1 | 866.2450 | Supplement for culture media. | N | N
Microbiology | LTW | Susceptibility Test Cards, Antimicrobial | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | JTN | Susceptibility Test Discs, Antimicrobial | 2 | 866.1620 | Antimicrobial susceptibility test disc. | N | N
Microbiology | NGZ | Susceptibility Test Plate, Antifungal | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | JTT | Susceptibility Test Powders, Antimicrobial | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | MJA | Susceptibility Test Powders, Antimycobacterial | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | MDB | System, Blood Culturing | 1 | 866.2560 | Microbial growth monitor. | N | N
Microbiology | MZC | System, Detection, Bacterial, For Platelet Transfusion Products | 1 | 866.2560 | Microbial growth monitor. | N | N
Microbiology | PEX | System, Mass Spectrometry, Maldi Tof, Microorganism Identification, Cultured Isolates | 2 | 866.3361 | Mass spectrometer system for clinical use for the identification of microorganisms. | N | N
Microbiology | QBQ | System, Microbial Growth Monitor Of Normally Sterile Body Fluid Culture | 1 | 866.2560 | Microbial growth monitor. | N | N
Microbiology | NJO | System, Mycolic Acid Analysis, Mycobacterium Tuberculosis | 1 | 866.3370 | Mycobacterium tuberculosis immunofluorescent reagents. | N | N
Microbiology | POC | System, Nucleic Acid Amplification Test, Dna, Antimicrobial Resistance Marker, Direct Specimen | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | PMY | System, Nucleic Acid Amplification Test, Dna, Carbapenem Non-Susceptible Gram Negative Organism, Colony | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | NQX | System, Nucleic Acid Amplification Test, Dna, Methicillin Resistant Staphylococcus Aureus, Direct Specimen | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | MWA | System, Nucleic Acid Amplification, Mycobacterium Tuberculosis Complex | 2 | 866.3372 | Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis complex in respiratory specimens. | N | N
Microbiology | PEU | System, Nucleic Acid-Based, Mycobacterium Tuberculosis Complex, Resistance Marker, Direct Specimen | 2 | 866.3373 | Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis complex (MTB-complex) and the genetic mutations associated with MTB-complex antibiotic resistance in respiratory specimens. | N | N
Microbiology | LON | System, Test, Automated, Antimicrobial Susceptibility, Short Incubation | 2 | 866.1645 | Fully automated short-term incubation cycle antimicrobial susceptibility system. | N | N
Microbiology | NIJ | System, Test, Genotypic Detection, Resistant Markers, Enterococcus Species | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | MYI | System, Test, Genotypic Detection, Resistant Markers, Staphylococcus Colonies | 2 | 866.1640 | Antimicrobial susceptibility test powder. | N | N
Microbiology | JTW | System, Transport, Aerobic | 1 | 866.2900 | Microbiological specimen collection and transport device. | N | N
Microbiology | LTS | Test Reagents, Neisseria Gonorrhoeae Biochemical | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | LLM | Test, Antigen, Nuclear, Epstein-Barr Virus | 1 | 866.3235 | Epstein-Barr virus serological reagents. | N | N
Microbiology | MYR | Test, Donor, Syphilis, Antigens, Treponemal | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | GMW | Test, Sorbent, Fta-Abs | 2 | 866.3830 | Treponema pallidum treponemal test reagents. | N | N
Microbiology | MSQ | Test, Urea (Breath Or Blood) | 1 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | OZA | Test, Urea Adult And Pediatric (Breath), | 3 | 866.3110 | Campylobacter fetus serological reagents. | N | N
Microbiology | MXB | Test, Vaginal, Bacterial Sialidase | 1 | 866.2660 | Microorganism differentiation and identification device. | N | N
Microbiology | JTX | Transport Systems, Anaerobic | 1 | 866.2900 | Microbiological specimen collection and transport device. | N | N
Microbiology | PQA | Vaginitis And Bacterial Vaginosis Nucleic Acid Detection System | 2 | 866.3975 |  | N | N
Microbiology | PRA | Variola Virus Nucleic Acid-Based Detection Assay | 2 | 866.3316 |  | N | N
Microbiology | QFO | Zika Virus Serological Reagents | 2 | 866.3935 |  | N | N
Molecular Genetics | OYX | Bcr/Abl1 Monitoring Test | 2 | 866.6060 | BCR-ABL quantitation test. | N | N
Molecular Genetics | QAZ | Cancer Predisposition Risk Assessment System | 2 | 866.6090 |  | N | N
Molecular Genetics | QDJ | Direct-To-Consumer Access Pharmacogenetic Assessment System | 2 | 862.3364 |  | N | N
Molecular Genetics | QDI | Fish Based Detection Of Chromosomal Abnormalities From Patients With Hematologic Malignancies | 2 | 864.1880 |  | N | N
Molecular Genetics | PSU | Jak2 Gene Mutation Detection Test | 2 | 866.6070 |  | N | N
Neurology | MYU | Accessory, Barium Sulfate, Methyl Methacrylate For Cranioplasty | 2 | 882.5300 | Methyl methacrylate for cranioplasty. | N | Y
Neurology | GWL | Amplifier, Physiological Signal | 2 | 882.1835 | Physiological signal amplifier. | N | N
Neurology | OMA | Amplitude-Integrated Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GZM | Analyzer, Rigidity | 2 | 882.1020 | Rigidity analyzer. | N | N
Neurology | GWS | Analyzer, Spectrum, Electroencephalogram Signal | 1 | 882.1420 | Electroencephalogram (EEG) signal spectrum analyzer. | N | N
Neurology | GXM | Anvil, Skull Plate | 1 | 882.4030 | Skull plate anvil. | N | N
Neurology | HCI | Applier, Aneurysm Clip | 2 | 882.4175 | Aneurysm clip applier. | N | N
Neurology | GWW | Ataxiagraph | 1 | 882.1030 | Ataxiagraph. | N | N
Neurology | PTB | Ataxiagraph With Interpretive Software | 1 | 882.1030 | Ataxiagraph. | N | N
Neurology | OMB | Automatic Event Detection Software For Full-Montage Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | OLZ | Automatic Event Detection Software For Polysomnograph With Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | MZQ | Balloon, Detachable, For Neurovascular Occlusion | 2 | 882.5950 | Neurovascular embolization device. | Y | Y
Neurology | JXL | Block, Bite | 2 | 882.5070 | Bite block. | N | N
Neurology | GXX | Block, Test, Ultrasonic Scanner Calibration | 1 | 882.1925 | Ultrasonic scanner calibration test block. | N | N
Neurology | PIW | Brain Injury Adjunctive Interpretive Electroencephalograph Assessment Aid | 2 | 882.1450 | Brain injury adjunctive interpretive electroencephalograph assessment aid. | N | N
Neurology | QEA | Brain Injury Adjunctive Interpretive Oculomotor Assessment Aid | 2 | 882.1455 |  | N | N
Neurology | ORT | Burst Suppression Detection Software For Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | HCD | Cannula, Ventricular | 1 | 882.4060 | Ventricular cannula. | N | N
Neurology | HBZ | Catheter, Intravascular Occluding | 3 | 882.5150 | Intravascular occluding catheter. | N | N
Neurology | NUF | Catheter, Neurovasculature, Occluding Balloon | 3 | 882.5150 | Intravascular occluding catheter. | Y | N
Neurology | PAV | Catheter, Neuro-Vasculature, Occluding Balloon | HDE | 882.5150 | Intravascular occluding catheter. | N | N
Neurology | HCA | Catheter, Ventricular | 2 | 882.4100 | Ventricular catheter. | N | N
Neurology | NHC | Catheter, Ventricular (Containing Antibiotic Or Antimicrobial Agents) | 2 | 882.4100 | Ventricular catheter. | N | Y
Neurology | HBN | Chair, Neurosurgical | 1 | 882.4125 | Neurosurgical chair. | N | N
Neurology | HCE | Clamp, Carotid Artery | 2 | 882.5175 | Carotid artery clamp. | N | Y
Neurology | HCH | Clip, Aneurysm | 2 | 882.5200 | Aneurysm clip. | N | Y
Neurology | HBP | Clip, Implanted Malleable | 2 | 882.5225 | Implanted malleable clip. | N | Y
Neurology | HBO | Clip, Scalp | 2 | 882.4150 | Scalp clip. | N | N
Neurology | ONA | Computer Assisted Hair Harvesting System | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | PWE | Computerized Behavioral Therapy Device For Psychiatric Disorders | 2 | 882.5801 | Computerized behavioral therapy device for psychiatric disorders. | N | N
Neurology | PKQ | Computerized Cognitive Assessment Aid | 2 | 882.1470 | Computerized cognitive assessment aid. | N | N
Neurology | POM | Computerized Cognitive Assessment Aid For Concussion | 2 | 882.1471 | Computerized cognitive assessment aid for concussion. | N | N
Neurology | PTY | Computerized Cognitive Assessment Aid, Exempt | 2 | 882.1470 | Computerized cognitive assessment aid. | N | N
Neurology | GWK | Conditioner, Signal, Physiological | 2 | 882.1845 | Physiological signal conditioner. | N | N
Neurology | GXR | Cover, Burr Hole | 2 | 882.5250 | Burr hole cover. | N | Y
Neurology | PBJ | Cranial Distraction System | 2 | 882.5330 | Preformed nonalterable cranioplasty plate. | Y | Y
Neurology | JXK | Cranial Electrotherapy Stimulator To Treat Depression | 3 | 882.5800 | Cranial electrotherapy stimulator. | N | N
Neurology | QJQ | Cranial Electrotherapy Stimulator To Treat Insomnia And/Or Anxiety | 2 | 882.5800 | Cranial electrotherapy stimulator. | N | N
Neurology | POP | Cranial Motion Measurement Device | 2 | 882.1630 | Cranial motion measurement device. | N | N
Neurology | PPT | Cranial Surgical Planning And Instrument Guides | 2 | 882.4310 | Powered simple cranial drills, burrs, trephines, and their accessories. | N | N
Neurology | JXI | Cuff, Nerve | 2 | 882.5275 | Nerve cuff. | N | Y
Neurology | HCB | Device, Aversive Conditioning | 2 | 882.5235 | Aversive conditioning device. | N | N
Neurology | HCC | Device, Biofeedback | 2 | 882.5050 | Biofeedback device. | N | N
Neurology | GXC | Device, Electroconvulsive Therapy | 3 | 882.5940 | Electroconvulsive therapy device. | N | N
Neurology | GZO | Device, Galvanic Skin Response Measurement | 2 | 882.1540 | Galvanic skin response measurement device. | N | N
Neurology | GWM | Device, Monitoring, Intracranial Pressure | 2 | 882.1620 | Intracranial pressure monitoring device. | N | N
Neurology | JXE | Device, Nerve Conduction Velocity Measurement | 2 | 882.1550 | Nerve conduction velocity measurement device. | N | N
Neurology | HCG | Device, Neurovascular Embolization | 2 | 882.5950 | Neurovascular embolization device. | N | Y
Neurology | HCJ | Device, Skin Potential Measurement | 2 | 882.1560 | Skin potential measurement device. | N | N
Neurology | LEL | Device, Sleep Assessment | 2 | 882.5050 | Biofeedback device. | N | N
Neurology | GXH | Device, Surgical, Cryogenic | 2 | 882.4250 | Cryogenic surgical device. | N | N
Neurology | HCS | Device, Temperature Measurement, Direct Contact, Powered | 2 | 882.1570 | Powered direct-contact temperature measurement device. | N | N
Neurology | LLN | Device, Vibration Threshold Measurement | 1 | 882.1200 | Two-point discriminator. | N | N
Neurology | QFX | Diagnostic Neurosurgical Microscope Filter | 2 | 882.4950 |  | N | N
Neurology | GWI | Discriminator, Two-Point | 1 | 882.1200 | Two-point discriminator. | N | N
Neurology | QGT | Distal Transcutaneous Electrical Stimulator For Treatment Of Acute Migraine | 2 | 882.5899 |  | N | N
Neurology | HBF | Drills, Burrs, Trephines & Accessories (Compound, Powered) | 2 | 882.4305 | Powered compound cranial drills, burrs, trephines, and their accessories. | N | N
Neurology | HBG | Drills, Burrs, Trephines & Accessories (Manual) | 2 | 882.4300 | Manual cranial drills, burrs, trephines, and their accessories | N | N
Neurology | HBE | Drills, Burrs, Trephines & Accessories (Simple, Powered) | 2 | 882.4310 | Powered simple cranial drills, burrs, trephines, and their accessories. | N | N
Neurology | NLO | Drills, Burrs, Trephines And Accessories (Manual), Reprocessed | 2 | 882.4300 | Manual cranial drills, burrs, trephines, and their accessories | N | N
Neurology | NLN | Drills, Burrs, Trephines And Accessories (Simple Powered), Reprocessed | 2 | 882.4310 | Powered simple cranial drills, burrs, trephines, and their accessories. | N | N
Neurology | NLP | Drills, Burrs, Trephines, And Accessories (Powered Compound), Reprocessed | 2 | 882.4305 | Powered compound cranial drills, burrs, trephines, and their accessories. | N | N
Neurology | GXQ | Dura Substitute | 2 | 882.5910 | Dura substitute. | N | Y
Neurology | PGW | Ear, Nose, And Throat Stereotaxic Instrument | 2 | 882.4560 | Stereotaxic instrument. | N | N | Acclarent
Neurology | OBV | Eating Disorder Conditioning Tool | 2 | 882.5060 |  | N | N
Neurology | GXW | Echoencephalograph | 2 | 882.1240 | Echoencephalograph. | N | N
Neurology | QGH | Electroconvulsive Therapy Device For Catatonia, Major Depressive Disorder, And Bipolar Disorder | 2 | 882.5940 | Electroconvulsive therapy device. | N | N
Neurology | GYC | Electrode, Cortical | 2 | 882.1310 | Cortical electrode. | N | N
Neurology | GXY | Electrode, Cutaneous | 2 | 882.1320 | Cutaneous electrode. | N | N
Neurology | GZL | Electrode, Depth | 2 | 882.1330 | Depth electrode. | N | N
Neurology | GZK | Electrode, Nasopharyngeal | 2 | 882.1340 | Nasopharyngeal electrode. | N | N
Neurology | GXZ | Electrode, Needle | 2 | 882.1350 | Needle electrode. | N | N
Neurology | LHG | Electrode, Spinal Epidural | 2 | 882.5880 | Implanted spinal cord stimulator for pain relief. | N | Y
Neurology | QFF | Electromechanical Arm For A Transcranial Magnetic Stimulation System | 2 | 882.5805 | Repetitive transcranial magnetic stimulation system. | N | N
Neurology | GYE | Encephalogram Telemetry System | 2 | 882.1855 | Electroencephalogram (EEG) telemetry system. | N | N
Neurology | GWG | Endoscope, Neurological | 2 | 882.1480 | Neurological endoscope. | N | N
Neurology | PVL | Endoscopic Shunt Placement Kit | 2 | 882.5550 | Central nervous system fluid shunt and components. | N | N
Neurology | PVM | Endoscopic Ventricular Catheter Placement Kit | 2 | 882.4100 | Ventricular catheter. | N | N
Neurology | GXB | Esthesiometer | 1 | 882.1500 | Esthesiometer. | N | N
Neurology | PNA | Evoked Photon Image Capture Device | 1 | 882.1561 | Evoked photon image capture device. | N | N
Neurology | NTU | Evoked Potential Stimulator, Thermal | 2 | 882.1870 | Evoked response electrical stimulator. | N | N
Neurology | PCB | External Cerebrospinal Fluid (Csf) Diversion | 2 | 882.5560 |  | N | N
Neurology | QBC | External Upper Limb Tremor Stimulator | 2 | 882.5897 | External upper limb tremor stimulator. | N | N
Neurology | HBW | Fastener, Plate, Cranioplasty | 2 | 882.5360 | Cranioplasty plate fastener. | N | Y
Neurology | PJM | Filler, Bone Void, Alterable Compound For Cranioplasty | 2 | 882.5320 | Preformed alterable cranioplasty plate. | N | Y
Neurology | PJN | Filler, Bone Void, Non-Alterable Compound For Cranioplasty | 2 | 882.5330 | Preformed nonalterable cranioplasty plate. | N | Y
Neurology | GWX | Fork, Tuning | 1 | 882.1525 | Tuning fork. | N | N
Neurology | GWQ | Full-Montage Standard Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GWR | Generator, Electroencephalograph Test Signal | 1 | 882.1430 | Electroencephalograph test signal generator. | N | N
Neurology | GXD | Generator, Lesion, Radiofrequency | 2 | 882.4400 | Radiofrequency lesion generator. | N | N
Neurology | HBD | Handpiece (Brace), Drill | 1 | 882.4325 | Cranial drill handpiece (brace). | N | N
Neurology | HBM | Headrest, Neurosurgical | 1 | 882.4440 | Neurosurgical headrests. | N | N
Neurology | HBL | Holder, Head, Neurosurgical (Skull Clamp) | 2 | 882.4460 | Neurosurgical head holder (skull clamp). | N | N
Neurology | LEM | Human Lyophilized Dura Mater | 2 | 882.5975 | Human dura mater. | N | Y
Neurology | GZA | Implanted Cerebellar Stimulator | 3 | 882.5820 | Implanted cerebellar stimulator. | N | Y
Neurology | GZE | Implanted Diaphragmatic/Phrenic Nerve Stimulator | 3 | 882.5830 | Implanted diaphragmatic/phrenic nerve stimulator. | N | Y
Neurology | OLW | Index-Generating Electroencephalograph Software | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | HBS | Instrument, Clip Forming/Cutting | 1 | 882.4190 | Clip forming/cutting instrument. | N | N
Neurology | HBQ | Instrument, Clip Removal | 1 | 882.4200 | Clip removal instrument. | N | N
Neurology | NMN | Instrument, Clip, Forming/Cutting, Reprocessed | 1 | 882.4190 | Clip forming/cutting instrument. | N | N
Neurology | HBX | Instrument, Cranioplasty Material Forming | 1 | 882.4500 | Cranioplasty material forming instrument. | N | N
Neurology | GZQ | Instrument, Dowel Cutting | 2 | 882.4275 | Dowel cutting instrument. | N | N
Neurology | GZX | Instrument, Microsurgical | 1 | 882.4525 | Microsurgical instrument. | N | N
Neurology | GYK | Instrument, Shunt System Implantation | 1 | 882.4545 | Shunt system implantation instrument. | N | N
Neurology | HAO | Instrument, Surgical, Non-Powered | 1 | 882.4535 | Nonpowered neurosurgical instrument. | N | N
Neurology | LIH | Interferential Current Therapy | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | PSP | Intracranial Pressure & Temperature Monitoring Kit | 2 | 882.1620 | Intracranial pressure monitoring device. | N | N
Neurology | ONN | Intraoperative Orthopedic Joint Assessment Aid | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | GXE | Leukotome | 1 | 882.4600 | Leukotome. | N | N
Neurology | PVO | Lumbar Drainage Catheter Kit | 2 | 882.5550 | Central nervous system fluid shunt and components. | N | N
Neurology | OLY | Magnetoencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GYB | Media, Electroconductive | 2 | 882.1275 | Electroconductive media. | N | N
Neurology | JXH | Methyl Methacrylate For Aneurysmorrhaphy | 2 | 882.5030 | Methyl methacrylate for aneurysmorrhaphy. | N | Y
Neurology | GXP | Methyl Methacrylate For Cranioplasty | 2 | 882.5300 | Methyl methacrylate for cranioplasty. | N | Y
Neurology | GXS | Monitor, Alpha | 2 | 882.1610 | Alpha monitor. | N | N
Neurology | GXT | Monitor, Lesion Temperature | 2 | 882.5500 | Lesion temperature monitor. | N | N
Neurology | HBC | Motor, Drill, Electric | 2 | 882.4360 | Electric cranial drill motor. | N | N
Neurology | HBB | Motor, Drill, Pneumatic | 2 | 882.4370 | Pneumatic cranial drill motor. | N | N
Neurology | HAS | Needle, Neurosurgical Suture | 1 | 882.4650 | Neurosurgical suture needle. | N | N
Neurology | HAW | Neurological Stereotaxic Instrument | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | ORR | Neurological Stereotaxic Instrument, Real-Time Intraoperative Mri | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | PVP | Neurological Test Kit | 1 | 882.1750 | Pinwheel. | N | N
Neurology | OJG | Neurological Tray | 1 | 882.4525 | Microsurgical instrument. | N | N
Neurology | NCG | Neuropsychiatric Interpretative Electroencephalograph Assessment Aid | 2 | 882.1440 | Neuropsychiatric interpretive electroencephalograph assessment aid. | N | N
Neurology | HBA | Neurosurgical Paddie | 2 | 882.4700 | Neurosurgical paddie. | N | N
Neurology | POL | Neurovascular Mechanical Thrombectomy Device For Acute Ischemic Stroke Treatment | 2 | 882.5600 | Neurovascular mechanical thrombectomy device for acute ischemic stroke treatment. | N | N
Neurology | PKR | Non-Invasive Vagus Nerve Stimulator - Headache | 2 | 882.5892 | External vagal nerve stimulator for headache. | N | N
Neurology | QAK | Non-Invasive Vagus Nerve Stimulator For Migraine Headache | 2 | 882.5892 | External vagal nerve stimulator for headache. | N | N
Neurology | OLT | Non-Normalizing Quantitative Electroencephalograph Software | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | OLU | Normalizing Quantitative Electroencephalograph Software | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GWN | Nystagmograph | 2 | 882.1460 | Nystagmograph. | N | N
Neurology | OJP | Orthopedic Computer Controlled Surgical System | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | OLO | Orthopedic Stereotaxic Instrument | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | MVA | Orthosis, Cranial | 2 | 882.5970 | Cranial orthosis. | N | N
Neurology | OAN | Orthosis, Cranial, Laser Scan | 2 | 882.5970 | Cranial orthosis. | N | N
Neurology | OSE | Patient Specific Manual Navigation System | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | OSF | Patient Specific Manual Orthopedic Stereotaxic System | 2 | 882.4560 | Stereotaxic instrument. | N | N
Neurology | GWZ | Percussor | 1 | 882.1700 | Percussor. | N | N
Neurology | PZR | Percutaneous Nerve Stimulator For Opioid Withdrawal | 2 | 882.5896 | Percutaneous nerve stimulator for substance use disorders. | N | N
Neurology | POS | Physiological Signal Based Seizure Monitoring System | 2 | 882.1580 | Non-electroencephalogram (EEG) physiological signal based seizure monitoring system. | N | N
Neurology | OSG | Piezo-Electric Stimulator For Relief Of Mosquito Bite Itch | 2 | 882.5894 | Limited output transcutaneous piezoelectric stimulator for skin reactions associated with insect bites. | N | N
Neurology | GWY | Pinwheel | 1 | 882.1750 | Pinwheel. | N | N
Neurology | GWO | Plate, Cranioplasty, Preformed, Alterable | 2 | 882.5320 | Preformed alterable cranioplasty plate. | N | Y
Neurology | GXN | Plate, Cranioplasty, Preformed, Non-Alterable | 2 | 882.5330 | Preformed nonalterable cranioplasty plate. | N | Y
Neurology | JXF | Plethysmograph, Ocular | 3 | 882.1790 | Ocular plethysmograph. | N | N
Neurology | LKG | Port & Catheter, Implanted, Subcutaneous, Intraventricular | 2 | 882.5550 | Central nervous system fluid shunt and components. | N | Y
Neurology | GXI | Probe, Radiofrequency Lesion | 2 | 882.4725 | Radiofrequency lesion probe. | N | N
Neurology | GXJ | Punch, Skull | 1 | 882.4750 | Skull punch. | N | N
Neurology | HBR | Rack, Clip | 1 | 882.4215 | Clip rack. | N | N
Neurology | OMC | Reduced- Montage Standard Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GZT | Retractor, Self-Retaining, For Neurosurgery | 2 | 882.4800 | Self-retaining retractor for neurosurgery. | N | N
Neurology | GZN | Rheoencephalograph | 3 | 882.1825 | Rheoencephalograph. | N | N
Neurology | HAE | Rongeur, Manual | 2 | 882.4840 | Manual rongeur. | N | N
Neurology | HAD | Rongeur, Powered | 2 | 882.4845 | Powered rongeur. | N | N
Neurology | GXL | Screwdriver, Skullplate | 1 | 882.4900 | Skullplate screwdriver. | N | N
Neurology | JXG | Shunt, Central Nervous System And Components | 2 | 882.5550 | Central nervous system fluid shunt and components. | N | Y
Neurology | OLX | Source Localization Software For Electroencephalograph Or Magnetoencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | PIY | Stainless Steel Instrument, Shunt System Implantation | 1 | 882.4545 | Shunt system implantation instrument. | N | N
Neurology | OLV | Standard Polysomnograph With Electroencephalograph | 2 | 882.1400 | Electroencephalograph. | N | N
Neurology | GWJ | Stimulator, Auditory, Evoked Response | 2 | 882.1900 | Evoked response auditory stimulator. | N | N
Neurology | GWF | Stimulator, Electrical, Evoked Response | 2 | 882.1870 | Evoked response electrical stimulator. | N | N
Neurology | NYN | Stimulator, Electrical, Transcutaneous, For Arthritis | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | NYX | Stimulator, Electrical, Transcutaneous, With Limited Output, For Aesthetic Purposes | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | GYZ | Stimulator, Intracerebral/Subcortical, Implanted | 3 | 882.5840 | Implanted intracerebral/subcortical stimulator for pain relief. | N | Y
Neurology | GZP | Stimulator, Mechanical, Evoked Response | 2 | 882.1880 | Evoked response mechanical stimulator. | N | N
Neurology | NHI | Stimulator, Nerve, Electrical, Percutaneous (Pens), For Pain Relief | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | PCC | Stimulator, Nerve, Electrical, Transcutaneous, For Migraine | 2 | 882.5891 | Transcutaneous electrical nerve stimulator to treat headache. | N | N
Neurology | OCF | Stimulator, Nerve, Electrical, Transcutaneous, Limited Output, Arthritis Pain Relief | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | NYW | Stimulator, Nerve, Electrical, Transcutaneous, With Limited Output, For Pain Relief | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | GZJ | Stimulator, Nerve, Transcutaneous, For Pain Relief | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | NUH | Stimulator, Nerve, Transcutaneous, Over-The-Counter | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | GZI | Stimulator, Neuromuscular, External Functional | 2 | 882.5810 | External functional neuromuscular stimulator. | N | N
Neurology | GZC | Stimulator, Neuromuscular, Implanted | 3 | 882.5860 | Implanted neuromuscular stimulator. | N | Y
Neurology | GZF | Stimulator, Peripheral Nerve, Implanted (Pain Relief) | 2 | 882.5870 | Implanted peripheral nerve stimulator for pain relief. | N | Y
Neurology | GWE | Stimulator, Photic, Evoked Response | 2 | 882.1890 | Evoked response photic stimulator. | N | N
Neurology | GZB | Stimulator, Spinal-Cord, Implanted (Pain Relief) | 2 | 882.5880 | Implanted spinal cord stimulator for pain relief. | N | Y
Neurology | GZD | Stimulator, Spinal-Cord, Implanted, For Bladder Evacuation | 3 | 882.5850 | Implanted spinal cord stimulator for bladder evacuation. | N | Y
Neurology | NFO | Stimulator, Transcutaneous Electrical, Aesthetic Purposes | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | GXO | Strip, Craniosynostosis, Preformed | 2 | 882.5900 | Preformed craniosynostosis strip. | N | Y
Neurology | PUU | Temporary Coil Embolization Assist Device | 2 | 882.5955 |  | N | N
Neurology | QFD | Tens For Sinus Pain | 2 | 882.5890 | Transcutaneous electrical nerve stimulator for pain relief. | N | N
Neurology | LQW | Test, Temperature Discrimination | 1 | 882.1200 | Two-point discriminator. | N | N
Neurology | GYA | Tester, Electrode/Lead, Electroencephalograph | 1 | 882.1410 | Electroencephalograph electrode/lead tester. | N | N
Neurology | PLU | Thermal System For Insomnia | 2 | 882.5700 | Thermal system for insomnia. | N | N
Neurology | QAR | Thermal Vestibular Stimulator For Headache | 2 | 882.5893 | Thermal vestibular stimulator for headache. | N | N
Neurology | HAX | Tong, Skull For Traction | 2 | 882.5960 | Skull tongs for traction. | N | N
Neurology | OEW | Tracking, Soft Tissue, Intraoperative | 2 | 882.4560 | Stereotaxic instrument. | Y | N
Neurology | QGL | Trancutaneous Nerve Stimulator For Adhd | 2 | 882.5898 |  | N | N
Neurology | QCI | Transcranial Magnetic Stimulation System For Obsessive-Compulsive Disorder | 2 | 882.5802 | Transcranial magnetic stimulation system for neurological and psychiatric disorders and conditions. | N | N
Neurology | OBP | Transcranial Magnetic Stimulator | 2 | 882.5805 | Repetitive transcranial magnetic stimulation system. | N | N
Neurology | OKP | Transcranial Magnetic Stimulator For The Treatment Of Migraine Headache | 2 | 882.5808 | Transcranial magnetic stimulator for headache. | N | N
Neurology | GYD | Transducer, Tremor | 2 | 882.1950 | Tremor transducer. | N | N
Obstetrics/Gynecology | MQG | Accessory, Assisted Reproduction | 2 | 884.6120 | Assisted reproduction accessories. | N | N
Obstetrics/Gynecology | PUB | Accessory, Assisted Reproduction, Exempt | 2 | 884.6120 | Assisted reproduction accessories. | N | N
Obstetrics/Gynecology | HFA | Amnioscope, Transabdominal (Fetoscope) (And Accessories) | HDE | 884.1600 | Transabdominal amnioscope (fetoscope) and accessories. | N | N
Obstetrics/Gynecology | HGE | Amniotome | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HEO | Analyzer, Data, Obstetric | 3 | 884.2050 | Obstetric data analyzer. | N | N
Obstetrics/Gynecology | HGD | Applicator, Vaginal | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HFC | Aspirator, Endocervical | 2 | 884.1050 | Endocervical aspirator. | N | N
Obstetrics/Gynecology | HFF | Aspirator, Endometrial | 2 | 884.1060 | Endometrial aspirator. | N | N
Obstetrics/Gynecology | HFY | Band, Tubal Occlusion | 3 | 884.5380 | Contraceptive tubal occlusion device (TOD) and introducer. | N | Y
Obstetrics/Gynecology | MSC | Barrier, Std, Oral Sex | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | FHG | Bell, Circumcision | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PWA | Breast Pump Accessory, Milk Expression Aid, Via Positive Pressure | 1 | 884.5150 | Nonpowered breast pump. | N | N
Obstetrics/Gynecology | OHH | Breast Pump Kit | 2 | 884.5160 | Powered breast pump. | N | N
Obstetrics/Gynecology | HFE | Brush, Endometrial | 2 | 884.1100 | Endometrial brush. | N | N
Obstetrics/Gynecology | HEN | Caliper, Fetal Head, Ultrasonic | 2 | 884.2225 | Obstetric-gynecologic ultrasonic imager. | N | N
Obstetrics/Gynecology | PGK | Cannula, Injector, Uterine, Endometrial Biopsy | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | MFD | Cannula, Intrauterine Insemination | 2 | 884.5250 | Cervical cap. | N | N
Obstetrics/Gynecology | LKF | Cannula, Manipulator/Injector, Uterine | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HGH | Cannula, Suction, Uterine | 2 | 884.5070 | Vacuum abortion system. | N | N
Obstetrics/Gynecology | HDR | Cap, Cervical | 2 | 884.5250 | Cervical cap. | N | N
Obstetrics/Gynecology | MQF | Catheter, Assisted Reproduction | 2 | 884.6110 | Assisted reproduction catheters. | N | N
Obstetrics/Gynecology | PON | Catheter, Balloon, Dilation Of Cervical Canal | 2 | 884.4260 | Hygroscopic Laminaria cervical dilator. | N | N
Obstetrics/Gynecology | PFJ | Catheter, Balloon, Dilation Of Cervical Canal Prior To Labor | 2 | 884.4260 | Hygroscopic Laminaria cervical dilator. | N | N
Obstetrics/Gynecology | HGS | Catheter, Intrauterine And Introducer | 2 | 884.2700 | Intrauterine pressure monitor and accessories. | N | N
Obstetrics/Gynecology | MOV | Catheters, Salpingography | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | OHM | Cesarean Section Tray | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | HEH | Chamber, Decompression, Abdominal | 3 | 884.5225 | Abdominal decompression chamber. | N | N
Obstetrics/Gynecology | OHG | Circumcision Tray | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | NBZ | Clamp And Cutter, Umbilical | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HFX | Clamp, Circumcision | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HFW | Clamp, Umbilical | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HGC | Clamp, Uterine | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HHN | Cleidoclast | 2 | 884.4500 | Obstetric fetal destructive instrument. | N | N
Obstetrics/Gynecology | HGB | Clip, Tubal Occlusion | 3 | 884.5380 | Contraceptive tubal occlusion device (TOD) and introducer. | N | Y
Obstetrics/Gynecology | HFI | Coagulator, Culdoscopic (And Accessories) | 2 | 884.4160 | Unipolar endoscopic coagulator-cutter and accessories. | N | N
Obstetrics/Gynecology | HFH | Coagulator, Hysteroscopic (And Accessories) | 2 | 884.4160 | Unipolar endoscopic coagulator-cutter and accessories. | N | N
Obstetrics/Gynecology | HFG | Coagulator, Laparoscopic, Unipolar (And Accessories) | 2 | 884.4160 | Unipolar endoscopic coagulator-cutter and accessories. | N | N
Obstetrics/Gynecology | HIN | Coagulator-Cutter, Endoscopic, Bipolar (And Accessories) | 2 | 884.4150 | Bipolar endoscopic coagulator-cutter and accessories. | N | N
Obstetrics/Gynecology | KNF | Coagulator-Cutter, Endoscopic, Unipolar (And Accessories) | 2 | 884.4160 | Unipolar endoscopic coagulator-cutter and accessories. | N | N
Obstetrics/Gynecology | HEX | Colposcope (And Colpomicroscope) | 2 | 884.1630 | Colposcope. | N | N
Obstetrics/Gynecology | PTZ | Colposcope (And Colpomicroscope), Exempt | 2 | 884.1630 | Colposcope. | N | N
Obstetrics/Gynecology | HIS | Condom | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | LTZ | Condom With Nonoxynol-9 | 2 | 884.5310 | Condom with spermicidal lubricant. | N | N
Obstetrics/Gynecology | OBY | Condom, Female, Animal Tissue | 3 | 884.5330 | Multiple-use female condom. | N | N
Obstetrics/Gynecology | MOL | Condom, Synthetic | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | PMU | Containment System, Laparoscopic Power Morcellation, With Instrument Port | 2 | 884.4050 | Gynecologic laparoscopic power morcellation containment system. | N | N
Obstetrics/Gynecology | HGG | Controller, Abortion Unit, Vacuum | 2 | 884.5070 | Vacuum abortion system. | N | N
Obstetrics/Gynecology | HGA | Cranioclast | 2 | 884.4500 | Obstetric fetal destructive instrument. | N | N
Obstetrics/Gynecology | HFZ | Craniotribe | 2 | 884.4500 | Obstetric fetal destructive instrument. | N | N
Obstetrics/Gynecology | OHF | Culdocentesis Kit | 2 | 884.1640 | Culdoscope and accessories. | N | N
Obstetrics/Gynecology | HEW | Culdoscope (And Accessories) | 2 | 884.1640 | Culdoscope and accessories. | N | N
Obstetrics/Gynecology | OYO | Culture, Intravaginal, Assisted Reproduction | 2 | 884.6165 | Intravaginal culture system. | N | N
Obstetrics/Gynecology | HHE | Cup, Menstrual | 2 | 884.5400 | Menstrual cup. | N | N
Obstetrics/Gynecology | HHK | Curette, Suction, Endometrial (And Accessories) | 2 | 884.1175 | Endometrial suction curette and accessories. | N | N
Obstetrics/Gynecology | HCY | Curette, Uterine | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | OKY | Cytology Kit | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PVR | D&C Tray | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | NBV | Device, Engorgement, Clitoral | 2 | 884.5970 | Clitoral engorgement device. | N | N
Obstetrics/Gynecology | PYT | Device, Fertility Diagnostic, Contraceptive, Software Application | 2 | 884.5370 | Software application for contraception. | N | N
Obstetrics/Gynecology | HDT | Device, Intrauterine, Contraceptive And Introducer | 3 | 884.5360 | Contraceptive intrauterine device (IUD) and introducer. | N | Y
Obstetrics/Gynecology | HDW | Diaphragm, Contraceptive (And Accessories) | 2 | 884.5350 | Contraceptive diaphragm and accessories. | N | N
Obstetrics/Gynecology | HDN | Dilator, Cervical, Expandable | 3 | 884.4250 | Expandable cervical dilator. | N | N
Obstetrics/Gynecology | HDQ | Dilator, Cervical, Fixed Size | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HDY | Dilator, Cervical, Hygroscopic-Laminaria | 2 | 884.4260 | Hygroscopic Laminaria cervical dilator. | N | N
Obstetrics/Gynecology | PKN | Dilator, Cervical, Synthetic, Osmotic | 2 | 884.4260 | Hygroscopic Laminaria cervical dilator. | N | N
Obstetrics/Gynecology | HHG | Dilator, Cervical, Vibratory | 3 | 884.4270 | Vibratory cervical dilators. | N | N
Obstetrics/Gynecology | HDX | Dilator, Vaginal | 2 | 884.3900 | Vaginal stent. | N | N
Obstetrics/Gynecology | LXE | Doppler, Fetal, Ultrasound | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HED | Douche Apparatus, Vaginal, Therapeutic | 1 | 884.5900 | Therapeutic vaginal douche apparatus. | N | N
Obstetrics/Gynecology | HFL | Drain, Cervical | 2 | 884.3200 | Cervical drain. | N | N
Obstetrics/Gynecology | HIM | Electrocautery, Endoscopic And Accessories | 2 | 884.4100 | Endoscopic electrocautery and accessories. | N | N
Obstetrics/Gynecology | HGI | Electrocautery, Gynecologic (And Accessories) | 2 | 884.4120 | Gynecologic electrocautery and accessories. | N | N
Obstetrics/Gynecology | HGP | Electrode, Circular (Spiral), Scalp And Applicator | 2 | 884.2675 | Fetal scalp circular (spiral) electrode and applicator. | N | N
Obstetrics/Gynecology | HIQ | Electrode, Clip, Fetal Scalp (And Applicator) | 3 | 884.2685 | Fetal scalp clip electrode and applicator. | N | N
Obstetrics/Gynecology | HDP | Elevator, Uterine | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PBH | Embryo Image Assessment System, Assisted Reproduction | 2 | 884.6195 | Assisted Reproduction Embryo Image Assessment System. | N | N
Obstetrics/Gynecology | OKU | Emergency Obstetrical Kit | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | OKZ | Endometrial Sampling Kit | 2 | 884.1175 | Endometrial suction curette and accessories. | N | N
Obstetrics/Gynecology | HGK | Endoscope, Fetal Blood Sampling (And Accessories) | 2 | 884.1660 | Transcervical endoscope (amnioscope) and accessories. | N | N
Obstetrics/Gynecology | HEZ | Endoscope, Transcervical (Amnioscope)(And Accessories) | 2 | 884.1660 | Transcervical endoscope (amnioscope) and accessories. | N | N
Obstetrics/Gynecology | QAG | Endoscopic Electrosurgical Clip Cutting System | 2 | 876.4310 | Endoscopic electrosurgical clip cutting system. | N | N
Obstetrics/Gynecology | PTE | Endoscopic Morcellator Gastroenterology | 2 | 884.1690 | Hysteroscope and accessories. | N | N
Obstetrics/Gynecology | HDB | Extractor, Vacuum, Fetal | 2 | 884.4340 | Fetal vacuum extractor. | N | N
Obstetrics/Gynecology | MKO | Falloposcope | 2 | 884.1690 | Hysteroscope and accessories. | N | N
Obstetrics/Gynecology | PVS | Fetal Blood Sampling Kit (Excludes Hiv Testing) | 2 | 884.1560 | Fetal blood sampler. | N | N
Obstetrics/Gynecology | PWB | Fetal Head Elevator | 2 | 884.4350 | Fetal head elevator. | N | N
Obstetrics/Gynecology | PBQ | Fixation, Non-Absorbable Or Absorbable, For Pelvic Use | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | Y
Obstetrics/Gynecology | HFB | Forceps, Biopsy, Gynecological | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | NMG | Forceps, Biopsy, Gynecological, Reprocessed | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HDA | Forceps, Obstetrical | 2 | 884.4400 | Obstetric forceps. | N | N
Obstetrics/Gynecology | HCZ | Forceps, Surgical, Gynecological | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | OHE | Forensic Evidence Sexual Assault Kit | 2 | 884.1630 | Colposcope. | N | N
Obstetrics/Gynecology | OHD | Gynecological Laparoscopic Kit | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | KND | Heater, Perineal | 2 | 884.5390 | Perineal heater. | N | N
Obstetrics/Gynecology | HGZ | Heater, Perineal, Direct Contact | 2 | 884.5390 | Perineal heater. | N | N
Obstetrics/Gynecology | HHA | Heater, Perineal, Radiant, Non-Contact | 2 | 884.5390 | Perineal heater. | N | N
Obstetrics/Gynecology | OOA | Hemorrhoid Prevention Pressure Wedge | 2 | 884.5200 | Hemorrhoid prevention pressure wedge. | N | N
Obstetrics/Gynecology | HHX | Holograph, Fetal, Acoustical | 2 | 884.2225 | Obstetric-gynecologic ultrasonic imager. | N | N
Obstetrics/Gynecology | LQK | Home Uterine Activity Monitor | 2 | 884.2730 | Home uterine activity monitor. | N | N
Obstetrics/Gynecology | HDI | Hook, Destructive, Obstetrical | 2 | 884.4500 | Obstetric fetal destructive instrument. | N | N
Obstetrics/Gynecology | HDE | Hook, Fibroid, Gynecological | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | OJF | Hysterectomy Kit | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HIH | Hysteroscope (And Accessories) | 2 | 884.1690 | Hysteroscope and accessories. | N | N
Obstetrics/Gynecology | NWW | Hysteroscope Accessories | 1 | 884.1690 | Hysteroscope and accessories. | N | N
Obstetrics/Gynecology | HEM | Imager, Ultrasonic Obstetric-Gynecologic | 2 | 884.2225 | Obstetric-gynecologic ultrasonic imager. | N | N
Obstetrics/Gynecology | KNB | Instrument, Destructive, Fetal, Obstetric | 2 | 884.4500 | Obstetric fetal destructive instrument. | N | N
Obstetrics/Gynecology | KOH | Instrument, Manual, General Obstetric-Gynecologic | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | KNA | Instrument, Manual, Specialized Obstetric-Gynecologic | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PWJ | Instrumentation, Surgical Mesh, Urogynecologic, Stress Urinary Incontinence | 2 | 884.4910 | Specialized surgical instrumentation for use with urogynecologic surgical mesh. | N | N
Obstetrics/Gynecology | PWK | Instrumentation, Surgical Mesh, Urogynecologic, Transabdominal Repair Of Pelvic Organ Prolapse | 2 | 884.4910 | Specialized surgical instrumentation for use with urogynecologic surgical mesh. | N | N
Obstetrics/Gynecology | PWI | Instrumentation, Surgical Mesh, Urogynecologic, Transvaginal Repair Of Pelvic Organ Prolapse | 2 | 884.4910 | Specialized surgical instrumentation for use with urogynecologic surgical mesh. | N | N
Obstetrics/Gynecology | HES | Insufflator, Carbon-Dioxide, Uterotubal (And Accessories) | 2 | 884.1300 | Uterotubal carbon dioxide insufflator and accessories. | N | N
Obstetrics/Gynecology | OSV | Insufflator, Endoscopic Vessel Harvesting | 2 | 884.1730 | Laparoscopic insufflator. | N | N
Obstetrics/Gynecology | HIG | Insufflator, Hysteroscopic | 2 | 884.1700 | Hysteroscopic insufflator. | N | N
Obstetrics/Gynecology | PGT | Insufflator, Hysteroscopic, Fluid, Closed-Loop Recirculation With Cutter-Coagulator, Endoscopic, Bipolar | 2 | 884.1710 | Closed loop hysteroscopic insufflator with cutter-coagulator. | N | N
Obstetrics/Gynecology | HIF | Insufflator, Laparoscopic | 2 | 884.1730 | Laparoscopic insufflator. | N | N
Obstetrics/Gynecology | HEC | Insufflator, Vaginal | 1 | 884.5920 | Vaginal insufflator. | N | N
Obstetrics/Gynecology | OQY | Intrauterine Tamponade Balloon | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | OBB | Kit, Conception-Assist, Home Use | 2 | 884.5250 | Cervical cap. | N | N
Obstetrics/Gynecology | HDZ | Knife, Cervical Cone | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | MLS | Labor And Delivery Kit | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | MQK | Labware, Assisted Reproduction | 2 | 884.6160 | Assisted reproduction labware. | N | N
Obstetrics/Gynecology | PUD | Labware, Assisted Reproduction, Exempt | 2 | 884.6160 | Assisted reproduction labware. | N | N
Obstetrics/Gynecology | NMH | Laparoscope And Accessories, Gynecologic, Reprocessed | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | HET | Laparoscope, Gynecologic (And Accessories) | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | NWV | Laparoscopic Accessories, Gynecologic | 1 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | OCQ | Laparoscopic Bladder-Neck Suspension Instrument, Stress Urinary Incontinence | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | KNH | Laparoscopic Contraceptive Tubal Occlusion Device | 3 | 884.5380 | Contraceptive tubal occlusion device (TOD) and introducer. | N | Y
Obstetrics/Gynecology | NMI | Laparoscopic Insufflator And Accessories, Reprocessed | 1 | 884.1730 | Laparoscopic insufflator. | N | N
Obstetrics/Gynecology | LLW | Laser, Neodymium:Yag For Gynecologic Use | 2 | 884.4550 | Gynecologic surgical laser. | N | N
Obstetrics/Gynecology | HHR | Laser, Surgical, Gynecologic | 2 | 884.4550 | Gynecologic surgical laser. | N | N
Obstetrics/Gynecology | HHJ | Locator, Intracorporeal Device, Ultrasonic, Obstetric-Gynecologic | 2 | 884.2225 | Obstetric-gynecologic ultrasonic imager. | N | N
Obstetrics/Gynecology | NUC | Lubricant, Personal | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | PEB | Lubricant, Personal, Gamete, Fertilization, And Embryo Compatible | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | PVT | Maternity Kit | 1 | 884.5435 | Unscented menstrual pad. | N | N
Obstetrics/Gynecology | MQL | Media, Reproductive | 2 | 884.6180 | Reproductive media and supplements. | N | N
Obstetrics/Gynecology | PAI | Mesh, Surgical, Non-Synthetic, Urogynecologic, For Pelvic Organ Prolapse, Transvaginally Placed | 3 | 884.5980 | Surgical mesh for transvaginal pelvic organ prolapse repair. | N | Y
Obstetrics/Gynecology | OTP | Mesh, Surgical, Synthetic, Urogynecologic, For Pelvic Organ Prolapse, Transvaginally Placed | 3 | 884.5980 | Surgical mesh for transvaginal pelvic organ prolapse repair. | N | Y
Obstetrics/Gynecology | LZL | Micro-Condom | 3 | 884.5320 | Glans sheath. | N | N
Obstetrics/Gynecology | MQJ | Micromanipulators And Microinjectors, Assisted Reproduction | 2 | 884.6150 | Assisted reproduction micromanipulators and microinjectors. | N | N
Obstetrics/Gynecology | MTX | Microscope And Microscope Accessories, Reproduction, Assisted | 1 | 884.6190 | Assisted reproductive microscopes and microscope accessories. | N | N
Obstetrics/Gynecology | MQI | Microtool Fabrication, Assisted Reproduction | 2 | 884.6140 | Assisted reproduction micropipette fabrication instruments. | N | N
Obstetrics/Gynecology | MQH | Microtools, Assisted Reproduction (Pipettes) | 2 | 884.6130 | Assisted reproduction microtools. | N | N
Obstetrics/Gynecology | PUC | Microtools, Assisted Reproduction (Pipettes), Exempt | 2 | 884.6130 | Assisted reproduction microtools. | N | N
Obstetrics/Gynecology | HFK | Mold, Vaginal | 2 | 884.3900 | Vaginal stent. | N | N
Obstetrics/Gynecology | HEP | Monitor, Blood-Flow, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | KXN | Monitor, Cardiac, Fetal | 2 | 884.2600 | Fetal cardiac monitor. | N | N
Obstetrics/Gynecology | HGO | Monitor, Electroencephalographic, Fetal (And Accessories) | 3 | 884.2620 | Fetal electroencephalographic monitor. | N | N
Obstetrics/Gynecology | MAA | Monitor, Fetal Doppler Ultrasound | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | MOH | Monitor, Heart Rate, Fetal, Non-Stress Test (Home Use) | 2 | 884.2730 | Home uterine activity monitor. | N | N
Obstetrics/Gynecology | HEL | Monitor, Heart Rate, Fetal, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HEK | Monitor, Heart Sound, Fetal, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HEI | Monitor, Heart-Valve Movement, Fetal, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HEJ | Monitor, Hemic Sound, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HFP | Monitor, Phonocardiographic, Fetal | 2 | 884.2640 | Fetal phonocardiographic monitor and accessories. | N | N
Obstetrics/Gynecology | HEQ | Monitor, Pressure, Arterial, Fetal, Ultrasonic | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | KXO | Monitor, Pressure, Intrauterine | 2 | 884.2700 | Intrauterine pressure monitor and accessories. | N | N
Obstetrics/Gynecology | KNG | Monitor, Ultrasonic, Fetal | 2 | 884.2660 | Fetal ultrasonic monitor and accessories. | N | N
Obstetrics/Gynecology | HFM | Monitor, Uterine Contraction, External (For Use In Clinic) | 2 | 884.2720 | External uterine contraction monitor and accessories. | N | N
Obstetrics/Gynecology | MQE | Needle, Assisted Reproduction | 2 | 884.6100 | Assisted reproduction needles. | N | N
Obstetrics/Gynecology | HDH | Needle, Cerclage, Gynecological | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | NNB | Needle, Reproduction, Assisted, Reprocessed | 2 | 884.6100 | Assisted reproduction needles. | N | N
Obstetrics/Gynecology | PVU | Obstetrical Anesthesia Kit | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | OKV | Obstetrical Kit | 2 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | OHC | Obstetrical Vacuum Delivery Kit | 2 | 884.4340 | Fetal vacuum extractor. | N | N
Obstetrics/Gynecology | ONX | Ovarian Adnexal Mass Assessment Score Test System | 2 | 866.6050 | Ovarian adnexal mass assessment score test system. | N | N
Obstetrics/Gynecology | HDM | Packer, Uterine | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | NUR | Pad, Interlabial | 1 | 884.5435 | Unscented menstrual pad. | N | N
Obstetrics/Gynecology | NUQ | Pad, Menstrual, Reusable | 1 | 884.5435 | Unscented menstrual pad. | N | N
Obstetrics/Gynecology | HHL | Pad, Menstrual, Scented | 2 | 884.5425 | Scented or scented deodorized menstrual pad. | N | N
Obstetrics/Gynecology | HHD | Pad, Menstrual, Unscented | 1 | 884.5435 | Unscented menstrual pad. | N | N
Obstetrics/Gynecology | NRC | Pads, Menstrual, Scented-Deodorized | 1 | 884.5425 | Scented or scented deodorized menstrual pad. | N | N
Obstetrics/Gynecology | OHL | Pap Smear Kit | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PVV | Paracervical Anesthesia Kit | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | MLT | Pelvic Exam Kit | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HER | Pelvimeter, External | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | LBX | Pelvimeter, Internal | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HIR | Perineometer | 2 | 884.1425 | Perineometer. | N | N
Obstetrics/Gynecology | HHW | Pessary, Vaginal | 2 | 884.3575 | Vaginal pessary. | N | N
Obstetrics/Gynecology | PNU | Pressure Wedge, Perianal, For Reduction Of Cesarean Delivery | 2 | 884.5210 | Pressure wedge for the reduction of cesarean delivery. | N | N
Obstetrics/Gynecology | HFJ | Prosthesis, Fallopian Tube | 2 | 884.3650 | Fallopian tube prosthesis. | N | Y
Obstetrics/Gynecology | PVW | Pudendal Anesthesia Kit | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | HGY | Pump, Breast, Non-Powered | 1 | 884.5150 | Nonpowered breast pump. | N | N
Obstetrics/Gynecology | HGX | Pump, Breast, Powered | 2 | 884.5160 | Powered breast pump. | N | N
Obstetrics/Gynecology | HFO | Recorder, Pressure, Intrauterine | 2 | 884.2700 | Intrauterine pressure monitor and accessories. | N | N
Obstetrics/Gynecology | HHF | Remover, Intrauterine Device, Contraceptive, Hook-Type | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HDL | Retractor, Vaginal | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HIO | Sampler, Amniotic Fluid (Amniocentesis Tray) | 1 | 884.1550 | Amniotic fluid sampler (amniocentesis tray). | N | N
Obstetrics/Gynecology | HGW | Sampler, Blood, Fetal | 2 | 884.1560 | Fetal blood sampler. | N | N
Obstetrics/Gynecology | PCF | Sampler, Endocervical | 2 | 884.1050 | Endocervical aspirator. | N | N
Obstetrics/Gynecology | HDK | Scissors, Episiotomy | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HDJ | Scissors, Umbilical | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HHO | Screw, Fibroid, Gynecological | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | QHZ | Seal, Cervical, Mechanical, Unpowered | 1 | 884.1690 | Hysteroscope and accessories. | N | N
Obstetrics/Gynecology | OKW | Seminal Fluid Collection Kit | 2 | 884.5300 | Condom. | N | N
Obstetrics/Gynecology | KNE | Set, Anesthesia, Obstetric | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | HEE | Set, Anesthesia, Paracervical | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | HEG | Set, Anesthesia, Pudendal | 2 | 884.5100 | Obstetric anesthesia set. | N | N
Obstetrics/Gynecology | NHM | Sheet, Recording, Breast Examination | 2 | 884.2990 | Breast lesion documentation system. | N | N
Obstetrics/Gynecology | FHJ | Shield, Circumcision | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | MBU | Single-Use Internal Condom | 2 | 884.5340 | Single-use internal condom. | N | N
Obstetrics/Gynecology | HHM | Sound, Uterine | 1 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HGF | Source, Abortion Unit, Vacuum | 2 | 884.5070 | Vacuum abortion system. | N | N
Obstetrics/Gynecology | MPU | Source, Chemiluminescent Light | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HHT | Spatula, Cervical, Cytological | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HDF | Speculum, Vaginal, Metal | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HDG | Speculum, Vaginal, Metal, Fiberoptic | 1 | 884.4520 | Obstetric-gynecologic general manual instrument. | N | N
Obstetrics/Gynecology | HIB | Speculum, Vaginal, Nonmetal | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HIC | Speculum, Vaginal, Nonmetal, Fiberoptic | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | KXP | Stent, Vaginal | 2 | 884.3900 | Vaginal stent. | N | N
Obstetrics/Gynecology | HGN | Stethoscope, Fetal | 1 | 884.2900 | Fetal stethoscope. | N | N
Obstetrics/Gynecology | HII | Stimulator, Vaginal, Muscle, Powered, For Therapeutic Use | 3 | 884.5940 | Powered vaginal muscle stimulator for therapeutic use. | N | N
Obstetrics/Gynecology | HHQ | System, Abortion, Metreurynter-Balloon | 3 | 884.5050 | Metreurynter-balloon abortion system. | N | N
Obstetrics/Gynecology | HHI | System, Abortion, Vacuum | 2 | 884.5070 | Vacuum abortion system. | N | N
Obstetrics/Gynecology | MRX | System, Assisted Reproduction Laser | 2 | 884.6200 | Assisted reproduction laser system. | N | N
Obstetrics/Gynecology | NKA | System, Documentation, Breast Lesion | 2 | 884.2990 | Breast lesion documentation system. | N | N
Obstetrics/Gynecology | MDG | System, Intrafallopian Cannula | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HGM | System, Monitoring, Perinatal | 2 | 884.2740 | Perinatal monitoring system and accessories. | N | N
Obstetrics/Gynecology | LHQ | System, Telethermographic (Adjunctive Use) | 1 | 884.2980 | Telethermographic system. | N | N
Obstetrics/Gynecology | LHP | System, Telethermographic (Sole Diagnostic Screen) | 3 | 884.2980 | Telethermographic system. | N | N
Obstetrics/Gynecology | IYM | System, Telethermographic, Infrared | 3 | 884.2980 | Telethermographic system. | N | N
Obstetrics/Gynecology | LHM | System, Thermographic, Liquid Crystal | 1 | 884.2982 | Liquid crystal thermographic system. | N | N
Obstetrics/Gynecology | LHR | System, Thermographic, Liquid Crystal (Sole Diagnostic Screen) | 3 | 884.2982 | Liquid crystal thermographic system. | N | N
Obstetrics/Gynecology | KYA | System, Thermographic, Liquid Crystal, Nonpowered (Adjunctive Use) | 1 | 884.2982 | Liquid crystal thermographic system. | N | N
Obstetrics/Gynecology | KXZ | System, Thermographic, Liquid Crystal, Powered (Adjunctive Use) | 3 | 884.2982 | Liquid crystal thermographic system. | N | N
Obstetrics/Gynecology | MTW | System, Water, Reproduction, Assisted, And Purification | 2 | 884.6170 | Assisted reproduction water and water purification systems. | N | N
Obstetrics/Gynecology | KNC | Table, Obstetric (And Accessories) | 2 | 884.4900 | Obstetric table and accessories. | N | N
Obstetrics/Gynecology | HDD | Table, Obstetrical, Ac-Powered (And Accessories) | 2 | 884.4900 | Obstetric table and accessories. | N | N
Obstetrics/Gynecology | HHP | Table, Obstetrical, Manual (And Accessories) | 2 | 884.4900 | Obstetric table and accessories. | N | N
Obstetrics/Gynecology | HIL | Tampon, Menstrual, Scented, Deodorized | 2 | 884.5460 | Scented or scented deodorized menstrual tampon. | N | N
Obstetrics/Gynecology | HEB | Tampon, Menstrual, Unscented | 2 | 884.5470 | Unscented menstrual tampon. | N | N
Obstetrics/Gynecology | HDC | Tenaculum, Uterine | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | PUA | Tenaculum, Uterine, Exempt | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | HHS | Transcervical Contraceptive Tubal Occlusion Device | 3 | 884.5380 | Contraceptive tubal occlusion device (TOD) and introducer. | N | Y
Obstetrics/Gynecology | HFN | Transducer, Pressure, Intrauterine | 2 | 884.2700 | Intrauterine pressure monitor and accessories. | N | N
Obstetrics/Gynecology | HGL | Transducer, Ultrasonic, Obstetric | 2 | 884.2960 | Obstetric ultrasonic transducer and accessories. | N | N
Obstetrics/Gynecology | OKX | Trocar Kit | 2 | 884.1720 | Gynecologic laparoscope and accessories. | N | N
Obstetrics/Gynecology | NVT | Tubing And Tubing/Filter Fits | 1 | 884.1700 | Hysteroscopic insufflator. | N | N
Obstetrics/Gynecology | NKC | Tubing/Tubing With Filter, Insufflation, Laparoscopic | 1 | 884.1730 | Laparoscopic insufflator. | N | N
Obstetrics/Gynecology | OSP | Uterine Electromyographic Monitor | 2 | 884.2720 | External uterine contraction monitor and accessories. | N | N
Obstetrics/Gynecology | PVX | Vaginal Examination Tray | 2 | 884.4530 | Obstetric-gynecologic specialized manual instrument. | N | N
Obstetrics/Gynecology | MOK | Vaginoscope And Accessories | 2 | 884.1630 | Colposcope. | N | N
Obstetrics/Gynecology | HDO | Valve, Tubal Occlusion | 3 | 884.5380 | Contraceptive tubal occlusion device (TOD) and introducer. | N | Y
Obstetrics/Gynecology | KXQ | Vibrator For Therapeutic Use, Genital | 2 | 884.5960 | Genital vibrator for therapeutic use. | N | N
Obstetrics/Gynecology | LHZ | Viscometer, Mucus, Cervical | 1 | 884.1040 | Viscometer for cervical mucus. | N | N
Obstetrics/Gynecology | HFD | Washer, Endometrial | 2 | 884.1185 | Endometrial washer. | N | N
ophthalmic | NCF | Aberrometer, Ophthalmic | 1 | 886.1760 | Ophthalmic refractometer. | N | N
ophthalmic | LPN | Accessories, Soft Lens Products | 2 | 886.5928 | Soft (hydrophilic) contact lens care products. | N | N
ophthalmic | LYL | Accessories, Solution, Ultrasonic Cleaners For Lenses | 2 | 886.5928 | Soft (hydrophilic) contact lens care products. | N | N
ophthalmic | HJW | Adaptometer (Biophotometer) | 1 | 886.1050 | Adaptometer (biophotometer). | N | N
ophthalmic | OUM | Adaptometer (Biophotometer), Software-Based Data Acquisition And Stimulus Generation | 1 | 886.1050 | Adaptometer (biophotometer). | N | N
ophthalmic | LZP | Aid, Surgical, Viscoelastic | 3 | 886.4275 | Intraocular fluid. | N | N
ophthalmic | HOT | Aid, Vision Image Intensification | 1 | 886.5910 | Image intensification vision aid. | N | N
ophthalmic | HPF | Aid, Vision, Electronic, Ac-Powered | 1 | 886.5900 | Electronic vision aid. | N | N
ophthalmic | HPG | Aid, Vision, Electronic, Battery-Powered | 1 | 886.5900 | Electronic vision aid. | N | N
ophthalmic | HPI | Aid, Vision, Optical, Ac-Powered | 1 | 886.5915 | Optical vision aid. | N | N
ophthalmic | HPE | Aid, Vision, Optical, Battery-Powered | 1 | 886.5915 | Optical vision aid. | N | N
ophthalmic | HIW | Anomaloscope | 1 | 886.1070 | Anomaloscope. | N | N
ophthalmic | HQR | Apparatus, Cautery, Radiofrequency, Ac-Powered | 2 | 886.4100 | Radiofrequency electrosurgical cautery apparatus. | N | N
ophthalmic | HQQ | Apparatus, Cautery, Radiofrequency, Battery-Powered | 2 | 886.4100 | Radiofrequency electrosurgical cautery apparatus. | N | N
ophthalmic | PUL | Apparatus, Cutting, Radiofrequency, Electrosurgical, Ac-Powered | 2 | 886.4100 | Radiofrequency electrosurgical cautery apparatus. | N | N
ophthalmic | NCR | Apparatus, Cutting, Radiofrequency, Electrosurgical, Battery-Powered | 2 | 886.4100 | Radiofrequency electrosurgical cautery apparatus. | N | N
ophthalmic | LCC | Applicator, Ocular Pressure | 2 | 886.4610 | Ocular pressure applicator. | N | N
ophthalmic | HJO | Biomicroscope, Slit-Lamp, Ac-Powered | 2 | 886.1850 | AC-powered slitlamp biomicroscope. | N | N
ophthalmic | PUE | Biomicroscope, Slit-Lamp, Ac-Powered, Exempt | 2 | 886.1850 | AC-powered slitlamp biomicroscope. | N | N
ophthalmic | NKY | Blade, Keratome, Reprocessed | 1 | 886.4370 | Keratome. | N | N
ophthalmic | HIY | Brush, Haidinger, (Including Macular Integrity) | 1 | 886.1090 | Haidinger brush. | N | N
ophthalmic | HQS | Burr, Corneal, Ac-Powered | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | OHX | Burr, Corneal, Ac-Powered, Rust Ring Removal | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | OHW | Burr, Corneal, Battery Powered, Rust Ring Removal | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | HOG | Burr, Corneal, Battery-Powered | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | HOF | Burr, Corneal, Manual | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | NCK | Button, Iris, Eye, Artificial | 1 | 886.3200 | Artificial eye. | N | N
ophthalmic | HRE | Cabinet, Instrument, Ac-Powered, Ophthalmic | 1 | 886.1860 | Ophthalmic instrument stand. | N | N
ophthalmic | HLA | Calibrator, Tonometer | 2 | 886.1930 | Tonometer and accessories. | N | N
ophthalmic | HOE | Caliper, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HKI | Camera, Ophthalmic, Ac-Powered | 2 | 886.1120 | Ophthalmic camera. | N | N
ophthalmic | PJZ | Camera, Ophthalmic, General-Use | 2 | 886.1120 | Ophthalmic camera. | N | N
ophthalmic | QER | Camera, Ophthalmic, Slit-Scanning | 2 | 886.1120 | Ophthalmic camera. | N | N
ophthalmic | HOP | Campimeter, Stereo, Battery-Powered | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HMX | Cannula, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | MXY | Cannula, Ophthalmic, Posterior Capsular Polishing, Polyvinyl Acetal | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | NGY | Cannula, Trocar, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | LRX | Case, Contact Lens | 2 | 886.5928 | Soft (hydrophilic) contact lens care products. | N | N
ophthalmic | HME | Chair, Ophthalmic, Ac-Powered | 1 | 886.1140 | Ophthalmic chair. | N | N
ophthalmic | HMD | Chair, Ophthalmic, Manual | 1 | 886.1140 | Ophthalmic chair. | N | N
ophthalmic | HOX | Chart, Visual Acuity | 1 | 886.1150 | Visual acuity chart. | N | N
ophthalmic | HOD | Clamp, Eyelid, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HOB | Clamp, Muscle, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HOC | Clip, Iris Retractor | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HPB | Clip, Lens, Trial, Ophthalmic | 1 | 886.1410 | Ophthalmic trial lens clip. | N | N
ophthalmic | HQW | Clip, Tantalum, Ophthalmic | 2 | 886.3100 | Ophthalmic tantalum clip. | N | Y
ophthalmic | MOE | Collagen Corneal Shield | 1 | 886.4750 | Ophthalmic eye shield. | N | N
ophthalmic | HOA | Compressor, Orbital | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HQN | Conformer, Ophthalmic | 2 | 886.3130 | Ophthalmic conformer. | N | N
ophthalmic | NQB | Conformer, Ophthalmic, Biological Tissue | 2 | 886.3130 | Ophthalmic conformer. | N | Y
ophthalmic | QGJ | Corneal Inlay Inserter Handle | 1 | 886.4355 | Corneal inlay inserter handle. | N | N
ophthalmic | HJZ | Crutch, Ptosis | 1 | 886.5600 | Ptosis crutch. | N | N
ophthalmic | HNZ | Curette, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNY | Cystotome | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNX | Depressor, Orbital | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | MXK | Device, Analysis, Anterior Segment | 2 | 886.1850 | AC-powered slitlamp biomicroscope. | N | N
ophthalmic | HPL | Device, Fixation, Ac-Powered, Ophthalmic | 1 | 886.1290 | Fixation device. | N | N
ophthalmic | KYG | Device, Irrigation, Ocular Surgery | 1 | 886.4360 | Ocular surgery irrigation device. | N | N
ophthalmic | HLF | Device, Measuring, Lens Radius, Ophthalmic | 1 | 886.1430 | Ophthalmic contact lens radius measuring device. | N | N
ophthalmic | PIB | Diabetic Retinopathy Detection Device | 2 | 886.1100 |  | N | N
ophthalmic | MMC | Dilator, Expansive Iris (Accessory) | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | HNW | Dilator, Lachrymal | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HMM | Distometer | 1 | 886.1190 | Distometer. | N | N
ophthalmic | HMS | Drum, Eye Knife Test | 1 | 886.4230 | Ophthalmic knife test drum. | N | N
ophthalmic | HOW | Drum, Opticokinetic | 1 | 886.1200 | Optokinetic drum. | N | N
ophthalmic | HLZ | Electrode, Corneal | 2 | 886.1220 | Corneal electrode. | N | N
ophthalmic | QKV | Electromechanical Tear Stimulator | 2 | 886.5305 |  | N | N
ophthalmic | HRG | Engine, Trephine, Accessories, Ac-Powered | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | HRF | Engine, Trephine, Accessories, Battery-Powered | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | HLD | Engine, Trephine, Accessories, Gas-Powered | 1 | 886.4070 | Powered corneal burr. | N | N
ophthalmic | HNT | Erisophake | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HJC | Esthesiometer, Ocular | 1 | 886.1040 | Ocular esthesiometer. | N | N
ophthalmic | HMK | Euthyscope, Ac-Powered | 2 | 886.1250 | Euthyscope. | N | N
ophthalmic | HML | Euthyscope, Battery-Powered | 1 | 886.1250 | Euthyscope. | N | N
ophthalmic | HLS | Exophthalmometer | 1 | 886.1270 | Exophthalmometer. | N | N
ophthalmic | NFM | Expander, Tissue, Orbital | 2 | 886.3320 | Eye sphere implant. | N | Y
ophthalmic | HNS | Expressor | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | POE | Extended Depth Of Focus Intraocular Lens | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | OJK | Eye Tray | 2 | 886.4790 | Ophthalmic sponge. | N | N
ophthalmic | HQH | Eye, Artificial, Non-Custom | 1 | 886.3200 | Artificial eye. | N | N
ophthalmic | ORZ | Eyelid Thermal Pulsation System | 2 | 886.5200 | Eyelid thermal pulsation system. | N | N
ophthalmic | NXR | Fixation, Vision Training, Binocular | 1 | 886.1290 | Fixation device. | N | N
ophthalmic | NJG | Flasher, Afterimage, Ophthalmic | 2 | 886.1300 | Afterimage flasher. | N | N
ophthalmic | LWL | Fluid, Intraocular | 3 | 886.4275 | Intraocular fluid. | N | N
ophthalmic | MUS | Fluidic, Phacoemulsification/Phacofragmentation | 2 | 886.4670 | Phacofragmentation system. | N | N
ophthalmic | MSS | Folders And Injectors, Intraocular Lens (Iol) | 1 | 886.4300 | Intraocular lens guide. | N | N
ophthalmic | HNR | Forceps, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HKG | Fornixscope | 1 | 886.1320 | Fornixscope. | N | N
ophthalmic | HQZ | Frame, Spectacle | 1 | 886.5842 | Spectacle frame. | N | N
ophthalmic | HPA | Frame, Trial, Ophthalmic | 1 | 886.1415 | Ophthalmic trial lens frame. | N | N
ophthalmic | LPO | Gases Used Within Eye To Place Pressure On Detached Retina | 3 | 886.4270 | Intraocular gas. | N | N
ophthalmic | HLN | Gauge, Lens, Ophthalmic | 1 | 886.1420 | Ophthalmic lens gauge. | N | N
ophthalmic | OTZ | Graft Insertion Instrument For Endothelial Keratoplasty | 1 | 886.4300 | Intraocular lens guide. | N | N
ophthalmic | HOQ | Grid, Amsler | 1 | 886.1330 | Amsler grid. | N | N
ophthalmic | HJT | Haploscope | 1 | 886.1340 | Haploscope. | N | N
ophthalmic | HPQ | Headlamp, Operating, Ac-Powered | 2 | 886.4335 | Operating headlamp. | N | N
ophthalmic | HPP | Headlamp, Operating, Battery-Operated | 1 | 886.4335 | Operating headlamp. | N | N
ophthalmic | FCT | Headlight, Fiberoptic Focusing | 2 | 886.4335 | Operating headlamp. | N | N
ophthalmic | HNQ | Hook, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HJD | Illuminator, Color Vision Plate | 1 | 886.1160 | Color vision plate illuminator. | N | N
ophthalmic | HQJ | Implant, Absorbable, (Scleral Buckling Methods) | 2 | 886.3300 | Absorbable implant (scleral buckling method). | N | Y
ophthalmic | HPZ | Implant, Eye Sphere | 2 | 886.3320 | Eye sphere implant. | N | Y
ophthalmic | KYF | Implant, Eye Valve | 2 | 886.3920 | Aqueous shunt. | N | Y
ophthalmic | HQX | Implant, Orbital, Extra-Ocular | 2 | 886.3340 | Extraocular orbital implant. | N | Y
ophthalmic | OPT | Infrared Hematoma Detector | 2 | 882.1935 | Near Infrared (NIR) Brain Hematoma Detector. | N | N
ophthalmic | NCE | Injector, Capsular Tension Ring | 1 | 886.4300 | Intraocular lens guide. | N | N
ophthalmic | KYE | Inserter/Remover Contact Lens | 1 | 886.5420 | Contact lens inserter/remover. | N | N
ophthalmic | HJB | Instrument, Measuring, Corneal Radius | 1 | 886.1450 | Corneal radius measuring device. | N | N
ophthalmic | HLM | Instrument, Measuring, Lens, Ac-Powered | 1 | 886.1425 | Lens measuring instrument. | N | N
ophthalmic | HLC | Instrument, Measuring, Stereopsis | 1 | 886.1460 | Stereopsis measuring instrument. | N | N
ophthalmic | HPJ | Instrument, Visual Field, Laser | 2 | 886.1360 | Visual field laser instrument. | N | N
ophthalmic | HQE | Instrument, Vitreous Aspiration And Cutting, Ac-Powered | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | HKP | Instrument, Vitreous Aspiration And Cutting, Battery-Powered | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | PQJ | Intranasal Electrostimulation Device | 2 | 886.5300 | Tear electrostimulation device. | N | N
ophthalmic | QBR | Intranasal Electrostimulation Device For Dry Eye Symptoms | 2 | 886.5310 | Intranasal electrostimulation device for dry eye symptoms. | N | N
ophthalmic | HQL | Intraocular Lens | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | NVB | Intraocular, Pressure Measuring Device | 3 | 886.4280 | Intraocular pressure measuring device. | N | N
ophthalmic | HNP | Introducer, Sphere | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | QDV | Iol Tray | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNO | Keratome, Ac-Powered | 1 | 886.4370 | Keratome. | N | N
ophthalmic | HMY | Keratome, Battery-Powered | 1 | 886.4370 | Keratome. | N | N
ophthalmic | MYD | Keratome, Water Jet | 1 | 886.4370 | Keratome. | N | N
ophthalmic | HQM | Keratoprosthesis, Permanent Implant | 2 | 886.3400 | Keratoprosthesis. | N | Y
ophthalmic | MLP | Keratoprosthesis, Temporary Implant, Surgical Use | 2 | 886.3400 | Keratoprosthesis. | N | Y
ophthalmic | HLQ | Keratoscope, Ac-Powered | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | HLR | Keratoscope, Battery-Powered | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | HNN | Knife, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | NLA | Knife, Ophthalmic, Reprocessed | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | LXS | Laser, Neodymium:Yag, Ophthalmic For Posterior Capsulotomy And Cutting Pupilla | 2 | 886.4392 | Nd:YAG laser for posterior capsulotomy and peripheral iridotomy. | N | N
ophthalmic | HQF | Laser, Ophthalmic | 2 | 886.4390 | Ophthalmic laser. | N | N
ophthalmic | MXO | Laser, System, Phacolysis | 2 | 886.4670 | Phacofragmentation system. | N | N
ophthalmic | HPD | Lens, Bagolini | 1 | 886.1375 | Bagolini lens. | N | N
ophthalmic | HJL | Lens, Condensing, Diagnostic | 1 | 886.1380 | Diagnostic condensing lens. | N | N
ophthalmic | MUW | Lens, Contact (Orthokeratology) | 2 | 886.5916 | Rigid gas permeable contact lens. | N | N
ophthalmic | HQD | Lens, Contact (Other Material) - Daily | 2 | 886.5916 | Rigid gas permeable contact lens. | N | N
ophthalmic | MWL | Lens, Contact (Rigid Gas Permeable), Extended Wear | 3 | 886.5916 | Rigid gas permeable contact lens. | N | N
ophthalmic | MVN | Lens, Contact, (Disposable) | 2 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | NCZ | Lens, Contact, For Color Vision Deficiency | 2 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | NIC | Lens, Contact, For Reading Discomfort | 2 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | NUU | Lens, Contact, Orthokeratology, Overnight | 3 | 886.5916 | Rigid gas permeable contact lens. | N | N
ophthalmic | HJK | Lens, Contact, Polymethylmethacrylate, Diagnostic | 2 | 886.1385 | Polymethylmethacrylate (PMMA) diagnostic contact lens. | N | N
ophthalmic | HJJ | Lens, Fresnel, Flexible, Diagnostic | 1 | 886.1390 | Flexible diagnostic Fresnel lens. | N | N
ophthalmic | HJI | Lens, Fundus, Hruby, Diagnostic | 1 | 886.1395 | Diagnostic Hruby fundus lens. | N | N
ophthalmic | KYB | Lens, Guide, Intraocular | 1 | 886.4300 | Intraocular lens guide. | N | N
ophthalmic | NAA | Lens, Intraocular, Accommodative | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | MJP | Lens, Intraocular, Toric Optics | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | NIZ | Lens, Iris Reconstruction | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | HKR | Lens, Maddox | 1 | 886.1400 | Maddox lens. | N | N
ophthalmic | MFK | Lens, Multifocal Intraocular | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | NAI | Lens, Prescription, Color Deficiency | 1 | 886.5844 | Prescription spectacle lens. | N | N
ophthalmic | NJH | Lens, Spectacle (Prescription), For Reading Discomfort | 1 | 886.5844 | Prescription spectacle lens. | N | N
ophthalmic | HQG | Lens, Spectacle, Non-Custom (Prescription) | 1 | 886.5844 | Prescription spectacle lens. | N | N
ophthalmic | LQJ | Lens, Surgical, Laser, Accesssory, Ophthalmic Laser | 2 | 886.4390 | Ophthalmic laser. | N | N
ophthalmic | LPL | Lenses, Soft Contact, Daily Wear | 2 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | LPM | Lenses, Soft Contact, Extended Wear | 3 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | PZK | Light Adjustable Lens (Lal) And Light Delivery Device (Ldd) | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | FSR | Light, Headband, Surgical | 2 | 886.4335 | Operating headlamp. | N | N
ophthalmic | FTZ | Locator, Magnetic | 1 | 886.4445 | Permanent magnet. | N | N
ophthalmic | HPM | Locator, Metal, Electronic | 2 | 886.4400 | Electronic metal locator. | N | N
ophthalmic | HJH | Loupe, Binocular, Low Power | 1 | 886.5120 | Low-power binocular loupe. | N | N
ophthalmic | HPO | Magnet, Ac-Powered | 2 | 886.4440 | AC-powered magnet. | N | N
ophthalmic | HPN | Magnet, Permanent | 1 | 886.4445 | Permanent magnet. | N | N
ophthalmic | HJF | Magnifier, Hand-Held, Low-Vision | 1 | 886.5540 | Low-vision magnifier. | N | N
ophthalmic | HMR | Marker, Ocular | 1 | 886.4570 | Ophthalmic surgical marker. | N | N
ophthalmic | HMQ | Marker, Sclera | 1 | 886.4570 | Ophthalmic surgical marker. | N | N
ophthalmic | HIX | Maxwell Spot, Ac-Powered | 1 | 886.1435 | Maxwell spot. | N | N
ophthalmic | NQE | Microscope, Specular | 2 | 886.1850 | AC-powered slitlamp biomicroscope. | N | N
ophthalmic | HKF | Mirror, Headband, Ophthalmic | 1 | 886.1500 | Headband mirror. | N | N
ophthalmic | HLL | Monitor, Eye Movement | 2 | 886.1510 | Eye movement monitor. | N | N
ophthalmic | HMC | Monitor, Eye Movement, Diagnostic | 2 | 886.1510 | Eye movement monitor. | N | N
ophthalmic | HNM | Needle, Ophthalmic Suturing | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | NKX | Needle, Phacoemulsification, Reprocessed | 2 | 886.4670 | Phacofragmentation system. | N | N
ophthalmic | PLZ | Ocular Pattern Recorder | 2 | 886.1925 | Diurnal pattern recorder system. | N | N
ophthalmic | MQU | Ocular Peg | 2 | 886.3320 | Eye sphere implant. | N | Y
ophthalmic | PUF | Ocular Peg, Exempt | 2 | 886.3320 | Eye sphere implant. | N | Y
ophthalmic | OOE | Ophthalmic Femtosecond Laser | 2 | 886.4390 | Ophthalmic laser. | N | N
ophthalmic | NYK | Ophthalmodynamometer, Diagnostic Contact Lens, Polymethylmethacrylate (Pmma) | 2 | 886.1385 | Polymethylmethacrylate (PMMA) diagnostic contact lens. | N | N
ophthalmic | HLI | Ophthalmoscope, Ac-Powered | 2 | 886.1570 | Ophthalmoscope. | N | N
ophthalmic | HLJ | Ophthalmoscope, Battery-Powered | 2 | 886.1570 | Ophthalmoscope. | N | N
ophthalmic | MYC | Ophthalmoscope, Laser, Scanning | 2 | 886.1570 | Ophthalmoscope. | N | N
ophthalmic | MSG | Ophthalmoscopes, Replacement Batteries, Hand-Held | 2 | 886.1570 | Ophthalmoscope. | N | N
ophthalmic | PIC | Oral Electronic Vision Aid | 2 | 886.5905 | Oral electronic vision aid. | N | N
ophthalmic | HRP | Pen, Marking, Surgical | 1 | 886.4570 | Ophthalmic surgical marker. | N | N
ophthalmic | HOO | Perimeter, Ac-Powered | 1 | 886.1605 | Perimeter. | N | N
ophthalmic | HPT | Perimeter, Automatic, Ac-Powered | 1 | 886.1605 | Perimeter. | N | N
ophthalmic | HON | Perimeter, Manual | 1 | 886.1605 | Perimeter. | N | N
ophthalmic | QDW | Phaco Kit | 2 | 886.4670 | Phacofragmentation system. | N | N
ophthalmic | QCB | Phakic Toric Intraocular Lens | 3 | 886.3600 | Intraocular lens. | N | Y
ophthalmic | HQB | Photocoagulator And Accessories | 2 | 886.4690 | Ophthalmic photocoagulator. | N | N
ophthalmic | HJA | Photokeratoscope | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | NOH | Photokeratoscope, Battery Powered | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | MMF | Photorefractor | 2 | 886.1120 | Ophthalmic camera. | N | N
ophthalmic | HLX | Photostimulator, Ac-Powered | 2 | 886.1630 | AC-powered photostimulator. | N | N
ophthalmic | LXP | Plug, Scleral | 2 | 886.4155 | Scleral plug. | N | N
ophthalmic | HLT | Preamplifier, Ac-Powered, Ophthalmic | 2 | 886.1640 | Ophthalmic preamplifier. | N | N
ophthalmic | HLW | Preamplifier, Battery-Powered, Ophthalmic | 2 | 886.1640 | Ophthalmic preamplifier. | N | N
ophthalmic | HKW | Prism, Bar, Ophthalmic | 1 | 886.1650 | Ophthalmic bar prism. | N | N
ophthalmic | HKT | Prism, Fresnel, Ophthalmic | 1 | 886.1655 | Ophthalmic Fresnel prism. | N | N
ophthalmic | HKS | Prism, Gonioscopic | 1 | 886.1660 | Gonioscopic prism. | N | N
ophthalmic | HKQ | Prism, Rotary, Ophthalmic | 1 | 886.1665 | Ophthalmic rotary prism. | N | N
ophthalmic | HPW | Probe And Counter, Isotope, For Phosphorus 32 | 2 | 886.1670 | Ophthalmic isotope uptake probe. | N | N
ophthalmic | HNL | Probe, Lachrymal | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNK | Probe, Trabeculotomy | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | MRC | Products, Contact Lens Care, Rigid Gas Permeable | 2 | 886.5918 | Rigid gas permeable contact lens care products. | N | N
ophthalmic | HOS | Projector, Ophthalmic | 1 | 886.1680 | Ophthalmic projector. | N | N
ophthalmic | NXM | Prosthesis, Eyelid Spacer/Graft, Biologic | 2 | 886.3130 | Ophthalmic conformer. | N | Y
ophthalmic | HNJ | Punch, Corneo-Scleral | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HRI | Pupillograph | 1 | 886.1690 | Pupillograph. | N | N
ophthalmic | HLG | Pupillometer, Ac-Powered | 1 | 886.1700 | Pupillometer. | N | N
ophthalmic | HLH | Pupillometer, Manual | 1 | 886.1700 | Pupillometer. | N | N
ophthalmic | HMH | Rack, Skiascopic | 1 | 886.1750 | Skiascopic rack. | N | N
ophthalmic | HJY | Reader, Bar, Ophthalmic | 1 | 886.5800 | Ophthalmic bar reader. | N | N
ophthalmic | HJX | Reader, Prism, Ophthalmic | 1 | 886.5810 | Ophthalmic prism reader. | N | N
ophthalmic | HKO | Refractometer, Ophthalmic | 1 | 886.1760 | Ophthalmic refractometer. | N | N
ophthalmic | HKN | Refractor, Manual, Non-Powered, Including Phoropter | 1 | 886.1770 | Manual refractor. | N | N
ophthalmic | HKL | Retinoscope, Ac-Powered | 2 | 886.1780 | Retinoscope. | N | N
ophthalmic | HKM | Retinoscope, Battery-Powered | 1 | 886.1780 | Retinoscope. | N | N
ophthalmic | HNI | Retractor, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNH | Ring, Ophthalmic (Flieringa) | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNG | Rongeur, Lachrymal Sac | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HLE | Ruler, Nearpoint (Punctometer) | 1 | 886.1790 | Nearpoint ruler. | N | N
ophthalmic | HNF | Scissors, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HOM | Screen, Tangent, Ac-Powered (Campimeter) | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HOL | Screen, Tangent, Felt (Campimeter) | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HOK | Screen, Tangent, Projection, Ac-Powered | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HMJ | Screen, Tangent, Projection, Battery-Powered | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HOJ | Screen, Tangent, Target | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | HLK | Screen, Tangent, Target, Battery-Powered | 1 | 886.1810 | Tangent screen (campimeter). | N | N
ophthalmic | NFD | Selector, Color, Spectacle, Lens, Tint | 1 | 886.1160 | Color vision plate illuminator. | N | N
ophthalmic | HPC | Set, Lens, Trial, Ophthalmic | 1 | 886.1405 | Ophthalmic trial lens set. | N | N
ophthalmic | HQT | Shell, Scleral | 2 | 886.3800 | Scleral shell. | N | Y
ophthalmic | HOY | Shield, Eye, Ophthalmic (Including Sunlamp Protective Eyewear And Post-Mydriatic Eyewear) | 1 | 886.4750 | Ophthalmic eye shield. | N | N
ophthalmic | HOR | Simulatan (Including Crossed Cylinder) | 1 | 886.1840 | Simulatan (including crossed cylinder). | N | N
ophthalmic | HNE | Snare, Enucleating | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HND | Spatula, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HKC | Spectacle Microscope, Low-Vision | 1 | 886.5540 | Low-vision magnifier. | N | N
ophthalmic | HOI | Spectacle, Magnifying | 1 | 886.5840 | Magnifying spectacles. | N | N
ophthalmic | HOH | Spectacle, Operating (Loupe), Ophthalmic | 1 | 886.4770 | Ophthalmic operating spectacles (loupes). | N | N
ophthalmic | HNC | Specula, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HOZ | Sponge, Ophthalmic | 2 | 886.4790 | Ophthalmic sponge. | N | N
ophthalmic | HNB | Spoon, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HNA | Spud, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HMF | Stand, Instrument, Ac-Powered, Ophthalmic | 1 | 886.1860 | Ophthalmic instrument stand. | N | N
ophthalmic | NOI | Stand, Instrument, Ophthalmic, Battery Powered | 1 | 886.1860 | Ophthalmic instrument stand. | N | N
ophthalmic | HMG | Stand, Instrument, Ophthalmic, Non-Powered | 1 | 886.1860 | Ophthalmic instrument stand. | N | N
ophthalmic | HJQ | Stereoscope, Ac-Powered | 1 | 886.1870 | Stereoscope. | N | N
ophthalmic | HJR | Stereoscope, Battery-Powered | 1 | 886.1870 | Stereoscope. | N | N
ophthalmic | HRD | Sterilizer, Soft-Lens, Thermal, Ac-Powered | 2 | 886.5928 | Soft (hydrophilic) contact lens care products. | N | N
ophthalmic | HRC | Sterilizer, Soft-Lens, Thermal, Battery-Powered | 2 | 886.5928 | Soft (hydrophilic) contact lens care products. | N | N
ophthalmic | HKZ | Sterilizer, Tonometer | 1 | 886.1940 | Tonometer sterilizer. | N | N
ophthalmic | PMW | Strabismus Detection Device | 2 | 886.1342 | Strabismus detection device. | N | N
ophthalmic | KYD | Strip, Schirmer | 1 | 886.1800 | Schirmer strip. | N | N
ophthalmic | HQY | Sunglasses (Non-Prescription Including Photosensitive) | 1 | 886.5850 | Sunglasses (nonprescription). | N | N
ophthalmic | PZG | Surgical Eye Tray | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | MZD | System, In-Office Tinting, Contact Lenses | 2 | 886.5925 | Soft (hydrophilic) contact lens. | N | N
ophthalmic | NPB | System, Monitoring, For Progress Of Labor | 2 | 884.2800 | Computerized Labor Monitoring System. | N | N
ophthalmic | HJG | System, Reading, Television, Closed-Circuit | 1 | 886.5820 | Closed-circuit television reading system. | N | N
ophthalmic | HRK | Table, Instrument, Manual, Ophthalmic | 1 | 886.4855 | Ophthalmic instrument table. | N | N
ophthalmic | HRJ | Table, Instrument, Powered, Ophthalmic | 1 | 886.4855 | Ophthalmic instrument table. | N | N
ophthalmic | HKD | Tape, Nystagmus | 1 | 886.1905 | Nystagmus tape. | N | N
ophthalmic | HLP | Target, Fusion And Stereoscopic | 1 | 886.1880 | Fusion and stereoscopic target. | N | N
ophthalmic | PLX | Tear Duct Occluder | 1 | 886.5838 | Nasolacrimal compression device. | N | N
ophthalmic | HKB | Telescope, Hand-Held, Low-Vision | 1 | 886.5870 | Low-vision telescope. | N | N
ophthalmic | HKK | Telescope, Spectacle, Low-Vision | 1 | 886.5870 | Low-vision telescope. | N | N
ophthalmic | HJS | Test, Spectacle Dissociation, Ac-Powered (Lancaster) | 1 | 886.1910 | Spectacle dissociation test system. | N | N
ophthalmic | HLO | Test, Spectacle Dissociation, Battery-Powered (Lancaster) | 1 | 886.1910 | Spectacle dissociation test system. | N | N
ophthalmic | HIT | Tester, Color Vision | 1 | 886.1170 | Color vision tester. | N | N
ophthalmic | OBO | Tomography, Optical Coherence | 2 | 886.1570 | Ophthalmoscope. | N | N
ophthalmic | HPK | Tonograph | 2 | 886.1930 | Tonometer and accessories. | N | N
ophthalmic | HKX | Tonometer, Ac-Powered | 2 | 886.1930 | Tonometer and accessories. | N | N
ophthalmic | NJJ | Tonometer, Analyzer, Ocular Blood Flow | 2 | 886.1930 | Tonometer and accessories. | N | N
ophthalmic | HKY | Tonometer, Manual | 2 | 886.1930 | Tonometer and accessories. | N | N
ophthalmic | MMQ | Topographer, Corneal, Ac-Powered | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | NOG | Topographer, Corneal, Battery Powered | 1 | 886.1350 | Keratoscope. | N | N
ophthalmic | HMZ | Trabeculotome | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | HJM | Transilluminator, Ac-Powered | 2 | 886.1945 | Transilluminator. | N | N
ophthalmic | HJN | Transilluminator, Battery-Powered | 1 | 886.1945 | Transilluminator. | N | N
ophthalmic | HRH | Trephine, Manual, Ophthalmic | 1 | 886.4350 | Manual ophthalmic surgical instrument. | N | N
ophthalmic | MSR | Tubing, Replacement, Phacofragmentation Unit | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | HRL | Unit, Beta, Radiation, Ophthalmic | 2 | 886.5100 | Ophthalmic beta radiation source. | N | N
ophthalmic | HQO | Unit, Cautery, Thermal, Ac-Powered | 2 | 886.4115 | Thermal cautery unit. | N | N
ophthalmic | HQP | Unit, Cautery, Thermal, Battery-Powered | 2 | 886.4115 | Thermal cautery unit. | N | N
ophthalmic | HPS | Unit, Cryophthalmic | 2 | 886.4170 | Cryophthalmic unit. | N | N
ophthalmic | HRN | Unit, Cryophthalmic, Ac-Powered | 2 | 886.4170 | Cryophthalmic unit. | N | N
ophthalmic | HQA | Unit, Cryotherapy, Ophthalmic | 2 | 886.4170 | Cryophthalmic unit. | N | N
ophthalmic | HRO | Unit, Electrolysis, Ac-Powered, Ophthalmic | 2 | 886.4250 | Ophthalmic electrolysis unit. | N | N
ophthalmic | HPY | Unit, Electrolysis, Battery-Powered, Ophthalmic | 1 | 886.4250 | Ophthalmic electrolysis unit. | N | N
ophthalmic | HQC | Unit, Phacofragmentation | 2 | 886.4670 | Phacofragmentation system. | N | N
ophthalmic | QDY | Vitrectomy Kit | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | MLZ | Vitrectomy, Instrument Cutter | 2 | 886.4150 | Vitreous aspiration and cutting instrument. | N | N
ophthalmic | MML | Weights, Eyelid, External | 2 | 886.5700 | Eyelid weight. | N | N
ophthalmic | MTZ | Wrap, Implant, Orbital | 2 | 886.3320 | Eye sphere implant. | N | Y
orthopedic | NBH | Accessories, Arthroscopic | 1 | 888.1100 | Arthroscope. | N | N
orthopedic | LYP | Accessories, Fixation, Spinal Interlaminal | 2 | 888.3050 | Spinal interlaminal fixation orthosis. | N | Y
orthopedic | LYQ | Accessories, Fixation, Spinal Intervertebral Body | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | NJA | Accessory, System, External Fixator, Containing Antimicrobial Agent | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | PLW | Acetabular Cup Orientation System | 2 | 888.3350 | Hip joint metal/polymer semi-constrained cemented prosthesis. | N | N
orthopedic | NOV | Anchor, Suture, Bone Fixation, Metallic | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | OYK | Ankle Arthroplasty Implantation System | 2 | 888.3110 | Ankle joint metal/polymer semi-constrained cemented prosthesis. | N | N
orthopedic | PHQ | Anterior Staple As Supplemental Fixation For Fusion | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | HST | Apparatus, Traction, Non-Powered | 1 | 888.5850 | Nonpowered orthopedic traction apparatus and accessories. | N | N
orthopedic | KTT | Appliance, Fixation, Nail/Blade/Plate Combination, Multiple Component | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | LXT | Appliance, Fixation, Nail/Blade/Plate Combination, Multiple Component, Metal Composite | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | KTW | Appliance, Fixation, Nail/Blade/Plate Combination, Single Component | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | KWP | Appliance, Fixation, Spinal Interlaminal | 2 | 888.3050 | Spinal interlaminal fixation orthosis. | N | Y
orthopedic | KWQ | Appliance, Fixation, Spinal Intervertebral Body | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | KWK | Appliance, Nail/Blade/Plate Combination, Single Component | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | MRY | Appliances And Accessories, Fixation, Bone, Absorbable Single/Multiple Component | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HXN | Applier, Cerclage | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HRX | Arthroscope | 2 | 888.1100 | Arthroscope. | N | N
orthopedic | HWJ | Awl | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HSQ | Belt, Pelvic, Traction | 1 | 888.5890 | Noninvasive traction component. | N | N
orthopedic | HXW | Bender | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HTW | Bit, Drill | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | LOD | Bone Cement | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | MBB | Bone Cement, Antibiotic | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | PML | Bone Cement, Posterior Screw Augmentation | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | OWI | Bone Fixation Cerclage, Sublaminar | 2 | 888.3010 | Bone fixation cerclage. | N | Y
orthopedic | LYS | Bone Mill | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | PLF | Bone Wedge | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | NVC | Bone, Heterograft | 3 | 888.3015 | Bone heterograft. | N | Y
orthopedic | HXY | Brace, Drill | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HTQ | Broach | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HTT | Burr, Orthopedic | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | OIS | Calcium Salt Bone Void Filler, Drillable, Non-Screw Augmentation | 2 | 888.3045 | Resorbable calcium salt bone void filler device. | N | Y
orthopedic | KTZ | Caliper | 1 | 888.4150 | Calipers for clinical use. | N | N
orthopedic | JDT | Cap, Bone | 1 | 888.3000 | Bone cap. | N | Y
orthopedic | OBL | Cement, Bone, Pre-Formed, Modular, Polymeric, Vertebroplasty | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | NDN | Cement, Bone, Vertebroplasty | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | OBM | Cement, Bone, Vertebroplasty, Pre-Formed, Modular | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | JDQ | Cerclage, Fixation | 2 | 888.3010 | Bone fixation cerclage. | N | Y
orthopedic | NDD | Cerclage, Fixation, Metallic | 2 | 888.3010 | Bone fixation cerclage. | N | Y
orthopedic | LGF | Component, Cast | 1 | 888.5940 | Cast component. | N | N
orthopedic | JEC | Component, Traction, Invasive | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | KQZ | Component, Traction, Non-Invasive | 1 | 888.5890 | Noninvasive traction component. | N | N
orthopedic | JDP | Condylar Plate Fixation Implant | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | HWI | Corkscrew | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HWW | Countersink | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HXQ | Crimper, Pin | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HXZ | Cutter, Wire | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | QHG | Defect Measurement Tool | 2 | 888.4510 |  | N | N
orthopedic | JDO | Device, Fixation, Proximal Femoral, Implant | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | KIH | Dispenser, Cement | 1 | 888.4200 | Cement dispenser. | N | N
orthopedic | HWR | Driver, Prosthesis | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | LBB | Dynamometer, Ac-Powered | 2 | 888.1240 | AC-powered dynamometer. | N | N
orthopedic | HRW | Dynamometer, Nonpowered | 1 | 888.1250 | Nonpowered dynamometer. | N | N
orthopedic | OOL | Electronic Depth Gauge | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | N
orthopedic | JDY | Evacuator, Vapor, Cement Monomer | 1 | 888.4220 | Cement monomer vapor evacuator. | N | N
orthopedic | HWB | Extractor | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MAI | Fastener, Fixation, Biodegradable, Soft Tissue | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | MBI | Fastener, Fixation, Nondegradable, Soft Tissue | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HTP | File | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MQV | Filler, Bone Void, Calcium Compound | 2 | 888.3045 | Resorbable calcium salt bone void filler device. | N | Y
orthopedic | MBP | Filler, Bone Void, Osteoinduction (W/O Human Growth Factor) | 2 | 888.3045 | Resorbable calcium salt bone void filler device. | N | Y
orthopedic | LYT | Fixation Accessory | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | HXE | Fork | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HTJ | Gauge, Depth | 1 | 888.4300 | Depth gauge for clinical use. | N | N
orthopedic | NKI | Goniometer With Electrodes | 2 | 888.1500 | Goniometer. | N | N
orthopedic | KQX | Goniometer, Ac-Powered | 1 | 888.1500 | Goniometer. | N | N
orthopedic | KQW | Goniometer, Nonpowered | 1 | 888.1520 | Nonpowered goniometer. | N | N
orthopedic | HSS | Halter, Head, Traction | 1 | 888.5890 | Noninvasive traction component. | N | N
orthopedic | OQG | Hip Prosthesis, Semi-Constrained, Cemented, Metal/Polymer, + Additive, Porous, Uncemented | 2 | 888.3358 | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | N | Y
orthopedic | OQI | Hip, Semi-Constrained, Cemented, Metal/Ceramic/Polymer + Additive, Porous Uncemented | 2 | 888.3353 | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | N | Y
orthopedic | OQH | Hip, Semi-Constrained, Cemented, Metal/Polymer + Additive, Cemented | 2 | 888.3350 | Hip joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | HXK | Holder, Needle; Orthopedic | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HWA | Impactor | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | JDN | Implant, Fixation Device, Spinal | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | OAR | Injector, Vertebroplasty (Does Not Contain Cement) | 1 | 888.4200 | Cement dispenser. | N | N
orthopedic | HXP | Instrument, Bending Or Contouring | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | LGG | Instrument, Cast Application/Removal, Manual | 1 | 888.5980 | Manual cast application and removal instrument. | N | N
orthopedic | LGH | Instrument, Cast Removal, Ac-Powered | 1 | 888.5960 | Cast removal instrument. | N | N
orthopedic | HWN | Instrument, Compression | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | JDX | Instrument, Surgical, Sonic And Accessory/Attachment | 2 | 888.4580 | Sonic surgical instrument and accessories/attachments. | N | N
orthopedic | QBO | Instruments Designed For Press-Fit Osteochondral Implants | 2 | 888.4505 |  | N | N
orthopedic | OZI | Internal Hinged Elbow Fixator | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | ODP | Intervertebral Fusion Device With Bone Graft, Cervical | 2 | 888.3080 | Intervertebral body fusion device. | N | Y
orthopedic | MAX | Intervertebral Fusion Device With Bone Graft, Lumbar | 2 | 888.3080 | Intervertebral body fusion device. | N | Y
orthopedic | PHM | Intervertebral Fusion Device With Bone Graft, Thoracic | 2 | 888.3080 | Intervertebral body fusion device. | N | Y
orthopedic | OVE | Intervertebral Fusion Device With Integrated Fixation, Cervical | 2 | 888.3080 | Intervertebral body fusion device. | N | Y
orthopedic | OVD | Intervertebral Fusion Device With Integrated Fixation, Lumbar | 2 | 888.3080 | Intervertebral body fusion device. | N | Y
orthopedic | QFP | Intraoperative Orthopedic Strain Sensor | 2 | 888.3090 |  | N | N
orthopedic | LZE | Isometer | 1 | 888.1520 | Nonpowered goniometer. | N | N
orthopedic | OOG | Knee Arthroplasty Implantation System | 2 | 888.3560 | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | N | N
orthopedic | HTS | Knife, Orthopedic | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | NWN | Laparoscopic Bone Anchor Urethropexy Instrument Kit | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | N
orthopedic | QLQ | Manual Instruments Designed For Total Disc Replacement Device | 2 | 888.4515 |  | N | N
orthopedic | KRN | Metal Cemented Constrained Femorotibial Knee Prosthesis | 3 | 888.3480 | Knee joint femorotibial metallic constrained cemented prosthesis. | N | Y
orthopedic | KYM | Metallic Cemented Glenoid Hemi-Shoulder Prosthesis | 3 | 888.3680 | Shoulder joint glenoid (hemi-shoulder) metallic cemented prosthesis. | N | Y
orthopedic | JDZ | Mixer, Cement, For Clinical Use | 1 | 888.4210 | Cement mixer for clinical use. | N | N
orthopedic | JDS | Nail, Fixation, Bone | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | NDH | Nail, Fixation, Bone, Metallic | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | PBF | Orthopaedic Surgical Planning And Instrument Guides | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | N
orthopedic | LXH | Orthopedic Manual Surgical Instrument | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MNI | Orthosis, Spinal Pedicle Fixation | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | Y
orthopedic | NQW | Orthosis, Spine, Plate, Laminoplasty, Metal | 2 | 888.3050 | Spinal interlaminal fixation orthosis. | N | Y
orthopedic | MNH | Orthosis, Spondylolisthesis Spinal Fixation | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | Y
orthopedic | HWQ | Passer | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HXI | Passer, Wire, Orthopedic | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | PQC | Pedicle Screw Placement Guide | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | N
orthopedic | OSH | Pedicle Screw Spinal System, Adolescent Idiopathic Scoliosis | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | Y
orthopedic | OVZ | Pin, Fixation, Resorbable, Hard Tissue | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HTY | Pin, Fixation, Smooth | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | NDL | Pin, Fixation, Smooth, Metallic | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | JDW | Pin, Fixation, Threaded | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | NDM | Pin, Fixation, Threaded, Metallic | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | OBT | Plate, Bone, Growth Control, Pediatric, Epiphysiodesis | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | Y | Y
orthopedic | HRS | Plate, Fixation, Bone | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | NDF | Plate, Fixation, Bone, Non-Spinal, Metallic | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | KIL | Positioner, Socket | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | NKG | Posterior Cervical Screw System | 2 | 888.3075 | Posterior cervical screw system. | N | Y
orthopedic | NQP | Posterior Metal/Polymer Spinal System, Fusion | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | Y
orthopedic | HXB | Probe | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MAY | Prosthesis, Hip, Semi-Constrained, Metal/Ceramic/Polymer, Cemented Or Non-Porous Cemented, Osteophilic Finish | 2 | 888.3353 | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | N | Y
orthopedic | KXC | Prosthesis, Ankle, Cemented, Non-Constrained | 3 | 888.3120 | Ankle joint metal/polymer non-constrained cemented prosthesis. | N | Y
orthopedic | KMD | Prosthesis, Ankle, Semi-Constrained, Cemented, Metal/Composite | 2 | 888.3100 | Ankle joint metal/composite semi-constrained cemented prosthesis. | N | Y
orthopedic | HSN | Prosthesis, Ankle, Semi-Constrained, Cemented, Metal/Polymer | 2 | 888.3110 | Ankle joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | JDC | Prosthesis, Elbow, Constrained, Cemented | 2 | 888.3150 | Elbow joint metal/polymer constrained cemented prosthesis. | N | Y
orthopedic | KWJ | Prosthesis, Elbow, Hemi-, Humeral, Metal | 3 | 888.3180 | Elbow joint humeral (hemi-elbow) metallic uncemented prosthesis. | N | Y
orthopedic | KWI | Prosthesis, Elbow, Hemi-, Radial, Polymer | 2 | 888.3170 | Elbow joint radial (hemi-elbow) polymer prosthesis. | N | Y
orthopedic | JDB | Prosthesis, Elbow, Semi-Constrained, Cemented | 2 | 888.3160 | Elbow joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | LBC | Prosthesis, Finger, Constrained, Metal, Cemented | 3 | 888.3210 | Finger joint metal/metal constrained cemented prosthesis. | N | Y
orthopedic | KWE | Prosthesis, Finger, Constrained, Metal, Uncemented | 3 | 888.3200 | Finger joint metal/metal constrained uncemented prosthesis. | N | Y
orthopedic | KWG | Prosthesis, Finger, Constrained, Metal/Polymer | 3 | 888.3220 | Finger joint metal/polymer constrained cemented prosthesis. | N | Y
orthopedic | KYJ | Prosthesis, Finger, Constrained, Polymer | 2 | 888.3230 | Finger joint polymer constrained prosthesis. | N | Y
orthopedic | KWF | Prosthesis, Finger, Polymer | 2 | 888.3230 | Finger joint polymer constrained prosthesis. | N | Y
orthopedic | KWZ | Prosthesis, Hip, Constrained, Cemented Or Uncemented, Metal/Polymer | 2 | 888.3310 | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | N | Y
orthopedic | PBI | Prosthesis, Hip, Constrained, Cemented Or Uncemented, Metal/Polymer, + Additive | 2 | 888.3310 | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | N | Y
orthopedic | KXD | Prosthesis, Hip, Constrained, Metal | 3 | 888.3300 | Hip joint metal constrained cemented or uncemented prosthesis. | N | Y
orthopedic | JDG | Prosthesis, Hip, Femoral Component, Cemented, Metal | 2 | 888.3360 | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | N | Y
orthopedic | KXA | Prosthesis, Hip, Femoral, Resurfacing | 2 | 888.3400 | Hip joint femoral (hemi-hip) metallic resurfacing prosthesis. | N | Y
orthopedic | KWB | Prosthesis, Hip, Hemi-, Acetabular, Cemented, Metal | 3 | 888.3370 | Hip joint (hemi-hip) acetabular metal cemented prosthesis. | N | Y
orthopedic | KWL | Prosthesis, Hip, Hemi-, Femoral, Metal | 2 | 888.3360 | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | N | Y
orthopedic | LZY | Prosthesis, Hip, Hemi-, Femoral, Metal Ball | 2 | 888.3360 | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | N | Y
orthopedic | KWY | Prosthesis, Hip, Hemi-, Femoral, Metal/Polymer, Cemented Or Uncemented | 2 | 888.3390 | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | N | Y
orthopedic | JDH | Prosthesis, Hip, Hemi-, Trunnion-Bearing, Femoral, Metal/Polyacetal | 3 | 888.3380 | Hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal cemented prosthesis. | N | Y
orthopedic | KXB | Prosthesis, Hip, Pelvifemoral Resurfacing, Metal/Polymer | 3 | 888.3410 | Hip joint metal/polymer or ceramic/polymer semiconstrained resurfacing cemented prosthesis. | N | Y
orthopedic | OCG | Prosthesis, Hip, Pelvifemoral Resurfacing, Metal/Polymer, Uncemented | 3 | 888.3410 | Hip joint metal/polymer or ceramic/polymer semiconstrained resurfacing cemented prosthesis. | N | Y
orthopedic | JDL | Prosthesis, Hip, Semi-Constrained (Metal Cemented Acetabular Component) | 3 | 888.3320 | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | N | Y
orthopedic | KWA | Prosthesis, Hip, Semi-Constrained (Metal Uncemented Acetabular Component) | 3 | 888.3330 | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | N | Y
orthopedic | KMC | Prosthesis, Hip, Semi-Constrained, Composite/Metal | 2 | 888.3340 | Hip joint metal/composite semi-constrained cemented prosthesis. | N | Y
orthopedic | LZO | Prosthesis, Hip, Semi-Constrained, Metal/Ceramic/Polymer, Cemented Or Non-Porous, Uncemented | 2 | 888.3353 | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | N | Y
orthopedic | JDI | Prosthesis, Hip, Semi-Constrained, Metal/Polymer, Cemented | 2 | 888.3350 | Hip joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | LPH | Prosthesis, Hip, Semi-Constrained, Metal/Polymer, Porous Uncemented | 2 | 888.3358 | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | N | Y
orthopedic | LWJ | Prosthesis, Hip, Semi-Constrained, Metal/Polymer, Uncemented | 2 | 888.3360 | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | N | Y
orthopedic | MEH | Prosthesis, Hip, Semi-Constrained, Uncemented, Metal / Polymer, Non-Porous, Calcium Phosphate | 2 | 888.3353 | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | N | Y
orthopedic | MBL | Prosthesis, Hip, Semi-Constrained, Uncemented, Metal/Polymer, Porous | 2 | 888.3358 | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | N | Y
orthopedic | NPJ | Prosthesis, Knee Patellofemorotibial, Partial, Semi-Constrained, Cemented, Polymer/Metal/Polymer | 2 | 888.3560 | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | KRO | Prosthesis, Knee, Femorotibial, Constrained, Cemented, Metal/Polymer | 2 | 888.3510 | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | N | Y
orthopedic | HSX | Prosthesis, Knee, Femorotibial, Non-Constrained, Cemented, Metal/Polymer | 2 | 888.3520 | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | N | Y
orthopedic | KTX | Prosthesis, Knee, Femorotibial, Non-Constrained, Metal/Composite Cemented | 2 | 888.3490 | Knee joint femorotibial metal/composite non-constrained cemented prosthesis. | N | Y
orthopedic | KYK | Prosthesis, Knee, Femorotibial, Semi-Constrained, Cemented, Metal/Composite | 2 | 888.3500 | Knee joint femorotibial metal/composite semi-constrained cemented prosthesis. | N | Y
orthopedic | HRY | Prosthesis, Knee, Femorotibial, Semi-Constrained, Cemented, Metal/Polymer | 2 | 888.3530 | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | LGE | Prosthesis, Knee, Femorotibial, Semi-Constrained, Cemented, Trunnion-Bearing | 2 | 888.3530 | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | NJD | Prosthesis, Knee, Femorotibial, Unicompartmental/Unicondylar, Uncemented, Porous-Coated, Metal/Polymer | 2 | 888.3535 | Knee joint femorotibial (uni-compartmental) metal/polymer porous-coated uncemented prosthesis. | N | Y
orthopedic | HSA | Prosthesis, Knee, Hemi-, Femoral | 3 | 888.3570 | Knee joint femoral (hemi-knee) metallic uncemented prosthesis. | N | Y
orthopedic | KRS | Prosthesis, Knee, Hemi-, Femoral (Uncemented) | 3 | 888.3570 | Knee joint femoral (hemi-knee) metallic uncemented prosthesis. | N | Y
orthopedic | HTG | Prosthesis, Knee, Hemi-, Patellar Resurfacing, Uncemented | 2 | 888.3580 | Knee joint patellar (hemi-knee) metallic resurfacing uncemented prosthesis. | N | Y
orthopedic | HSH | Prosthesis, Knee, Hemi-, Tibial, Resurfacing (Uncemented) | 2 | 888.3590 | Knee joint tibial (hemi-knee) metallic resurfacing uncemented prosthesis. | N | Y
orthopedic | HRZ | Prosthesis, Knee, Hinged (Metal-Metal) | 3 | 888.3480 | Knee joint femorotibial metallic constrained cemented prosthesis. | N | Y
orthopedic | KMB | Prosthesis, Knee, Non-Constrained (Metal-Carbon Reinforced Polyethylene) Cemented | 2 | 888.3490 | Knee joint femorotibial metal/composite non-constrained cemented prosthesis. | N | Y
orthopedic | KRR | Prosthesis, Knee, Patello/Femoral, Semi-Constrained, Cemented, Metal/Polymer | 2 | 888.3540 | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | N | Y
orthopedic | KRP | Prosthesis, Knee, Patello/Femorotibial, Constrained, Cemented, Polymer/Metal/Metal | 3 | 888.3550 | Knee joint patellofemorotibial polymer/metal/metal constrained cemented prosthesis. | N | Y
orthopedic | KRQ | Prosthesis, Knee, Patello/Femorotibial, Constrained, Cemented, Polymer/Metal/Polymer | 2 | 888.3550 | Knee joint patellofemorotibial polymer/metal/metal constrained cemented prosthesis. | N | Y
orthopedic | MBV | Prosthesis, Knee, Patello/Femorotibial, Semi-Constrained, Uhmwpe, Pegged, Cemented, Polymer/Metal/Polymer | 2 | 888.3560 | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | MBH | Prosthesis, Knee, Patello/Femorotibial, Semi-Constrained, Uncemented, Porous, Coated, Polymer/Metal/Polymer | 2 | 888.3565 | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | N | Y
orthopedic | OIY | Prosthesis, Knee, Patellofemorotibial, Semi-Constrained, Cemented, Polymer + Additive/Metal/Polymer + Additive | 2 | 888.3560 | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | JWH | Prosthesis, Knee, Patellofemorotibial, Semi-Constrained, Cemented, Polymer/Metal/Polymer | 2 | 888.3560 | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | KWR | Prosthesis, Shoulder, Constrained, Metal/Metal Or Metal/Polymer Cemented | 3 | 888.3640 | Shoulder joint metal/metal or metal/polymer constrained cemented prosthesis. | N | Y
orthopedic | HSD | Prosthesis, Shoulder, Hemi-, Humeral, Metallic Uncemented | 2 | 888.3690 | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | N | Y
orthopedic | KWT | Prosthesis, Shoulder, Non-Constrained, Metal/Polymer Cemented | 2 | 888.3650 | Shoulder joint metal/polymer non-constrained cemented prosthesis. | N | Y
orthopedic | PAO | Prosthesis, Shoulder, Semi-Constrained, Metal/Polymer + Additive, Cemented | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | KWS | Prosthesis, Shoulder, Semi-Constrained, Metal/Polymer Cemented | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | MBF | Prosthesis, Shoulder, Semi-Constrained, Metal/Polymer, Uncemented | 2 | 888.3670 | Shoulder joint metal/polymer/metal nonconstrained or semi-constrained porous-coated uncemented prosthesis. | N | Y
orthopedic | MJW | Prosthesis, Subtalar, Plug, Polymer | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HXA | Prosthesis, Tendon, Passive | 2 | 888.3025 | Passive tendon prosthesis. | N | Y
orthopedic | KWH | Prosthesis, Toe, Constrained, Polymer | 2 | 888.3720 | Toe joint polymer constrained prosthesis. | N | Y
orthopedic | KWD | Prosthesis, Toe, Hemi-, Phalangeal | 2 | 888.3730 | Toe joint phalangeal (hemi-toe) polymer prosthesis. | N | Y
orthopedic | PKC | Prosthesis, Total Anatomic Shoulder, Uncemented Metaphyseal Humeral Stem With No Diaphyseal Incursion, Semi-Constrained | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | JDD | Prosthesis, Upper Femoral | 2 | 888.3360 | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | N | Y
orthopedic | JWI | Prosthesis, Wrist, 2 Part Metal-Plastic Articulation, Semi-Constrained | 2 | 888.3800 | Wrist joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | JWJ | Prosthesis, Wrist, 3 Part Metal-Plastic-Metal Articulation, Semi-Constrained | 2 | 888.3800 | Wrist joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | KWN | Prosthesis, Wrist, Carpal Lunate | 2 | 888.3750 | Wrist joint carpal lunate polymer prosthesis. | N | Y
orthopedic | KWO | Prosthesis, Wrist, Carpal Scaphoid | 2 | 888.3760 | Wrist joint carpal scaphoid polymer prosthesis. | N | Y
orthopedic | KYI | Prosthesis, Wrist, Carpal Trapezium | 2 | 888.3770 | Wrist joint carpal trapezium polymer prosthesis. | N | Y
orthopedic | KYN | Prosthesis, Wrist, Constrained, Metal | 3 | 888.3790 | Wrist joint metal constrained cemented prosthesis. | N | Y
orthopedic | KIG | Prosthesis, Wrist, Constrained, Polymer | 2 | 888.3780 | Wrist joint polymer constrained prosthesis. | N | Y
orthopedic | KXE | Prosthesis, Wrist, Hemi-, Ulnar | 2 | 888.3810 | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | N | Y
orthopedic | KWM | Prosthesis, Wrist, Semi-Constrained | 2 | 888.3800 | Wrist joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | HTH | Protractor | 1 | 888.4600 | Protractor for clinical use. | N | N
orthopedic | HWP | Punch, Femoral Neck | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HXO | Pusher, Socket | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HTO | Reamer | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | OJM | Resorbable Spinal Intervertebral Body Fixation Orthosis - Mesh | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | OJB | Resorbable Spinal Intervertebral Body Fixation Orthosis - Plate | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | HSB | Rod, Fixation, Intramedullary And Accessories | 2 | 888.3020 | Intramedullary fixation rod. | N | Y
orthopedic | QAD | Rod, Fixation, Intramedullary And Accessories, In-Vivo Cured, Light-Activated | 2 | 888.3023 | In vivo cured intramedullary fixation rod. | N | Y
orthopedic | NDE | Rod, Fixation, Intramedullary And Accessories, Metallic And Non-Collapsible | 2 | 888.3020 | Intramedullary fixation rod. | N | Y
orthopedic | HTX | Rongeur | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | OUR | Sacroiliac Joint Fixation | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HRR | Scissors, Orthopedic, Surgical | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | QAC | Screw Sleeve Bone Fixation Device | 2 | 888.3043 |  | N | Y
orthopedic | HWC | Screw, Fixation, Bone | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | NDJ | Screw, Fixation, Bone, Non-Spinal, Metallic | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HXX | Screwdriver | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HWL | Set, Hollow Mill | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | QHE | Shoulder Arthroplasty Implantation System | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | N
orthopedic | PHX | Shoulder Prosthesis, Reverse Configuration | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | HWO | Skid, Bone | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | OSN | Software For Diagnosis/Treatment | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | LTO | Spacer, Cement | 3 | 888.3320 | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | N | Y
orthopedic | OCJ | Spinal Channeling Instrument, Vertebroplasty | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MQP | Spinal Vertebral Body Replacement Device | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | PLR | Spinal Vertebral Body Replacement Device - Cervical | 2 | 888.3060 | Spinal intervertebral body fixation orthosis. | N | Y
orthopedic | PEK | Spinous Process Plate | 2 | 888.3050 | Spinal interlaminal fixation orthosis. | N | Y
orthopedic | HSP | Splint, Traction | 1 | 888.5890 | Noninvasive traction component. | N | N
orthopedic | HXJ | Staple Driver | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | MNU | Staple, Absorbable | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | JDR | Staple, Fixation, Bone | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | NDI | Staple, Fixation, Bone, Metallic | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | HWD | Starter, Bone Screw | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HRT | Stripper, Surgical | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | PPR | Surgical Cannulaes & Depth Gauge Kit For Arthroscopic Procedures | 2 | 888.1100 | Arthroscope. | N | N
orthopedic | LZV | System, Cement Removal Extraction | 2 | 888.4580 | Sonic surgical instrument and accessories/attachments. | N | N
orthopedic | NDK | System, External Fixator (With Metallic Invasive Components) | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | Y
orthopedic | HXG | Tamp | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HWX | Tap, Bone | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HWT | Template | 1 | 888.4800 | Template for clinical use. | N | N
orthopedic | NGR | Tester, Stiffness, Cartilage, Arthroscopic | 2 | 888.1100 | Arthroscope. | N | N
orthopedic | NKB | Thoracolumbosacral Pedicle Screw System | 2 | 888.3070 | Thoracolumbosacral pedicle screw system. | N | Y
orthopedic | QHQ | Total Shoulder Arthroplasty System | 2 | 888.3660 | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | N | Y
orthopedic | PPQ | Transtibial Anterior Cruciate Ligament Kit | 2 | 888.3040 | Smooth or threaded metallic bone fixation fastener. | N | N
orthopedic | HWK | Trephine | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | KII | Tube, Cement Ventilation | 1 | 888.4230 | Cement ventilation tube. | N | N
orthopedic | HXS | Twister, Wire | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
orthopedic | HSR | Unit, Traction, Hip, Non-Powered, Non-Penetrating | 1 | 888.5890 | Noninvasive traction component. | N | N
orthopedic | NUD | Vertebroplasty Compound, Calcium Based | 2 | 888.3027 | Polymethylmethacrylate (PMMA) bone cement. | N | Y
orthopedic | HTN | Washer, Bolt Nut | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | NDG | Washer, Bolt, Nut, Non-Spinal, Metallic | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | LRN | Wire, Surgical | 2 | 888.3030 | Single/multiple component metallic bone fixation appliances and accessories. | N | Y
orthopedic | HXC | Wrench | 1 | 888.4540 | Orthopedic manual surgical instrument. | N | N
Physical medicine | ILZ | Accessories, Traction | 1 | 890.5925 | Traction accessory. | N | N
Physical medicine | KNO | Accessories, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | ILD | Adaptor, Dressing | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | ILW | Adaptor, Grooming | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | IQG | Adaptor, Holder, Syringe | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | ILS | Adaptor, Hygiene | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | ILT | Adaptor, Recreational | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | OMW | Adhesive Backed Thermal Skin Patches | 1 | 890.5710 | Hot or cold disposable pack. | N | N
Physical medicine | IKX | Aid, Transfer | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | IMY | Armboard, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | IML | Armrest, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | LYH | Arthrometer | 1 | 890.1615 | Miniature pressure transducer. | N | N
Physical medicine | ISW | Assembly, Knee/Shank/Ankle/Foot, External | 2 | 890.3500 | External assembled lower limb prosthesis. | N | N
Physical medicine | KFT | Assembly, Shoulder/Elbow/Forearm/Wrist/Hand, Mechanical | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | KFW | Assembly, Shoulder/Elbow/Forearm/Wrist/Hand, Powered | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | KFX | Assembly, Thigh/Knee/Shank/Ankle/Foot, External | 2 | 890.3500 | External assembled lower limb prosthesis. | N | N
Physical medicine | INB | Attachment, Commode, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMQ | Attachment, Narrowing, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | ISR | Band Or Belt, Pelvic Support | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | ITG | Bandage, Cast | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | IOE | Bars, Parallel, Exercise | 1 | 890.5370 | Nonmeasuring exercise equipment. | N | N
Physical medicine | IRR | Bars, Parallel, Powered | 1 | 890.5380 | Powered exercise equipment. | N | N
Physical medicine | ILJ | Bath, Hydro-Massage | 2 | 890.5100 | Immersion hydrobath. | N | N
Physical medicine | IMC | Bath, Paraffin | 2 | 890.5110 | Paraffin bath. | N | N
Physical medicine | KTC | Bath, Sitz, Nonpowered | 1 | 890.5125 | Nonpowered sitz bath. | N | N
Physical medicine | ILM | Bath, Sitz, Powered | 2 | 890.5100 | Immersion hydrobath. | N | N
Physical medicine | INX | Bed, Air Fluidized | 2 | 890.5160 | Air-fluidized bed. | N | N
Physical medicine | IOQ | Bed, Flotation Therapy, Powered | 2 | 890.5170 | Powered flotation therapy bed. | N | N
Physical medicine | INY | Bed, Patient Rotation, Manual | 1 | 890.5180 | Manual patient rotation bed. | N | N
Physical medicine | IKZ | Bed, Patient Rotation, Powered | 2 | 890.5225 | Powered patient rotation bed. | N | N
Physical medicine | IQB | Belt, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMX | Board, Lap, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | KNL | Board, Scooter, Prone | 1 | 890.5370 | Nonmeasuring exercise equipment. | N | N
Physical medicine | IMW | Brake, Extension, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMB | Cabinet, Moist Steam | 2 | 890.5250 | Moist steam cabinet. | N | N
Physical medicine | ISN | Cable | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IKD | Cable, Electrode | 2 | 890.1175 | Electrode cable. | N | N
Physical medicine | ITM | Cage, Knee | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IPS | Cane | 1 | 890.3075 | Cane. | N | N
Physical medicine | KHY | Cane, Safety Walk | 1 | 890.3075 | Cane. | N | N
Physical medicine | INN | Chair, Adjustable, Mechanical | 1 | 890.3100 | Mechanical chair. | N | N
Physical medicine | INO | Chair, Positioning, Electric | 2 | 890.3110 | Electric positioning chair. | N | N
Physical medicine | INM | Chair, With Casters | 1 | 890.3100 | Mechanical chair. | N | N
Physical medicine | IKP | Chronaximeter | 2 | 890.1225 | Chronaximeter. | N | N
Physical medicine | IMN | Climber, Curb, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | PHW | Cold Pack And Vibrating Massager | 1 | 890.5975 | Therapeutic vibrator. | N | N
Physical medicine | ISH | Component, External, Limb, Ankle/Foot | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IOD | Components, Exercise | 1 | 890.5350 | Exercise component. | N | N
Physical medicine | KNN | Components, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IPM | Cover, Limb | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | IPR | Crutch | 1 | 890.3150 | Crutch. | N | N
Physical medicine | INC | Cuff, Pusher, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | KIC | Cushion, Flotation | 1 | 890.3175 | Flotation cushion. | N | N
Physical medicine | MOC | Cushion, Flotation, Therapeutic | 2 | 890.5170 | Powered flotation therapy bed. | N | N
Physical medicine | IMP | Cushion, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMR | Device, Anti-Tip, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | NHH | Device, Discharge, Electrostatic (For Pain Relief) | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | LQX | Device, Finger-Sucking | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | LBG | Device, Fluidized Therapy, Dry Heat | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | EGJ | Device, Iontophoresis, Other Uses | 2 | 890.5525 | Iontophoresis device. | N | N
Physical medicine | KTB | Device, Iontophoresis, Specific Uses | 2 | 890.5525 | Iontophoresis device. | N | N
Physical medicine | KZM | Device, Muscle Monitoring | 2 | 890.1375 | Diagnostic electromyograph. | N | N
Physical medicine | KNM | Device, Pressure Applying | 1 | 890.5765 | Pressure-applying device. | N | N
Physical medicine | IQO | Device, Prosthesis Alignment | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | IRN | Device, Warning, Overload, External Limb, Powered | 2 | 890.5575 | Powered external limb overload warning device. | N | N
Physical medicine | IOA | Diathermy, Microwave, For Use In Applying Therapeutic Deep Heat | 2 | 890.5275 | Microwave diathermy. | N | N
Physical medicine | HPH | Diathermy, Microwave, For Use Other Than Applying Therapeutic Deep Heat | 3 | 890.5275 | Microwave diathermy. | N | N
Physical medicine | IMJ | Diathermy, Shortwave, For Use In Applying Therapeutic Deep Heat | 2 | 890.5290 | Shortwave diathermy. | N | N
Physical medicine | ILX | Diathermy, Shortwave, For Use Other Than Applying Therapeutic Deep Heat | 2 | 890.5290 | Shortwave diathermy. | N | N
Physical medicine | LXF | Diathermy, Ultrasonic, For Use Other Than Applying Therapeutic Deep Heat | 3 | 890.5300 | Ultrasonic diathermy. | N | N
Physical medicine | NXB | Dispenser, Solid Medication | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | IKT | Electrode, Needle, Diagnostic Electromyograph | 2 | 890.1385 | Diagnostic electromyograph needle electrode. | N | N
Physical medicine | IKN | Electromyograph, Diagnostic | 2 | 890.1375 | Diagnostic electromyograph. | N | N
Physical medicine | ING | Elevator, Wheelchair, Portable | 2 | 890.3930 | Wheelchair elevator. | N | N
Physical medicine | ITH | Equipment, Traction, Powered | 2 | 890.5900 | Power traction equipment. | N | N
Physical medicine | OAL | Exercise Equipment, Powered, Emg-Triggered | 2 | 890.1375 | Diagnostic electromyograph. | N | N
Physical medicine | JFA | Exerciser, Finger, Powered | 1 | 890.5410 | Powered finger exerciser. | N | N
Physical medicine | ISD | Exerciser, Measuring | 2 | 890.5360 | Measuring exercise equipment. | N | N
Physical medicine | ION | Exerciser, Non-Measuring | 1 | 890.5370 | Nonmeasuring exercise equipment. | N | N
Physical medicine | PKS | Exerciser, Non-Measuring For Stroke Rehabilitation | 1 | 890.5370 | Nonmeasuring exercise equipment. | N | N
Physical medicine | BXB | Exerciser, Powered | 1 | 890.5380 | Powered exercise equipment. | N | N
Physical medicine | OIZ | Foot Examination Tool For Inflammatory Changes | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | OTX | Foot Wrap For Treating Restless Leg Syndrome Symptoms | 1 | 890.5760 | Nonpowered lower extremity pressure wrap. | N | N
Physical medicine | IMM | Footrest, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | OAG | Garment, Storage, Peritoneal Dialysis Catheter | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | IKO | Hammer, Reflex, Powered | 2 | 890.1450 | Powered reflex hammer. | N | N
Physical medicine | IRA | Hand, External Limb Component, Mechanical | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IQZ | Hand, External Limb Component, Powered | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IMO | Handrim, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IRS | Head Halter, Traction | 1 | 890.5925 | Traction accessory. | N | N
Physical medicine | NDA | Helmet, Cranial, For Protective Use | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | INA | Hill Holder, Wheelchair | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMZ | Holder, Crutch And Cane, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | IQX | Hook, External Limb Component, Mechanical | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IQW | Hook, External Limb Component, Powered | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | MPO | Infant Heel Warmer (Chemical Heat Pack) | 1 | 890.5710 | Hot or cold disposable pack. | N | N
Physical medicine | OJL | Insulin Vial Protector/Holder | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | LXJ | Interactive Rehabilitation Exercise Devices | 2 | 890.5360 | Measuring exercise equipment. | N | N
Physical medicine | ITW | Joint, Ankle, External Brace | 1 | 890.3475 | Limb orthosis. | N | N | DJO
Physical medicine | IRD | Joint, Elbow, External Limb Component, Mechanical | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IRE | Joint, Elbow, External Limb Component, Powered | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | ITS | Joint, Hip, External Brace | 1 | 890.3475 | Limb orthosis. | N | N | DJO
Physical medicine | ISL | Joint, Hip, External Limb Component | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | ITQ | Joint, Knee, External Brace | 1 | 890.3475 | Limb orthosis. | N | N | DJO
Physical medicine | ISY | Joint, Knee, External Limb Component | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IQQ | Joint, Shoulder, External Limb Component | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | IOB | Lamp, Infrared, Non Heating | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | ILY | Lamp, Infrared, Therapeutic Heating | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | OAP | Laser, Comb, Hair | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | NZY | Light, Lymphedema Reduction, Low Energy | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | IRP | Massager, Powered Inflatable Tube | 2 | 890.5650 | Powered inflatable tube massager. | N | N
Physical medicine | ISA | Massager, Therapeutic, Electric | 1 | 890.5660 | Therapeutic massager. | N | N
Physical medicine | LYG | Massager, Therapeutic, Manual | 1 | 890.5660 | Therapeutic massager. | N | N
Physical medicine | OSD | Massager, Therapeutic, To Internally Massage Trigger Points In The Pelvic Floor Musculature | 2 | 890.5670 |  | N | N
Physical medicine | NHF | Monitor, Orthotic/Prosthetic | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | PQY | Nonthermal Shortwave Therapy Device Indicated For Over The Counter Use For The Treatment Of Pain | 2 | 890.5290 | Shortwave diathermy. | N | N
Physical medicine | KTD | Orthosis, Abdominal | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | IQK | Orthosis, Cervical | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | IQF | Orthosis, Cervical-Thoracic, Rigid | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | KNP | Orthosis, Corrective Shoe | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | OHI | Orthosis, Foot Drop | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IQI | Orthosis, Limb Brace | 1 | 890.3475 | Limb orthosis. | N | N | DJO
Physical medicine | PMV | Orthosis, Limb, For Back Pain | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IQE | Orthosis, Lumbar | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | IPY | Orthosis, Lumbo-Sacral | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | MNE | Orthosis, Moldable, Supportive, Skin Protective | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IPO | Orthosis, Pneumatic Structure, Rigid | 3 | 890.3610 | Rigid pneumatic structure orthosis. | N | N
Physical medicine | IPX | Orthosis, Rib Fracture, Soft | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | IPW | Orthosis, Sacroiliac, Soft | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | IPT | Orthosis, Thoracic | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | NJB | Orthosis, Truncal, For Dysmenorrhea | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | MRI | Orthosis, Truncal/Orthosis, Limb | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | OZC | Over The Counter Thermal Treatment For Acne. | 2 | 890.5740 | Powered heating pad. | N | N
Physical medicine | IMA | Pack, Heat, Moist | 1 | 890.5730 | Moist heat pack. | N | N
Physical medicine | IMD | Pack, Hot Or Cold, Disposable | 1 | 890.5710 | Hot or cold disposable pack. | N | N
Physical medicine | IME | Pack, Hot Or Cold, Reusable | 1 | 890.5700 | Cold pack. | N | N
Physical medicine | ILO | Pack, Hot Or Cold, Water Circulating | 2 | 890.5720 | Water circulating hot or cold pack. | N | N
Physical medicine | IRT | Pad, Heating, Powered | 2 | 890.5740 | Powered heating pad. | N | N
Physical medicine | PCE | Permanently Mounted Wheelchair Platform Lift | 2 | 890.3930 | Wheelchair elevator. | N | N
Physical medicine | KHX | Platform, Force-Measuring | 1 | 890.1575 | Force-measuring platform. | N | N
Physical medicine | INT | Plinth | 1 | 890.3520 | Plinth. | N | N
Physical medicine | PCM | Portable Liquid Oxygen Unit Stand | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | PHL | Powered Exoskeleton | 2 | 890.3480 | Powered lower extremity exoskeleton. | N | N
Physical medicine | NHN | Powered Light Based Laser Non-Thermal Instrument With Non-Heating Effect For Adjunctive Use In Pain Therapy | 2 | 890.5500 | Infrared lamp. | N | N
Physical medicine | PCD | Powered Stairway Chair Lift | 2 | 890.5150 | Powered patient transport. | N | N
Physical medicine | ISS | Prosthesis, Thigh Socket, External Component | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | ISM | Pylon, Post Surgical | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | NXQ | Reminder, Medication | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | KID | Restraint, Wheelchair, Non-Protective | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | IQP | Rotator, Transverse | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | INF | Scale, Platform, Wheelchair | 1 | 890.3940 | Wheelchair platform scale. | N | N
Physical medicine | QDT | Shoe With Adjustable Sole Units | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IPG | Shoe, Cast | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | ILI | Sling, Arm | 1 | 890.3640 | Arm sling. | N | N
Physical medicine | ILE | Sling, Arm, Overhead Supported | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | INE | Sling, Overhead Suspension, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | IOZ | Splint, Abduction, Congenital Hip Dislocation | 1 | 890.3665 | Congenital hip dislocation abduction splint. | N | N
Physical medicine | IQJ | Splint, Clavicle | 1 | 890.3490 | Truncal orthosis. | N | N
Physical medicine | ITN | Splint, Denis Brown | 1 | 890.3675 | Denis Brown splint. | N | N
Physical medicine | ILH | Splint, Hand, And Components | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IQM | Splint, Temporary Training | 1 | 890.3025 | Prosthetic and orthotic accessory. | N | N
Physical medicine | PFW | Stationary Ultrasonic Diathermy Device For Use In Applying Therapeutic Deep Heat | 2 | 890.5300 | Ultrasonic diathermy. | N | N
Physical medicine | ISB | Stimulator, Muscle, Diagnostic | 2 | 890.1850 | Diagnostic muscle stimulator. | N | N
Physical medicine | IPF | Stimulator, Muscle, Powered | 2 | 890.5850 | Powered muscle stimulator. | N | N
Physical medicine | NUW | Stimulator, Muscle, Powered, Dental | 2 | 890.5850 | Powered muscle stimulator. | N | N
Physical medicine | NGX | Stimulator, Muscle, Powered, For Muscle Conditioning | 2 | 890.5850 | Powered muscle stimulator. | N | N
Physical medicine | NYZ | Stimulator, Muscle, Powered, With Limited Output, For Muscle Conditioning | 2 | 890.5850 | Powered muscle stimulator. | N | N
Physical medicine | NYY | Stimulator, Muscle, Powered, With Limited Output, For Rehabilitation | 2 | 890.5850 | Powered muscle stimulator. | N | N
Physical medicine | IMG | Stimulator, Ultrasound And Muscle, For Use In Applying Therapeutic Deep Heat | 2 | 890.5860 | Ultrasound and muscle stimulator. | N | N
Physical medicine | ITC | Stirrup, External Brace Component | 1 | 890.3410 | External limb orthotic component. | N | N | DJO
Physical medicine | ILG | Stocking, Elastic | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | INK | Stretcher, Wheeled, Powered | 2 | 890.3690 | Powered wheeled stretcher. | N | N
Physical medicine | LBE | Stroller, Adaptive | 1 | 890.3850 | Mechanical wheelchair. | N | N
Physical medicine | IOY | Support, Arm | 1 | 890.3475 | Limb orthosis. | N | N
Physical medicine | IMS | Support, Head And Trunk, Wheelchair | 1 | 890.3910 | Wheelchair accessory. | N | N
Physical medicine | ILP | System, Communication, Non-Powered | 1 | 890.3700 | Nonpowered communication system. | N | N
Physical medicine | ILQ | System, Communication, Powered | 2 | 890.3710 | Powered communication system. | N | N
Physical medicine | IQA | System, Environmental Control, Powered | 2 | 890.3725 | Powered environmental control system. | N | N
Physical medicine | IKK | System, Isokinetic Testing And Evaluation | 2 | 890.1925 | Isokinetic testing and evaluation system. | N | N
Physical medicine | JFC | System, Pressure Measurement, Intermittent | 1 | 890.1600 | Intermittent pressure measurement system. | N | N
Physical medicine | INW | Table, Mechanical | 1 | 890.3750 | Mechanical table. | N | N
Physical medicine | JFB | Table, Physical Therapy, Multi Function | 2 | 890.5880 | Multi-function physical therapy table. | N | N
Physical medicine | INQ | Table, Powered | 1 | 890.3760 | Powered table. | N | N
Physical medicine | INP | Tips And Pads, Cane, Crutch And Walker | 1 | 890.3790 | Cane, crutch, and walker tips and pads. | N | N
Physical medicine | IKE | Transducer, Miniature Pressure | 1 | 890.1615 | Miniature pressure transducer. | N | N
Physical medicine | ILK | Transport, Patient, Powered | 2 | 890.5150 | Powered patient transport. | N | N
Physical medicine | IOG | Treadmill, Mechanical | 1 | 890.5370 | Nonmeasuring exercise equipment. | N | N
Physical medicine | IOL | Treadmill, Powered | 1 | 890.5380 | Powered exercise equipment. | N | N
Physical medicine | ITO | Twister, Brace Setting | 1 | 890.3410 | External limb orthotic component. | N | N | DJO
Physical medicine | IMI | Ultrasonic Diathermy For Use In Applying Therapeutic Deep Heat | 2 | 890.5300 | Ultrasonic diathermy. | N | N
Physical medicine | LFF | Unit, Activation, Breath Control | 1 | 890.3920 | Wheelchair component. | N | N
Physical medicine | IMF | Unit, Chilling | 1 | 890.5940 | Chilling unit. | N | N
Physical medicine | LSB | Unit, Fluidotherapy | 2 | 890.5100 | Immersion hydrobath. | N | N
Physical medicine | IRQ | Unit, Heating, Powered | 1 | 890.5950 | Powered heating unit. | N | N
Physical medicine | ISZ | Unit, Wrist, External Limb Component, Mechanical | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | KGH | Unit, Wrist, External Limb Component, Powered | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | PAE | Upper Extremity Prosthesis With Multiple Simultaneous Degrees Of Freedom And Controlled Via Cutaneous Electromyography | 2 | 890.3450 | Upper extremity prosthesis including a simultaneously powered elbow and/or shoulder with greater than two simultaneous powered degrees of freedom and controlled by non-implanted electrical components. | N | N
Physical medicine | ILC | Utensil, Eating | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | IKW | Utensil, Homemaking | 1 | 890.5050 | Daily activity assist device. | N | N
Physical medicine | ISP | Valve, Prosthesis | 1 | 890.3420 | External limb prosthetic component. | N | N
Physical medicine | INI | Vehicle, Motorized 3-Wheeled | 2 | 890.3800 | Motorized three-wheeled vehicle. | N | N
Physical medicine | IRO | Vibrator, Therapeutic | 1 | 890.5975 | Therapeutic vibrator. | N | N
Physical medicine | OVP | Vibratory Counter-Stimulation | 2 | 882.5895 | Vibratory counter-stimulation device. | N | N
Physical medicine | ITJ | Walker, Mechanical | 1 | 890.3825 | Mechanical walker. | N | N
Physical medicine | NXE | Walker, Mechanical, Poly Vinyl Chloride (Pvc) | 1 | 890.3825 | Mechanical walker. | N | N
Physical medicine | LHC | Warmer, Irrigation Solution | 1 | 890.5950 | Powered heating unit. | N | N
Physical medicine | IOR | Wheelchair, Mechanical | 1 | 890.3850 | Mechanical wheelchair. | N | N
Physical medicine | ITI | Wheelchair, Powered | 2 | 890.3860 | Powered wheelchair. | N | N
Physical medicine | IQC | Wheelchair, Special Grade | 2 | 890.3880 | Special grade wheelchair. | N | N
Physical medicine | IMK | Wheelchair, Stair Climbing | 2 | 890.3890 | Stair-climbing wheelchair. | N | N
Physical medicine | IPL | Wheelchair, Standup | 2 | 890.3900 | Standup wheelchair. | N | N
Radiology | IYE | Accelerator, Linear, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | MWW | Accessory - Film Dosimetry System | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | MYN | Analyzer, Medical Image | 2 | 892.2070 |  | N | N
Radiology | IZS | Aperture, Radiographic | 2 | 892.1610 | Diagnostic x-ray beam-limiting device. | N | N
Radiology | EAJ | Apron, Leaded | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IWO | Apron, Protective | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | OII | Arthrogram Tray | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | ITY | Assembly, Tube Housing, X-Ray, Diagnostic | 1 | 892.1760 | Diagnostic x-ray tube housing assembly. | N | N
Radiology | ITZ | Assembly, Tube Housing, X-Ray, Therapeutic | 2 | 892.5930 | Therapeutic x-ray tube housing assembly. | N | N
Radiology | QIH | Automated Radiological Image Processing Software | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | IWX | Barrier, Control Panel, X-Ray, Movable | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IYZ | Bed, Scanning, Nuclear | 1 | 892.1350 | Nuclear scanning bed. | N | N
Radiology | IYG | Betatron, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | OIJ | Biopsy Needle Guide Kit | 2 | 892.1560 | Ultrasonic pulsed echo imaging system. | N | N
Radiology | IXI | Block, Beam-Shaping, Radiation Therapy | 2 | 892.5710 | Radiation therapy beam-shaping block. | N | N
Radiology | MUA | Bone Sonometer | 2 | 892.1180 | Bone sonometer. | N | N
Radiology | MWP | Cabinet, X-Ray System | 2 | 892.1680 | Stationary x-ray system. | N | N
Radiology | KPT | Calibrator, Dose, Radionuclide | 2 | 892.1360 | Radionuclide dose calibrator. | N | N
Radiology | IXH | Camera, Focal Spot, Radiographic | 1 | 892.1940 | Radiologic quality assurance instrument. | N | N
Radiology | LMC | Camera, Multi Format, Radiological | 2 | 892.2040 | Medical image hardcopy device. | N | N
Radiology | IZC | Camera, Positron | 1 | 892.1110 | Positron camera. | N | N
Radiology | IYX | Camera, Scintillation (Gamma) | 1 | 892.1100 | Scintillation (gamma) camera. | N | N
Radiology | IZJ | Camera, X-Ray, Fluorographic, Cine Or Spot | 2 | 892.1620 | Cine or spot fluorographic x-ray camera. | N | N
Radiology | RCC | C-Arm Fluoroscopic X-Ray System | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | IWN | Cassette, Measurement, Ardran-Crooks | 1 | 892.1940 | Radiologic quality assurance instrument. | N | N
Radiology | IXA | Cassette, Radiographic Film | 2 | 892.1850 | Radiographic film cassette. | N | N
Radiology | HBK | Chair, Pneumoencephalographic | 2 | 892.1820 | Pneumoencephalographic chair. | N | N
Radiology | KPX | Changer, Radiographic Film/Cassette | 2 | 892.1860 | Radiographic film/cassette changer. | N | N
Radiology | OMJ | Chest X-Ray Computer Aided Detection | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | MOS | Coil, Magnetic Resonance, Specialty | 2 | 892.1000 | Magnetic resonance diagnostic device. | N | N
Radiology | IZW | Collimator, Automatic, Radiographic | 2 | 892.1610 | Diagnostic x-ray beam-limiting device. | N | N
Radiology | IYL | Collimator, Dermatological, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IYK | Collimator, High Voltage, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IYJ | Collimator, Low Voltage, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IZX | Collimator, Manual, Radiographic | 2 | 892.1610 | Diagnostic x-ray beam-limiting device. | N | N
Radiology | IYI | Collimator, Orthovoltage, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | NWE | Colon Computed Tomography System, Computer Aided Detection | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | POK | Computer-Assisted Diagnostic Software For Lesions Suspicious For Cancer | 2 | 892.2060 |  | N | N
Radiology | IZT | Cone, Radiographic | 2 | 892.1610 | Diagnostic x-ray beam-limiting device. | N | N
Radiology | ONL | Conformal Brachytherapy Source | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | EGT | Controller, Temperature, Radiographic | 2 | 892.1900 | Automatic radiographic film processor. | N | N
Radiology | JAI | Couch, Radiation Therapy, Powered | 2 | 892.5770 | Powered radiation therapy patient support assembly. | N | N
Radiology | JAN | Counter, Whole Body, Nuclear | 1 | 892.1130 | Nuclear whole body counter. | N | N
Radiology | KXH | Cradle, Patient, Radiologic | 1 | 892.1830 | Radiologic patient cradle. | N | N
Radiology | PDK | Cream For X-Ray Attenuation | 2 | 892.6510 |  | N | N
Radiology | IWQ | Curtain, Protective, Radiographic | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IWK | Cyclotron, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | KGI | Densitometer, Bone | 2 | 892.1170 | Bone densitometer. | N | N
Radiology | QJG | Device System, Imaging, Tympanic Membrane And Middle Ear | 2 | 892.1560 | Ultrasonic pulsed echo imaging system. | N | N
Radiology | IWD | Device, Beam Limiting, Teletherapy, Radionuclide | 2 | 892.5750 | Radionuclide radiation therapy system. | N | N
Radiology | KPW | Device, Beam Limiting, X-Ray, Diagnostic | 2 | 892.1610 | Diagnostic x-ray beam-limiting device. | N | N
Radiology | KQA | Device, Beam Limiting, X-Ray, Therapeutic | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | NFG | Device, Communications, Images, Ophthalmic | 1 | 892.2020 | Medical image communications device. | N | N
Radiology | LMB | Device, Digital Image Storage, Radiological | 1 | 892.2010 | Medical image storage device. | N | N
Radiology | NFI | Device, Hardcopy, Images, Ophthalmic | 2 | 892.2040 | Medical image hardcopy device. | N | N
Radiology | IXL | Device, Spot-Film | 2 | 892.1670 | Spot-film device. | N | N
Radiology | NFF | Device, Storage, Images, Ophthalmic | 1 | 892.2010 | Medical image storage device. | N | N
Radiology | OQQ | Diagnostic Ultrasonic Transducer, Robotic | 2 | 892.1570 | Diagnostic ultrasonic transducer. | N | N
Radiology | LMA | Digitizer, Image, Radiological | 2 | 892.2030 | Medical image digitizer. | N | N
Radiology | NFH | Digitizer, Images, Ophthalmic | 2 | 892.2030 | Medical image digitizer. | N | N
Radiology | OIL | Discography Kit | 2 | 892.1730 | Photofluorographic x-ray system. | N | N
Radiology | PGY | Display, Diagnostic Radiology | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | NZT | Dosimeter, Ionizing Radiation, Implanted | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | Y
Radiology | EGW | Dryer, Film, Radiographic | 2 | 892.1900 | Automatic radiographic film processor. | N | N
Radiology | IWZ | Film, Radiographic | 1 | 892.1840 | Radiographic film. | N | N
Radiology | QHX | Fluoroscopic X-Ray Image Intensifier | 1 | 892.1960 | Radiographic intensifying screen. | N | N
Radiology | MUE | Full Field Digital, System, X-Ray, Mammographic | 2 | 892.1715 | Full-field digital mammography system. | N | N
Radiology | IYH | Generator, Dermatological (Grenz Ray), Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | KPZ | Generator, High Voltage, X-Ray, Therapeutic | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IZO | Generator, High-Voltage, X-Ray, Diagnostic | 1 | 892.1700 | Diagnostic x-ray high voltage generator. | N | N
Radiology | IYD | Generator, Low Voltage, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IYC | Generator, Orthovoltage, Therapeutic X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | IXJ | Grid, Radiographic | 1 | 892.1910 | Radiographic grid. | N | N
Radiology | PSW | High Level Disinfection Reprocessing Instrument For Ultrasonic Transducers, Liquid | 2 | 892.1570 | Diagnostic ultrasonic transducer. | N | N
Radiology | OUJ | High Level Disinfection Reprocessing Instrument For Ultrasonic Transducers, Mist | 2 | 892.1570 | Diagnostic ultrasonic transducer. | N | N
Radiology | IWY | Holder, Head, Radiographic | 1 | 892.1920 | Radiographic head holder. | N | N
Radiology | IXY | Holder, Radiographic Cassette, Wall-Mounted | 1 | 892.1880 | Wall-mounted radiographic cassette holder. | N | N
Radiology | IWR | Holder, Syringe, Lead | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | OVB | Hydrogel Spacer | 2 | 892.5725 | Absorbable perirectal spacer. | N | N
Radiology | IXC | Illuminator, Radiographic-Film | 1 | 892.1890 | Radiographic film illuminator. | N | N
Radiology | JAG | Illuminator, Radiographic-Film, Explosion-Proof | 1 | 892.1890 | Radiographic film illuminator. | N | N
Radiology | QJU | Image Acquisition And/Or Optimization Guided By Artificial Intelligence | 2 | 892.2100 |  | N | N
Radiology | OXO | Image-Intensified Fluoroscopic X-Ray System, Mobile | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | LHO | Instrument, Quality-Assurance, Radiologic | 1 | 892.1940 | Radiologic quality assurance instrument. | N | N
Radiology | OWB | Interventional Fluoroscopic X-Ray System | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | PCS | Liver Iron Concentration Imaging Companion Diagnostic For Deferasirox | 2 | 892.1001 |  | N | N
Radiology | OEB | Lung Computed Tomography System, Computer-Aided Detection | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | MUI | Media, Coupling, Ultrasound | 2 | 892.1570 | Diagnostic ultrasonic transducer. | N | N
Radiology | LQA | Media, Reusable Image | 1 | 892.1840 | Radiographic film. | N | N
Radiology | JAE | Microtron, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | IWE | Monitor, Patient Position, Light-Beam | 1 | 892.5780 | Light beam patient position indicator. | N | N
Radiology | JAF | Monitor, Ultrasonic, Nonfetal | 2 | 892.1540 | Nonfetal ultrasonic monitor. | N | N
Radiology | OIM | Mri Disposable Kit | 2 | 892.1000 | Magnetic resonance diagnostic device. | N | N
Radiology | IWF | Needle, Isotope, Gold, Titanium, Platinum | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | NMP | Needle, Isotope, Reprocessed | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | IYP | Phantom, Anthropomorphic, Nuclear | 1 | 892.1370 | Nuclear anthropomorphic phantom. | N | N
Radiology | IXG | Phantom, Anthropomorphic, Radiographic | 1 | 892.1950 | Radiographic anthropomorphic phantom. | N | N
Radiology | IYQ | Phantom, Flood Source, Nuclear | 1 | 892.1380 | Nuclear flood source phantom. | N | N
Radiology | JAR | Phantom, Test-Pattern, Radionuclide | 1 | 892.1420 | Radionuclide test pattern phantom. | N | N
Radiology | POY | Post Breast Biopsy Hemostatic Breast Compression Device | 2 | 892.1710 | Mammographic x-ray system. | N | N
Radiology | IZD | Probe, Uptake, Nuclear | 1 | 892.1320 | Nuclear uptake probe. | N | N
Radiology | IXX | Processor, Cine Film | 2 | 892.1900 | Automatic radiographic film processor. | N | N
Radiology | IXW | Processor, Radiographic-Film, Automatic | 2 | 892.1900 | Automatic radiographic film processor. | N | N
Radiology | EGY | Processor, Radiographic-Film, Automatic, Dental | 2 | 892.1900 | Automatic radiographic film processor. | N | N
Radiology | IZP | Programmer, Changer, Film/Cassette, Radiographic | 2 | 892.1870 | Radiographic film/cassette changer programmer. | N | N
Radiology | PCT | Prostate Immobilizer Rectal Balloon | 2 | 892.5720 | Rectal balloon for prostate immobilization. | N | N
Radiology | OIN | Prostate Seeding Kit | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | OIO | Radiographic Contrast Tray | 2 | 892.1680 | Stationary x-ray system. | N | N
Radiology | IWP | Radiographic Protective Glove | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | QBS | Radiological Computer Assisted Detection/Diagnosis Software For Fracture | 2 | 892.2090 |  | N | N
Radiology | QDQ | Radiological Computer Assisted Detection/Diagnosis Software For Lesions Suspicious For Cancer | 2 | 892.2090 |  | N | N
Radiology | QFM | Radiological Computer-Assisted Prioritization Software For Lesions | 2 | 892.2080 |  | N | N
Radiology | QAS | Radiological Computer-Assisted Triage And Notification Software | 2 | 892.2080 |  | N | N
Radiology | QKB | Radiological Image Processing Software For Radiation Therapy | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | OIK | Radiology Dental Tray | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | OIP | Radiology-Diagnostic Kit | 2 | 892.1680 | Stationary x-ray system. | N | N
Radiology | JAO | Scanner, Fluorescent | 2 | 892.1220 | Fluorescent scanner. | N | N
Radiology | IYW | Scanner, Rectilinear, Nuclear | 1 | 892.1300 | Nuclear rectilinear scanner. | N | N
Radiology | JAM | Scanner, Whole Body, Nuclear | 1 | 892.1330 | Nuclear whole body scanner. | N | N
Radiology | EAM | Screen, Intensifying, Radiographic | 1 | 892.1960 | Radiographic intensifying screen. | N | N
Radiology | EAK | Screen, Leaded, Operator Radiation Protector | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IWG | Seed, Isotope, Gold, Titanium, Platinum | 2 | 892.5730 | Radionuclide brachytherapy source. | N | Y
Radiology | IWS | Shield, Eye, Radiological | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IWT | Shield, Gonadal | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | KPY | Shield, Protective, Personnel | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | IWW | Shield, Vial | 1 | 892.6500 | Personnel protective shield. | N | N
Radiology | PZO | Software For Visualization Of Vascular Anatomy And Intravascular Devices | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | QHY | Solid State Fluoroscopic X-Ray Imager | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | MQB | Solid State X-Ray Imager (Flat Panel/Digital Imager) | 2 | 892.1680 | Stationary x-ray system. | N | N
Radiology | KXK | Source, Brachytherapy, Radionuclide | 2 | 892.5730 | Radionuclide brachytherapy source. | N | Y
Radiology | IXD | Source, Calibration, Sealed, Nuclear | 1 | 892.1400 | Nuclear sealed calibration source. | N | N
Radiology | IWI | Source, Isotope, Sealed, Gold, Titanium, Platinum | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | IWH | Source, Teletherapy, Radionuclide | 1 | 892.5740 | Radionuclide teletherapy source. | N | N
Radiology | IWA | Source, Wire, Iridium, Radioactive | 2 | 892.5730 | Radionuclide brachytherapy source. | N | N
Radiology | IXO | Synchronizer, Ecg / Respirator, Radiographic | 1 | 892.1970 | Radiographic ECG/respirator synchronizer. | N | N
Radiology | IYY | Synchronizer, Electrocardiograph, Nuclear | 1 | 892.1410 | Nuclear electrocardiograph synchronizer. | N | N
Radiology | IWM | Synchrotron, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | IWJ | System, Applicator, Radionuclide, Manual | 1 | 892.5650 | Manual radionuclide applicator system. | N | N
Radiology | JAQ | System, Applicator, Radionuclide, Remote-Controlled | 2 | 892.5700 | Remote controlled radionuclide applicator system. | N | N
Radiology | LMD | System, Digital Image Communications, Radiological | 1 | 892.2020 | Medical image communications device. | N | N
Radiology | NFJ | System, Image Management, Ophthalmic | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | LLZ | System, Image Processing, Radiological | 2 | 892.2050 | Picture archiving and communications system. | N | N
Radiology | NCS | System, Imaging, Holography, Acoustic | 2 | 892.1550 | Ultrasonic pulsed doppler imaging system. | N | N
Radiology | NQQ | System, Imaging, Optical Coherence Tomography (Oct) | 2 | 892.1560 | Ultrasonic pulsed echo imaging system. | N | N
Radiology | IYN | System, Imaging, Pulsed Doppler, Ultrasonic | 2 | 892.1550 | Ultrasonic pulsed doppler imaging system. | N | N
Radiology | IYO | System, Imaging, Pulsed Echo, Ultrasonic | 2 | 892.1560 | Ultrasonic pulsed echo imaging system. | N | N
Radiology | IXK | System, Imaging, X-Ray, Electrostatic | 2 | 892.1630 | Electrostatic x-ray imaging system. | N | N
Radiology | LNH | System, Nuclear Magnetic Resonance Imaging | 2 | 892.1000 | Magnetic resonance diagnostic device. | N | N
Radiology | LNI | System, Nuclear Magnetic Resonance Spectroscopic | 2 | 892.1000 | Magnetic resonance diagnostic device. | N | N
Radiology | MUJ | System, Planning, Radiation Therapy Treatment | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | LHN | System, Radiation Therapy, Charged-Particle, Medical | 2 | 892.5050 | Medical charged-particle radiation therapy system. | N | N
Radiology | IWL | System, Radiation Therapy, Neutron, Medical | 2 | 892.5300 | Medical neutron radiation therapy system. | N | N
Radiology | IWB | System, Radiation Therapy, Radionuclide | 2 | 892.5750 | Radionuclide radiation therapy system. | N | N
Radiology | IYT | System, Rebreathing, Radionuclide | 2 | 892.1390 | Radionuclide rebreathing system. | N | N
Radiology | KPQ | System, Simulation, Radiation Therapy | 2 | 892.5840 | Radiation therapy simulation system. | N | N
Radiology | JAD | System, Therapeutic, X-Ray | 2 | 892.5900 | X-ray radiation therapy system. | N | N
Radiology | JWM | System, Tomographic, Nuclear | 2 | 892.1310 | Nuclear tomography system. | N | N
Radiology | KPS | System, Tomography, Computed, Emission | 2 | 892.1200 | Emission computed tomography system. | N | N
Radiology | IZI | System, X-Ray, Angiographic | 2 | 892.1600 | Angiographic x-ray system. | N | N
Radiology | JAC | System, X-Ray, Film Marking, Radiographic | 1 | 892.1640 | Radiographic film marking system. | N | N
Radiology | JAA | System, X-Ray, Fluoroscopic, Image-Intensified | 2 | 892.1650 | Image-intensified fluoroscopic x-ray system. | N | N
Radiology | JAB | System, X-Ray, Fluoroscopic, Non-Image-Intensified | 2 | 892.1660 | Non-image-intensified fluoroscopic x-ray system. | N | N
Radiology | IZH | System, X-Ray, Mammographic | 2 | 892.1710 | Mammographic x-ray system. | N | N
Radiology | IZL | System, X-Ray, Mobile | 2 | 892.1720 | Mobile x-ray system. | N | N
Radiology | IZG | System, X-Ray, Photofluorographic | 2 | 892.1730 | Photofluorographic x-ray system. | N | N
Radiology | KPR | System, X-Ray, Stationary | 2 | 892.1680 | Stationary x-ray system. | N | N
Radiology | IZF | System, X-Ray, Tomographic | 2 | 892.1740 | Tomographic x-ray system. | N | N
Radiology | JAK | System, X-Ray, Tomography, Computed | 2 | 892.1750 | Computed tomography x-ray system. | N | N
Radiology | IZZ | Table, Radiographic, Non-Tilting, Powered | 2 | 892.1980 | Radiologic table. | N | N
Radiology | IXQ | Table, Radiographic, Stationary Top | 2 | 892.1980 | Radiologic table. | N | N
Radiology | IXR | Table, Radiographic, Tilting | 2 | 892.1980 | Radiologic table. | N | N
Radiology | KXJ | Table, Radiologic | 2 | 892.1980 | Radiologic table. | N | N
Radiology | IXF | Test-Pattern, Radiographic | 1 | 892.1940 | Radiologic quality assurance instrument. | N | N
Radiology | OUO | Tomographic Imager Combining Emission Computed Tomography With Nuclear Magnetic Resonance | 2 | 892.1200 | Emission computed tomography system. | N | N
Radiology | ITX | Transducer, Ultrasonic, Diagnostic | 2 | 892.1570 | Diagnostic ultrasonic transducer. | N | N
Radiology | LEK | Transilluminator (Diaphanoscope) | 3 | 892.1990 | Transilluminator for breast evaluation. | N | N
Radiology | IYB | Tube Mount, X-Ray, Diagnostic | 1 | 892.1770 | Diagnostic x-ray tube mount. | N | N
Radiology | PSV | Ultrasound Bronchoscope | 2 | 892.1550 | Ultrasonic pulsed doppler imaging system. | N | N
Radiology | LWI | Ultrasound, Sinus | 2 | 892.1560 | Ultrasonic pulsed echo imaging system. | N | N
Radiology | QHA | X-Ray Angiographic Imaging Based Coronary Vascular Simulation Software Device | 2 | 892.1600 | Angiographic x-ray system. | N | N
Radiology | RGC | X-Ray Field Indicator Light (Laser) | 1 | 892.5780 | Light beam patient position indicator. | N | N
Radiology | OAS | X-Ray, Tomography, Computed, Dental | 2 | 892.1750 | Computed tomography x-ray system. | N | N
Toxicology | DLI | Acetylcholine Chloride, Specific Reagent For Pseudo Cholinesterase | 1 | 862.3240 | Cholinesterase test system. | N | N
Toxicology | DKC | Alcohol Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DIC | Alcohol Dehydrogenase, Specific Reagent For Ethanol Enzyme Method | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | KLP | Amikacin Serum Assay | 2 | 862.3035 | Amikacin test system. | N | N
Toxicology | DKQ | Antiserum, Digitoxin | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DKA | Antiserum, Digoxin | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DJI | Antiserum, Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | NDW | Assay, Porphyrin, Spectrophotometry, Lithium | 2 | 862.3560 | Lithium test system. | N | N
Toxicology | DNE | Atomic Absorption, Antimony | 1 | 862.3110 | Antimony test system. | N | N
Toxicology | DNZ | Atomic Absorption, Arsenic | 1 | 862.3120 | Arsenic test system. | N | N
Toxicology | JII | Atomic Absorption, Lithium | 2 | 862.3560 | Lithium test system. | N | N
Toxicology | DID | Bacillus Subtilis Microbiology Assay, Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | DKD | Bacillus Subtilis, Microbiological Assay, Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | OIC | Blood Alcohol Kit (Excludes Hiv Testing) | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | ORA | Breast Milk-Alcohol Test System | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DKB | Calibrators, Drug Mixture | 2 | 862.3200 | Clinical toxicology calibrator. | N | N
Toxicology | DLJ | Calibrators, Drug Specific | 2 | 862.3200 | Clinical toxicology calibrator. | N | N
Toxicology | DNN | Calibrators, Ethyl Alcohol | 2 | 862.3200 | Clinical toxicology calibrator. | N | N
Toxicology | DIG | Cholinesterase Test Paper | 1 | 862.3240 | Cholinesterase test system. | N | N
Toxicology | LDP | Colorimetry, Acetaminophen | 2 | 862.3030 | Acetaminophen test system. | N | N
Toxicology | DIH | Colorimetry, Cholinesterase | 1 | 862.3240 | Cholinesterase test system. | N | N
Toxicology | DKJ | Colorimetry, Salicylate | 2 | 862.3830 | Salicylate test system. | N | N
Toxicology | ODW | Cytochrome P450 2c9 (Cyp450 2c9) Drug Metabolizing Enzyme Genotyping System | 2 | 862.3360 | Drug metabolizing enzyme genotyping system. | N | N
Toxicology | DIZ | Delayed Analysis, Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DJZ | Devices, Breath Trapping, Alcohol | 1 | 862.3050 | Breath-alcohol test system. | N | N
Toxicology | NVI | Diagnostic Software, K-Nearest Neighbor Algorithm, Autoimmune Disease | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DJK | Digitoxin Control Serum, Ria | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DMP | Digoxin Control Serum, Ria | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | NTI | Drug Metabolizing Enzyme Genotyping Systems | 2 | 862.3360 | Drug metabolizing enzyme genotyping system. | N | N
Toxicology | DIF | Drug Mixture Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | LAS | Drug Specific Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | LAL | Eas Chromatography, Quinine | 1 | 862.3750 | Quinine test system. | N | N
Toxicology | DOH | Electrometry, Cholinesterase | 1 | 862.3240 | Cholinesterase test system. | N | N
Toxicology | DMT | Enzymatic Method, Alcohol Dehydrogenase, Ultraviolet | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | LDN | Enzymatic Radiochemical Assay, Amikacin | 2 | 862.3035 | Amikacin test system. | N | N
Toxicology | KLO | Enzymatic Radiochemical Assay, Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | LDO | Enzymatic Radiochemical Assay, Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | DKZ | Enzyme Immunoassay, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DIS | Enzyme Immunoassay, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | JXM | Enzyme Immunoassay, Benzodiazepine | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | LDJ | Enzyme Immunoassay, Cannabinoids | 2 | 862.3870 | Cannabinoid test system. | N | N
Toxicology | KLT | Enzyme Immunoassay, Carbamazepine | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | JXO | Enzyme Immunoassay, Cocaine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | DIO | Enzyme Immunoassay, Cocaine And Cocaine Metabolites | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | LFM | Enzyme Immunoassay, Digitoxin | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | KXT | Enzyme Immunoassay, Digoxin | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DIP | Enzyme Immunoassay, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DLF | Enzyme Immunoassay, Ethosuximide | 2 | 862.3380 | Ethosuximide test system. | N | N
Toxicology | LCD | Enzyme Immunoassay, Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | KLR | Enzyme Immunoassay, Lidocaine | 2 | 862.3555 | Lidocaine test system. | N | N
Toxicology | DJR | Enzyme Immunoassay, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | LAN | Enzyme Immunoassay, N-Acetylprocainamide | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | MKU | Enzyme Immunoassay, Nicotine And Nicotine Metabolites | 1 | 862.3220 | Carbon monoxide test system. | N | N
Toxicology | DJG | Enzyme Immunoassay, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | POX | Enzyme Immunoassay, Oxcarbazepine Metabolite | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | DLZ | Enzyme Immunoassay, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | DJD | Enzyme Immunoassay, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | LAR | Enzyme Immunoassay, Procainamide | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | JXN | Enzyme Immunoassay, Propoxyphene | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | LBZ | Enzyme Immunoassay, Quinidine | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | KLS | Enzyme Immunoassay, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | LEG | Enzyme Immunoassay, Valproic Acid | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | OUF | Everolimus Immunoassay | 2 | 862.3840 | Sirolimus test system. | N | N
Toxicology | JIH | Flame Photometry, Lithium | 2 | 862.3560 | Lithium test system. | N | N
Toxicology | LGJ | Fluorescence Polarization Immunoassay, Amibacin | 2 | 862.3035 | Amikacin test system. | N | N
Toxicology | LGI | Fluorescence Polarization Immunoassay, Carbamazepine | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | MOJ | Fluorescence Polarization Immunoassay, Diphenylhydantoin (Free) | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LGR | Fluorescence Polarization Immunoassay, Diphenylhydantoin (Total) | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LGQ | Fluorescence Polarization Immunoassay, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | LGS | Fluorescence Polarization Immunoassay, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | LFW | Fluorescence Polarization Immunoassay, Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | LCQ | Fluorescent Immunoassay Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | LES | Fluorescent Immunoassay, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LET | Fluorescent Immunoassay, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | LFT | Fluorescent Immunoassay, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | LER | Fluorescent Immunoassay, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | LCR | Fluorescent Immunoassay, Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | DOX | Fluorometer, Lead (Dedicated Instruments) | 2 | 862.3550 | Lead test system. | N | N
Toxicology | DJJ | Fluorometry, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DJL | Free Radical Assay, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DIR | Free Radical Assay, Cocaine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | DOL | Free Radical Assay, Lsd | 2 | 862.3580 | Lysergic acid diethylamide (LSD) test system. | N | N
Toxicology | DPP | Free Radical Assay, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | DOK | Free Radical Assay, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DKT | Free Radical Assay, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | DNG | Free Radical, Benzoylecgnonine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | OTF | Gabapentin Assay | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DLS | Gas Chromatograph, Alcohol (Dedicated Instruments) | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | JKT | Gas Chromatograph, Carbon-Monoxide | 1 | 862.3220 | Carbon monoxide test system. | N | N
Toxicology | DLH | Gas Chromatography, Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DOD | Gas Chromatography, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | KZZ | Gas Chromatography, Benzodiazepine | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | DIN | Gas Chromatography, Cocaine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | LAD | Gas Chromatography, Codeine | 2 | 862.3270 | Codeine test system. | N | N
Toxicology | DKH | Gas Chromatography, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DIY | Gas Chromatography, Ethosuximide | 2 | 862.3380 | Ethosuximide test system. | N | N
Toxicology | DMB | Gas Chromatography, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | LAF | Gas Chromatography, Methamphetamine | 2 | 862.3610 | Methamphetamine test system. | N | N
Toxicology | DMY | Gas Chromatography, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DJF | Gas Chromatography, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | DJH | Gas Chromatography, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | DMQ | Gas Chromatography, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | LAJ | Gas Chromatography, Propoxyphene | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | DIM | Gas Chromatography, Salicylate | 2 | 862.3830 | Salicylate test system. | N | N
Toxicology | DMF | Gas Liquid Chromatography, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | JKQ | Gold Chloride (Colorimetric), Bromide | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | DIE | Heavy Metals Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DLX | Hemagglutination Inhibition, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | DNO | Hemagglutination Inhibition, Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | DIW | Hemagglutination Inhibition, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | DLR | Hemagglutination Inhibition, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DLN | Hemagglutination, Cocaine Metabolites (Benzoylecgnonine) | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | DLT | Hemagglutination, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | KZY | High Pressure Liquid Chromatography, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | LAA | High Pressure Liquid Chromatography, Benzodiazepine | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | LAC | High Pressure Liquid Chromatography, Cocaine And Cocaine Metabolites | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | LAE | High Pressure Liquid Chromatography, Codeine | 2 | 862.3270 | Codeine test system. | N | N
Toxicology | LAG | High Pressure Liquid Chromatography, Methamphetamine | 2 | 862.3610 | Methamphetamine test system. | N | N
Toxicology | LAH | High Pressure Liquid Chromatography, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | LAK | High Pressure Liquid Chromatography, Propoxyphene | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | LAM | High Pressure Liquid Chromatography, Quinine | 1 | 862.3750 | Quinine test system. | N | N
Toxicology | LFI | High Pressure Liquid Chromatography, Tricyclic Antidepressant Drugs | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | NWM | Immunoassay, Anti-Seizure Drug | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | KZT | Kit, Disc, Agar Gel Diffusion, Determine Serum Levels Of Gentamicin | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | KZW | Kit, Disc, Agar Gel Diffusion, Determine Serum Levels Of Kanamycin | 2 | 862.3520 | Kanamycin test system. | N | N
Toxicology | KZX | Kit, Disc, Agar Gel Diffusion, Determine Serum Levels Of Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | ORH | Lamotrigine Assay | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DOF | Lead, Atomic Absorption | 2 | 862.3550 | Lead test system. | N | N
Toxicology | DIJ | Lead, Delta Amino Levulinic Acid | 2 | 862.3550 | Lead test system. | N | N
Toxicology | DNX | Lead, Protoporphyrin Zinc Method, Fluorometric | 2 | 862.3550 | Lead test system. | N | N
Toxicology | DMK | Lead, Protoporphyrin, Fluorometric | 2 | 862.3550 | Lead test system. | N | N
Toxicology | ORI | Levetiracetam Assay | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LAX | Lidocaine Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DNI | Liquid Chromatography, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DOZ | Liquid Chromatography, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DNF | Liquid Chromatography, Ethosuximide | 2 | 862.3380 | Ethosuximide test system. | N | N
Toxicology | DNT | Liquid Chromatography, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | DPK | Liquid Chromatography, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DOO | Liquid Chromatography, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | DPQ | Liquid Chromatography, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | DMX | Liquid Chromatography, Salicylate | 2 | 862.3830 | Salicylate test system. | N | N
Toxicology | QBK | Meprobamate Test System | 2 | 862.3590 | Meprobamate test system. | N | N
Toxicology | DJN | Mercury Dithiazone, Colorimetry, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | DPH | Mercury, Atomic Absorption | 1 | 862.3600 | Mercury test system. | N | N
Toxicology | LAY | Methotrexate Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | OAV | Mycophenolic Acid Test System | 2 | 862.3840 | Sirolimus test system. | N | N
Toxicology | LAZ | N-Acetylprocainamide Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DML | Nad-Nadh, Specific Reagent For Alcohol Enzyme Method | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | LFO | Nephelometric Inhibition Immunoassay, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LFN | Nephelometric Inhibition Immunoassay, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | PDY | Organophosphate Test System | 2 | 862.3652 | Organophosphate test system. | N | N
Toxicology | QDR | Plazomicin Test System, Immunoassay | 2 | 862.3460 |  | N | N
Toxicology | DMI | Potassium Dichromate Specific Reagent For Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DOJ | Potassium Dichromate, Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | LBA | Procainamide Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | LPX | Radioceptor Assay, Neuroleptic Drugs | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | KLQ | Radioimmunoassay, Amikacin | 2 | 862.3035 | Amikacin test system. | N | N
Toxicology | DJP | Radioimmunoassay, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DPJ | Radioimmunoassay, Amphetamine (125-I), Goat Antibody, Ammonium Sulfate Sep. | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DKN | Radioimmunoassay, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | LAT | Radioimmunoassay, Cannabinoid(S) | 2 | 862.3870 | Cannabinoid test system. | N | N
Toxicology | KLN | Radioimmunoassay, Cocaine Metabolite | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | LCW | Radioimmunoassay, Digitoxin (125-I) | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DPL | Radioimmunoassay, Digitoxin (125-I), Bovine Antibody, Charcoal Sep. | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DPG | Radioimmunoassay, Digitoxin (125-I), Rabbit Antibody, Coated Tube Sep. | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DND | Radioimmunoassay, Digitoxin (125-I), Rabbit Antibody, Solid Phase Sep. | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DOB | Radioimmunoassay, Digitoxin (3-H) | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DNW | Radioimmunoassay, Digitoxin (3-H), Rabbit Antibody, Charcoal Sep. | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | DOW | Radioimmunoassay, Digitoxin (3-H), Rabbit Antibody, Coated Tube | 2 | 862.3300 | Digitoxin test system. | N | N
Toxicology | LCS | Radioimmunoassay, Digoxin (125-I) | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DPI | Radioimmunoassay, Digoxin (125-I), Bovine Antibody, Charcoal Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DNJ | Radioimmunoassay, Digoxin (125-I), Goat Antibody, 2nd Antibody Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DOA | Radioimmunoassay, Digoxin (125-I), Goat Antibody, Anion Exchange, Resin Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DPB | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Charcoal Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DPO | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Coated Tube Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DNS | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Double Label Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DOG | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Polyethylene Glycol | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DNL | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Second Antibody Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DON | Radioimmunoassay, Digoxin (125-I), Rabbit Antibody, Solid Phase Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | LCT | Radioimmunoassay, Digoxin (3-H) | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DOR | Radioimmunoassay, Digoxin (3-H), Bovine Antibody, Charcoal Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DOY | Radioimmunoassay, Digoxin (3-H), Goat Antibody, 2nd Antibody Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DPD | Radioimmunoassay, Digoxin (3-H), Rabbit Antibody, Charcoal Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DNQ | Radioimmunoassay, Digoxin (3-H), Rabbit Antibody, Coated Tubes Sep. | 2 | 862.3320 | Digoxin test system. | N | N
Toxicology | DLP | Radioimmunoassay, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DJX | Radioimmunoassay, Ethosuximide | 2 | 862.3380 | Ethosuximide test system. | N | N
Toxicology | DJB | Radioimmunoassay, Gentamicin (125-I), Second Antibody Sep. | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | KJI | Radioimmunoassay, Kanamycin | 2 | 862.3520 | Kanamycin test system. | N | N
Toxicology | DLB | Radioimmunoassay, Lsd (125-I) | 2 | 862.3580 | Lysergic acid diethylamide (LSD) test system. | N | N
Toxicology | KXS | Radioimmunoassay, Methaqualone | 2 | 862.3630 | Methaqualone test system. | N | N
Toxicology | DOE | Radioimmunoassay, Morphine (125-I), Goat Antibody Ammonium Sulfate Sep. | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DIQ | Radioimmunoassay, Morphine (3-H), Goat Antibody Ammonium Sulfate Sep. | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | DNA | Radioimmunoassay, Morphine-Barbiturate (125-I), Goat Antibody | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | LCE | Radioimmunoassay, Netilmicin (I-125) | 2 | 862.3450 | Gentamicin test system. | N | N
Toxicology | DKP | Radioimmunoassay, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | DIL | Radioimmunoassay, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | LCA | Radioimmunoassay, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | KLB | Radioimmunoassay, Tobramycin | 2 | 862.3900 | Tobramycin test system. | N | N
Toxicology | LFG | Radioimmunoassay, Tricyclic Antidepressant Drugs | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | LEH | Radioimmunoassay, Vancomycin | 2 | 862.3950 | Vancomycin test system. | N | N
Toxicology | DLE | Reagents, Test, Bromides | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | DKE | Reagents, Test, Tetrahydrocannabinol | 2 | 862.3870 | Cannabinoid test system. | N | N
Toxicology | NRP | Sirolimus Test System | 2 | 862.3840 | Sirolimus test system. | N | N
Toxicology | OIW | Software, Similarity Score Algorithm, Tissue Of Origin For Malignant Tumor Types | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DMR | Solution, M-Nitrophenol, Specific Reagent For Cholinesterase | 1 | 862.3240 | Cholinesterase test system. | N | N
Toxicology | JKS | Spectral Absorb. Curve, Oxyhemoglobin, Carboxyhemoglobin, Carbon-Monoxide | 1 | 862.3220 | Carbon monoxide test system. | N | N
Toxicology | DNC | Strip, Paper, Salicylate | 2 | 862.3830 | Salicylate test system. | N | N
Toxicology | MXA | System, Test, Breath Nitric Oxide | 2 | 862.3080 | Breath nitric oxide test system. | N | N
Toxicology | MSL | System, Test, Topiramatee | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DMJ | Test Reagents For Ethyl Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DLA | Test Reagents For Methyl Alcohol | 2 | 862.3040 | Alcohol test system. | N | N
Toxicology | DJQ | Test Reagents For Phenothiazine | 2 | 862.3670 | Phenothiazine test system. | N | N
Toxicology | DKW | Test Reagents For Sulphanimide Derivatives | 1 | 862.3850 | Sulfonamide test system. | N | N
Toxicology | MRS | Test System, Nicotine, Cotinine, Metabolites | 1 | 862.3220 | Carbon monoxide test system. | N | N
Toxicology | PUX | Test, Amphetamine, Employment And Insurance Testing, Exempt | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | NFT | Test, Amphetamine, Over The Counter | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | PUY | Test, Barbiturate, Employment And Insurance Testing, Exempt | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | PTH | Test, Barbiturate, Over The Counter | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | PUZ | Test, Benzodiazepine, Employment And Insurance Testing, Exempt | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | NFV | Test, Benzodiazepine, Over The Counter | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | PVJ | Test, Cannabinoid, Employment And Insurance Testing, Exempt | 2 | 862.3870 | Cannabinoid test system. | N | N
Toxicology | NFW | Test, Cannabinoid, Over The Counter | 2 | 862.3870 | Cannabinoid test system. | N | N
Toxicology | PVA | Test, Cocaine And Cocaine Metabolite, Employment And Insurance Testing, Exempt | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | NFY | Test, Cocaine And Cocaine Metabolites, Over The Counter | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | PVB | Test, Codeine, Employment And Insurance Testing, Exempt | 2 | 862.3270 | Codeine test system. | N | N
Toxicology | PVE | Test, Methadone, Employment And Insurance Testing, Exempt | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | PTG | Test, Methadone, Over The Counter | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | PVD | Test, Methamphetamine, Employment And Insurance Testing, Exempt | 2 | 862.3610 | Methamphetamine test system. | N | N
Toxicology | NGG | Test, Methamphetamine, Over The Counter | 2 | 862.3610 | Methamphetamine test system. | N | N
Toxicology | PVF | Test, Methaqualone, Employment And Insurance Testing, Exempt | 2 | 862.3630 | Methaqualone test system. | N | N
Toxicology | PVG | Test, Morphine, Employment And Insurance Testing, Exempt | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | NGI | Test, Morphine, Over The Counter | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | PVH | Test, Opiates, Employment And Insurance Testing, Exempt | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | NGL | Test, Opiates, Over The Counter | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | PVI | Test, Propoxyphene, Employment And Insurance Testing, Exempt | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | QBF | Test, Propoxyphene, Over The Counter | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | PVK | Test, Tricyclic Antidepressant Drugs, Employment And Insurance Testing, Exempt | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | QAW | Test, Tricyclic Antidepressants, Over The Counter | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | LAW | Theophylline Control Materials | 1 | 862.3280 | Clinical toxicology control material. | N | N
Toxicology | DIT | Thin Layer Chromatography, Amphetamine | 2 | 862.3100 | Amphetamine test system. | N | N
Toxicology | DKX | Thin Layer Chromatography, Barbiturate | 2 | 862.3150 | Barbiturate test system. | N | N
Toxicology | LAB | Thin Layer Chromatography, Benzodiazepine | 2 | 862.3170 | Benzodiazepine test system. | N | N
Toxicology | DOM | Thin Layer Chromatography, Benzoylecgnonine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | DMN | Thin Layer Chromatography, Cocaine | 2 | 862.3250 | Cocaine and cocaine metabolite test system. | N | N
Toxicology | DLD | Thin Layer Chromatography, Codeine | 2 | 862.3270 | Codeine test system. | N | N
Toxicology | DPE | Thin Layer Chromatography, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | DNP | Thin Layer Chromatography, Ethosuximide | 2 | 862.3380 | Ethosuximide test system. | N | N
Toxicology | DKR | Thin Layer Chromatography, Methadone | 2 | 862.3620 | Methadone test system. | N | N
Toxicology | DJC | Thin Layer Chromatography, Methamphetamine | 2 | 862.3610 | Methamphetamine test system. | N | N
Toxicology | DNK | Thin Layer Chromatography, Morphine | 2 | 862.3640 | Morphine test system. | N | N
Toxicology | LAI | Thin Layer Chromatography, Opiates | 2 | 862.3650 | Opiate test system. | N | N
Toxicology | DIX | Thin Layer Chromatography, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | DOS | Thin Layer Chromatography, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | DPN | Thin Layer Chromatography, Propoxyphene | 2 | 862.3700 | Propoxyphene test system. | N | N
Toxicology | DIK | Thin Layer Chromatography, Quinine | 1 | 862.3750 | Quinine test system. | N | N
Toxicology | DJE | Thin Layer Chromatography, Salicylate | 2 | 862.3830 | Salicylate test system. | N | N
Toxicology | LEW | Thin Layer Chromatography, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | MLK | Thin Layer Chromatography, Tricyclic Antidepressant Drugs | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | LDB | U.V. Spectrometry, Carbamazepine | 2 | 862.3645 | Neuroleptic drugs radioreceptor assay test system. | N | N
Toxicology | LDC | U.V. Spectrometry, Diphenylhydantoin | 2 | 862.3350 | Diphenylhydantoin test system. | N | N
Toxicology | LDA | U.V. Spectrometry, Phenobarbital | 2 | 862.3660 | Phenobarbital test system. | N | N
Toxicology | LCZ | U.V. Spectrometry, Primidone | 2 | 862.3680 | Primidone test system. | N | N
Toxicology | LCY | U.V. Spectrometry, Theophylline | 2 | 862.3880 | Theophylline test system. | N | N
Toxicology | LFH | U.V. Spectrometry, Tricyclic Antidepressant Drugs | 2 | 862.3910 | Tricyclic antidepressant drugs test system. | N | N
Toxicology | PUJ | Voriconazole Test System | 2 | 862.3970 |  | N | N
General Hospital | QBP | Cap, Device Disinfectant | 2 |  | A molded cap with a disinfectant pad to fit on a luer access device; Disinfect needleless access valves and may act as a physical barrier to contamination if not removed for a set period of time | N | N | ICU Medical | CAP, Swabtip
Hematology | JPA | System, Multipurpose For In Vitro Coagulation Studies | 2 | 864.5425 | Multipurpose system for in vitro coagulation studies. | N | N | Accriva | hemochron
Hematology | JOZ | System, Automated Platelet Aggregation | 2 | 864.57oo | utomated platelet aggregation system | N | N | Accriva | verifyNow
Toxicology | JKS | Spectral Absorb. Curve, Oxyhemoglobin, Carboxyhemoglobin, Carbon-Monoxide | 1 | 862.322o | Carbon monoxide test system | N | N | Accriva | AVOXImeter
Hematology | JCA | Device, Bleeding Time |  | 864.61oo |  | N | N | Accriva | Tenderfoot, surgicutt
General & Plastic Surgery | FMK | Lancet, Blood |  | 878.48oo |  | N | N | Accriva | Tenderlett

## HS_collection
Pacer_HS | Brace_HS | Stent Graft(X) | Stent Graft(EN) | Viewer
Abnormal Cardiac Pacing |  |  | Aspiration Issue | Treatment
Abnormal Defibrillation Therapy | Abnormal Therapy |  | Balloon Deflation Issue | Abnormal Treatment
Abnormal Misuse | Misuse | Misuse | Balloon Inflation Issue | Loss of Treatment
Application of Excessive Force | Application of Excessive Force | Application of Excessive Force | Application of Excessive Force | Application of Excessive Energy
Biocompatibility Issue | Skin Allergy or Irritation | Cytoxic | Cytoxic | Application of Undesired Energy
Confusion | User Confused | User Confused | User Confused | Application of Inadequate Energy
Device Heating | Device Heating |  |  | Incorrect Energy Delivery Area
Device Tissue Pressure |  |  | Blood Leak through Device | Application of Excessive Force at Incision
Extravascular Foreign Body | Foreign Materials |  | Component Attachment / Detachment Issue | Application of Excessive Force to Tissue
Inaccessible Diagnostic Data | Inaccessible Data |  | Compromised Graft Integrity | Procedure Delayed
Inappropriate Defibrillation Therapy | Inappropriate Treatment | Inappropriate Treatment |  | Procedure Stopped
Inappropriate Programmed Settings | Inappropriate Settings |  | Damaged / Broken Component | Alternative Procedure Applied
Incomplete | Device Reused |  | Detachment of Components (Implant) | Surgical Complications
Incomplete or Misleading Diagnostic Data | Incomplete or Misleading Diagnostic Data |  | Detachment of Components (Accessories) | Loss of Session
Inconvenience | Inconvenience | Inconvenience | Detachment of Components (Catheter) | Image, Model, Data, Pre-op Plan
Lack of Sterility |  | Compromised Sterility | Compromised Sterility | Poor Patient Image during pre-operative
Loss of Cardiac Pacing | Loss of Material Integrity | Loss of Material Integrity | Loss/Deterioration of function: Delivery system integrity | Poor Patient Image during intra-operative
Loss of Defibrillation Therapy | Loss of Structural Integrity | Loss of Structural Integrity | Loss/Deterioration of function: Prosthesis Integrity | Loss of Patient Image during pre-operative
Loss of Patient Notification | Loss of Mechanical Integrity | Loss of Mechanical Integrity |  | Loss of Patient Image during intra-operative
Surgical Complications |  | Surgical Complications | Difficult to Position Implant | Inaccessible Patient Image during pre-operative
Unintentional Extra Cardiac Stimulation | Unintend Motion | Unintend Motion | Unintend Motion | Inaccessible Patient Image during intra-operative
| Material Degraded | Material Degraded | Difficult to Retract / Unsheath | unable to Manipulate Patient Image during pre-operative
| Inadquate IFU or Labelling |  | Difficult to Separate (Demate) | unable to Manipulate Patient Image during intra-operative
| Brace Immobilizated | Unintend Stop | Difficult to Withdraw | unable to Display Patient Image during pre-operative
| Fail to Operate |  | Difficult to Dock/Undock | unable to Display Patient Image during intra-operative
| Product Malfunction |  | Endobag Compromised | Inappropriate Patient 3D Model during pre-operative
| Improper Fit or adjustment | Improper Fit or Adjustment | Endobag Protrusion | Inappropriate Patient 3D Model during intra-operative
| Inadequate Support |  | Failure to Unfold (Graft) | Loss of Patient 3D Model during pre-operative
| Brace Applied Inappropriately |  | Flaked | Loss of Patient 3D Model during intra-operative
| Product Contiminated |  | Fluid Leak | Inaccessible Patient 3D Model during pre-operative
| Abrasion to User |  | Fracture (Implant) | Inaccessible Patient 3D Model during intra-operative
| Bacteria Growth |  | Implant Migration, Cuff | unable to Manipulate 3D Model during pre-operative
| Chocking on Small Parts |  | Implant Migration, Limb | unable to Manipulate 3D Model during intra-operative
| Exposure to Chemicals |  | Implant Migration, Main Body
|  | Blood Leak through Device | Implant movement | Loss of Patient Data during pre-operative
|  | Inadequate Stent-Graft Patency | Implant Separation | Loss of Patient Data during intra-operative
|  | Visualization or Visibility Issue | Visualization or Visibility Issue | Corrupted Patient Data during pre-operative
|  |  | Loss/Deterioration of function: Visualization | Corrupted Patient Data during intra-operative
|  | Unable to Withdraw | Unable to Withdraw | Incorrect Patient Data during pre-operative
|  | Unable to Advance | Unable to Advance | Incorrect Patient Data during intra-operative
|  | Unable to Deploy | Unable to Deploy | Inaccessible Patient Data during pre-operative
|  |  | Loss of Seal | Inaccessible Patient Data during intra-operative
Material Invagination (Graft)
|  |  | MRI Compatibility Issue | Loss of Surgical Records during pre-operative
|  |  | Obstruction/Blockage within Device | Loss of Surgical Records during intra-operative
|  |  | Occlusion within Device (Implant) | Inappropriate Pre-Operative Plan
|  |  | Polymer Cartridge Issue | Corrupted Pre-Operative Plan
|  |  | Size Incorrect for Patient (Oversized) | Delayed Pre-Operative Plan
|  |  | Size Incorrect for Patient (Undersized) | Software, System
|  |  | Surepass Wire Issues | Software Frozen or locked Up during pre-operative
|  |  | Unable to Dock/Undock | Software Frozen or locked Up during intra-operative
|  |  | Unable to hold pressure in endobag | Software Unresponsive during pre-operative
|  |  | Unable to Inflate | Software Unresponsive during intra-operative
|  |  | Unable to Prepare Device | Hardware (Mechanical, Electronical)
|  |  | Unable to Retract / Unsheath | Device or Subassemblies Collision
|  |  | Unable to Dispense Polymer | Unintend Motion
|  |  | Unable to Remove | Unintend Stop
|  |  | Inadequate Stent/Graft Patency | Loss of Jaw Motion or Articulation
|  |  | Incorrect Product for intended use (off-label) | Low Grip Force
|  |  | Kinked Component | Vibration of Instrument
|  |  | Balloon Rupture | Bicomaptibility
|  |  | Difficult to Dispense Polymer | Biocompatibility Issue
|  |  | Inadequate endobag-wall apposition | Breach in Sterility
|  |  | Blood in Polymer Line (major) | User or Patient Experience
|  |  | Blood in Polymer Line (minor) | User Confusion
|  |  | Fragment in body | User Annoyance
|  |  | Material Degraded | User Inconvenience
User Discomfort
No Hazardous Situation
Electrical Shock
Device Heating
User or Patient Collision

## reference
| battery | 1.- R.P. Ramasamy, et al., "Discharge characteristics of silver vanadium oxide cathodes," Journal of Applied Electrochemistry (2006) 487–497
2.- Christine A. Frysz et al., "Selection of cathode current collector materials using electrochemical techniques" 0-7803-0552-3/92.03.00 IEEE 1992
3. - Esther S. Takeuchi "Reliability systems for implantable cardiac defibrillator batteries" Journal of Power Sources 54 (1995) 115-119
9/1/5/2020: add multiple outputs
Standards
1 | ISO 10555-1 2013 | intravascular catheters - sterile and single-use catheters
2 | ISO 11070 2014 | intravascular introducers - dilators - guidewires
3 | ISO 11135-2014 | sterilization EO
4 | ISO 11137-1 2016 | sterilization validation
5 | ISO 11137-2 2013 | sterilzation dose
6 | ISO 11607-1 2019 | packaging for sterile barrier
7 | ISO 13408 2008 | aseptic processing of health care products
8 | ISO 14630 2012 | non-active surgical implants
9 | ISO 14971 2019 | risk management for medical devices
10 | ISO 17664 2017 | resterilization medical device
11 | ISO 22442-1 2015 | utlizing animal tissues
12 | ISO 25539-1 2017 | endovascular prostheses
13 | ISO 62304 | software for medical devices
14 | ISO 62366 | usability engineering for medical devices
15 | ISO 7198 2016 | vascular protheses
16 | ISO 80369-7 2016 | luer conector
| Device learned | Keywords | Training sentences | Model included | Evaluation
1 | brace
2 | muscle stimulator
3 | pacer
4 | ICD
5 | lead
6 | ventilator
7 | robotic surgical
8 | infusion pump
9 | diabete pump
10 | pneumatical pump
11 | battery
12 | capacitor
| Stent-Graft | standard only

## Brace_shop1
Failure_Mode | Hazard | Hazardous_Situation | Harm | Severity
accessory unattached | Detachment of Device Component | Fail to Operate | Loss of Treatment | 2
adding stability failed | Unstable | Loss of Structural Integrity | Loss of Treatment | 2
adjustment difficult | Device Difficult to Setup or Prepare | Fail to Operate | Delay of Treatment | 1
adjustment difficult | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
adjustment failed | Device Difficult to Setup or Prepare | Fail to Operate | Delay of Treatment | 1
adjustment failed | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
air flow blocked | Blockage within Device or Device Component | Inadequate Support | Discomfort | 2
allergic composition | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
allergic reaction | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
anatomical shape failed | Patient-Device Incompatibility | Improper Fit or adjustment | Loss of Treatment | 2
articulating bar low rigidity | Material Rigid or Stiff | Loss of Mechanical Integrity | Loss of Treatment | 2
articulating bar low rigidity | support failed | Inadequate Support | Delay of Treatment | 1
attach loosely | Detachment of Device Component | Fail to Operate | Loss of Treatment | 2
authentication failed | Unauthorized Access to Computer System | Misuse | Injury to Patient | 2
belt mismatched | Component Incompatible | Fail to Operate | Delay of Treatment | 1
belt size out of specification | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
Bio-contamination | Device Contamination with Biological Material | Skin Allergy or Irritation | Skin Irriation | 2
biological substance | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
biomechanics testing failed | Incorrect or Inadequate Result | Loss of Mechanical Integrity | Loss of Treatment | 2
brace control stop failed | Device Operates Differently than Expected | Inappropriate Treatment | Injury to Patient | 2
brace mismatched | Component Incompatible | Fail to Operate | Delay of Treatment | 1
brace misused | Use of Device Issue | Misuse | Loss of Treatment | 2
brace uncomfortable | Human-Device Interface Issue | Inadequate Support | Discomfort | 2
breathable failed | Use of Device Issue | Misuse | Loss of Treatment | 2
breathable material compromised | Material Integrity Issue | Material Degraded | Loss of Treatment | 2
buckle broken | Break | Loss of Mechanical Integrity | Injury to Patient | 2
CFR regulation violated | Compliance Issue | Loss of Compliance | No Harm | 1
clinical setting changed | Use of Incorrect Control Settings | Abnormal Treatment | Injury to Patient | 2
COG out of range | Market Issue | Inconvenience | No Harm | 1
color deviation | Use of Device Issue | Inconvenience | No Harm | 1
connector mismatched | Component Incompatible | Fail to Operate | Delay of Treatment | 1
covering ankle undesired | Human Factors Issue | Abnormal Treatment | Tissue Necrosis | 3
cross-border selling | Market Issue | Inconvenience | No Harm | 1
design lack of intuitively implemented | Inadequate User Interface | Misuse | Delay of Treatment | 1
design out of date | Inadequate User Interface | Misuse | Delay of Treatment | 1
device damaged by inappropriate cleaning | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
device malfunction | Device Operates Differently than Expected | Product Malfunction | Injury to Patient | 2
difficult to apply brace | Human-Device Interface Issue | Brace Applied Inappropriately | Loss of Treatment | 2
difficult to apply device | Use of Device Issue | Brace Applied Inappropriately | Injury to Patient | 2
difficult to position | Difficult to Position | Fail to Operate | Delay of Treatment | 1
difficult to remove | Difficult to Remove | Fail to Operate | Delay of Treatment | 1
difficult to setup | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
download failed | Problem with Software Installation | Fail to Operate | Loss of Treatment | 2
download failed | Use of Device Issue | Inconvenience | No Harm | 1
duty cycle low | Device Operates Differently than Expected | Abnormal Treatment | Injury to Patient | 2
electrode out of condition | Environmental Control or Utility Issue | Material Degraded | Loss of Treatment | 2
electrode washed with belt | Device Cleaning Issue | Fail to Operate | Loss of Treatment | 2
exercise over activity limit | Incorrect Measurement | Inappropriate Treatment | Injury to Patient | 2
expired electrode | Device Expiration Issue | Material Degraded | Loss of Treatment | 2
fail to access | Human-Device Interface Issue | Fail to Operate | Delay of Treatment | 1
fail to access | Human-Device Interface Issue | Fail to Operate | Loss of Treatment | 2
fail to access | Human-Device Interface Issue | Inconvenience | Delay of Treatment | 1
fail to adjust size | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
fail to air dry | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
fail to avoid inversion movement | Unintended Movement | Unintend Motion | Injury to Patient | 2
fail to be washed | Device Cleaning Issue | Abrasion to User | Skin Irriation | 2
fail to be washed | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
fail to downgrade stability along with treatment | Unstable | Loss of Structural Integrity | Injury to Patient | 2
fail to email privacy policy | Use of Device Issue | Inconvenience | No Harm | 1
fail to email Terms & Conditions | Use of Device Issue | Inconvenience | No Harm | 1
fail to fit into shoes | Fitting Problem | Improper Fit or adjustment | Loss of Treatment | 2
fail to hand wash | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
fail to identify the user | Application Security Issue | Misuse | Injury to Patient | 2
fail to indicate battery level | Battery Issue | Loss of Data | Delay of Treatment | 1
fail to log out | Use of Device Issue | Inconvenience | No Harm | 1
fail to monitor health and status | Protective Measure Issue | Misuse | Injury to Patient | 2
fail to replace component | Device Maintenance Issue | Fail to Operate | Delay of Treatment | 1
fail to respect reimbursement | Market Issue | Inconvenience | No Harm | 1
fail to support light sport | Use of Device Issue | Inadequate Support | Delay of Treatment | 1
fail to take photo | Use of Device Issue | Inconvenience | No Harm | 1
flexion degree out of range | Defective Component | Material Degraded | Injury to Patient | 2
flexion degree out of range | Defective Component | Unintend Motion | Injury to Patient | 2
grey-trade | Market Issue | Inconvenience | No Harm | 1
group electrodes wrong | Component Incompatible | Fail to Operate | Delay of Treatment | 1
healthcare data breached | Application Security Issue | Product Malfunction | Injury to Patient | 2
heavy weight | Patient-Device Incompatibility | Improper Fit or adjustment | Discomfort | 2
high cost | Market Issue | Inconvenience | No Harm | 1
high resistance | Physical Resistance | Inconvenience | No Harm | 1
IEC regulation violated | Compliance Issue | Loss of Compliance | No Harm | 1
immobilization compromised | Unstable | Patient mobilizated | Loss of Treatment | 2
improper or incorrect use of device | Use of Device Issue | Misuse | Loss of Treatment | 2
inability to fully insert | Difficult to Insert | Fail to Operate | Delay of Treatment | 1
inaccurate accelerometer | Incorrect Measurement | User Confused | No Harm | 1
inaccurate counter number | Incorrect Measurement | Inappropriate Treatment | Injury to Patient | 2
inaccurate height adjustment | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
inaccurate measurement | Incorrect Measurement | Inappropriate Treatment | Injury to Patient | 2
inadequate caution | Device Markings Issue | Inadquate IFU or Labeling | Injury to Patient | 2
inadequate device information | Device Markings Issue | Misuse | Injury to Patient | 2
inadequate DHF | Compliance Issue | Loss of Compliance | No Harm | 1
Inadequate extension | Mechanical Issue | Inadequate Support | Delay of Treatment | 2
inadequate IFU | Instruction for Use Issue | Brace Applied Inappropriately | Loss of Treatment | 2
inadequate IFU | Instruction for Use Issue | Inadequate IFU or Labeling | No Harm | 1
inadequate IFU | Instruction for Use Issue | Inadequate IFU or Labeling | Injury to Patient | 2
inadequate instruction | Inadequate instructions for healthcare professional | Inadequate IFU or Labeling | Delay of Treatment | 1
inadequate instruction | Inadequate instructions for healthcare professional | Misuse | No Harm | 1
inadequate instruction | Inadequate instructions for non-healthcare professional | Inadequate IFU or Labeling | Delay of Treatment | 1
inadequate instruction | Inadequate instructions for non-healthcare professional | Inadequate IFU or Labeling | Injury to Patient | 2
inadequate instruction | Inadequate instructions for non-healthcare professional | Misuse | Burn, Thermal | 2
inadequate labeling | Device Markings Issue | Fail to Operate | Delay of Treatment | 1
inadequate labeling | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
inadequate labeling | Device Markings Issue | Misuse | Injury to Patient | 2
inadequate screen graph | Incorrect Display | User Confused | No Harm | 1
inadequate size | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
inadequate size range to meet market demand | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
inadequate training | Inadequate Training | Misuse | Injury to Patient | 2
inadequate user manual | Use of Device Issue | Brace Applied Inappropriately | Injury to Patient | 2
inadequate user manual | Use of Device Issue | Misuse | Skin Irriation | 2
inappropriate clinical evaluation | Compliance Issue | Loss of Compliance | No Harm | 1
inappropriate color scheme | Use of Device Issue | Inconvenience | No Harm | 1
inappropriate distribute weight | Patient-Device Incompatibility | Patient immobilizated | Restricted Blood Circulation | 2
inappropriate environment condition | Environmental Control or Utility Issue | Product Malfunction | Injury to Patient | 2
inappropriate fit | Fitting Problem | Improper Fit or adjustment | No harm | 1
inappropriate fit | Fitting Problem | Improper Fit or adjustment | Discomfort | 2
inappropriate fit | Fitting Problem | Improper Fit or adjustment | Loss of Treatment | 2
inappropriate interface design | Inadequate User Interface | Misuse | Delay of Treatment | 1
inappropriate operation environment | Use of Device Issue | Fail to Operate | Loss of Treatment | 2
inappropriate stability change | Unstable | Loss of Structural Integrity | Loss of Treatment | 2
inappropriate storage environment or condition | Environmental Control or Utility Issue | Material Degraded | Loss of Treatment | 2
inappropriate symbol | Device Markings Issue | User Confused | No Harm | 1
inappropriate user | Use of Device Issue | Misuse | Loss of Treatment | 2
inappropriate user | User used incorrect product for intended use | Abnormal Treatment | Tissue Necrosis | 3
incompatible accessory | Accessory Incompatible | Fail to Operate | Delay of Treatment | 1
incompatible software versions | Application Program Version or Upgrade Problem | Abnormal Treatment | Injury to Patient | 2
incorrect brace size | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
incorrect brand | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
incorrect labeling | Device Markings Issue | Misuse | Injury to Patient | 2
incorrect logo | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
incorrect manufacturer address | Device Markings Issue | User Confused | Delay of Treatment | 1
incorrect offset value | No Display or Display Failure | Inconvenience | No Harm | 1
incorrect phone number | Device Markings Issue | User Confused | Delay of Treatment | 1
incorrect prescribed exercise | Inadequate instructions for healthcare professional | Inadequate IFU or Labeling | Delay of Treatment | 1
incorrect product classification | regulatory Issue | Loss of Compliance | No Harm | 1
incorrect software reference number | Application Program Version or Upgrade Problem | Abnormal Treatment | Injury to Patient | 2
insufficient fatigue cycle | Material Integrity Issue | Fail to Operate | Loss of Treatment | 2
insufficient package | Device Packaging Compromised | Fail to Operate | Delay of Treatment | 1
insufficient package | Device Packaging Compromised | Inadequate Support | Loss of Treatment | 2
insufficient package | Device Packaging Compromised | Material Degraded | Loss of Treatment | 2
insufficient sample size | Use of Incorrect Control Settings | Loss of Product Safety | Injury to Patient | 2
insufficient stability | Unstable | Loss of Structural Integrity | Loss of Treatment | 2
invalid pin number | Device Markings Issue | Fail to Operate | Loss of Treatment | 2
invalid UDI | Device Markings Issue | User Confused | Delay of Treatment | 1
invisible indication | No Display or Display Failure | Inconvenience | No Harm | 1
ISO regulation violated | Compliance Issue | Loss of Compliance | No Harm | 1
knee joint motion out of range | Unintended System Motion | Unintended Motion | Injury to Patient | 2
lack of care instructions | Inadequate instructions for non-healthcare professional | Material Degraded | Loss of Treatment | 2
lack of CE mark | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
lack of CE mark | Instruction for Use Issue | Inadequate IFU or Labeling | No Harm | 1
lack of contraindications | Instruction for Use Issue | Loss of Notification | Injury to Patient | 2
lack of contraindications | Instruction for Use Issue | Loss of Notification | Swelling | 2
lack of ductile | Material Rigid or Stiff | Loss of Mechanical Integrity | Loss of Treatment | 2
lack of indication | Instruction for Use Issue | Loss of Notification | Injury to Patient | 2
lack of labeling | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
lack of local language | Instruction for Use Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
lack of position indicator | Positioning Issue | Improper Fit or adjustment | Loss of Treatment | 2
lack of warning | Instruction for Use Issue | Loss of Notification | Injury to Patient | 2
lack of warning | Instruction for Use Issue | Loss of Notification | Swelling | 2
launch delayed | Market Issue | Inconvenience | No Harm | 1
less flexible | Material Rigid or Stiff | Brace Applied Inappropriately | Loss of Treatment | 2
less flexible | Material Rigid or Stiff | Loss of Mechanical Integrity | Loss of Treatment | 2
less size | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
less sticky | Detachment of Device Component | Fail to Operate | Delay of Treatment | 1
loss of comfort | Human-Device Interface Issue | Inadequate Support | Discomfort | 2
loss of connection | Connection Issue | Fail to Operate | Delay of Treatment | 1
loss of control | Calibration Issue | Abnormal Treatment | Injury to Patient | 2
Loss of elasticity | Device Cleaning Issue | Material Degraded | Injury to Patient | 2
loss of image | Image Display Error | User Confused | No Harm | 1
loss of indication | Application Program Version or Upgrade Problem | Fail to Operate | Loss of Treatment | 2
loss of indication | Instruction for Use Issue | Inadequate IFU or Labeling | No Harm | 1
loss of indication | No Display or Display Failure | Inconvenience | Delay of Treatment | 1
loss of indication | No Display or Display Failure | Inconvenience | No Harm | 1
loss of indication | No Display or Display Failure | User Confused | Delay of Treatment | 1
loss of message | Device Alarm System Issue | Inconvenience | No Harm | 1
loss of reimbursement | Market Issue | Inconvenience | No Harm | 1
loss of sensor data | Loss of Data | Inconvenience | No Harm | 1
loss of support | Use of Device Issue | Inadequate Support | Loss of Treatment | 2
loss of wireless communication | Wireless Communication Issue | Loss of Data | Delay of Treatment | 1
low battery | Battery Issue | Fail to Operate | Loss of Treatment | 2
low battery | Low Battery | Loss of Data | Delay of Treatment | 1
low level package type | Device Packaging Compromised | Material Degraded | Loss of Treatment | 2
low motion resolution | Imprecision | Unintend Motion | Injury to Patient | 2
low stiffness | Material Rigid or Stiff | Loss of Mechanical Integrity | Loss of Treatment | 2
low strength | Mechanical Issue | Loss of Mechanical Integrity | Loss of Treatment | 2
market share low | Market Issue | Inconvenience | No Harm | 1
material biocompatibility compromised | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
material bio-incompatibility | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
material degradation by inappropriate cleaning | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
material deviation | Material Integrity Issue | Loss of Material Integrity | Injury to Patient | 2
material deviation | Material Integrity Issue | Material Degraded | Loss of Treatment | 2
material flexibility compromised | Material Integrity Issue | Material Degraded | Loss of Treatment | 2
material life short | Material Integrity Issue | Material Degraded | Injury to Patient | 2
measure activity failed | Incorrect Measurement | Inappropriate Treatment | Injury to Patient | 2
missing brand | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
missing brand marking | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
missing contraindication | Inadequate instructions for healthcare professional | Loss of Notification | Tissue Necrosis | 3
missing IFU in package | Instruction for Use Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
missing legal manufacturer | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
missing moisture wicking material | Material Integrity Issue | Abrasion to User | Discomfort | 2
missing part | Component Missing | Inadequate Support | Loss of Treatment | 2
missing part number | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
missing replacement tool | Device Maintenance Issue | Fail to Operate | Delay of Treatment | 1
missing side effect notes | Use of Device Issue | Skin Allergy or Irritation | Swelling | 2
missing soft goods | Component Missing | Inadequate Support | Loss of Treatment | 2
missing splint | Component Missing | Inadequate Support | Loss of Treatment | 2
missing symbol | Device Markings Issue | Inadequate IFU or Labeling | Delay of Treatment | 1
misuse | Use of Device Issue | Misuse | Loss of Treatment | 2
modular support failed | Use of Device Issue | Inadequate Support | Discomfort | 2
out of dimension | Defective Component | Improper Fit or adjustment | Loss of Treatment | 2
over bending | Bent | Unintend Motion | Loss of Treatment | 2
over load | Device Damaged Prior to Use | Fail to Operate | Loss of Treatment | 2
over utilization | Use of Device Issue | Misuse | Injury to Patient | 2
part mismatched | Component Incompatible | Fail to Operate | Delay of Treatment | 1
part torn | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
part torn | Torn Material | Loss of Material Integrity | Loss of Treatment | 2
part visible undesired | Human-Device Interface Issue | Inconvenience | No Harm | 1
part washing damaged | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
patient data breached | Application Security Issue | Product Malfunction | Injury to Patient | 2
position adjustment difficult | Difficult to Position | Fail to Operate | Delay of Treatment | 1
power out of range | Electrical Power Problem | Loss of Data | Delay of Treatment | 1
prescribed exercise confused | Improper or Incorrect Procedure or Method | Inadequate IFU or Labeling | Injury to Patient | 2
product damaged by inappropriate washing | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
product life short | Device Operates Differently than Expected | Misuse | Injury to Patient | 2
product life short | Material Integrity Issue | Material Degraded | Injury to Patient | 2
prompt 911 message failed | Device Alarm System Issue | Loss of Notification | Injury to Patient | 2
prompt failed | Inappropriate Prompts | User Confused | No Harm | 1
pull strength low | Mechanical Issue | Loss of Mechanical Integrity | No Harm | 1
pull strength low | Mechanical Issue | Loss of Mechanical Integrity | Loss of Treatment | 2
recover delayed | Mechanical Issue | Inadequate Support | Delay of Treatment | 1
recovery status confused | Inadequate User Interface | User Confused | No Harm | 1
reduced life cycle | Device Maintenance Issue | Fail to Operate | Loss of Treatment | 2
reduced service life | Device Maintenance Issue | Fail to Operate | Loss of Treatment | 2
reduced washable | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
reduced wearing time | Use of Device Issue | Inadequate Support | Discomfort | 2
reducing inversion or eversion movement failed | Unintended Movement | Unintend Motion | Injury to Patient | 2
release delayed | Market Issue | Inconvenience | No Harm | 1
remove difficult | Difficult to Remove | Fail to Operate | Delay of Treatment | 1
scratched parts | Device Handling Issue | Loss of Material Integrity | Injury to Patient | 2
self-locking mode failed | Fail-Safe Mechanism Issue | Inappropriate Treatment | Injury to Patient | 2
sensor damaged | Break | Loss of Data | Delay of Treatment | 1
sensor data breached | Application Security Issue | Misuse | Injury to Patient | 2
sensor measurement interfered | Incorrect Measurement | Abnormal Treatment | Injury to Patient | 2
set time failed | Use of Device Issue | Inconvenience | No Harm | 1
shipping test failed | Device Packaging Compromised | Loss of Material Integrity | Injury to Patient | 2
single hand adjustment failed | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
single use violated | Use of Device Issue | Device Reused | Skin Irriation | 2
size chart confused | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
size chart confused | Size Incorrect for patient | User Confused | No Harm | 1
skin scratched | Human-Device Interface Issue | Abrasion to User | Bleed | 2
skip failed | Use of Device Issue | User Confused | No Harm | 1
slip coating | Split | Loss of Material Integrity | Loss of Treatment | 2
sticky button | Human-Device Interface Issue | Product Malfunction | Injury to Patient | 2
strap shift | Migration of Device or Device Component | Unintend Motion | Injury to Patient | 2
support compromised | Use of Device Issue | Inadequate Support | Loss of Treatment | 2
support failed | Use of Device Issue | Inadequate Support | Loss of Treatment | 2
symbol loss of explanation | Device Markings Issue | User Confused | No Harm | 1
temperature out of range | Temperature Issue | Fail to Operate | Loss of Treatment | 2
TENS accessory incompatible | Accessory Incompatible | Fail to Operate | Delay of Treatment | 1
TENS accessory incompatible | Accessory Incompatible | Improper Fit or adjustment | Loss of Treatment | 2
thickness out of specification | Size Incorrect for patient | Improper Fit or adjustment | Discomfort | 2
tightening adjustment difficult | Device Difficult to Setup or Prepare | Improper Fit or adjustment | Discomfort | 2
unclear symbol | Device Markings Issue | User Confused | No Harm | 1
uncomfortable | Human-Device Interface Issue | Inadequate Support | Discomfort | 2
undesired talus movement | Unintended Movement | Improper Fit or adjustment | Loss of Treatment | 2
undesired talus movement | Unintended Movement | Unintend Motion | Injury to Patient | 2
unique number violated | Device Markings Issue | Misuse | Delay of Treatment | 1
unmatched belt set | Patient-Device Incompatibility | Improper Fit or adjustment | Loss of Treatment | 2
unreadable message | Device Alarm System Issue | Inconvenience | No Harm | 1
unsecured data | Application Security Issue | Misuse | Injury to Patient | 2
unsecured hinge | Detachment of Device or Device Component | Unintend Motion | Injury to Patient | 2
unstable ankle brace | Unstable | Loss of Structural Integrity | Loss of Treatment | 2
unsticking gel | Failure to adhere or bond | Fail to Operate | Loss of Treatment | 2
untimely recovery status | Unable to Access | Loss of Data | Delay of Treatment | 1
usability test failed | Use of Device Issue | Brace Applied Inappropriately | Loss of Treatment | 2
usage misaligned | Use of Device Issue | Misuse | Loss of Treatment | 2
user inflammation or allergy | Biocompatibility Issue | Skin Allergy or Irritation | Skin Irriation | 2
user uncomfortable | Human-Device Interface Issue | Inadequate Support | Discomfort | 2
void warranty | Market Issue | Inconvenience | No Harm | 1
wash cycle high | Device Cleaning Issue | Loss of Structural Integrity | Discomfort | 2
wash cycle high | Device Cleaning Issue | Material Degraded | Loss of Treatment | 2
weak immobilization | Unstable | Patient mobilizated | Loss of Treatment | 2
weak interface strength | Mechanical Issue | Loss of Mechanical Integrity | Loss of Treatment | 2
wear for prolong periods | Use of Device Issue | Bacteria Growth | Infection | 3
wear test failed | Use of Device Issue | Improper Fit or adjustment | Discomfort | 2
weight out of range | Patient-Device Incompatibility | Improper Fit or adjustment | Discomfort | 2
weight out of specification | Patient-Device Incompatibility | Improper Fit or adjustment | Discomfort | 2

## unknown_categorylabel
Dental | teeth prosthesis
hip prosthesis
ankle prosthesis
elbow prosthesis
Ear, Nose, and Throat | hearing aid
Nasal dilator
Otoscope
Neurological | EEG telemetry system
Thermal system for insomnia
Neurosurgical headrests
Ophthalmic | Fornixscope
Stereoscope
Eye sphere implant
Ophthalmic laser
Electronic vision aid
Radiology | X-ray radiation therapy system
Nuclear tomography system
Radiologic table
General hospital | Ice bag
Steam sterilizer
Ultrasonic cleaner
Neonatal incubator
Cardiovascular | Trocar
Apex cardiographic transducer
Vascular clip
Gastroenterology and Urology | Stomach pH electrode
Urine flow or volume measuring system
Urological table and accessories
Ostomy irrigator
Rectal control system

## IV_Set
Failure_Mode | Failure_Cause | Hazard | Hazardous_Situation | Harm | Severity | Mitigation
filter leakage; connector leakage | Loss of the hydrophobic properties due to infusate interactions; filter torn | fluid leak | inadequate medication | inadequate pain relief
| administer at a high rate | fluid leak | inadequate medication | hypoglycemia
|  | fluid leak | inadequate medication | dehydration
|  | fluid leak | inadequate medication | medication delayed
|  | fluid leak | hazardous drug exposed | skin reaction
|  | fluid leak | skin contact | skin reaction
|  | fluid leak | hazardous drug spilled on patient | Contamination During Use
bag leakage |  | fluid leak | odor | discomfort, dizziness, nausea, headache, burning eyes |  | double bag
bag spike dropped | [5 empty] | clip system

## Viewer
Category | Inputs | Failure_Modes | Hazard | Hazardous Situation | Harm | Severity
Market | Pre-op planning is a critical step to ensure safety, efficacy and efficiency of resection procedures. | inadequate pre-op planning | application program issue | Inappropriate Pre-Operative Plan | Prolonged Surgery | S2 | Harm | Severity
Market | Pre-op planning is intended to inform 3 clinical decisions as framed by surgeons: “where to enter”, “how to move” and the “amount of resection”. | incorrect clinical decision | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Major hepatic hemorrhage | serious | S3
Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 1. Size, location and vascularity of the tumor(s) | poor tumor image | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Postoperative liver failure | serious | S5
Market | Larger tumors are more concerning for spread. Anything above 6 or 7 cm might mean a full removal and how it sits on the kidney in relationship to other organs. | tumor removed inappropriately | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Central Nervous System Event (CNS) | serious | S4
Market | Where is the cancer? Has it invaded the outside wall? Where has it invaded? What structures have been invaded? | fail to visualize cancer location | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Systemic Sepsis | serious | S5
Market | My focus is to see the proper anatomy of the mass and its vascular anatomy | poor anatomy image | optical issue | Poor Patient Image during pre-operative | Pro-operative Delayed | S1 | Nerve Injury | serious | S4
Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 2. Tumor relationship to surrounding critical structures (i.e. vasculature, ureter, chest wall, etc.) | fail to visualize surrounding critical structures | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Peripheral Neuropathy | serious | S4
Market | I need to avoid critical structures like the ureter. Knowing where it is can speed things up for us. | fail to visualize surrounding critical structures | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Acute Bleeding | serious | S3
Market | I’m looking for the relationship of tumor to surrounding structures, particularly the vasculature. The hepatic veins and vena cava are the bigger ones for liver | fail to visualize surrounding critical structures | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Major Vascular Injury | serious | S3
Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 3. Patient-specific anatomical variations or aberrant anatomy. | fail to visualize anatomical variation | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Bronchial Injury | serious | S3
Market | There is a great deal of variety within the anatomy. This is what we're looking for in the CT | fail to visualize anatomy variety | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Acute Respiratory Failure | serious | S5
Market | Because of the anatomical variability, we are often operating blindly. We don't know how many bronchi we will come into.” | fail to visualize anatomy variety | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2 | Additional Surgical Procedure | serious | S3
Market | By integrating 3D models and patient simulation from Visible Patient into J&J ecosystem, we’ll provide the following value to the clinicians: | inadequate 3D model | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2 | Prolonged Surgery | serious | S2
Market | we’ll provide the following value to the clinicians: Highly detailed, virtual 3D model, using patients’ pre-op images (CT, MRI) | fail to develop detailed 3D model | application program issue | Inappropriate Patient 3D Model during pre-operative | Pro-operative Delayed | S1 | Pro-operative Delayed | Minor | S1
Market | we’ll provide the following value to the clinicians: Color-coded, interactive experience that allows you to explore and visualize spatial relationships between tumors, critical structures and surrounding anatomy of your patient | fail to visualize spatial relationship | image display error | unable to Manipulate 3D Model during pre-operative | Pro-operative Delayed | S1 | No harm | Negligible | S0
Market | we’ll provide the following value to the clinicians: Easy-to-use interface that is accessible on your team’s PC/Mac, tablet, and mobile devices for reference in the OR, at home, and on-the-go | difficult to use | human-device interface issue | No Hazardous Situation | No harm | S0
Market | we’ll provide the following value to the clinicians: Allows you to explore anatomy on your schedule, without reliance on interpretation by your radiology team | fail to explore anatomy independently | human-device interface issue | User Inconvenience | No harm | S0
Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to provide virtual clip-applying tool | human-device interface issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to real time assess organ | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Market | Surgeons use 2D images to “envision a 3D model” in their mind when planning their approach. More experienced surgeons are more likely to be confident in this ability. However, even experienced surgeons acknowledge inefficiency and communication challenges with the surgical team when developing a surgical plan, e.g. reliance on the radiologist to develop a surgical plan. | fail to develop surgical plan independently | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Market | In addition to 3D models, 2D visualization, with color, of a 3D model (Anatomical Atlas) is also a valuable tool for clinicians as it is “what I’m used to and know how to interpret, but can see things better”. | inappropriate color code | use of device issue | Inappropriate Patient 3D Model during pre-operative | Prolonged Surgery | S2
Market | The patient simulation (Clip Applying) tool provides significant value when planning painstaking, tissue-sparing procedures. For many, the patient simulation tool is considered a “breakthrough” because of high-quality image and instant volume calculations that allow for simulation of different surgical plans pre-operatively (e.g. segmentectomies vs a lobectomy). | fail to provide real time organ volume calculation | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | inefficacy pre-op planning | application program issue | Inappropriate Pre-Operative Plan | Prolonged Surgery | S2
Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | unclear simulation | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Market | Surgeons see the flexibility in access across devices and locations (at home, at the office, on the go, in the OR) as a critical attribute of the solution | fail to access across devices and locations | human-device interface issue | Inaccessible Patient Data during pre-operative | Pro-operative Delayed | S1
Market | All surgeons for the target specialties (thoracic, HBP, colorectal and urology) are procedurally required to review pre-operative patient images before the procedure. However, only a subset of surgeons refers back to pre-operative patient images intra-operatively, mostly due to an unexpected finding or uncertainty. Hence, the primary use case for pre-operative planning is before the procedure begins; intraoperative reference of the preoperative planning is secondary. | missing review of pre-operative patient images | use of device issue | Inaccessible Patient Image during intra-operative | Prolonged Surgery | S2
Market | The primary clinical goals of this project are to design, develop and deploy a solution that will enable Digital Solution to ● launch surgical decision support, especially medical imaging software application products | incorrect surgical decision | application program issue | Delayed Pre-Operative Plan | Pro-operative Delayed | S1
Market | enable surgeons to order, manage, and review patient-specific 2D and 3D anatomical models | fail to manage 3D model | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Market | enable surgeons to view pre-operative images and 2D and 3D anatomical models online | fail to manage 2D images online | application program issue | Inaccessible Patient Image during intra-operative | Prolonged Surgery | S2
Market | enable surgeons to perform a simulated resection by applying virtual clips to approximated segments in order to simulate post-operative organ volume | fail to provide real time organ volume calculation | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Market | enable surgeons to save a viewing and simulation session and continue at a later time | fail to continue the paused session | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Market | enable surgeons to continue their viewing and simulation session in the OR using the Digital Hub hardware device after the commercial release of the Digital Hub (viewing of data sets on the Digital Hub will be outlined in a future Charter), | insufficient memory size | data back-up problem | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Market | enable surgeons to continue viewing their images and simulation in the OR on the web portal | fail to access through web portal | human-device interface issue | Inaccessible Patient Image during pre-operative | Pro-operative Delayed | S1
Market | Integrated solution should be comprised of the following major components: surgeon portal, digital solutions viewer, digital surgery enablement platform. | fail to access through web portal | application program issue | Inaccessible Patient Image during pre-operative | Pro-operative Delayed | S1
Market | Surgeon Portal - a web portal for surgeons to log into, select the appropriate language, order the modeling service, review the status of the modeling service, upload DICOM images, start the Digital Solutions Viewer. | insecure login | unauthorized access to computer system | Corrupted Pre-Operative Plan | Prolonged Surgery | S2
Market | Digital Solutions Viewer - a DICOM viewer that allows for interactions with the anatomical model, allow for simulation of resection by placing a virtual clip. | fail to provide simulation resection | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Market | Digital Solutions Viewer - a DICOM viewer that allows for viewing a DICOM image and/or the 3D anatomical model selected by the surgeon through the surgeon portal. | fail to view images through web portal | application program issue | Inaccessible Patient Image during intra-operative | Prolonged Surgery | S2
Market | Digital Solutions Viewer - a DICOM viewer allows for placement of annotations and measurements, show the anatomical model, allow for interactions with the anatomical model | fail to place measurements | application program issue | unable to Manipulate 3D Model during pre-operative | Pro-operative Delayed | S1
Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables user management, data storage | fail to manage users | application program issue | User Confusion | No harm | S0
Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables data storage, and provides DICOM studies to the external 3D model creation service as well as to the Digital Solutions Viewer. | fail to provide external 3D model creation | application program issue | Inappropriate Patient 3D Model during pre-operative | Pro-operative Delayed | S1
Market | Bandit Phase 2 will deliver an integrated experiences for the customer, which can be categorized in three groups: ● 3D Model Submission and Delivery ● Digital Solution Viewer ● Patient Modeling and Outcome Simulation | fail to deliver 3D model | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
Market | AI-enhanced visualization and personalized insights to help surgeons get a clearer roadmap for surgery | incorrect surgical decision | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Market | 3D model submission and delivery: 3D model order submission by clinical user; 3D model order management by order admin; 3D model order billing; PACS integration | inappropriate user | application program issue | User Confusion | No harm | S0
Market | Image viewer: image viewer; meaurement tools | lack of measurement tool | application program issue | unable to Manipulate 3D Model during pre-operative | Pro-operative Delayed | S1
Market | patient modeling and outcome simulation: integrated 3D models; patient specific simulation of resection | fail to provide simulation resection | application program issue | Inappropriate Patient 3D Model during pre-operative | Prolonged Surgery | S2
Market | Target for the Digital Solutions Viewer is the EU. The Digital Solutions Viewer will be the first CE marked product and first product under the new EU Medical Device Regulation (MDR) for Auris. | fail to comply with CE Mark | compliance issue | No Hazardous Situation | No harm | S0
Market | The Digital Solutions Viewer is considered software as a medical device (SaMD) and will need to conform to the requirements of the Medical Device Regulation (EU) 2017/745. The Viewer will need to provide Technical Documentation (TD) covering the design, development, and state of the art with respect to safety and performance and demonstrate compliance to the General safety and Performance requirements in Annex I. | IEC 62304 violated | computer software issue | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Market | The proposed classification of the Digital Solutions Viewer is a Class IIa Rule 11 (Software intended to provide information which is used to make decisions with diagnosis or therapeutic purposes). | inappropriate medical device classification | compliance issue | No Hazardous Situation | No harm | S0
Market | As the Digital Solutions viewer will be the first MDR product for Auris, the Auris QMS must be MDR ready to support certification. The proposed certificate scope will be for image visualization and planning software. | inappropriate certificate scope | compliance issue | No Hazardous Situation | No harm | S0
Market | The MDR stipulates that the manufacturer must provide sufficient clinical evidence (Article 61) to justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR. | insufficient clinical evidence | data issue | No Hazardous Situation | No harm | S0
Market | The MDR provides a possibility to use clinical data related to a marketed equivalent device (found in the clinical evaluation required for a device) under conformity assessment. | lack of clinical data from equivalent device | data issue | No Hazardous Situation | No harm | S0
Market | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes of the software are intended to be similar for the Auris Viewer. | inconsistent algorithm | programming issue | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Market | The equivalence requirement to demonstrate sufficient levels of access to comparator data should be met as Visible Patient is collaborating with J&J and the Auris Viewer is intended to interface with the Visible Patient Lab 3D recreated models. | fail to access across devices and locations | human-device interface issue | Inaccessible Patient 3D Model during pre-operative | Pro-operative Delayed | S1
Market | Primary User Targets: Surgeons in Thoracic, Urology, Colorectal, and Hepato-biliary | inappropriate user | user used incorrect product for intended use | User Confusion | No harm | S0
Market | Secondary User Targets (secondary expansion within primary user hospitals): Lower volume surgeons in Primary User Target specialties, Surgeons in GYN | inappropriate user | user used incorrect product for intended use | User Confusion | No harm | S0
Market | Builds Loyal Digital Surgery Customer Base: Increases the value of the online Ethicon Digital Solutions video, image & case library, creating network effects & switching costs | lack of case library | inadequate instructions for healthcare professional | User Inconvenience | No harm | S0
Market | Provides Critical Data Pipeline: Adds pre-op image and 3D model data to feed the Ethicon AI Insights platform | incompatible accessory | component or accessory incompatibility | Inaccessible Patient Image during pre-operative | Pro-operative Delayed | S1
Market | Pre-operative 3D planning software is in-market today, but most focus on orthopaedics and neurosurgery. Bandit has three main competitors in the US (Fujifilm Synapse, TeraRecon iNtuition, and Intuitive Surgical IRIS™), but only Fujifilm and TeraRecon have a significant presence in EMEA, Bandit’s initial regional foothold market. Only Fujifilm is used today in APAC. | inadequate pre-op planning | application program issue | Inappropriate Pre-Operative Plan | Prolonged Surgery | S2
Market | Fujifilm and TeraRecon require radiologist creation of 3D models and involvement with interpretation and dictation, whereas Bandit eliminates this step, giving surgeons a streamlined pre-op planning sequence that puts the control in their hands on their schedule. | fail to develop surgical plan independently | application program issue | Inappropriate Pre-Operative Plan | Pro-operative Delayed | S1
Market | Intuitive IRIS™ is currently cleared for use with CT only, whereas the others, including Bandit, are applicable for MRI as well. | fail to use 2D MRI images to envision a 3D model | application program issue | Loss of Patient 3D Model during pre-operative | Pro-operative Delayed | S1
Market | The major project milestones and estimated timing are listed below. This schedule assumes a fully staffed development team and approved program budget by the end of Phase 1. An updated schedule and list of major milestones along with program risks will be provided in the Design and Development Plan (DDP) prior to exiting the Planning and Definition Phase. | fail to identify risk | market issue | No Hazardous Situation | No harm | S0
Market | The following assumptions were made as basis of this Project Authorization Charter: ● CE Mark is prioritized (EMEA launch) and 510K Approval is performed afterwards | inappropriate authorization lists | market issue | No Hazardous Situation | No harm | S0
Market | Expedited Review (MDR) has minimal questions and no major revisions needed | prolonged Eu MDR | compliance issue | No Hazardous Situation | No harm | S0
Market | FDA acceptance of substantial equivalence with minimal questions | prolonged FDR approval | compliance issue | No Hazardous Situation | No harm | S0
Market | Code the complete the Digital Solutions Viewer is solid with few bugs/issues by mid-March 2021 | major bug existed | computer software issue | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Market | Web Portal & DSEP are not part of MDR submission | MDR submission delayed | compliance issue | No Hazardous Situation | No harm | S0
Market | QMS audit can completed by regulatory body by end of May (earlier than planned) | audit observations | compliance issue | No Hazardous Situation | No harm | S0
Market | QMS documents can be reviewed and released in Agile within 2 weeks (documents complete to released) | inadequate QMS document lists | compliance issue | No Hazardous Situation | No harm | S0
Clinical_Workflow | The intended use of the device is an aid in pre-surgical planning for procedures requiring spatial understanding of complex patient anatomy. | fail to provide spatial relationship | application program issue | Inappropriate Pre-Operative Plan | Pro-operative Delayed | S1
Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | inadequate surgical plan | improper or incorrect procedure or method | Delayed Pre-Operative Plan | Pro-operative Delayed | S1
Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor defects | application program issue | Inappropriate Pre-Operative Plan | Prolonged Surgery | S2
Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor abnormalities | application program issue | Inappropriate Pre-Operative Plan | Prolonged Surgery | S2
Clinical_Workflow | Patients undergoing procedures in the general, bariatric, HPN, gynecology, thoracic, and urology specialties, who require medical images to be taken to plan for their surgical procedure. | poor medical image | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | If the medical device fails while being used to reference medical images or 3D models during a procedure, the user may need to access the patient images on another PACS viewer to proceed. | lack of backup viewer | fail-safe mechanism issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | Measurements are taken by the medical device. The medical device can take linear measurements between two points displayed on the DICOM or 3D model | inappropriate measurement algorithm | measurement system incompatibility | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | Measurements are taken by the medical device. The medical device can calculate a designated area shown on a 2D DICOM slice | fail to calculate slice area | incorrect software programming calculations | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | Measurements are taken by the medical device. The medical device can through the clip apply feature, calculate volume of a designated portion of an organ to be resected | fail to calculate volume of resected organ | incorrect software programming calculations | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | Measurements taken on a 2D device should be accurate for images acquired with high resolution. | low resolution image | Image Resolution Poor | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | fail to reconstruct 3D model | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | inappropriate segnmentation alogorithm applied | application program issue | Inappropriate Patient 3D Model during pre-operative | Prolonged Surgery | S2
Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | incorrect volume calculation | parameter calculation error due to software problem | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The medical device may be used with a radiology monitor | fail to compatible with radiology monitor | Incompatibility Problem | unable to Display Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | software installation problem | problem with software installation | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | obsolete software used | application program version or upgrade problem | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | fail to compatible with operating system | Incompatibility Problem | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Clinical | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | short software life cycle | computer software issue | Software Frozen or locked Up during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | When using a tablet to interface with the device during a procedure, there is a risk that the procedure would coincide with battery depletion. | premature battery depletion | Premature Discharge of Battery | User Inconvenience | No harm | S0
Clinical_Workflow | successful application of the medical device is critically dependent on human factors such as the user interface. The user interface is fundamental to interacting with and displaying measurements of the patient images and anatomic models. | inappropriate interface design | inadequate user interface | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | use errors can arise from not understanding or properly navigating the user interface. Symbols for the user interface have been based on those common in other PACS viewer products, but may be arranged differently or applied slightly differently for this medical device. | unclear symbol | device markings issue | User Confusion | Prolonged Surgery | S2
Clinical_Workflow | In general, the medical device is used in the office environment for planning prior to a case. When using the medical device in the operating room, the user may be more distracted and could make use errors such as loading the incorrect patient’s data, or ignoring the orientation indicators. | improper or incorrect use of device | improper or incorrect procedure or method | User Confusion | Prolonged Surgery | S2
Clinical_Workflow | The medical device displays information. The display relies on text, colors, sizing, etc to relay spatial information and distinguish tissue types. These can be adjusted for accessibility but must take colorblind users into account | fail to consider colorblind user | incorrect display | User Discomfort | No harm | S0
Clinical_Workflow | the medical device must be used with a display large enough and with high enough resolution so that the user can clearly see details of the patient images and anatomic models. | low resolution | Image Resolution Poor | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | The device must render the patient models taking all these factors into account. | inappropriate model rendered | Poor Quality Image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to navigate functions | inadequate user interface | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to access settings | inadequate user interface | Inaccessible Patient Data during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall launch program by openning Aurora viewer during the selecting an order phase | fail to launch program | application program issue | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to navigate tab | human-device interface issue | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to select order | human-device interface issue | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall find patient(s) of interest by searching for patient by name, MRN, surgeon during the selecting an order phase | fail to search patient name | application program issue | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall find patient(s) of interest through sorting lists by specialty, patient, surgeon, delivery date, or status to find patient of interest during the selecting an order phase | fail to sort patient interest list | application program issue | Inaccessible Patient Data during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall find patient(s) of interest through filters list using "FILTER BY" tool in order to find patient during the selecting an order phase | fail to filter patient interest list | application program issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall select patient and launch image viewer by touching/clicking on row and row will highlight when selected during the selecting an order phase | unresponsive touch | failure to deploy | User Confusion | No harm | S0
Clinical_Workflow | User shall select patient and launch image viewer by clicking "NEXT" to load patient images and 3D model during the selecting an order phase | unresponsive click | failure to deploy | Software Unresponsive during intra-operative | No harm | S0
Clinical_Workflow | User shall verify images are correct by double checking patient informaiton during the selecting an order phase | fail to verify image | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall view 2D image stacks by monitoring loading of the 2D images on the "Viewing" panel during the load and view desired image stacks phase | fail to monitor loading images | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks. | fail to view image stack | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas during setting screen layout phase | inadequate image types | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select whether they would like the segmented anatomic regions to be displayed by infill or contour when turning on and off segmentation layers | fail to display anatomic contour | application program issue | unable to Manipulate 3D Model during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall adjust view of desired image stacks by operating pan, rotate, zoom, and scroll to select stack they would like to manipulate during the adjust view of desired image stacks phase | unable to adjust view of image stack | application program issue | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall use touchscreen, mouse, or side menu to select "Pan, Rotate, Zoom or Scroll" | insensitive touch screen | decreased sensitivity | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall undo a single adjustment with the "Undo" button during the adjust view of desired image stacks phase | fail to undo an adjustment | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall reset the view altogether with the "Re-Center" button during the adjust view of desired image stacks phase | fail to reset view | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall navigate to the Windowing menu on the sidebar to adjust the contrast or brightness of the image during the adjust view of desired image stacks phase | fail to adjust image contrast | application program issue | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select from presets to auto-adjust window width and level based on target anatomy the image was acquired to show during the adjust view of desired image stacks phase | inappropriate window width | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall choose to select "Width/Level" button to make customize adjustments to the brightness and contrast during the adjust view of desired image stacks phase | fail to adjust image brightness | application program issue | Poor Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall drag slider(s) to desire width and level percentages when adjust windowing during the adjust view of desired image stacks phase | fail to drag slider | human-device interface issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall reset the view using the "Reset" button when adjust windowing during the adjust view of desired image stacks phase | fail to reset view | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select the annotation menu by navigate to tools from the sidebar during the annotation tool phase | inappropriate annotation tool | human-device interface issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall select "distance" tool to measure the straight-line distance between two points on the image during the annotation tool phase | incorrect distance measurement | incorrect measurement | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall click on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen during the annotation tool phase | fail to pick measurement point | human-device interface issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to drag and drop line | human-device interface issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to resize line | human-device interface issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu by using distance tool during the annotation tool phase | fail to add line | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall add and delete arrows in a similar fashion after clicking the "Arrow" menu option. Arrows added to the screen will also appear in the Annotation sidebar menu list by arrow tool during the annotation tool phase | fail to delete arrow | human-device interface issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall add a note by selecting the "note" annotation option and clicking anywhere on the image during the annotation tool phase | fail to add notes at anywhere | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | when a textbox appears, user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu during the annotation tool phase | fail to add any notes | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox during the annotation tool phase | fail to edit notes | application program issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall remove any added note using the trash icon in the Annotation menu or on the note itself during the annotation tool phase | fail to remove any notes | application program issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename arrow | human-device interface issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename distance | human-device interface issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall reset all annotations using the "Reset" button during the annotation tool phase | fail to reset annotations | application program issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle on 3D model | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle off 2D images | application program issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall confirm they are viewing the correct 3D model by checking the information shown onscreen. They toggle on "Patient Info" if not already on when openning 3D Model View during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select "3D Model" menu to see which anatomic structures are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select/deselect the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to deselect 3D model | application program issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall scan the volumes of the different models to check against expectations when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to scan model volume | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships when setting desired orientation during the visualizing and manipulating 3D model phase | inappropriate model orientation | image orientation incorrect | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | Alternatively, user shall select from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view when setting desired orientation during the visualizing and manipulating 3D model phase | fail to auto-orient model | image orientation incorrect | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall enter fullscreen at times for a better view of the model when setting desired orientation during the visualizing and manipulating 3D model phase | fail to enter fullscreen | application program issue | User Inconvenience | No harm | S0
Clinical_Workflow | User shall select an individual structure when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure visibility | optical issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall adjust the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure opacity | optical issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select "Done" button to complete the adjustment when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to done an adjustment | failure to deploy | User Inconvenience | No harm | S0
Clinical_Workflow | User shall select "Clip Apply Model" menu and allows module to load when navigating to Clip Apply feature during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | human-device interface issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | inappropriate model orientation | image orientation incorrect | Inappropriate Patient 3D Model during intra-operative | Pro-operative Delayed | S1
Clinical_Workflow | User shall right click on the model the apply the clip when adding and removing Clips during applying clips to simulate resection volume phase | wrong mouse click | use of device issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall observe the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to place next clip | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | If the user is unhappy with the clip placement they chose, they shall either remove a single clip, or remove all clips when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | At any time, the user shall select any of the clips in the menu in order to review the volume simulation details for that clip when adding and removing Clips when saving the plan during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons when saving the plan during applying clips to simulate resection volume phase | fail to save surgical plan | data back-up problem | Delayed Pre-Operative Plan | Pro-operative Delayed | S1
Clinical_Workflow | If they choose to save the plan, user names the plans and confirms. (Need to check) when saving the plan during applying clips to simulate resection volume phase | fail to name surgical plan | data back-up problem | Delayed Pre-Operative Plan | Pro-operative Delayed | S1
Clinical_Workflow | User shall navigate to "My Views" menu when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall select "SAVE NEW VIEW" and pop-up appears when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" when saving current view for later access during saving current view phase | fail to replace current view | data back-up problem | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
Clinical_Workflow | User shall press "CONFIRM" a second time and the view is saved when saving current view for later access during saving current view phase | fail to confirm current view | application program issue | User Inconvenience | No harm | S0
DIR | Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer. | fail to compatible with regular computer | Incompatibility Problem | unable to Display Patient Image during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | fail to open images | image display error | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | inappropriate image format | data issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
DIR | The system shall have the capability to render 2D views of the preoperative patient images. | unable to render 2D views | application program issue | unable to Manipulate Patient Image during pre-operative | Pro-operative Delayed | S1
DIR | The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study. | unable to render 3D views | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall have the capability to render a view of 3D model segmentations | unable to render 3D views | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view | unable to render 2D overlay on 3D segmentation | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide a method to zoom in and out on the views. | unable to zoom in views | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to pan a viewing panel. | unable to pan viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to rotate a 2D viewing panel. | unable to rotate 2D viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to rotate a 3D viewing panel. | unable to rotate 3D viewing panel | application program issue | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel. | fail to toggle image visibility | application program issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide a method to rotate the cut plane within an MPR view. | unable to rotate cut plane | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to undo and redo user interactions. | fail to undo user interaction | use of device issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to undo and redo user interactions. | fail to redo user interaction | use of device issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size. | unable to resize viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method for hiding a viewing panel. | unable to hide viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to adjust the Window Center and Window Width adjustments using presets. | unable to adjust window center | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method for manually adjusting window width and center values for the intensity range | unable to adjust window width | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method for a reset all viewing panels simultaneously | unable to reset viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to select the layout for the viewing panels | unable to select layout for viewing panel | application program issue | User Inconvenience | No harm | S0
DIR | The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume. | unable to advance MPR cut plane | application program issue | User Inconvenience | No harm | S0
DIR | The system shall be able to display metadata contained in the DICOM data. | fail to display DICOM data | image display error | User Inconvenience | No harm | S0
DIR | The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects. | fail to list 3D models | application program issue | Inaccessible Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model. | fail to display volumetric sizing information | no display or display failure | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall be able to display metadata contained in the 3D model container | fail to display metadata | no display or display failure | User Inconvenience | No harm | S0
DIR | The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels | fail to display semi-transparent overlay | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall use SI units and their derivatives | inappropriate measurement unit | incorrect measurement | Incorrect Patient Data during pre-operative | Pro-operative Delayed | S1
DIR | The system shall provide a method for performing one or more measurements within one of the viewing panels | unable to perform multiple measurements | application program issue | User Inconvenience | No harm | S0
DIR | The system shall label and list measurements automatically | fail to label measurement automatically | application program issue | Incorrect Patient Data during pre-operative | Pro-operative Delayed | S1
DIR | The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object. | fail to place annotation on 3D model | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel. | fail to edit notes | application program issue | User Inconvenience | No harm | S0
DIR | The system shall allow for simulation of resection by applying a clip to approximated segments. | fail to simulate resection | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip. | fail to display volumetric sizing information | no display or display failure | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
DIR | The system shall provide the ability to support individual language packs. | inadequate language pack | inadequate user interface | User Annoyance | No harm | S0
DIR | The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language. | inadequate language pack | inadequate user interface | User Annoyance | No harm | S0
DIR | The system shall, when appropriate and available, use standard internationally recognized symbols. | unstandardized symbol | device markings issue | User Annoyance | No harm | S0
DIR | The system shall, when appropriate and available, use standard internationally recognized icons. | unstandardized icon | device markings issue | User Annoyance | No harm | S0
DIR | The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System. | fail to record operative data | data back-up problem | Loss of Surgical Records during intra-operative | Prolonged Surgery | S2
DIR | The System shall record all functions performed by the System pre, intra, and post operatively. | fail to record operative data | loss of data | Loss of Surgical Records during intra-operative | Prolonged Surgery | S2
DIR | The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel. | fail to record monitoring data | data back-up problem | Loss of Surgical Records during intra-operative | Prolonged Surgery | S2
DIR | The System shall record system functions and operational events generated by the System. | fail to record operational events | data back-up problem | Loss of Surgical Records during intra-operative | Prolonged Surgery | S2
CIR | The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning. | fail to create pre-op planning | application program issue | Delayed Pre-Operative Plan | Pro-operative Delayed | S1
CIR | The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images. | incomplete 3D model | application program issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
CIR | All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated. | incorrect revision | device markings issue | User Annoyance | No harm | S0
CIR | 3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator. | fail to create 3D Model | poor quality image | Poor Patient Image during intra-operative | Prolonged Surgery | S2
CIR | 3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon. | fail to provide patient diagnosis info | use of device issue | Loss of Surgical Records during intra-operative | Prolonged Surgery | S2
CIR | Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components. | software corrupted | computer software issue | User Inconvenience | No harm | S0
CIR | DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data. | image not in DICOM format | image display error | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
CIR | PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images. | fail to process medical imaging | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format. | image not in DICOM format | data issue | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively | fail to refer pre-op plan intraoperatively | application program issue | Delayed Pre-Operative Plan | Prolonged Surgery | S2
CIR | 3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models | fail to upload 3D model | application program issue | Inaccessible Patient 3D Model during intra-operative | Prolonged Surgery | S2
CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | inadequate IFU | inadequate instructions for healthcare professional | User Inconvenience | No harm | S0
CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | lack of video | nadequate instructions for non-healthcare professional | User Inconvenience | No harm | S0
CIR | Marketing and Sales Representative: Internal users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies. | fail to monitor product use | Market Issue | No Hazardous Situation | No harm | S0
CIR | View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to manipulate visualization parameters | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings. | unable to measure point to point | application program issue | User Inconvenience | No harm | S0
CIR | 2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings | unable to develop surgical plan | application program issue | Inappropriate Pre-Operative Plan | Pro-operative Delayed | S1
CIR | View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to review 3D model of pathology | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
CIR | Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings | fail to manipulate visualization parameters | application program issue | unable to Manipulate Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | fail to provide simulation resection | application program issue | unable to Manipulate 3D Model during pre-operative | Pro-operative Delayed | S1
CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | incorrect volume calculation | parameter calculation error due to software problem | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure image display | application program issue | Poor Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure 3D model display | application program issue | unable to Manipulate 3D Model during intra-operative | Prolonged Surgery | S2
CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display 3D model at OR monitor | incompatibility problem | Loss of Patient 3D Model during intra-operative | Prolonged Surgery | S2
CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display patient image at OR monitor | incompatibility problem | Loss of Patient Image during intra-operative | Prolonged Surgery | S2
CIR | Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device. | clinical workflow interrupted | disinfection or sterilization issue at user location | User Inconvenience | Prolonged Surgery | S2
CIR | Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice. | patient data breached | unauthorized access to computer system | Loss of Patient Data during intra-operative | Prolonged Surgery | S2
CIR | Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently. | lack of local language | human-device interface issue | User Inconvenience | No harm | S0
CIR | Indications for Use for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party. | Indications for Use misalignment | failure to align | User Confusion | No harm | S0
CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain regulatory clerance | compliance issue | No Hazardous Situation | No harm | S0
CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain registration | compliance issue | No Hazardous Situation | No harm | S0
CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to capture data | data issue | Loss of Patient Data during intra-operative | Prolonged Surgery | S2
CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to store data | data issue | Loss of Patient Data during intra-operative | Prolonged Surgery | S2
CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to follow privacy regulation | compliance issue | No Hazardous Situation | No harm | S0
CIR | Branding: As company marketing representative, I need the application to follow company defined branding, so that I can maximize consistency in the look and feel of company digital products that may be used by the same user base and reduce cognitive load on the users when using company digital products. | branding misaligned | compliance issue | No Hazardous Situation | No harm | S0
CIR | Demo Capabilities: As company marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively. | unclear workflow | data issue | Inappropriate Patient 3D Model during intra-operative | Prolonged Surgery | S2
Clinical_HA | Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Pro-operative Delayed | S1
Clinical_HA | Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Postoperative liver failure | S5
Clinical_HA | Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Central Nervous System Event (CNS) | S4
Clinical_HA | Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body. | incorrect procedure | bacterial contamination of device | Inappropriate Pre-Operative Plan | Systemic Sepsis | S5
Clinical_HA | Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Nerve Injury | S4
Clinical_HA | Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Peripheral Neuropathy | S4
Clinical_HA | Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Acute Bleeding | S3
Clinical_HA | Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Major Vascular Injury | S3
Clinical_HA | Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Bronchial Injury | S3
Clinical_HA | Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Inappropriate Pre-Operative Plan | Acute Respiratory Failure | S5

## new training data
text | failure mode | hazard
Bag Spike – A bag piercing device that complies with ISO 8536-4. The bag spike product may also have a dry spike adaptor, vent port, one or more access ports, or an IV line |  |  | Bag clips
The Clip, Bag Spike is a family of plastic clips that add an extra protection to secure bag spikes in IV bags, specifically for use in oncology
The clips will be used in pharmacy preparation and transported to the point of care in nursing administration.
The clips will be installed in pharmacy by pharmacists or pharmacy technicians, transported and received on nursing floor for administration to patients. The pharmacists need to know how to secure clip onto the specific IV bags, nurses are just made aware of the clip and its intended use to secure bag spike in the specific IV bags.
All populations receiving IV oncology treatment.
The pharmacist or pharmacy technician fully inserts the ICU Medical bag spike into a specific bag.
The spike is rotated to the user’s preferred orientation.
The spike is used to prepare the oncology drugs/compounds.
The pharmacist or pharmacy technician selects the appropriate clip and snaps it onto the IV bag and Bag Spike. Once the clip is in place, the spike, bag, and clip are a secure unit and remain together until disposal.
The bag is prepared for transport in an individual transport bag that is placed in a bin or carried by hand.
Transport: The bin with multiple bags is transported to the point of care on a cart.
At the point of care, the nurse removes the bag/spike/clip from the transport bag.
At the point of care, The nurse connects a patient line to the bag spike.
At the point of care, The drug is delivered using an IV pump.
At the point of care, The nurse disconnects patient line from the spike.
At the point of care, The bag/spike/clip is disposed.
The bag could be stored at room temperature or in a refrigerator at 4°C for up to 48 hours before it is administered
The transportation may involve a car if pharmacy and point of care are in different buildings.
Clip must secure bag spike in IV bag throughout preparation, storage, transportation, and therapy
Clip must not interfere with any function of the Bag Spike such as additive ports or vents.
Clip must not fall off
Clip must be compatible with ICU Universal Bag Spike.
Clip must be compatible with ICU 2-down bag.
Clip must be compatible with ICU Belly-button bag.
Clip must be compatible with ICU Vis-IV bag.
The product must be biocompatible, based on use case.
Clip must have a 3 year shelf life.
Product must be Reach/RoHS compliant
Clip shall be easily installed.
Spike and bag will meet ISO 15747 retention force (15N for 15 seconds) after worst-case time, temperature, transportation and use conditions
The clip must be free of material in an area of a Ø.55” cylinder, starting .27” from the central axis of the clip and extending outward, centered .275” away from the surface that contacts the spike flange.
Clip shall remain connected after 10 drops from 6-12” above table.
Clip shall not damage the bag or spike.
Device shall function until expiry date
The product shall be subjected to a validated sterilization process
Patient contacting materials shall comply with applicable parts of EN ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process
Product shall be Reach/RoHS compliant.
Product shall remain sterile and work until expiry date.