## impella_IFU
Statements | Failure_Mode | Hazard
INTENDED AUDIENCE: This manual is intended for use by cardiologists, surgeons, nurses, perfusionists, and cardiac catheterization lab technicians who have been trained in the use of the Circulatory Support System. | incorrect intended user | use of device issue
Warnings: Warnings alert you to situations that can cause death or serious injury. The dark gray symbol appears before warning messages. | low level risk added into warning list | protective measure issue
Use of the Circulatory Support System by trained and experienced practitioners has been associated with improved outcomes. Consequently, the first use System should be preceded by the completion of a contemporary training program and include on-site proctoring during the first use by clinical support personnel certified in the use of system. | lack of on-site training | inadequate training
Fluoroscopy is required for the insertion of the guidewire and Impella 5.0 Catheter. | poor fluoroscopic visibility | poor quality image
Be sure that the stopcock on the repositioning sheath is always kept in the closed position. Significant bleed back can result if the stopcock is open. | fail to close stopcock | bleed back
During insertion, avoid manual compression of the inlet, outlet, or sensor areas of the cannula assembly on the Impella 5.0 Catheter. | cannula assembly compressed | device damaged prior to use
During insertion, avoid manual compression of the inlet, outlet, or sensor areas of the cannula assembly on the Impella 5.0 Catheter. | sensor compressed | device handling issue
The sterile components of the Impella 5.0 System can be used only if the sterilization indicators show that the contents have been sterilized, the packaging is not damaged, and the expiration date has not elapsed. | product expired | device disinfection or sterilization issue
The sterile components of the Impella 5.0 System can be used only if the sterilization indicators show that the contents have been sterilized, the packaging is not damaged, and the expiration date has not elapsed. | package damaged | delivered as unsterile product
The sterile components of the Impella 5.0 System can be used only if the sterilization indicators show that the contents have been sterilized, the packaging is not damaged, and the expiration date has not elapsed. | sterilization indicator not visible | disinfection or sterilization issue at user location
Do NOT resterilize or reuse the Impella 5.0 Catheter. It is a disposable device and is intended for single use only. Reuse, reprocessing, or resterilization may compromise the structural integrity of the catheter and/or lead to catheter failure which, in turn, may result in patient injury, illness, or death. | catheter resterilized | material integrity issue
Do NOT resterilize or reuse the Impella 5.0 Catheter. It is a disposable device and is intended for single use only. Reuse, reprocessing, or resterilization may compromise the structural integrity of the catheter and/or lead to catheter failure which, in turn, may result in patient injury, illness, or death. | catheter reprocessed | structural problem
Do NOT resterilize or reuse the Impella 5.0 Catheter. It is a disposable device and is intended for single use only. Reuse, reprocessing, or resterilization may compromise the structural integrity of the catheter and/or lead to catheter failure which, in turn, may result in patient injury, illness, or death. | catheter reused | delivered as unsterile product
Retrograde flow will occur across the aortic valve if the Impella 5.0 Catheter is set at P-level P-0. | retrograde flow | bleed back
Retrograde flow will occur across the aortic valve if the Impella 5.0 Catheter is set at P-level P-0. | retrograde flow | pressure issue
Do NOT use saline in the purge system. | inappropriate purge solution | device cleaning issue
Do NOT use alcohol-based fluids for lubrication | inappropriate lubrication material used | device maintenance issue
Do NOT use an Impella 5.0 System if any part of the system is damaged | damaged device in use | user used incorrect product for intended use
To prevent the risk of explosion, do NOT operate the Impella 5.0 System near flammable anesthetics. | device used near flammable anesthetics | burst
To prevent the risk of explosion, do NOT operate the Impella 5.0 System near flammable anesthetics. | catheter exploded | burst
If at any time during the course of support with the Impella 5.0 Catheter, the Automated Impella Controller alarms “Purge Pressure Low” or “Purge System Open,” follow the instructions presented in this manual. | purge system alarm | device alarm system issue
MR Unsafe - Do NOT subject a patient who has been implanted with an Impella 5.0 Catheter to magnetic resonance imaging (MRI). The strong magnetic energy produced by an MRI machine may cause the Impella 5.0 System components to stop working, and result in injuries to the patient. An MRI may also damage the electronics of the Impella 5.0 System. | catheter malfunctioned | electromagnetic interference (EMI)
MR Unsafe - Do NOT subject a patient who has been implanted with an Impella 5.0 Catheter to magnetic resonance imaging (MRI). The strong magnetic energy produced by an MRI machine may cause the Impella 5.0 System components to stop working, and result in injuries to the patient. An MRI may also damage the electronics of the Impella 5.0 System. | catheter interfered | electromagnetic interference (EMI)
During defibrillation, do NOT touch the Impella 5.0 Catheter, cables, or Automated Impella Controller. | catheter touched inappropriately | use of device issue
During defibrillation, do NOT touch the Impella 5.0 Catheter, cables, or Automated Impella Controller. | cable touched inappropriately | inappropriate shock
Avoid overinserting the Impella 5.0 Catheter and possibly impinging the catheter tip against the walls of the vasculature, atria, or ventricles. | vasculature wall impinged | material perforation
Avoid overinserting the Impella 5.0 Catheter and possibly impinging the catheter tip against the walls of the vasculature, atria, or ventricles. | ventricle wall impinged | material perforation
Avoid overinserting the Impella 5.0 Catheter and possibly impinging the catheter tip against the walls of the vasculature, atria, or ventricles. | catheter overinserted | use of device issue
Do NOT advance or withdraw the Impella 5.0 Catheter against resistance without using fluoroscopy to determine the cause of the resistance. Doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or ventricular perforation. | catheter inserted without fluoroscopy | device handling issue
Do NOT advance or withdraw the Impella 5.0 Catheter against resistance without using fluoroscopy to determine the cause of the resistance. Doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or ventricular perforation. | catheter withdrawed without fluoroscopy | use of device issue
Do NOT advance or withdraw the Impella 5.0 Catheter against resistance without using fluoroscopy to determine the cause of the resistance. Doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or ventricular perforation. | resistance encountered | physical resistance
Cautions indicate situations in which equipment may malfunction, be damaged, or cease to operate. The light gray symbol appears before caution messages. | catheter ceased to operate | device operates differently than expected
Cautions indicate situations in which equipment may malfunction, be damaged, or cease to operate. The light gray symbol appears before caution messages. | inadequate caution | device alarm system issue
Handle with care. The Impella 5.0 Catheter can be damaged during removal from packaging, preparation, insertion, and removal. Do NOT bend, pull, or place excess pressure on the catheter or mechanical components at any time. | catheter damaged from packaging | packaging issue
Handle with care. The Impella 5.0 Catheter can be damaged during removal from packaging, preparation, insertion, and removal. Do NOT bend, pull, or place excess pressure on the catheter or mechanical components at any time. | catheter damaged from preparation | device handling issue
Handle with care. The Impella 5.0 Catheter can be damaged during removal from packaging, preparation, insertion, and removal. Do NOT bend, pull, or place excess pressure on the catheter or mechanical components at any time. | catheter damaged from removal | device handling issue
Patients with aortic stenosis or other abnormal aortic valve performance may be compromised by the use of the Impella 5.0 Catheter. Patients with aortic valve disease should be observed for aortic insufficiency. | catheter compromised | device operates differently than expected
Patients with aortic stenosis or other abnormal aortic valve performance may be compromised by the use of the Impella 5.0 Catheter. Patients with aortic valve disease should be observed for aortic insufficiency. | aortic stenosis | calcified
Use only original accessories and replacement parts supplied by Abiomed. | incompatible replacement | accessory incompatible
Do NOT use damaged or contaminated connector cables | damaged cable in use | use of device issue
To prevent device failure, do NOT start the Impella 5.0 Catheter until the placement guidewire has been removed. | catheter started inappropriately | use of device issue
Do NOT remove the Impella 5.0 Catheter over the length of the placement guidewire. | catheter removed inappropriately | use of device issue
When replacing the purge cassette, the replacement process should be completed within 90 seconds of luer disconnection. The Impella 5.0 Catheter may be damaged if replacement takes longer than 90 seconds of luer disconnection. | replacement time long | device difficult to setup or prepare
Do NOT kink or clamp any part of the Impella 5.0 Catheter | catheter clamped | kinked
Have a backup Automated Impella Controller, purge cassette, connector cable, and Impella RP Catheter available in the unlikely event of a device failure. | fail to provide backup | failure to back-up
Do NOT use the bed mount as a handle | inappropriate handle used | use of device issue
INTENDED USE IN CANADA: The Impella 5.0 Circulatory Support System is intended to provide hemodynamic support of the left ventricle in situations where a patient is hemodynamically impaired, or where hemodynamic instability is expected, in order to prevent the patient from experiencing hemodynamic collapse and shock. | hemodynamic support compromised | therapeutic or diagnostic output failure
INTENDED USE IN CANADA: The Impella 5.0 Circulatory Support System is intended to provide hemodynamic support of the left ventricle in situations where a patient is hemodynamically impaired, or where hemodynamic instability is expected, in order to prevent the patient from experiencing hemodynamic collapse and shock. | patient hemodynamic impaired | adverse event without identified device or use problem
The Impella 5.0 Circulatory Support System is approved for usage of up to 10 days. | device over usage | unexpected therapeutic results
CONTRAINDICATIONS • Prosthetic aortic valves, severe aortic valvular stenosis or valvular regurgitation | valvular regurgitation | peri-valvular leak
CONTRAINDICATIONS • Prosthetic aortic valves, severe aortic valvular stenosis or valvular regurgitation | aortic valvular stenosis | calcified
CONTRAINDICATIONS • Hematological disorder causing fragility of the blood cells or hemolysis | patient hematological disorder | adverse event without identified device or use problem
CONTRAINDICATIONS • Hematological disorder causing fragility of the blood cells or hemolysis | hemolysis fragility | adverse event without identified device or use problem
CONTRAINDICATIONS • Hypertrophic obstructive cardiomyopathy (HOCM) | patient hypertrophic obstructive cardiomyopathy | adverse event without identified device or use problem
CONTRAINDICATIONS • Aneurysm or necrotomy or severe anomaly of the ascending aorta and/or the aortic arch | patient aorta anomaly | adverse event without identified device or use problem
CONTRAINDICATIONS • Mural thrombus in the left ventricle | patient mural thrombus | adverse event without identified device or use problem
CONTRAINDICATIONS • Ventricular septal defect (VSD) after myocardial infarction | patient septal defect | adverse event without identified device or use problem
CONTRAINDICATIONS • Anatomic conditions precluding insertion of the pump | difficult to insert catheter | patient-device incompatibility
CONTRAINDICATIONS • Other illnesses or therapy requirements precluding use of the pump • Peripheral arterial occlusion disease (PAOD) | patient contraindication | adverse event without identified device or use problem
There are risks of complications with every procedure using a blood pump. These include among others: • Embolism, thrombosis | patient embolism | patient device interaction problem
There are risks of complications with every procedure using a blood pump. These include among others: • Hemolysis, • Bleeding | patient bleeding | patient device interaction problem
There are risks of complications with every procedure using a blood pump. These include among others: • Immune reaction | patient immune reaction | biocompatibility issue
There are risks of complications with every procedure using a blood pump. These include among others: • Vascular injury through to angionecrotomy | patient vascular injury | patient device interaction problem
There are risks of complications with every procedure using a blood pump. These include among others: • Infection and septicemia | patient infection | bacterial contamination of device
There are risks of complications with every procedure using a blood pump. These include among others: • Patient dependency on the pump after use for support | patient dependency on pump | adverse event without identified device or use problem
There are risks of complications with every procedure using a blood pump. These include among others: Cardiovalvular injuries due to extreme movement of the suction cannula in relation to the cardiac valve or as a result of attachment by suction of the pump to the valve system following incorrect positioning | cannula suction movement | unintended movement
There are risks of complications with every procedure using a blood pump. These include among others: Cardiovalvular injuries due to extreme movement of the suction cannula in relation to the cardiac valve or as a result of attachment by suction of the pump to the valve system following incorrect positioning | catheter misplaced | dislodged or dislocated
There are risks of complications with every procedure using a blood pump. These include among others: • Endocardiac injuries as a result of attachment of the pump due to suction | patient endocardiac injury | patient device interaction problem
There are risks of complications with every procedure using a blood pump. These include among others: • Dislocation of the pump | patient complication | adverse event without identified device or use problem
The Impella 5.0 Catheter is an intravascular microaxial blood pump that supports a patient’s circulatory system. The Impella 5.0 Catheter is inserted via femoral or axillary artery cut-down through the artery and into the left ventricle | circulatory system support compromised | unexpected therapeutic results
When properly positioned, the Impella 5.0 Catheter delivers blood from the inlet area, which sits inside the left ventricle, through the cannula, to the outlet opening in the ascending aorta. Physicians and device operators monitor the correct positioning and functioning of the Impella 5.0 Catheter on the display screen of the Automated Impella Controller. | blood delivery compromised | device operates differently than expected
REUSABLE SYSTEM COMPONENTS The Impella 5.0 System consists of the following reusable components: • Automated Impella Controller—provides the user interface, alarm indications, and portable battery • Automated Impella Controller cart—for easy transport of the Automated Impella Controller | resuable controller compromised | device operates differently than expected
REUSABLE SYSTEM COMPONENTS The Impella 5.0 System consists of the following reusable components: • Automated Impella Controller—provides the user interface, alarm indications, and portable battery • Automated Impella Controller cart—for easy transport of the Automated Impella Controller | device damaged by inappropriate cleaning | device cleaning issue
SINGLE-USE SYSTEM COMPONENTS The Impella 5.0 System also includes the following single-use components: • Impella 5.0 Catheter • Purge cassette • 0.018 inch, 260 cm placement guidewire • Connector cable • Impella Axillary Insertion kit | single use violated | use of device issue
The Impella 5.0 Catheter is an intravascular microaxial blood pump that delivers up to 5.0 liters of blood per minute from the left ventricle into the aorta. | fail to deliver sufficient blood | device operates differently than expected
The pressure sensor is a flexible membrane integrated into the cannula. One side of the sensor is exposed to the blood pressure on the outside of the cannula and the other side is exposed to the pressure of the blood inside of the cannula. | sensor detached | detachment of device or device component
The sensor generates an electrical signal proportional to the difference between the pressure outside the cannula and the pressure inside. This signal is displayed on the Automated Impella Controller as the placement signal. | fail to generate signal | electronic property issue
The sensor generates an electrical signal proportional to the difference between the pressure outside the cannula and the pressure inside. This signal is displayed on the Automated Impella Controller as the placement signal. | fail to display signal | no display or display failure
The pressure sensor membrane flexes when the pressure on one side is different from the pressure on the other side, and the electrical properties of the membrane change when it flexes. This allows the sensor to generate an electrical signal proportional to how much the membrane is flexed, and thus proportional to the difference between the pressure on the outside of the cannula and the pressure inside. | membrane torn | torn material
The pressure sensor membrane flexes when the pressure on one side is different from the pressure on the other side, and the electrical properties of the membrane change when it flexes. This allows the sensor to generate an electrical signal proportional to how much the membrane is flexed, and thus proportional to the difference between the pressure on the outside of the cannula and the pressure inside. | membrane flexed unproportionally | material rigid or stiff
When the Impella 5.0 Catheter is placed in the correct position across the aortic valve, the top (outer surface) of the sensor is exposed to the aortic pressure and the bottom (inner surface) of the sensor is exposed to the ventricular pressure. Therefore, the placement signal is approximately equal to the difference between the aortic pressure and the ventricular pressure. | inconsistent measurement | incorrect measurement
When the Impella 5.0 Catheter is correctly positioned across the aortic valve, the changes in pressure associated with the cardiac cycle result in a pulsatile placement signal. | fail to position catheter valve | malposition of device
During diastole, the large pressure difference between the aorta and the left ventricle creates a large electrical signal. | fail to create large electrical signal | electronic property issue
Then at the peak of systole, when the aortic valve opens, the pressure difference between the aorta and the left ventricle—and thus the electrical signal—is zero. Thus, the continual pressure changes associated with the cardiac cycle produce the pulsatile (up and down) waveform seen on the Automated Impella Controller display | fail to produce waveform | electronic property issue
When the Impella 5.0 Catheter is not properly placed across the aortic valve, or when it is fully in the aorta or fully in the ventricle, the pressures outside and inside the cannula are the same throughout the cardiac cycle. As a result, the pressure on either side of the sensor membrane is the same, resulting in a flat placement signal. | flat signal | failure to read input signal
When the Impella 5.0 Catheter is not properly placed across the aortic valve, or when it is fully in the aorta or fully in the ventricle, the pressures outside and inside the cannula are the same throughout the cardiac cycle. As a result, the pressure on either side of the sensor membrane is the same, resulting in a flat placement signal. | catheter fully in ventricle | malposition of device
When the Impella 5.0 Catheter is not properly placed across the aortic valve, or when it is fully in the aorta or fully in the ventricle, the pressures outside and inside the cannula are the same throughout the cardiac cycle. As a result, the pressure on either side of the sensor membrane is the same, resulting in a flat placement signal. | catheter fully in aorta | malposition of device
The purge cassette delivers rinsing fluid to the Impella 5.0 Catheter. The purge fluid (typically 5% dextrose solution) flows from the purge cassette through the catheter to the microaxial blood pump to prevent blood from entering the motor. | purge pressure low | bleed back
The purge cassette delivers rinsing fluid to the Impella 5.0 Catheter. The purge fluid (typically 5% dextrose solution) flows from the purge cassette through the catheter to the microaxial blood pump to prevent blood from entering the motor. | fail to deliver rinsing fluid | no flow
ACCESSORIES: The white connector cable connects the Impella 5.0 Catheter to the Automated Impella Controller. Clips on the cable are used to secure the purge tubing to the cable. The socket at the black end of the cable connects to the red Impella plug. The white plug at the opposite end of the cable is inserted into the blue catheter plug on the front of the Automated Impella Controller. | fail to clip connector | loose or intermittent connection
ACCESSORIES: The silicone plugs may be attached to the shaft of the Impella 5.0 Catheter to aid in inserting the catheter. | fail to aid in inserting catheter | difficult to insert
The 0.018 inch, 260 cm guidewire is used for the placement of the Impella 5.0 Catheter only. | fail to place catheter | difficult to advance
The Impella Axillary Insertion kit facilitates placement of the Impella 5.0 Catheter via the axillary artery. It contains a 23 Fr diameter x 6 cm length peel-away introducer and two (2) graft locks used to attach a graft onto the introducer. (Note: Only one graft lock is required when used with the recommended Hemashield Platinum graft; a back-up is provided.) | fail to insert catheter | difficult to advance
The Impella Axillary Insertion kit facilitates placement of the Impella 5.0 Catheter via the axillary artery. It contains a 23 Fr diameter x 6 cm length peel-away introducer and two (2) graft locks used to attach a graft onto the introducer. (Note: Only one graft lock is required when used with the recommended Hemashield Platinum graft; a back-up is provided.) | part missing | component missing
The kit is packaged with an 8 Fr silicone-coated lubrication dilator. It is recommended to be used in conjunction with a 10 mm diameter x 20 cm length Hemashield Platinum graft. | incompatible kit | component or accessory incompatibility
Hospital Provided: Dextrose solution (typically 5% dextrose in water with 50 IU/mL of heparin) is used as the purge fluid through the Impella 5.0 Catheter. | inappropriate purge solution | chemical issue
The Automated Impella Controller cart holds the Automated Impella Controller. The cart has wheels for easy transport of the controller and a storage basket. (For more information, including assembly instructions, refer to the Automated Impella Controller cart instructions for use.) | cart tipped over | tip-over
Startup: SUPPLIES NEEDED • Automated Impella Controller • Impella 5.0 Catheter and accessories • Diagnostic catheter (AL1 or MP without side holes or pigtail without side holes) • 500 cc bag of dextrose solution for purge solution (5% recommended; 5% to 40% acceptable) with 50 IU heparin/mL • 10 mm x 20 cm Dacron vascular graft (if using Axillary Insertion kit; Hemashield Platinum graft recommended) • 8 mm x 20 cm Dacron vascular graft (optional for femoral insertion of Impella 5.0) | fail to startup | installation-related problem
CASE START 1. Press the START NEW CASE soft button from the startup screen or plug in a new Impella Catheter. “Case Start” can also be selected by pressing the MENU soft key. 2. The controller displays the screen shown below. | fail to startup | difficult or delayed activation
INSERT PURGE CASSETTE 1. Open the purge cassette package. Disconnect and discard the Y connector. Secure the Yellow luer to the sterile field. | fail to secure part | installation-related problem
INSERT PURGE CASSETTE 2. Pass the purge cassette and spike off the sterile field. | difficult to spike off | installation-related problem
INSERT PURGE CASSETTE 3. Press the Next soft button to continue. | fail to press enough | installation-related problem
INSERT PURGE CASSETTE 4. Open the purge cassette door by pressing the release on the left side of the controller. Insert the purge cassette into the Automated Impella Controller. | fail to release door | installation-related problem
INSERT PURGE CASSETTE 5. The purge cassette snaps into a molded compartment on the front of the controller. Follow the diagram on the inside of the purge cassette door for proper placement. | fail to snap | installation-related problem
INSERT PURGE CASSETTE 5. The purge cassette snaps into a molded compartment on the front of the controller. Follow the diagram on the inside of the purge cassette door for proper placement. | incorrect placement | installation-related problem
INSERT PURGE CASSETTE 6. Slide the purge disc into the slot to the right of the purge cassette until it snaps into place. The controller will automatically begin priming the purge cassette. | fail to slide into slot | installation-related problem
INSERT PURGE CASSETTE 6. Slide the purge disc into the slot to the right of the purge cassette until it snaps into place. The controller will automatically begin priming the purge cassette. | prime slow | failure to prime
INSERT PURGE CASSETTE 7. Extend the purge tubing and close the purge cassette door. There is sufficient room around the edges of the purge cassette door so that it will not pinch the purge tubing as it exits. | material puncture | installation-related problem
INSERT PURGE CASSETTE 8. The Automated Impella Controller automatically begins priming the purge cassette after it is inserted. | priming stopped | failure to prime
CONNECT THE IMPELLA CATHETER 1. Remove the Impella 5.0 Catheter from its package using sterile technique and inspect the catheter, including its connector, for damage. | device removed without using sterile technique | disinfection or sterilization issue at user location
CONNECT THE IMPELLA CATHETER 2. Remove the white connector cable from its package using sterile technique. | fail to use sterile technique | installation-related problem
CONNECT THE IMPELLA CATHETER 3. Inspect the cable for damage, including damage to the connector pins at the controller end. | forget to inspect | installation-related problem
CONNECT THE IMPELLA CATHETER 4. Secure the black end of the cable to the sterile field. | loose connection | installation-related problem
CONNECT THE IMPELLA CATHETER 5. Insert the catheter plug into the connector cable socket (grey end). The tab and the slot must be aligned during connection. | connector joining misaligned | installation-related problem
CONNECT THE IMPELLA CATHETER 6. Pull back on the connection to make sure that the plug has snapped into place. | fail to snap into place | positioning issue
CONNECT THE IMPELLA CATHETER 7. Snap the purge clip (located on the pressure reservoir of the clear sidearm) to the connector cable. | fail to snap into place | installation-related problem
CONNECT THE IMPELLA CATHETER Snapping the purge clip on the pressure reservoir to the connector cable is important to prevent the tube from kinking. | tubing kinked | kinked
CONNECT THE IMPELLA CATHETER 8. Pass the sterile connector cable from the Impella 5.0 Catheter off the sterile field. | sterile field compromised | disinfection or sterilization issue at user location
CONNECT THE IMPELLA CATHETER 9. Line up the notch on the connector cable with the notch in the blue catheter plug on the front of the Automated Impella Controller and plug the cable into the controller. | fail to line up notch | failure to align
CONNECT THE IMPELLA CATHETER 10. If you have not already done so, disconnect and discard the Y connector with the red and yellow luers from the purge tubing. | fail to discard part | installation-related problem
CONNECT THE IMPELLA CATHETER 11. Connect the yellow luer on the end of the purge tubing to the yellow luer on the clear sidearm of the Impella 5.0 Catheter. | wrong connection | installation-related problem
CONNECT THE IMPELLA CATHETER 12. When the controller detects that the luer is connected, it automatically begins priming the purge lumen. | fail to initialize priming | failure to prime
ENTER PURGE FLUID DATA 1. Enter the purge fluid information. | incorrect purge information entered | installation-related problem
ENTER PURGE FLUID DATA 2. To select the default values displayed on the screen, press the ACCEPT soft button. This will select those values and automatically advance to the next screen. NOTE: the Automated Impella Catheter will use the default values for the purge fluid unless changed. | fail to select default value | installation-related problem
3. To change the purge fluid information, press the EDIT soft button, scroll to the appropriate item and push the selector knob to select it or use the white arrow soft buttons. Then scroll through the values and push the selector knob to make a new selection. Press the DONE button to finish editing. The controller will use the default values if no other selections are made. • Purge fluid can be set to 50 mL, 100 mL, 250 mL, 500 mL (default), or 1000 mL. • Dextrose concentration can be set to 5% (default), 10%, 20%, 30%, or 40%. • Heparin concentration (anticoagulant in the treatment of thrombosis) can be set to 0 (default), 5, 6.25, 10, 12.5, 15, 20, 25, 40, or 50 IU/mL. | fail to change correct value | installation-related problem
INSERTING THE IMPELLA 5.0 CATHETER NOTE – Proper surgical procedures and techniques are the responsibility of the medical professional. The described procedure is furnished for information purposes only. Each physician must evaluate the appropriateness of the procedure based on his or her medical training and experience, the type of procedure, and the type of systems used. This describes the femoral insertion procedure. For using the Axillary Insertion Kit please consult the Axillary insertion Kit User Manual. | inappropriate surgical procedure | improper or incorrect procedure or method
INSERTING THE IMPELLA 5.0 CATHETER 1. Identify the femoral artery and perform a cut-down of 3 to 5 cm. | inappropriate insert technique | improper or incorrect procedure or method
INSERTING THE IMPELLA 5.0 CATHETER 1. Identify the femoral artery and perform a cut-down of 3 to 5 cm. | difficult to perform a cut | use of device issue
INSERTING THE IMPELLA 5.0 CATHETER 2. Expose the femoral artery. Wrap vessel loops, one distal and one proximal to the subsequent point of incision, one and a half times around the artery. Make the vessel loops as far apart as possible. | vessel loop wrapped inappropriately | use of device issue
3. To prepare the repositioning sheath, remove the luer plug at the end of the sidearm tube and flush the tube with 0.9% NaCl solution. Place the luer plug back in the sidearm tube and secure the plug. | inappropriate reposition procedure | installation-related problem
4. Make the incision as close as possible to the distal loop. Insert a 6 Fr diagnostic catheter with no side holes (Abiomed recommends a 6 Fr AL1 or Multipurpose without side holes or 5 Fr pigtail without side holes) over a diagnostic 0.035 inch or 0.038 inch guidewire into the left ventricle. | incorrect incision position | positioning issue
5. Remove the diagnostic guidewire and exchange it for the supplied 0.018 inch placement guidewire. | fail to exchange | installation-related problem
6. Hold tension on the proximal vessel loop to prevent bleeding. Straighten the blue pigtail and thread it over the 0.018 inch placement guidewire. Wet the cannula with sterile water and backload the catheter onto the placement guidewire. One or two people can load the catheter on the guidewire. | less tension on vessel loop | bleed back
6. Hold tension on the proximal vessel loop to prevent bleeding. Straighten the blue pigtail and thread it over the 0.018 inch placement guidewire. Wet the cannula with sterile water and backload the catheter onto the placement guidewire. One or two people can load the catheter on the guidewire. | fail to wet cannula with sterile water | installation-related problem
One-person technique a. Advance the placement guidewire into the Impella 5.0 Catheter and stabilize the cannula between the fingers. This prevents pinching of the inlet area. The placement guidewire must exit the outlet area on the inner radius of the cannula, and align with the straight black line on the catheter. The catheter can be hyperextended as necessary to ensure the placement guidewire exits on the inner radius of the cannula. | fail to stablize cannula | installation-related problem
One-person technique a. Advance the placement guidewire into the Impella 5.0 Catheter and stabilize the cannula between the fingers. This prevents pinching of the inlet area. The placement guidewire must exit the outlet area on the inner radius of the cannula, and align with the straight black line on the catheter. The catheter can be hyperextended as necessary to ensure the placement guidewire exits on the inner radius of the cannula. | catheter pinched | installation-related problem
One-person technique a. Advance the placement guidewire into the Impella 5.0 Catheter and stabilize the cannula between the fingers. This prevents pinching of the inlet area. The placement guidewire must exit the outlet area on the inner radius of the cannula, and align with the straight black line on the catheter. The catheter can be hyperextended as necessary to ensure the placement guidewire exits on the inner radius of the cannula. | inappropriate insert technique | improper or incorrect procedure or method
Two-person technique b. The scrub assistant can help stabilize the catheter by holding the catheter proximal to the motor. This will allow the implanting physician to visualize the inner radius. The placement guidewire must exit the outlet area on the inner radius of the catheter and align with the straight black line on the catheter. The physician can focus on advancing the placement guidewire and, if the cannula needs to be hyperextended, the scrub assistant is available to assist. | inappropriate insert technique | installation-related problem
7. Make a transverse incision at the guidewire for the 21 Fr catheter. Use U-stitches instead of purse string sutures to avoid stenosis of the vessel after explantation. | incorrect stitch material used | installation-related problem
7. Make a transverse incision at the guidewire for the 21 Fr catheter. Use U-stitches instead of purse string sutures to avoid stenosis of the vessel after explantation. | inappropriate stitch technique | improper or incorrect procedure or method
8. Administer heparin and achieve ACT of at least 250 seconds. | wrong infusion time | medication error
9. Insert the catheter into the vessel and advance along the 0.018 inch placement guidewire until resistance is met at the proximal vessel loop. | fail to advance along with guidewire | failure to advance
10. Loosen the proximal vessel loop and advance the catheter into the vessel. When the motor housing is entirely past the proximal vessel loop, temporarily tighten the loop to control bleeding. | forget to loosen vessel loop | installation-related problem
10. Loosen the proximal vessel loop and advance the catheter into the vessel. When the motor housing is entirely past the proximal vessel loop, temporarily tighten the loop to control bleeding. | forget to tighten vessel loop | use of device issue
11. Advance the repositioning sheath, located on the catheter shaft, through the incision and into the femoral artery until bleeding is controlled. Secure the sheath outside of the vessel using the supplied suture loop. | fail to control bleeding | adverse event without identified device or use problem
11. Advance the repositioning sheath, located on the catheter shaft, through the incision and into the femoral artery until bleeding is controlled. Secure the sheath outside of the vessel using the supplied suture loop. | fail to secure sheath | installation-related problem
12. Stabilize the guidewire and repositioning sheath and advance the catheter through the sheath. Follow the catheter under fluoroscopy as it is advanced into the left ventricle. (Refer to the following page of this manual for information about waveforms displayed on the controller during placement.) | fail to stablize sheath | installation-related problem
13. When the catheter is correctly positioned, slightly loosen the proximal vessel loop and remove the 0.018 inch guidewire. Leave at least 2 to 3 cm of the repositioning sheath inside the vessel. | fail to loosen vessel loop | device handling issue
14. Tighten the prepared U-stitches to seal the sheath. | fail to tighten | device handling issue
15. Loosen the distal vessel loop. Then loosen the proximal vessel loop. | fail to loosen vessel loop in order | device handling issue
Use Fluoroscopy for Placement Impella 5.0 Catheter performance will be compromised if correct placement cannot be confirmed. While other imaging techniques, such as transesophageal echocardiography (TEE), can the Impella 5.0 Catheter after placement, TEE does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the Impella 5.0 Catheter across the aortic valve. | improper imaging technique | improper or incorrect procedure or method
Use Fluoroscopy for Placement Impella 5.0 Catheter performance will be compromised if correct placement cannot be confirmed. While other imaging techniques, such as transesophageal echocardiography (TEE), can the Impella 5.0 Catheter after placement, TEE does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the Impella 5.0 Catheter across the aortic valve. | fail to visualize catheter | poor quality image
Use Fluoroscopy for Placement Impella 5.0 Catheter performance will be compromised if correct placement cannot be confirmed. While other imaging techniques, such as transesophageal echocardiography (TEE), can the Impella 5.0 Catheter after placement, TEE does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the Impella 5.0 Catheter across the aortic valve. | fail to place catheter correctly | use of device issue
Impella 5.0 Use in Open Heart Surgery If the Impella 5.0 Catheter is used in the OR as part of open heart surgery, manipulation may be performed only through the 9 Fr steering catheter. Direct manipulation of the catheter through the aorta or ventricle may result in serious damage to the Impella 5.0 Catheter and serious injury to the patient. | catheter manipulated inappropriately | device handling issue
GP IIb-IIIa Inhibitors. If the patient is receiving a GP IIb-IIIa inhibitor, the Impella 5.0 Catheter can be inserted when ACT is 200 or above. | fail to insert catheter | use of device issue
Do NOT Touch Inlet or Outlet Areas. While feeding the Impella 5.0 Catheter through the femoral artery, hold the device at the cannula or motor housing. Do NOT touch the inlet area or the outlet area. | catheter outlet area touched | device damaged prior to use
Do NOT Touch Inlet or Outlet Areas. While feeding the Impella 5.0 Catheter through the femoral artery, hold the device at the cannula or motor housing. Do NOT touch the inlet area or the outlet area. | device damaged prior to use | device damaged prior to use
Maintaining ACT. After insertion of the catheter (and until explant), ACT should be maintained at 160 to 180 seconds | fail to maintain time | improper device output
Importance of Proper Impella 5.0 Catheter Placement. When the Impella 5.0 Catheter is not correctly placed, there is no effective unloading of the ventricle (hydraulic short circuit). The patient may not be profiting from the flow rate shown on the controller. | ineffective ventricle unloading | unexpected therapeutic results
Importance of Proper Impella 5.0 Catheter Placement. When the Impella 5.0 Catheter is not correctly placed, there is no effective unloading of the ventricle (hydraulic short circuit). The patient may not be profiting from the flow rate shown on the controller. | hydraulic short circuit | therapeutic or diagnostic output failure
Confirming Differential Pressure Sensor Zero. If the lower value of the differential pressure signal is not within 6 mmHg of 0 when the Impella 5.0 Catheter is in place across the aortic valve, zero the differential pressure sensor. | fail to zero sensor | failure to reset
Check Positioning at P-9. When the P-level is increased to P-9, the Impella 5.0 Catheter has a tendency to be drawn into the ventricle. Check positioning at P9 to ensure proper placement throughout the P-level setting range. | catheter fully in ventricle | dislodged or dislocated
Check Positioning at P-9. When the P-level is increased to P-9, the Impella 5.0 Catheter has a tendency to be drawn into the ventricle. Check positioning at P9 to ensure proper placement throughout the P-level setting range. | incorrect pressure reading | pressure issue
POSITIONING AND STARTING THE IMPELLA 5.0 CATHETER. 1. Confirm purge fluid is exiting the Impella Catheter. | purge fluid blocked | obstruction within device
POSITIONING AND STARTING THE IMPELLA 5.0 CATHETER. 2. While the catheter is being advanced in the aorta, the initial placement signal is flat. The inlet area of the catheter has not passed the aortic valve. | flat signal | unable to obtain readings
POSITIONING AND STARTING THE IMPELLA 5.0 CATHETER. 3. Gently advance the catheter forward until a pulsatile waveform is present on the placement signal screen. This signal is generated when the inlet area of the catheter crosses the aortic valve. | fail to generate pulsatile waveform | malposition of device
POSITIONING AND STARTING THE IMPELLA 5.0 CATHETER. 4. Confirm that the controller displays a pulsatile waveform and the inlet area of the Impella 5.0 Catheter is approximately 3.5 cm below the aortic valve. | fail to display pulsatile waveform | malposition of device
5. Press the START IMPELLA soft button | unresponsive button | difficult or delayed activation
6. Turn the selector knob to increase the P-level from P-0 to P-2. | knob stuck | sticking
7. Press the selector knob to select the new P-level. | fail to spin knob | defective component
8. The catheter operation icon in the lower left corner of the screen begins rotating when the Impella 5.0 Catheter begins to operate. | icon rotation failed | application program issue
9. Increase the P-level to P-9 to confirm correct and stable placement. Evaluate the catheter position in the aortic arch and remove any excess slack. The catheter should align against the lesser curvature of the aorta rather than the greater curvature. Verify placement with fluoroscopy and with the placement signal screen. | catheter excess slack | installation-related problem
10. Make sure there is no bleeding at the transition from the repositioning sheath to the femoral artery. Close and dress the wound. Secure the repositioning sheath by suturing it to the skin using the blue suture pads or a StatLock® stabilization device on the sheath. | bleeding at transition site | improper or incorrect procedure or method
11. Attach the anticontamination sleeve to the sheath. Lock it in place by turning clockwise. Secure the catheter by tightening the connected anchoring ring. | fail to attach sleeve | installation-related problem
12. Carefully extend the anticontamination sleeve to maximum length and secure the end closest to the red Impella plug by tightening the anchoring ring. | fail to extend sleeve | installation-related problem
13. Reposition the catheter as necessary. | inefficient reposition | difficult to position
14. Select the lowest P-level that will enable you to achieve the flow rate necessary for patient support. You can select one of ten P-levels (P-0 to P-9) for the Impella 5.0 Catheter. | inappropriate pressure level selected | pressure issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 1. Press PURGE MENU and select “Change Purge Cassette and Bag” from the menu. | difficult to navigate menu | use of device issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 2. Select START to begin the cassette and fluid change process. | fail to change cassette | installation-related problem
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 3. When prompted by the controller, disconnect the luer(s) from the Impella catheter. | incorrect prompt | inappropriate prompts
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 4. Open the purge cassette door by pressing the button on the left side of the console. Remove and discard the old cassette and purge fluid bag. | door opened manually | use of device issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 5. Open the new purge cassette. Spike the new purge fluid bag with the new purge cassette tubing. Select NEXT to continue. | difficult to spike bag | use of device issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 6. Insert the new purge cassette into the controller. Be sure to slide the purge disc into place and extend the purge tubing through the gap in the purge cassette door when you close the door. | door opened inappropriately | use of device issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 7. Confirm the luer(s) are disconnected. Press NEXT to proceed to prime the purge cassette. | fail to disconnect luer | failure to disconnect
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 8. Update the purge fluid information. a. To select the default purge fluid values displayed on the screen, select CONFIRM. b. To change the purge fluid information, select EDIT. Then use the soft keys to navigate selections and edit values. Select DONE to complete editing. | fail to update purge fluid info | installation-related problem
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 8. Update the purge fluid information. a. To select the default purge fluid values displayed on the screen, select CONFIRM. b. To change the purge fluid information, select EDIT. Then use the soft keys to navigate selections and edit values. Select DONE to complete editing. | fluid value selected inappropriately | use of device issue
CHANGE PURGE CASSETTE AND BAG. Follow these steps to change both the purge cassette and purge fluid: 9. Connect the luer(s) from the new purge cassette to the Impella catheter. | fail to connect luer | disconnection
RAPID WEANING 1. Initiate rapid weaning by decreasing catheter P-level in 2-level steps at intervals of several minutes (for example, P-6 to P-4 to P-2). Do NOT decrease P-level to below P-2 until just before removing the catheter from the ventricle. | short pressure decreasing period | installation-related problem
RAPID WEANING 2. When the P-level has been reduced to P-2, maintain the patient on P-2 support for at least 10 minutes before discontinuing circulatory support. | fail to maintain pressure | installation-related problem
RAPID WEANING 2. When the P-level has been reduced to P-2, maintain the patient on P-2 support for at least 10 minutes before discontinuing circulatory support. | short pressure maintain period | installation-related problem
RAPID WEANING 3. If the patient’s hemodynamics remain stable, decrease the P-level to P-1, pull the catheter into the aorta, and stop the motor by decreasing the P-level to P-0. | difficult to pull back catheter | installation-related problem
RAPID WEANING 4. Explant the catheter. | fail to explant catheter | installation-related problem
RAPID WEANING 5. Follow institutional guidelines for arterial closure. | inadequate guideline | Inadequate instructions for healthcare professional
RAPID WEANING 6. Disconnect the connector cable from the Automated Impella Controller and turn the controller off by pressing the power switch on the side of the controller for 3 seconds. | short button pressing period | use of device issue
SLOW WEANING 1. Initiate slow weaning by decreasing catheter P-level in 2-level steps over time as cardiac function allows (for example, P-6 to P-4 to P-2). Do NOT decrease P-level to below P-2 until just before removing the catheter from the ventricle. | catheter pressure decreased rapidly | use of device issue
SLOW WEANING 2. When the P-level has been reduced to P-2, maintain the patient on P-2 support until the patient’s hemodynamics remain stable before discontinuing circulatory support. | short pressure maintain period | installation-related problem
SLOW WEANING 3. If the patient’s hemodynamics remain stable, decrease the P-level to P-1, pull the catheter into the aorta, and stop the motor by decreasing the P-level to P-0. | difficult to pull back catheter | installation-related problem
SLOW WEANING 4. Explant the catheter. | fail to explant catheter | installation-related problem
SLOW WEANING 5. Follow institutional guidelines for arterial closure. | confused guideline | inadequate instructions for healthcare professional
SLOW WEANING 6. Disconnect the connector cable from the Automated Impella Controller and turn the controller off by pressing the power switch on the side of the controller for 3 seconds. | short button pressing period | use of device issue
Remove the Impella 5.0 Catheter With Care. Removal of the Impella 5.0 Catheter must be completed with care to avoid damage to the catheter assembly. | catheter damaged from removal | use of device issue
The electrical signal produced by the differential pressure sensor may drift over time. This drifting is normal behavior and is characterized by an upward or downward shift in the placement signal on the controller display. The amplitude (peak-to-trough, or maximum to minimum) of the waveform will not change. In addition, when sensor drifting occurs, the displayed flow rate may not match the P-level. | signal drifted | electronic property issue
The electrical signal produced by the differential pressure sensor may drift over time. This drifting is normal behavior and is characterized by an upward or downward shift in the placement signal on the controller display. The amplitude (peak-to-trough, or maximum to minimum) of the waveform will not change. In addition, when sensor drifting occurs, the displayed flow rate may not match the P-level. | flow rate displayed inappropriately | electronic property issue
Flow rate can vary due to suction or incorrect positioning. If you observe that the placement waveform has shifted up or down on the display, or the expected flow does not match the current P-level setting, zero the differential pressure sensor by performing the following steps: 1. Press the MENU key and select “Start Manual Zero.” | waveform shifted | electronic property issue
Flow rate can vary due to suction or incorrect positioning. If you observe that the placement waveform has shifted up or down on the display, or the expected flow does not match the current P-level setting, zero the differential pressure sensor by performing the following steps: 1. Press the MENU key and select “Start Manual Zero.” | fail to zero sensor | failure to reset
Flow rate can vary due to suction or incorrect positioning. If you observe that the placement waveform has shifted up or down on the display, or the expected flow does not match the current P-level setting, zero the differential pressure sensor by performing the following steps: 2. Press OK to adjust flow level. | fail to adjust flow level | improper flow or infusion
3. The controller displays “Wait until new flow level is reached.” and then “Offset calculation is in progress.” | fail to calculate flow rate offset | application program issue
4. Select OK to accept the new setting when the controller displays the “Placement Signal Offset Adjust finished” message. | fail to provide signal offset | data issue
5. The Impella will automatically be reset to the previous P-level. | fail to reset pressure level | reset issue
Accuracy of Displayed Flow Rate. Under normal operating conditions, displayed flow rate can deviate from the actual flow rate by up to 0.5 L/min. | inaccurate flow rate displayed | inaccurate flowrate
The controller software contains a data table listing the expected differential pressure for a given motor current when the motor speed is set to a specific value. | inadequate differential pressure table | data issue
To zero the differential pressure sensor while the Impella 5.0 Catheter is running, the software sets the motor speed and measures the motor current. Using the data table, the software determines what the measured differential pressure should be, then adjusts the signal from the differential pressure sensor so that it matches the expected value. | fail to zero differential pressure | failure to reset
PLACEMENT SIGNAL NOT RELIABLE AND EFFECTS ON FLOW CALCULATIONS If the pressure sensor fails, the controller can no longer calculate the flow rate. The controller displays a triangle with “Flow Calculation Disabled.” The Home screen displays a yellow question mark over the heart icon and “Placement Monitoring Suspended.” | unreliable signal | electronic property issue
If the pressure sensor fails, the controller can no longer calculate the flow rate. The controller displays a triangle with “Flow Calculation Disabled.” The Home screen displays a yellow question mark over the heart icon and “Placement Monitoring Suspended.” | fail to calculate flow rate | failure to sense
PLACEMENT SIGNAL NOT RELIABLE AND EFFECTS ON POSITION CONTROL If the pressure sensor fails, placement monitoring is switched off because it is not possible to display the position of the catheter. In this case, the motor current signal or imaging procedures can be used for position control. As long as the motor current signal is pulsatile, the Impella 5.0 Catheter is correctly positioned across the valve. This signal must be monitored closely because the catheter can become dislodged (displaced) when moving the patient or changing the patient’s position. Therefore, if patient hemodynamics change—for example, if arterial pressure falls or there are signs of left ventricular failure—check the correct positioning of the catheter using imaging procedures (eg, TEE) and the motor current signal. | sensor dislodged | dislodged or dislocated
PLACEMENT SIGNAL NOT RELIABLE AND EFFECTS ON POSITION CONTROL If the pressure sensor fails, placement monitoring is switched off because it is not possible to display the position of the catheter. In this case, the motor current signal or imaging procedures can be used for position control. As long as the motor current signal is pulsatile, the Impella 5.0 Catheter is correctly positioned across the valve. This signal must be monitored closely because the catheter can become dislodged (displaced) when moving the patient or changing the patient’s position. Therefore, if patient hemodynamics change—for example, if arterial pressure falls or there are signs of left ventricular failure—check the correct positioning of the catheter using imaging procedures (eg, TEE) and the motor current signal. | sensor failed | failure to sense
SUCTION DETECTION DURING SENSOR DRIFT OR PLACEMENT SIGNAL NOT RELIABLE If sensor drift occurs or the pressure sensor fails, the controller can no longer detect suction. The effectiveness of Impella 5.0 Catheter support can only be assessed by monitoring patient hemodynamics, cardiac imaging, and the Impella 5.0 Catheter motor current. Signs of suction include: • A drop in the patient’s arterial pressure • Decreased output, if a cardiac monitor is in place • Dampened or flat motor current waveforms | fail to detect suction | suction issue
SUCTION DETECTION DURING SENSOR DRIFT OR PLACEMENT SIGNAL NOT RELIABLE If sensor drift occurs or the pressure sensor fails, the controller can no longer detect suction. The effectiveness of Impella 5.0 Catheter support can only be assessed by monitoring patient hemodynamics, cardiac imaging, and the Impella 5.0 Catheter motor current. Signs of suction include: • A drop in the patient’s arterial pressure • Decreased output, if a cardiac monitor is in place • Dampened or flat motor current waveforms | suction dropped | decrease in suction
SUCTION DETECTION DURING SENSOR DRIFT OR PLACEMENT SIGNAL NOT RELIABLE If imaging reveals that the suction is caused by the catheter inlet area in close proximity to the intraventricular wall, reposition the catheter. If hemodynamic parameters, such as low aortic pressure or high pulmonary artery pressure, indicate suction caused by inadequate filling volume, reduce performance level and follow the management strategies. | over suction | increase in suction
The Impella 5.0 Catheter is designed to be operated with a purge solution that contains heparin. Operation of the system without heparin in the purge solution has not been tested. In the event that a patient is intolerant to heparin, due to heparin-induced thrombocytopenia (HIT) or bleeding, physicians should use their clinical judgment to assess the risks versus benefits of operating the Impella System without heparin. | heparin side effect | unexpected therapeutic results
If it is in the best interest of the patient to operate the system without heparin, the dextrose solution is still required, and physicians should consider systemic delivery of an alternative anticoagulant. DO NOT add any alternative anticoagulant (such as a direct thrombin inhibitor) to the purge fluid. The Impella 5.0 Catheter has not been tested with any alternative anticoagulants in the purge solution. | alternative anticoagulant applied | medication error
OPERATING THE IMPELLA 5.0 CATHETER IN ELECTROMAGNETIC FIELDS The Impella 5.0 Catheter contains a permanent magnet motor that emits an electromagnetic field. This field may produce electromagnetic interference with other equipment. In addition, other equipment that emits a strong electromagnetic field may affect the operation of the Impella 5.0 Catheter motor. | electromagnetic field emitted | electromagnetic interference (EMI)
OPERATING THE IMPELLA 5.0 CATHETER IN ELECTROMAGNETIC FIELDS The Impella 5.0 Catheter contains a permanent magnet motor that emits an electromagnetic field. This field may produce electromagnetic interference with other equipment. In addition, other equipment that emits a strong electromagnetic field may affect the operation of the Impella 5.0 Catheter motor. | interference presented | electromagnetic interference (EMI)
ELECTROANATOMIC MAPPING (EAM) SYSTEMS The electromagnetic field emitted by the Impella 5.0 Catheter may produce interference with the magnetic location detection component of the electroanatomic mapping (EAM) system, particularly when the mapping catheter is close to the Impella 5.0 Catheter motor. For example, mapping in the right or left ventricular outflow tracts places the mapping catheter in close proximity to the Impella 5.0 Catheter motor in the ascending aorta. | interference introduced | electromagnetic interference (EMI)
ELECTROANATOMIC MAPPING (EAM) SYSTEMS. Electromagnetic interference may appear as: • Magnetic interference errors generated by the electroanatomic mapping system | magnetic interference error | electromagnetic compatibility issue
ELECTROANATOMIC MAPPING (EAM) SYSTEMS. Electromagnetic interference may appear as: • Instability in the displayed location of the mapping catheter | display instability | electromagnetic interference (EMI)
ELECTROANATOMIC MAPPING (EAM) SYSTEMS When operating the Impella Catheter in the presence of an EAM system, use P-level mode. Operate the Impella Catheter at P-1–P-5 or P-7. The motor speeds at these P-levels cause the least interference. Best performance is observed when the Impella Catheter motor is at least 3 cm from the sensors in the mapping catheter. If you suspect interference, follow the troubleshooting steps in the following table. | sensor interfered by motor | electromagnetic interference (EMI)
MAGNETIC NAVIGATION SYSTEMS (MNS) When initiating Impella 5.0 Catheter support in the presence of a magnetic navigation system (MNS), follow the steps below: 1. Insert the Impella 5.0 Catheter following the steps outlined in section INSERTING THE IMPELLA 5.0 CATHETER above. 2. Place the MNS magnets in the “Reduced” or “Stowed” position. 3. Start the Impella 5.0 Catheter in the manner described above. Increase performance level to at least P-5. 4. Place the MNS magnets in the “Navigate” position and proceed with magnetic navigation. Keep operating the Impella 5.0 Catheter at a performance level of at least P-3 when the MNS magnets are in the “Navigate” position. If the performance level falls below P-3, the Impella 5.0 Catheter may stop running. To resume operation, follow the steps in the table below. | catheter interfered by magnetic navigation system | electromagnetic interference (EMI)
MAGNETIC NAVIGATION SYSTEMS (MNS) During magnetic navigation of the mapping catheter, the motor current of the Impella 5.0 Catheter may temporarily increase to the point that the catheter stops running. The table below explains how to resume operation. When the MNS magnets are in the “Navigate” position, the displayed Impella 5.0 Catheter flow may be artificially elevated. To accurately assess the flow rate, note the displayed flow when the magnets are in the “Stowed” position. | catheter malfunctioned | electromagnetic interference (EMI)
MAGNETIC NAVIGATION SYSTEMS (MNS) During magnetic navigation of the mapping catheter, the motor current of the Impella 5.0 Catheter may temporarily increase to the point that the catheter stops running. The table below explains how to resume operation. When the MNS magnets are in the “Navigate” position, the displayed Impella 5.0 Catheter flow may be artificially elevated. To accurately assess the flow rate, note the displayed flow when the magnets are in the “Stowed” position. | motor current increased | electromagnetic interference (EMI)
MAGNETIC NAVIGATION SYSTEMS (MNS) During magnetic navigation of the mapping catheter, the motor current of the Impella 5.0 Catheter may temporarily increase to the point that the catheter stops running. The table below explains how to resume operation. When the MNS magnets are in the “Navigate” position, the displayed Impella 5.0 Catheter flow may be artificially elevated. To accurately assess the flow rate, note the displayed flow when the magnets are in the “Stowed” position. | catheter flow display elevated | electromagnetic interference (EMI)
TRANSFERRING FROM THE AUTOMATED IMPELLA CONTROLLER TO A NEW AUTOMATED IMPELLA CONTROLLER A backup Automated Impella Controller should be available at all times when a patient is on support. In the event that the controller fails, follow the steps below to transition the Impella 5.0 Catheter to the backup controller. 1. Confirm that the backup controller is powered on and ready. 2. Press PURGE MENU on the original controller, select Change Purge Fluid Bag, and complete the step to bolus the purge system. (Do NOT flush the purge fluid from the cassette.) 3. Disconnect the yellow luer connector from the Impella 5.0 Catheter to release the pressure in the purge cassette. (“Purge System Open” alarm will sound.) 4. Transfer the purge cassette and purge solution from the original controller to the backup controller. 5. Reconnect the yellow luer connector to the Impella 5.0 Catheter. 6. Remove the white connector cable from the original controller and plug it into the catheter plug on the front of the backup controller. 7. Once the Impella 5.0 Catheter is connected to the backup controller, wait for a message to appear on the screen asking you to confirm re-starting the Impella 5.0 Catheter at the previously set performance level. 8. Press OK within 10 seconds to confirm restarting the Impella 5.0 Catheter at the previously set performance level. 9. If the message to restart the Impella 5.0 Catheter does not appear within 30 seconds, restart the Impella 5.0 Catheter using the FLOW CONTROL soft button. | difficult to transit backup controller | failure to back-up
PATIENT MANAGEMENT CHECKLIST FOLLOWING TRANSFER OF SUPPORT After transferring patient support to or from the Automated Impella Controller, perform each of the following patient management checklist items: 1. Confirm Impella 5.0 Catheter placement using echocardiography. 2. Tighten the Tuohy-Borst valve (tighten all the way to the right) on the Impella 5.0 Catheter to prevent catheter migration. | catheter migrated | migration of device or device component
The Automated Impella Controller monitors various functions to determine whether specific operational parameters are within expected limits. When a parameter goes outside of its specified limits, the controller sounds an alarm tone and displays an alarm message that can be viewed on the display screen on the front of the controller. The alarm tone indicates the severity of the alarm. The alarm message on the display screen is color-coded for severity and provides details on the cause of the alarm and how to resolve the alarm. | inappropriate alarm tone | defective alarm
ALARM LEVELS Alarms are divided into three levels of severity: • Advisory (white), Notification, 1 beep every 5 minutes, Alarm header on white background • Serious (yellow), May become harmful or life-threatening if not addressed immediately, 3 beeps every 15 seconds, Alarm header on yellow background • Critical (red), Immediately harmful or life-threatening, 10 beeps every 6.7 seconds, Alarm header on red background | inappropriate alarm level | improper alarm
ALARM LEVELS Alarms are divided into three levels of severity: • Advisory (white), Notification, 1 beep every 5 minutes, Alarm header on white background | improper advisory alarm defined | device alarm system issue
ALARM LEVELS Alarms are divided into three levels of severity: • Critical (red), Immediately harmful or life-threatening, 10 beeps every 6.7 seconds, Alarm header on red background | improper critical alarm defined | improper alarm
ALARM LEVELS Alarms are divided into three levels of severity: • Serious (yellow), May become harmful or life-threatening if not addressed immediately, 3 beeps every 15 seconds, Alarm header on yellow background | improper serious alarm defined | device alarm system issue
Alarms That Resolve On Their Own The audible indicator will shut off if an alarm condition is resolved before you press MUTE ALARM. The visual message, however, will continue to be displayed, with the alarm header on a gray background, for 20 minutes or until you press MUTE ALARM. This allows you to identify the alarm that occurred. | fail to shut off alarm | device alarm system issue
MUTE ALARM FUNCTION Pressing the MUTE ALARM button on the upper right of the Automated Impella Controller display screen will silence the audible alarm indicator for 2 minutes (for red or yellow alarms) or 5 minutes (for white advisory alarms). When an alarm is silenced, the words “MUTE ALARM” next to the button are replaced by the mute alarm indicator, a crossed-out bell icon. | fail to mute alarm | device remains activated
alarm: Air in Purge System, severity: critical, action: The purge system has stopped. Follow the instructions in the De-air Tool to remove the air from the system, cause: There is air in the purge tubing | air in purge system | infusion or flow issue
alarm: Air in Purge System, severity: critical, action: The purge system has stopped. Follow the instructions in the De-air Tool to remove the air from the system, cause: There is air in the purge tubing | air in purge system alarm | device alarm system issue
alarm: Air in Purge System, severity: critical, action: The purge system has stopped. Follow the instructions in the De-air Tool to remove the air from the system, cause: There is air in the purge tubing | fail to expel air | blockage within device or device component
alarm: Battery Critically Low, severity: critical, action: Plug controller into AC power, cause: Battery power has 15% remaining capacity | low battery alarm | low battery
alarm: Battery Failure. severity: critical. action: 1. Plug controller into AC power, 2. Press switch located on the underside of the controller. 3. Switch to backup controller. cause: One of the batteries has failed. A battery switch is turned off or there is a malfunction of the switch. | battery failed | battery issue
alarm: Battery Temperature High severity: critical action: Switch to backup controller cause: Battery temperature is greater than 60°C. | battery temperature high | overheating of device or device component
alarm: Complete Procedure severity: critical action: 1. Follow the steps on the screen or 2. Exit the procedure cause: Complete Procedure serious alarm (yellow; see next page) is active and the user has not responded for an additional 2 minutes. | procedure completion alarm | use of device issue
alarm: Controller Failure severity: critical action: Switch to backup controller cause: There is a problem with the controller electronics | controller failure alarm | device inoperable
alarm: Emergency Shutdown Imminent severity: critical action: Release ON/OFF push button cause: Power switch pressed for 15 seconds while Impella Catheter still connected. | emergency shutdown alarm | electrical power problem
alarm: catheter Disconnected severity: critical action: 1. Check cable connection to console. 2. Check Impella connection to cable. cause: Running Impella Catheter disconnected | catheter disconnection alarm | connection issue
alarm: catheter Failure severity: critical action: Replace catheter cause: There is a problem with the Catheter motor | catheter failure alarm | device inoperable
alarm: catheter Position Wrong severity: critical, action: 1. Reduce the P-Level to P-2 until imaging is available. 2. Reposition with imaging. 3. Return to preferred P-Level and confirm positioning with imaging. cause: Controller has detected that Catheter is in the wrong position. | catheter position harm | malposition of device
alarm: Impella Stopped severity: critical action: 1. Restart Impella. 2. Replace Impella after 3rd unsuccessful restart attempt. 1. Replace white connector cable. 2. Switch to back-up Controller.3. Replace Impella Catheter. cause: There may be a mechanical or electrical problem in the Catheter. There is a problem with the Catheter motor | device stop alarm | device inoperable
alarm: Motor Current High severity: critical action: 1. Restart Impella. 2. Replace Impella after 3rd unsuccessful restart attempt. cause: There is a problem with the Impella Catheter motor | motor current high alarm | electrical issue
alarm: Retrograde Flow severity: critical action: To prevent retrograde flow, restart Impella or withdraw pump from ventricle. cause: Impella Catheter is not running; possible reverse flow through Impella Catheter | back flow alarm | bleed back
alarm: Purge Disc Not Detected severity: critical action: Re-insert purge disc cause: The controller is not detecting that the purge disc is clicked into the front of the controller | purge disc detection alarm | failure to sense
alarm: Purge Pressure High severity: critical action: 1. Check purge system tubing for kinks. 2. Decrease concentration of dextrose in the purge solution. cause: Purge pressure is ≥1100 mmHg with the purge flow <2 mL/hr. | purge pressure high alarm | medication error
alarm: Purge Pressure High severity: critical action: 1. Check purge system tubing for kinks. 2. Decrease concentration of dextrose in the purge solution. cause: Purge pressure is ≥1100 mmHg with the purge flow <2 mL/hr. | purge pressure high | increase in pressure
alarm: Purge Pressure Low severity: critical action: 1. Check purge system tubing for leaks. 2. Increase concentration of dextrose in the purge solution. 3. Replace purge cassette. cause: Purge pressure has dropped below 300 mmHg with the purge flow ≥30 mL/hr for 30 seconds or longer. | purge pressure low alarm | fluid leak
alarm: Purge Pressure Low severity: critical action: 1. Check purge system tubing for leaks. 2. Increase concentration of dextrose in the purge solution. 3. Replace purge cassette. cause: Purge pressure has dropped below 300 mmHg with the purge flow ≥30 mL/hr for 30 seconds or longer. | purge pressure low | medication error
alarm: Purge System Blocked severity: critical action: 1. Check the purge system tubing for kinks or blockages. 2. Decrease concentration of dextrose in the purge solution. cause: Purge flow has dropped below 1 mL/hr. Kinked or blocked purge connecting tube. Kinked or blocked purge lumen in Impella Catheter | purge system blockage alarm | no flow
alarm: Purge System Blocked severity: critical action: 1. Check the purge system tubing for kinks or blockages. 2. Decrease concentration of dextrose in the purge solution. cause: Purge flow has dropped below 1 mL/hr. Kinked or blocked purge connecting tube. Kinked or blocked purge lumen in Impella Catheter | purge system blocked | kinked
alarm: Purge System Failure severity: critical action: 1. Replace purge cassette. 2. Switch to back-up Controller. cause: There is a problem with the purger unit driver | purge system failure alarm | device inoperable
alarm: Purge System Open severity: critical action: 1. Check the purge system tubing for open connections or leaks.2. Replace purge cassette. cause: Purge pressure has dropped below 100 mmHg for 20 seconds or longer | purge system open alarm | decrease in pressure
alarm: Battery Communication Failure severity: serious action: Plug controller into AC power cause: Loss of communication to the battery | battery communication failure alarm | communication or transmission issue
alarm: Battery Level Low severity: serious action: Plug controller into AC power cause: Battery power has 50% remaining capacity. | low battery alarm | low battery
alarm: Battery Temperature High severity: serious action: 1. Check controller for blocked air vents. 2. Switch to backup controller cause: Battery temperature is greater than 50°C and less than or equal to 60°C. | battery temperature high alarm | battery issue
alarm: Controller Error severity: serious action: Switch to backup controller cause: There is a problem with the controller electronics
alarm: Catheter Not Supported severity: serious action: 1. Replace Impella with supported catheter (2.5, CP, 5.0, LD, RP). 2. Contact Abiomed Service to upgrade Impella Controller. cause: The Impella Catheter is not supported to operate with the current version of controller software and/or hardware.
alarm: Catheter Defective severity: serious action: Do not use Catheter. Replace Catheter. cause: There is a problem with the Catheter electronics
alarm: Catheter Position Wrong severity: serious action: 1. Confirm Catheter position with imaging. 2. Pull Catheter back 2 cm. cause: Controller has detected that the Catheter is in the wrong position, with the outlet area too close to the aortic valve. | catheter position alarm | malposition of device
alarm: Catheter Stopped severity: serious action: Attempting to restart catheter. Catheter expected to restart within 30 secs. cause: Locate back-up Controller | device stop alarm | device inoperable
alarm: Placement Signal Not Reliable severity: serious action: Placement and Suction Monitoring are Suspended. 1. Monitor patient hemodynamics. 2. Monitor Impella position with imaging cause: There is a problem with the Impella Catheter sensor signal. | sensor signal alarm | failure to read input signal
alarm: Purge Cassette Failure severity: serious action: Replace purge cassette. cause: There is a problem with the purge cassette hardware | purge cassette failure alarm | defective component
alarm: Purge Volume Critically Low severity: serious action: 1. Press the PURGE MENU soft key and select Change Purge Fluid Bag. 2. Follow the instructions to change the purge fluid. cause: There are 15 mL (in addition to 5% of the starting bag volume) or fewer remaining in the purge fluid bag. | purge volume low alarm | medication error
alarm: Reinstall Software failed severity: serious action: Software installation was unsuccessful. Reinstall software cause: Software was not installed successfully | software installation alarm | problem with software installation
alarm: Suction pesented severity: serious action: 1. Check filling and volume status. 2. Check Impella position. 3. Reduce P-Level. cause: Suction is detected. | suction detection alarm | suction issue
alarm: AC Power Disconnected severity: serious action: Controller is running on battery power cause: AC power was disconnected | AC power disconnection alarm | disconnection
alarm: AC Power Disconnected severity: serious action: Controller is running on battery power cause: AC power was disconnected | power disconnection alarm | disconnection
alarm: Audio Off severity: serious action: The audio for the following alarms has been disabled. cause: User has disabled audio for Placement Signal Not Reliable, Purge Pressure High, Suction, Purge System Blocked. | audio off alarm | use of device issue
alarm: Impella Position Unknown severity: serious action: Impella 5.0 Catheter position unknown due to low pulsatility. Assess cardiac function. Confirm Impella Position with imaging. cause: low pulsatility. Catheter position unknown detected by algorithm. | fail to detect position | positioning issue
alarm: Purge Cassette Incompatible severity: serious action: Contact Service to update Impella controller. cause: Incompatible purge cassette RFID version | incompatible purge cassette alarm | component or accessory incompatibility
alarm: Purge Flow Decreased severity: serious action: The purge flow has decreased by 2.5 mL/hr or more. This is a notification only; no action is required. cause: Purge flow has decreased by ≥2.5 mL/hr. | purge flow decrease alarm | insufficient flow or underinfusion
alarm: Purge Flow Increased severity: serious action: The purge flow has increased by 2.5 mL/hr or more. This is a notification only; no action is required. cause: Purge flow has increased by ≥2.5 mL/hr. | purge flow increase alarm | excess flow or over-infusion
alarm: Purge Volume Low severity: serious action: 1. Press the PURGE MENU soft key and select Change Purge Fluid Bag. 2. Follow the instructions to change the purge fluid. cause: There are 30 mL (in addition to 5% of the starting bag volume) or fewer remaining in the purge fluid bag. | purge volume low alarm | infusion or flow issue
alarm: Surgical Mode Enabled severity: serious action: Impella pump stopped. Purge system running. ‘Impella Stopped’ alarm disabled. cause: Surgical Mode has been enabled to silence ‘Impella Stopped’ alarm at P-0. | device stop alarm | device operates differently than expected
alarm: Unexpected Controller Shutdown severity: serious action: Switch to back-up Controller if condition persists. cause: Unexpected restart of controller due to software of hardware failures. | unexpected shutdown alarm | inappropriate or unexpected reset
Latex Free The Automated Controller and Catheter, including all accessories, are latex free. | latex free violated | environmental control or utility issue
The Catheter and connector cable are disposable items that must be disposed of in accordance with hospital regulations for blood contaminated materials. | inappropriate disposal of blood contaminated materials | device contamination with blood or blood product

## connect_IFU
Statements | Failure_Mode | Hazard
The Impella Connect System is comprised of the following components: 1) The Impella Connect Cloud Infrastructure/Portal; a web-based user portal that allows authorized users to remotely view the screen of the Automated Impella Controller (AIC) 2) Remote Link Module (RLM); a small hardware box (0042-4100) attached to the back and side of an Automated Impella Controller (AIC) | fail to connect cloud | wireless communication issue
The Impella Connect System is comprised of the following components: 1) The Impella Connect Cloud Infrastructure/Portal; a web-based user portal that allows authorized users to remotely view the screen of the Automated Impella Controller (AIC) | fail to connect portal | communication or transmission issue
The Impella Connect System is comprised of the following components: 2) Remote Link Module (RLM); a small hardware box (0042-4100) attached to the back and side of an Automated Impella Controller (AIC) | module malfunctioned | device inoperable
The RLM connects to the VGA output of the AIC and transmits the display of the AIC screen to a cloud-based server. | fail to transmit to server | wireless communication issue
The RLM connects to the VGA output of the AIC and transmits the display of the AIC screen to a cloud-based server. | intermittent connection | wireless communication issue
The transmitted image can be viewed by authorized remote users, which may include hospital clinicians, Abiomed local support staff, and Abiomed Customer Support Center (CSC) team members. | image viewed unauthorizedly | unauthorized access to computer system
The Impella Connect System only enables passive viewing of the AIC video screen. Changes to the controller settings can only be made using the physical controls on the AIC. | passive viewing violated | computer system security issue
The Impella Connect System only enables passive viewing of the AIC video screen. Changes to the controller settings can only be made using the physical controls on the AIC. | settings changed unauthorizedly | application security issue
Impella Connect may be configured and connected to the hospital’s secured Wi-Fi network or be plugged in directly via the Ethernet cable to the secured hospital network to transmit the video image to a web-based portal. | insecure WI-FI connected | computer system security issue
Impella Connect may be configured and connected to the hospital’s secured Wi-Fi network or be plugged in directly via the Ethernet cable to the secured hospital network to transmit the video image to a web-based portal. | insecure image transmitted | application security issue
Impella Connect transmits data via the hospital’s secured Wi-Fi network from the AIC to a HIPAA-compliant database (Abiomed Cloud hosted by Microsoft Azure) and enables hospital and Abiomed staff to view case information on demand. | insecure data transmitted | application security issue
Impella Connect transmits data via the hospital’s secured Wi-Fi network from the AIC to a HIPAA-compliant database (Abiomed Cloud hosted by Microsoft Azure) and enables hospital and Abiomed staff to view case information on demand. | fail to compliant with HIPAA database | data issue
Login to the Impella Connect customer portal is limited to only approved and registered users and is protected by login/password security for each account. | insecure login | unauthorized access to computer system
The Impella Connect user portal provides multiple views of ongoing cases. The Case Summary Dashboard provides a summary dashboard of all cases, distinctive colors to indicate status, and the Following option to identify a case. | dashboard layout confused | inadequate user interface
The Case Video Page provides a remote view of the AIC screen, a case history log that includes video and alert summaries, and the ability to replay any moment in the case. The high-level system architecture diagram for the Impella Connect System | fail to replay case | data issue
The Case Video Page provides a remote view of the AIC screen, a case history log that includes video and alert summaries, and the ability to replay any moment in the case. The high-level system architecture diagram for the Impella Connect System | fail to view video | image display error
As depicted, the Impella Connect Module (or RLM), transmits the display of the AIC, defined as telemetry data, to a cloud-based server which can be viewed remotely via the Impella Connect web application. Telemetry data is defined as data transmitted to the Cloud Server and Remote Viewer, which includes video of the AIC screen and pump log data. This functionality is intended to enable cloud-based, remote monitoring of the Impella device status. | inadequate telemetry data | loss of data
As depicted, the Impella Connect Module (or RLM), transmits the display of the AIC, defined as telemetry data, to a cloud-based server which can be viewed remotely via the Impella Connect web application. Telemetry data is defined as data transmitted to the Cloud Server and Remote Viewer, which includes video of the AIC screen and pump log data. This functionality is intended to enable cloud-based, remote monitoring of the Impella device status. | inadequate video data | data issue
Intended Use: Impella Connect is intended to be used to enable remote patient monitoring to provide users with passive viewing of the AIC’s display which includes information on alarms and hemodynamic data useful for troubleshooting and managing Impella devices to aid in patient management. | fail to provide alarm info | loss of data
Intended Use: Impella Connect is intended to be used to enable remote patient monitoring to provide users with passive viewing of the AIC’s display which includes information on alarms and hemodynamic data useful for troubleshooting and managing Impella devices to aid in patient management. | fail to provide hemodynamic data | data issue
Intended Use: Impella Connect is intended to be used to enable remote patient monitoring to provide users with passive viewing of the AIC’s display which includes information on alarms and hemodynamic data useful for troubleshooting and managing Impella devices to aid in patient management. Impella Connect transfers the video stream from the AIC (via the VGA output) to a cloud-based remote viewing portal. Communication between the AIC and Impella Connect is one-way (AIC to Impella Connect), and the streamed video is limited to Impella device operating parameters and alarm messages with no patient identifiable information. Impella Connect is powered by the AIC. | incorrect product for intended use | user used incorrect product for intended use
This section outlines the relevant regulatory pathway and associated requirements for the Impella Connect System in the United States. The regulatory pathway and requirements of Impella Connect are dependent on its intended use, design, and current scope of features incorporated within the system. This section will be updated as necessary when new system updates or features are introduced to the Impella Connect System. | fail to update regulatory pathway | no apparent adverse event
As supported by the Impella Connect intended use and current feature set, the system serves solely as a remote viewer without any device analysis or interpretation of the medical device data within the viewer. The regulatory classification of the individual components has been evaluated in accordance with relevant FDA regulations and guidance documents | incorrect regulatory classification | no apparent adverse event
Compoent: Cloud Infrastructure/Portal Classification: Non-Device CDS (clinical decision support that are excluded from the definition of device) and Non-Device-MDDS (Medical Device Data Systems, not subject to FDA regulatory requirements ) Software | incorrect regulatory classification | no apparent adverse event
Compoent: Remote Link Module (RLM) Classification: Hardware is Device-MDDS, Class I per 21 CFR 880.6310 & 86 FR 20278 under enforcement discretion; Software is exempt from the definition of device per § 3060 of the 21st Century Cures Act | incorrect regulatory classification | no apparent adverse event
To ensure diligence on regulatory classification and requirements of the Impella Connect Cloud Infrastructure/Portal component, the individual software features have been evaluated. The regulatory classification of the individual software features per webpage of the Cloud Infrastructure/Portal are described in Table 2 (Case Summary & Settings Pages) and Table 3 (Case Video Page). | fail to evaluate software | no apparent adverse event
Case Tile (Non-Device CDS Software Function) All active cases are displayed as a “tile”, which provides a concise visual representation of cases being followed. Each tile displays pertinent information for case identification including hospital, case number, Impella Pump type, AIC SN, as well as amount of time receiving Impella support. This feature solely displays case-specific information to enable case identification and provides a summary of active cases for user convenience in pump management. | confused case identification | human-device interface issue
Case Tile Color (Non-Device CDS Software Function) Each individual case tile includes a color (green, yellow, or red) which is indicative of the current state of the AIC/case. This feature solely displays a graphical display of the AIC’s current state, is not considered an alarm system, and is not intended to support clinical decision making. | confused case identification color | human-device interface issue
Case Tile Color (Non-Device CDS Software Function) Each individual case tile includes a color (green, yellow, or red) which is indicative of the current state of the AIC/case. This feature solely displays a graphical display of the AIC’s current state, is not considered an alarm system, and is not intended to support clinical decision making. | confused case status | human-device interface issue
Case Tile Pump Metrics (Non-Device CDS Software Function): All case tiles include the following case-specific metrics: Impella Flow (Min, Max, and Average), P-Level, and Purge System (Flow and Pressure). This feature solely displays information equivalent to what’s communicated on the AIC re-organized for user convenience, does not perform any data analysis, and is not intended to support any clinical decision making. | inappropriate case metrics | data issue
Notifications via Email (Non-Device CDS Software Function): This feature allows users to optionally subscribe to and receive email notifications based on the AIC alert trigger data that is streamed to the cloud. This feature is solely intended for ease of use to support pump management and is not intended for active patient monitoring | fail to receive email notification | loss of data
Live Video-Stream of AIC Screen (Non-Device-MDDS Software Function): This feature provides live passive viewing of the AIC video screen. Changes to the controller settings can only be made using the physical controls on the AIC. This feature solely displays information equivalent to what’s communicated on the AIC, does not perform any data analysis, and is not intended to support any clinical decision making. | fail to receive video stream | loss of data
Video Recording Navigation (Non-Device-MDDS Software Function): This feature provides additional navigation tools for the live stream of the AIC video screen, including ‘Previous 5’, ‘Live’, ‘Archive’, and ‘Bookmark’. These tools enable the user to navigate back to the previous 5 seconds of the recording, return to the live video stream, access archived video, and bookmark a specific time point in the recording. This feature is solely intended to display medical device data to support clinical education, does not perform any data analysis, and is not intended to support any clinical decision making. | fail to navigate video | application program issue
Video Recording Navigation (Non-Device-MDDS Software Function): This feature provides additional navigation tools for the live stream of the AIC video screen, including ‘Previous 5’, ‘Live’, ‘Archive’, and ‘Bookmark’. These tools enable the user to navigate back to the previous 5 seconds of the recording, return to the live video stream, access archived video, and bookmark a specific time point in the recording. This feature is solely intended to display medical device data to support clinical education, does not perform any data analysis, and is not intended to support any clinical decision making. | fail to bookmark video | application program issue
List of Alerts (Non-Device CDS Software Function): This feature provides a scrollable list of the recent alerts per case. This feature is solely intended to re-organize and display medical device data, equivalent to what’s provided on the AIC, for user convenience and to support clinical education. This feature does not perform any data analysis and is not intended to support any clinical decision making. | fail to provide alert list | no display or display failure
Archive (Non-Device-MDDS Software Function):: This feature allows a user to navigate to archived video. This feature is solely intended to display a record of medical device data and does not provide any form of data analysis or interpretation. | fail to navigate archive | data issue
In accordance with this guidance document, we have assessed the potential impact of the Impella Connect Non-Device Functions within the broader Impella system. Because the Impella Connect system is designed to receive only one-way data communications from the AIC and bi-directional communications are not possible, the Non-Device Software Functions do not have any impact on the AIC’s device functions or performance. | one-way communication violated | communication or transmission issue
Impella Connect Cloud Infrastructure/Portal. The commercialized software functions of the Infrastructure/Portal are all classified as a Non-Device-MDDS Software Function and/or a Non-Device CDS Software Function in the United States. In accordance with the referenced FDA Guidance documents, software functions in these classifications are not considered devices as per section 201(h) of the FD&C Act and are thus not subject to regulatory oversight. Therefore, there are no applicable United States regulatory requirements for the Impella Connect Cloud Infrastructure/Portal software functions. | incorrect regulatory requirement | no apparent adverse event
Impella Connect Remote Link Module (RLM). As described, the RLM is classified as a hardware Device-MDDS (CLASS 1) in the United States and is solely intended to transfer data from the AIC to the Cloud Infrastructure/Portal. In accordance with the FDA Guidance Document, Medical Device Data Systems, Medical Storage Devices, and Medical Image Communications Devices, issued September 28, 2022, “FDA does not intend to enforce the requirements under the FD&C Act for hardware functions that are considered to be Device MDDS, medical image storage devices, or medical image communications devices, provided that the hardware function is limited to assisting the following software functions: electronic transfer, storage, conversion of formats, or display of medical device data”. Therefore, the functionality of the RLM is not subject to regulatory oversight in the United States. | incorrect regulatory classification | no apparent adverse event
The modification to the AIC (Automated Impella Controller) to incorporate the Impella Connect RLM has been reviewed and approved by FDA under the following regulatory files: P140003/S025 and P140003/S073. The subject of these submissions were hardware (P140003/S025) and software (P140003/S073) modifications to the AIC to enable the one-way data streaming functionality of the RLM, thus substantiating potential performance differences of the AIC as it relates to EMC/EMI and Electrical Safety as well as cybersecurity mitigations due to installation of the RLM. The RLM itself was not the subject of these submissions as it meets the definition of a Class I hardware MDDS | cybersecurity subjected | unauthorized access to computer system
The modification to the AIC (Automated Impella Controller) to incorporate the Impella Connect RLM has been reviewed and approved by FDA under the following regulatory files: P140003/S025 and P140003/S073. The subject of these submissions were hardware (P140003/S025) and software (P140003/S073) modifications to the AIC to enable the one-way data streaming functionality of the RLM, thus substantiating potential performance differences of the AIC as it relates to EMC/EMI and Electrical Safety as well as cybersecurity mitigations due to installation of the RLM. The RLM itself was not the subject of these submissions as it meets the definition of a Class I hardware MDDS | EMI subjected | electromagnetic interference (EMI)
Device definition: Temporary Non-Roller Type Left Heart Support Blood Pump. A blood pump that provides temporary full or partial left heart support. Rotor or axial type pump manufactured of various materials. Single use, disposable. | inappropirate device definition | data issue
Technical Method: pumps part or all of the patient's blood circulation either during an adjunctive procedure where the patient requires hemodynamic stability, or a temporary use of the device to unload the heart to give it time to recover from an acute condition. Target area: heart review panel: cardiovascular device class: 3 | incorrect regulatory classification | no apparent adverse event
The Impella Connect system consists of two parts: • A web-based user portal that allows authorized users to remotely view the screen of the Automated Impella Controller (AIC). • A small box (0042-4100) attached to the back and side of an Automated Impella Controller (AIC) | device attachment loosened | detachment of device component
The box connects to the VGA output of the AIC and transmits the display of the AIC screen and data stored by the AIC to a cloud based server. The transmitted information can be viewed by authorized remote users, which may include hospital clinicians, Abiomed local support staff and Abiomed Clinical Support Center (CSC) team members. The Impella Connect enables passive viewing the AIC screen. Changes to the controller settings can only be made using the physical controls on the AIC. Impella Connect may be configured and connected to the hospital’s secured Wi-Fi network or be plugged in directly via Ethernet cable to the secured hospital network in order to transmit the video image to a web-based portal. | settings changed unauthorizedly | application security issue
INTENDED USE: Impella Connect is intended to be used to enable remote patient monitoring by providing authorized users with passive viewing of the AIC’s display which includes information on alarms and data useful for troubleshooting and managing Impella devices to aid in patient management. Impella Connect transfers the video stream from the AIC (via the VGA output) and data stored by the AIC to a cloud-based remote viewing portal. Communication between the AIC and Impella Connect is one-way (AIC to Impella Connect), and the streamed information is limited to Impella device operating parameters and alarm messages with no patient identifiable information. Impella Connect is powered directly by the AIC. | patient information leaked | unauthorized access to computer system
PRECAUTIONS: The Impella Connect is not intended to provide real-time information for monitoring patient status on the AIC. | fail to provide precaution info | user used incorrect product for intended use
PRECAUTIONS: During use of the Impella Connect, there will be a delay between when an image appears on the AIC screen and when it is displayed at a remote viewing location. | fail to caution display delay | output issue
PRECAUTIONS: The Impella Connect is not a source of patient alarms, nor is its use intended as a replacement for monitoring the AIC’s alarms. | fail to caution patient alarm | device alarm system issue
PRECAUTIONS: During use of Impella Connect, receipt of the displayed controller information is not confirmed by the AIC, nor is the delivery of the displayed AIC information to the authorized remote users guaranteed. | fail to caution display info | data issue
PRECAUTIONS: Impella Connect is not designed for use during transport. | fail to caution transport | protective measure issue
PRECAUTIONS: When Impella Connect is configured to use the hospital’s secured Wi-Fi network, signal quality and strength is dependent on the hospital’s Wi-Fi. | fail to caution network | radio signal problem
PRECAUTIONS: When Impella Connect is configured to use the hospital’s secured Wi-Fi network, connection may be lost in physical locations of low Wi-Fi signal. | fail to caution Wi-FI | disconnection
PRECAUTIONS: No modification of this equipment is permitted. | fail to caution device modification | device maintenance issue
PRECAUTIONS: Radiated and conducted electromagnetic interference can affect the performance of the Impella Connect, causing a temporary loss of connectivity. To clear interference, either increase the distance between system components and the EMI source or turn off the EMI source. Any electromagnetic interference related to the Impella Connect will have no impact on any of the AIC functional specifications. | fail to caution EMI | electromagnetic interference (EMI)
PRECAUTIONS: Portable and mobile RF communications equipment can affect medical electrical equipment | fail to caution RF | radiofrequency interference (RFI)
PRECAUTIONS: When in use, the Ethernet cables connecting the Impella Connect to the wall could represent a tripping hazard and should be placed where foot traffic is minimal. | fail to caution tripping hazard | device unsafe to use in environment
PRECAUTIONS: Keep excess cabling on the cable wrap when not needed. | fail to caution tripping hazard | use of device issue
PRECAUTIONS: Disconnect, when in use, the Ethernet cables from the wall outlet before moving the AIC. | fail to caution cable | use of device issue
PRECAUTIONS: Impella Connect is not interpretive | incorrect labeling interpretation | device markings issue
PRECAUTIONS: Impella Connect is not intended to control or interact with any drugs | fail to caution drug interaction | protective measure issue
The Impella Connect will automatically connect to the cloud when in range of an approved* Wi-Fi network. In locations where Wi-Fi is not available, the Impella Connect can be connected to the cloud via the Ethernet port located on the Impella Connect. | fail to connect to WI-FI | wireless communication issue
To configure a Wi-Fi network, a user must enter Maintenance Mode. To do so, firmly press the Flight Mode button until the Connection Indicator light flashes amber and continue holding for 15 seconds before release. Once released, after a few seconds, the Connection Indicator light will turn to a continuous flashing blue indicating the Impella Connect is now in Maintenance Mode. Note: The Connection Indicator for older versions of the Impella Connect module may not flash amber when pressed firmly. | fail to configure network | use of device issue
To exit Maintenance Mode, press the Flight Mode button for >1 second and release. | fail to exit mode | use of device issue
Approved network means SSID and password/credentials have been configured in the Impella Connect by an authorized IT administrator. | fail to approve network | computer system security issue
To disable the Wi-Fi transmitter prior to air transport, press the Flight Mode button for between 3 and 10 seconds and release to enter Flight Mode. The Flight Mode indicator will flash then turn solid amber. To exit Flight mode, press the Flight Mode button for >1 second and release. | fail to deactivate mode | use of device issue

## connect_dFMEA
Input | Failure Mode | Hazard
Component Description: Board Power Component Intended Function: Provide power to all components on RLM and external USB port Component Specification: Input voltage 5V, maximum current = 1.4 A; multiple power rails derived from this input voltage | component short | electrical shorting
Component Description: Board Power & USB Cable Assembly (Power) Component Intended Function: Provide power to all components on RLM and external USB port Component Specification: Input voltage 5V, maximum current = 1.4 A; multiple power rails derived from this input voltage | component short | electrical shorting
Component Description: Board Power & USB Cable Assembly (Power) Component Intended Function: Provide power to all components on RLM and external USB port Component Specification: Input voltage 5V, maximum current = 1.4 A; multiple power rails derived from this input voltage | Component Open | electrical issue
Component Description: USB Hub & USB Cable Assembly (Data) Component Intended Function: Provide connection between AIC USB, internal RLM USB and external USB Component Specification: USB2.0 | faulty component | disconnection
Component Description: VGA & FPGA Component Intended Function: Provide video feed from AIC to IC Component Specification: Programmable logic | faulty component | no display or display failure
Component Description: FPGA & Switch Component Intended Function: Toggle Switch (USB data to SOM) Component Specification: Programmable logic | faulty switch | electrical issue
Component Description: SOM + SD Card Component Intended Function: Main CPU, Memory and Data Storage Component Specification: ARM based System on Module, micro SD card | faulty component | electrical shorting
Component Description: SOM + SD Card Component Intended Function: Main CPU, Memory and Data Storage Component Specification: ARM based System on Module, micro SD card | software failure | computer software issue
Component Description: SOM + SD Card Component Intended Function: Main CPU, Memory and Data Storage Component Specification: ARM based System on Module, micro SD card | incompatible card | installation-related problem
Component Description: Firewall Component Intended Function: Power Management and Separation/ Monitoring of traffic between SOM & USB and I/O button & LED management Component Specification: PIC microcontroller | faulty component | electrical shorting
Component Description: Firewall Component Intended Function: Power Management and Separation/ Monitoring of traffic between SOM & USB and I/O button & LED management Component Specification: PIC microcontroller | firmware failure | computer software issue
Component Description: RLM user interface Component Intended Function: Informs User of connection Status and allows user the ability to enter Flight Mode/ Maintenance Mode. Component Specification: Blue LED, Orange LED, and 1*Push button | fail to indicate connection status | inadequate user interface
Component Description: LED Blink Codes Component Intended Function: Informs User of connection Status. Component Specification: Blue LED and Orange LED | fail to inform status | use of device issue
Component Description: LED Blink Codes Component Intended Function: Informs User of connection Status. Component Specification: Blue LED and Orange LED | fail to indicate status | inadequate user interface
Component Description: RLM Component Intended Function: Transmit the display of the AIC screen and data stored by the AIC to a cloud based server Component Specification: Wi-Fi b/g/n standard | device interfered with WiFi | radiofrequency interference (RFI)
Component Description: RLM Component Intended Function: Transmit the display of the AIC screen and data stored by the AIC to a cloud based server Component Specification: Wi-Fi b/g/n standard | device interfered with radio emission | radiofrequency interference (RFI)
Component Description: RLM Component Intended Function: Transmit the display of the AIC screen and data stored by the AIC to a cloud based server Component Specification: RLM mounts on the AIC | cart tipped over | device or device component damaged by another device
Component Description: RLR Component Intended Function: Heartbeat Communication w. IC Module, OCR Screenshots, JSON Data, Config. and Diagnostic Actions. Component Specification: Remote Link Router Communication Interface Protocol | loss of service | failure to service
Component Description: Wowza Component Intended Function: 3rd party software which handles recording video streams as well as playback. Component Specification: Service should be able to hosted on VM with Ubuntu | fail to record video | failure to service
Component Description: OCR Component Intended Function: Read values/ colors from images Component Specification: Google OCR & Coordinate Map | inaccurate information | data issue
Component Description: OCR Component Intended Function: Read values/ colors from images Component Specification: System Design | information missed | loss of data
Component Description: RLFS Component Intended Function: Capture logs Component Specification: 2TB Harddrive, VM using Ubuntu OS | disk drive full | loss of data
Component Description: Data Repository Component Intended Function: Store and Serve Data Component Specification: Azure SQL database | insufficient data capacity | loss of data
Component Description: Data Repository Component Intended Function: Store and Serve Data Component Specification: Azure SQL database | data capacity degraded | data issue
Component Description: Web Portal Component Intended Function: Provide authorized users with passive viewing of the AIC’s display and, when available, Abi output. Component Specification: Okta, Azure App Service, Azure Cloud | fail to access web portal | application program issue
Component Description: Web Portal Component Intended Function: Provide authorized users with passive viewing of the AIC’s display and, when available, Abi output. Component Specification: Okta, Azure App Service, Azure Cloud | unauthorized webportal access | unauthorized access to computer system
Component Description: Web Portal Component Intended Function: Provide authorized users with passive viewing of the AIC’s display and, when available, Abi output. Component Specification: Chrome, I/E | incompatible browser | computer software issue
Component Description: Data Pipeline (AIC display) Component Intended Function: Provide authorized users with passive viewing of the AIC’s display Component Specification: Communication Interface Protocol, Azure Services | incorrect display | incorrect display
Component Description: Data Pipeline (Abi display) Component Intended Function: Provide authorized users with Abi output. Component Specification: Communication Interface Protocol, Azure Services | fail to display | no display or display failure
Component Description: Data Pipeline (Care Escalation) Component Intended Function: Provide authorized users with avg. flow defict and minutes in suction Component Specification: PowerBI, Azure Services, Communication Interface Protocol | fail to display | communication or transmission issue

## connect_DIR
Statements | Failure_Mode | Hazard
Headers, Banners and Footers. CSC Contact Information. Every page must include the CSC phone number in the banner. The Case View page also includes additional troubleshooting options.
Headers, Banners and Footers. Navigate to Summary Page. There shall be a way for the user to navigate from any page back to the Summary Page.
Headers, Banners and Footers. Navigation links. Impella Connect shall provide hyperlinks for the user to navigate through different content
User Access - Login Screen The portal must include a login screen that contains relevant information for users to access the website
User Access - Login A username and password is required to access the portal. If the user belongs to an institution that has SSO enabled, the password will be provided to the portal from the user's institution.
User Access - Login for Validated Credentials The portal must be able to recognize whether a username matches an authorized user of the portal and must inform the user if there is not a match
User Access - Limited Incorrect Login Attempts The user must have a limited number of incorrect login attempts before the account is locked Unlocking a locked account must be done by waiting a specified amount of time or by resetting the account password
User Access - Password User password must be private and only be known or set by the user
User Access - HIPAA Passwords User password must meet HIPAA security standards
User Access - Terms and Conditions A single user must only be allowed to agree to the Terms and Conditions for their account, not for other accounts on the same enrollment form or at the same institution.
User Access - Terms and Conditions Agreement All users must agree to Terms and Conditions prior to accessing the website for the first time or accessing the website after any updates have been made to the Terms and Conditions.
User Access - Terms and Conditions Disagreement If a user does not agree to the Terms and Conditions, they must be blocked from accessing the website.
User Access - One Time Terms and Conditions Agreement After accepting the Terms and Conditions prior to first login or prior to first login after Terms and Conditions have been changed, a user must not have to agree again during subsequent login attempts, but must still be able to view the Terms and Conditions through a link on the login page.
User Access - Reset Password A user who forgets their password must be able to reset it using a link on the login screen. A link must be emailed to the user to set a new password.
User Access - Language Display The EU site must be able to be viewed in English, German, Danish, Spanish, French or Italian The JA site must be able to be viewed in English or Japanese.
User Access - Log Out The banner/header of the portal must contain a way to sign out of the site.
User Access - Email Test An Admin user can access the Email Testing page (for testing purposes)
User Access - Session Invalidation The portal must be able to detect inactivity (defined as a period of time in which no requests are made to the server) and invalidate the user's session/log the user out.
Device Management - Adding Devices Abiomed Administrators must be able to add new Remote Link Modules to the cloud, allowing new Modules to appear on the website.
Device Management - Assigning & Changing Institution When new Remote Link Modules are added to the website, the Abiomed Administrator must be able to assign the new Module to a specific institution. Abiomed Administrators must also be able to change the assignation of an existing Module (e.g. change from Manufacturing to a hospital).
Device Management - Assigning & Changing Console The Abiomed Administrator must be able to link Remote Link Modules to a specific Console. Abiomed Administrators must also be able to change the assignation of an existing Module
Device Management - Displaying Added RLMs By default, any new Remote Link Module assigned to an institution becomes visible to any users with access to that institution.
Device Management The website must include a configuration site for Abiomed Administrators to conduct portal configuration activities with active Remote Link Modules
User Management - Required Fields The user management form for adding a user must have required fields and not allow submission unless the required fields are completed
User Management - email on user creation When an Abiomed Administrator adds a new user and the new user account is created, the user must receive a registration email
User Management - Adding New Users Abiomed administrators must be able to add users
User Management - Username The username generated for a new user must be the same as the email address used to enroll. This username cannot be changed by the user
User Management - Assigned Institutions All Abiomed employees who interact with customers regarding cases must have access to the cases those customers are authorized to view (i.e. CSC and Field Sales are all required to have accounts)
User Management - Edit & Deactivating Users Abiomed Administrators must be able to edit or deactivate user accounts
User Management - Add, Edit, and Deactivate Users Abiomed administrators must be able to add, edit, and deactivate internal users without the need for a completed enrollment form
User Management - Internal User Types Different internal user groups must be created to determine level of administrative privileges, which accounts the user can see, and which notifications the user receives
User Management - External User Types Different external user groups must be created to determine level of site access, reporting capabilities, and role in account management
Institution Management - Institution List The Abiomed administrator must have access to a list of active accounts with Remote Link that can be used to match the institutions listed in an enrollment form to Abiomed records
Institution Management - Designation of SSO Institution The Impella Connect Administrator must be able to designate an institution as SSO, meaning that the user's password is managed by the institution rather than Impella Connect.
Receive Salesforce Territory Mapping Data Impella Connect shall receive updated Salesforce data that maps territories to institutions to users every 12 hours to stay synchronized between systems.
General Dashboards: Checking for Pump Connected As soon as a Remote Link Module connects to the site, the site must check to see if there is a pump connected to the AIC.
General Dashboards: Display of Pump Connected As soon as the connection of a pump is confirmed, the connection of the pump must be displayed on all applicable dashboards.
General Dashboards: Viewing All Cases (Summary & Map View) The site must contain a dashboard where all cases the user is authorized to view can be seen in a summary and map view.
General Dashboards: Search Function There shall be a way for a user to search for institutions and cases.
General Dashboards: Checking for Console Connection As soon as the Remote Link Module connects to the site, the site must check to see if there is a console connected
Summary View - Open Case In the Summary View, live cases must be able to be selected, and the user must be taken to the case view and with the live video of the case streaming.

## AIC_sFMEA
Software Requirement or Functional Description | fm_option | Failure Cause (specific) | hazard_option | Failure Mode | Hazard
Graphics Driver Communication | Faulty or Incompatible I/O | Graphics Driver Failure | AIC Startup Failure | incompatible I/O | component or accessory incompatibility
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | faulty or incompatible I/O | component or accessory incompatibility
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | fail to start up console | device inoperable
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | fail to power up | failure to power-up
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | loss of user interface | application interface becomes nonfunctional or program exits abnormally
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | graphics driver crashed | device inoperable
Pump Types & Sensor Types Supported Support Impella 2.5. The software must support the Impella 2.5. Corrupted message between EEPROM & EPM or between EPM &EPC -> EEPROM read fails | Faulty or Incompatible I/O | Corrupted message between EEPROM & EPM or between EPM &EPC -> EEPROM read fails | Programming / Software Error | fail to support pump type | computer software issue; component or accessory incompatibility
Pump Types & Sensor Types Supported Support Impella 2.5. The software must support the Impella 2.5. Incorrect PMD SW installed which does not support pump type | Faulty or Incompatible I/O | Incorrect PMD SW installed which does not support pump type | Programming / Software Error | fail to support sensor type | problem with software installation
Pump Types Supported Support Impella 2.5. The software must support the Impella 2.5. New pump type/sensor available at site without updated software. | Faulty or Incompatible I/O | New pump type/sensor available at site without updated software | Programming / Software Error | fail to upgrade software | application program version or upgrade problem
Pump Flow Control – P-Level Given a pump is connected, when the pump is set to a given P-Level, then the software shall adjust the target motor speed based on the performance values stored in the pump EEPROM (/A_170/). | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error | fail to adjust motor speed | communication or transmission issue
Pump Flow Control – P-Level Given a pump is connected, when the pump is set to a given P-Level, then the software shall adjust the target motor speed based on the performance values stored in the pump EEPROM (/A_170/). | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error | fail to adjust motor speed | application program issue
Step 3: Connect Impella Catheter - Display Next Soft Key Given the Step 3 - Connect Cable: Connect Impella Catheter screen is displayed, when 20 seconds have elapsed, then a NEXT soft key shall be displayed and the EPM shall be power cycled. | Faulty or Incompatible I/O | After power cycle, EPM fails to establish communication with EPC | Programming / Software Error | unable to display soft key | application program issue
Step 3: Connect Impella Catheter - Display Next Soft Key Given the Step 3 - Connect Cable: Connect Impella Catheter screen is displayed, when 20 seconds have elapsed, then a NEXT soft key shall be displayed and the EPM shall be power cycled. | Faulty or Incompatible I/O | After power cycle, EPM fails to establish communication with EPC | Programming / Software Error | fail to display soft key | incorrect display
Step 3: Connect Impella Catheter - Display Next Soft Key Given the Step 3 - Connect Cable: Connect Impella Catheter screen is displayed, when 20 seconds have elapsed, then a NEXT soft key shall be displayed and the EPM shall be power cycled. | Faulty or Incompatible I/O | After power cycle, EPM fails to establish communication with EPC | Programming / Software Error | blank screen | communication or transmission issue
Step 3: Connect Impella Catheter - Display Next Soft Key Given the Step 3 - Connect Cable: Connect Impella Catheter screen is displayed, when 20 seconds have elapsed, then a NEXT soft key shall be displayed and the EPM shall be power cycled. | Faulty or Incompatible I/O | After power cycle, EPM fails to establish communication with EPC | Programming / Software Error | no display | application program issue
Step 4 Completion Criteria Given an Impella with an optical sensor is detected and the purge cassette is primed and the luers are detected as connected, when Step 3 is complete, then Step 4 is skipped and Step 6: Purge Fluid Entry is displayed and the purge system is set to standard configuration. | False Positive Condition (SW detects the condition when there is none) | Kink in the purge line causes false detection of luers connected and cassette primed and kink is removed before kink error screen is displayed | Inadequate Priming Of Purge Line | luer connection condition detected | application program issue
Purge Lumen Prime Completion - Optical Impella Pumps Given an Impella with an optical sensor is detected and the purge lumen is priming, when the purge lumen has been primed for 30 seconds, then switch to normal purge mode. | Faulty or Incompatible I/O | Communication error between the EPC and AHP causes the purge system to not be set to normal purge mode | Rough Surface | incorrect purge mode applied | application program issue
Purge Lumen Prime Completion - Optical Impella Pumps Given an Impella with an optical sensor is detected and the purge lumen is priming, when the purge lumen has been primed for 30 seconds, then switch to normal purge mode. | Faulty or Incompatible I/O | Communication error between the EPC and AHP causes the purge system to not be set to normal purge mode | Rough Surface | incorrect purge mode | use of incorrect control settings
Purge Lumen Prime Completion - Optical Impella Pumps Given an Impella with an optical sensor is detected and the purge lumen is priming, when the purge lumen has been primed for 30 seconds, then switch to normal purge mode. | Faulty or Incompatible I/O | Communication error between the EPC and AHP causes the purge system to not be set to normal purge mode | Rough Surface | incorrect purge mode applied | communication or transmission issue
Placement Lumen Flushing Detected Given on Step 5: Flush Placement Lumen or Placement Lumen Not Flushed Error, when the Placement Signal is greater than 100 mmHg and less than 3000 mmHg, then the Placement Lumen is detected as being flushed. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal is not reliable or not available | Inadequate Priming Of Purge Line | data corrupted | failure to prime
Placement Lumen Flushing Detected Given on Step 5: Flush Placement Lumen or Placement Lumen Not Flushed Error, when the Placement Signal is greater than 100 mmHg and less than 3000 mmHg, then the Placement Lumen is detected as being flushed. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal is not reliable or not available | Inadequate Priming Of Purge Line | unreliable signal | data issue
Start Impella - P-Level Selection Given the Flow Control Input is displayed on the Step 7 Start Impella screen and a non-zero P-level is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must start running at the selected Flow Control Option and the Placement Screen must be displayed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | unresponsive soft key | application interface becomes nonfunctional or program exits abnormally
Start Impella - P-Level Selection Given the Flow Control Input is displayed on the Step 7 Start Impella screen and a non-zero P-level is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must start running at the selected Flow Control Option and the Placement Screen must be displayed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | unresponsive knob | application interface becomes nonfunctional or program exits abnormally
Start Impella - P-Level Selection Given the Flow Control Input is displayed on the Step 7 Start Impella screen and a non-zero P-level is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must start running at the selected Flow Control Option and the Placement Screen must be displayed. Given the Flow Control Input is displayed on the Step 7 Start Impella screen and P0 is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must not start running and the Flow Control Input must be removed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | inappropriate P-level highlighted | application program issue
Start Impella - P-Level Selection Given the Flow Control Input is displayed on the Step 7 Start Impella screen and P0 is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must not start running and the Flow Control Input must be removed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | pump activated inappropriately | application program issue
Start Impella Confirmation - YES Soft Key Given the Start Impella Confirmation screen, when the YES soft key is pressed, then the Impella Catheter shall be started in AUTO mode and the Placement Screen shall be displayed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | inappropriate mode displayed | computer software issue
Confirm Sensor Offset Given the Differential Sensor Calibration: Calibration Complete screen is displayed and the placement signal was valid during the zero calibration, when the user presses the YES soft key, then the new Manual Zero offset must be saved to the pump EEPROM and Start Impella Screen must be displayed. Given the Differential Sensor Calibration: Calibration Complete screen is displayed and the placement signal was not reliable during the zero calibration, when the user presses the YES soft key, then the new offset must be discarded, the original offset must be used and the Start Impella Screen must be displayed. Given the Differential Sensor Calibration: Calibration Complete screen is displayed, when the user presses the NO soft key, then the new offset must be | Faulty or Incompatible I/O | Communication failures with EPM | Incorrect Pump Performance Metrics
Confirm Sensor Offset Given the Differential Sensor Calibration: Calibration Complete screen is displayed and the placement signal was valid during the zero calibration, when the user presses the YES soft key, then the new Manual Zero offset must be saved to the pump EEPROM and Start Impella Screen must be displayed. Given the Differential Sensor Calibration: Calibration Complete screen is displayed and the placement signal was not reliable during the zero calibration, when the user presses the YES soft key, then the new offset must be discarded, the original offset must be used and the Start Impella Screen must be displayed. Given the Differential Sensor Calibration: Calibration Complete screen is displayed, when the user presses the NO soft key, then the new offset must be | Faulty or Incompatible I/O | Communication failures with EPM | Programming / Software Error
Motor Control Safety (Pump Stop - Motor Current) If the total motor current of the pump motor exceeds the maximum allowed current, then the motor is stopped and the PMD SW alerts the EPC SW. | Faulty or Incompatible I/O | PMD comm failure - Pump stop is not reported to EPC | Motor Vibrates during the usage of the device
Motor Control Safety (Pump Stop - Motor Speed) Given a pump is connected and running, when the pump motor driver (PMD) detects a motor speed deviation of greater than ±8000 rpm, then the motor is stopped and the Impella Stopped alarm (GID-ALRM-231716) must be triggered. The detection of this condition shall be delayed to allow the EPC SW to detect a motor speed deviation first. | Faulty or Incompatible I/O | PMD comm failure - Pump stop is not reported to EPC | Motor Vibrates during the usage of the device
Pump Pressure Out of Range Given the pump pressure out of range monitoring is active, when the pump pressure is less than the P-level dependent minimum limit defined in the EEPROM (/E_20/, /F_50/) for the monitoring period defined in the pump EEPROM (/E_20/, /F_40/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Sensor System Failure
Pump Pressure Out of Range Given the pump pressure out of range monitoring is active, when the pump pressure is more than the P-level dependent maximum limit defined in the EEPROM (/E_30/, /F_50/) for the monitoring period defined in the pump EEPROM (/E_30/, /F_40/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Rough Surfaces
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Components damaged during insertion/removal
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Misuse of pump metrics
Pump Flow Calculation Based on Placement Signal Given a pump is connected, when the Flow Calculation pump EEPROM value in /M_65/ is 1, then the Flow Calculation based on Placement Signal Algorithm is enabled. The algorithms used for pump flow calculation based on placement signal can be found in 0042-96256. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked Placement signal adjustment applied when placement signal is not reliable | Programming / Software Error
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Incorrect Algorithms | Changes in ventricular or arterial compliance. | Incorrect Pump Performance Metrics
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Basic Cardiac Output Calculation The Basic Cardiac Output must be the user entered Cardiac Output | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Native Cardiac Output Calculation The Native Cardiac Output shall be calculated using the equation below: NCO = ( NCO0 / PP0 ) * PP where • NCO0 is the Initial Native Cardiac Output (see GID-SRS-516275) • PP0 is the Initial Pulse Pressure (see GID-SRS-516274) • PP = PPmax - PPmin | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Native Cardiac Output Calculation The Native Cardiac Output shall be calculated using the equation below: NCO = ( NCO0 / PP0 ) * PP where • NCO0 is the Initial Native Cardiac Output (see GID-SRS-516275) • PP0 is the Initial Pulse Pressure (see GID-SRS-516274) • PP = PPmax - PPmin | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Advanced Cardiac Power Output Calculation Advanced Cardiac Power Output shall be calculated every 3 seconds using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-516272) MAP: Mean Arterial Pressure | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Advanced Cardiac Power Output Calculation Advanced Cardiac Power Output shall be calculated every 3 seconds using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-516272) MAP: Mean Arterial Pressure | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Basic Cardiac Power Output Calculation Basic Cardiac Power Output shall be calculated using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-688038) MAP: Mean Arterial Pressure at time of Cardiac Output entry | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Update CO Reminder - Triggering Criteria Given the Update CO Reminder Counter is incremented, when the Update CO Reminder Counter reaches 150, then the Update CO Reminder must be displayed. The value 150 was chosen to balance detecting true error while minimizing alarming for false error, see ES2020-074 for more information.The value of 150 was also selected because this is approximately 7.5 minutes, which is a round value for ease of training and communication. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Suction Impella position not ok | Incorrect Pump Performance Metrics
EDP Calculation Per Beat When the LVP signal is calculated, then the EDP calculated value shall be the inflection point of the LVP signal at the end of diastole. | Incorrect Algorithms | Suction causes the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
EDP Calculation Per Beat When the LVP signal is calculated, then the EDP calculated value shall be the inflection point of the LVP signal at the end of diastole. | Incorrect Algorithms | Pump position in the aorta causes the pump metrics to be inaccurate. | Incorrect Pump Performance Metrics
LVP Calculation Given LVP signal is calculated, when AoP and motor current values are available, then the AIC shall calculate LVP using the equation below: LVP = AoP - dP AoP: Placement Signal dP: differential pressure see GID-SRS-662382 | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Pump Placement Calculation – Open Pressure Measurement The open pressure measurement algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked | Sensor System Failure
Placement Signal Lumen Blocked Detection The placement signal lumen is blocked if the placement signal is valid, the placement signal pulsatility is ≤ 10 mmHg and the placement signal minimum is > 130 mmHg for 30 consecutive seconds. Alarm criteria time measurement accuracy: ±3 seconds | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Optical Sensor The optical sensor placement algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked Placement signal adjustment applied when placement signal is not reliable | Sensor System Failure
Pump Placement Calculation – Phase Shift Detection Algorithm The phase shift detection algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Pump Placement Calculation The Motor Current Transition Detection Algorithm shall detect if the aortic valve is closing on the pump outlet. If the motor current transition exceeds the threshold (/E_230/, /E_240/) more than once in 2 minutes then the algorithm shall report the valve is contacting the pump outlet area. (This flag is currently checked by the open pressure measurement pump position monitoring.) Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation The Motor Current Transition Detection Algorithm shall detect if the aortic valve is closing on the pump outlet. If the motor current transition exceeds the threshold (/E_230/, /E_240/) more than once in 2 minutes then the algorithm shall report the valve is contacting the pump outlet area. (This flag is currently checked by the open pressure measurement pump position monitoring.) Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Rough Surfaces
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Components damaged during insertion/removal
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Misuse of pump metrics
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Optical Sensor Failure
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Sensor System Failure
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | CPU/Memory Management (out of CPU or memory) | CPU/Memory Management (out of CPU or memory) | AIC Startup Failure
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Sensor System Failure
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Rough Surfaces
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Components damaged during insertion/removal
OPM - Signal to Noise Ratio (Case Start) Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the average SNR used by Case Start shall be bad. The average SNR used by Case Start must be greater than the limit contained in the pump EEPROM plus 50 before the SNR is good again. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Misuse of pump metrics
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | AIC Startup Failure
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Sensor System Failure
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Rough Surfaces
OPM - Ambient Pressure (Pump Connected) Given an Impella pump with an optical sensor is connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Components damaged during insertion/removal
OPM - Ambient Pressure (Pump Connected) Given an Impella pump with an optical sensor is connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Misuse of pump metrics
Data Streaming: Enabled Given the AIC Controller is powered on, when an Impella Connect Module is detected, then data streaming must be enabled. | Incorrect Algorithms | Cybersecurity Attack Over USB Bi-directional communication on USB connection allows entry for cybersecurity attack | Programming / Software Error
Data Streaming: Enabled Given the AIC Controller is powered on, when an Impella Connect Module is detected, then data streaming must be enabled. | CPU/Memory Management (out of CPU or memory) | Software Process Lockup Software process consumes too much CPU/Memory or blocks critical processes from running | AIC Startup Failure
Data Streaming: AIC System Information Given an AIC is powered on, when data streaming is first enabled, then the following AIC information must be sent to the Impella Connect Module: • System revision string • AIC serial number • EPC version string • AHP version string • EPM version string • PBM1 version string • PBM2 version string • KBD version string • PMD/CPLD version string • OPM version string • OPM serial number • IP address • Netmask • Gateway Refer to 0042-97042 Impella Connect Communication Interface Protocol. | Faulty or Incompatible I/O | For AIC SN, System Revision & network configuration: Software fails to trigger sending the data or data is lost in the transmission to the Impella Connect Module (ICM). Data is only sent once and never reaches the Impella Connect cloud.
Data Streaming: AIC System Information Given an AIC is powered on, when data streaming is first enabled, then the following AIC information must be sent to the Impella Connect Module: • System revision string • AIC serial number • EPC version string • AHP version string • EPM version string • PBM1 version string • PBM2 version string • KBD version string • PMD/CPLD version string • OPM version string • OPM serial number • IP address • Netmask • Gateway Refer to 0042-97042 Impella Connect Communication Interface Protocol. | Faulty or Incompatible I/O | For AIC software versions, OPM SN: Software fails to trigger sending the data or data is lost in the transmission to the Impella Connect Module (ICM). Data is only sent once and never reaches the Impella Connect cloud.
Data Streaming: Purge Fluid Concentration Given the AIC is powered on, when • data streaming is first enabled, or • a change in the purge fluid dextrose, heparin or bicarb concentration is accepted, or • one minute has passed since the last send of the purge fluid concentration, then the purge fluid concentration must be sent to the Impella Connect Module. | Faulty or Incompatible I/O | Software fails to trigger sending the data or data is lost in the transmission to the Impella Connect Module (ICM). Data is only sent once and never reaches the Impella Connect cloud.
Test Execution Cleanup | Faulty or Incompatible I/O | Temporary connection error causes test cleanup steps to not be executed | AIC Startup Failure
Test Execution Cleanup | Faulty or Incompatible I/O | Temporary connection error causes test cleanup steps to not be executed | Motor Vibrates during the usage of the device
Test Execution Setup | Faulty or Incompatible I/O | Temporary connection error causes test execution setup steps to not be executed | AIC Startup Failure
Test Execution Setup | Faulty or Incompatible I/O | Temporary connection error causes test execution setup steps to not be executed | Motor Vibrates during the usage of the device
Test Run Status | False Positive Condition (SW detects the condition when there is none) | Test Run passes all test steps, but unexpected alarm is triggered during test | AIC Startup Failure
Test Run Status | False Positive Condition (SW detects the condition when there is none) | Test Run passes all test steps, but unexpected alarm is triggered during test | Motor Vibrates during the usage of the device
Test Run Log Files | Faulty or Incompatible I/O | Missing log files due to connection or download error
Test Results Stored in Jama | Faulty or Incompatible I/O | Test execution data is not recorded in Jama due to connection error
Test Report Generation | Faulty or Incompatible I/O | Missing test report due to connection error
AIC Software and Hardware Versions | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Opsens firmware version is manually set incorrectly in AIC console configuration file
TCP Socket - Send & Receive | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | TCP command and/or data is incorrect during send/receive | AIC Startup Failure
TCP Socket - Send & Receive | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | TCP command and/or data is incorrect during send/receive | Motor Vibrates during the usage of the device
TCP Socket - Send & Receive | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | TCP command and/or data is corrupt during send/receive | AIC Startup Failure
TCP Socket - Send & Receive | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | TCP command and/or data is corrupt during send/receive | Motor Vibrates during the usage of the device
TCP Socket - Send & Receive | Faulty or Incompatible I/O | TCP command is lost during communication | AIC Startup Failure
TCP Socket - Send & Receive | Faulty or Incompatible I/O | TCP command is lost during communication | Motor Vibrates during the usage of the device
Communications with the EPC (Motor Current) • Motor current | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Missing two motor current samples per day | N/A
Communications with the EPC (Motor Current) • Motor current | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | PMD unable to handle/generate the samples at 250Hz | AIC Startup Failure
Communications with the EPC (Motor Current) • Motor current | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | AIC disk space not enough to store 14 days of logs | N/A
Communications with the EPC (Motor Current) • Motor current | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | Motor current shared memory data may out of sync in a packet by 4ms or more
Communications with the EPC (Motor Current) • Motor current | Faulty or Incompatible I/O | Dropped or corrupted message; EPC unable to handle/generate the samples at 250Hz | Blood Damage
Communications with the EPC (Motor Current) • Motor current | Faulty or Incompatible I/O | EPC unable to handle/generate the samples at 250Hz | AIC Startup Failure
Communications with the EPC (Motor Current) • Motor current | Faulty or Incompatible I/O | Dropped or corrupted message; EPC unable to handle/generate the samples at 250Hz | Blood Damage
Communications with the EPC (Motor Current) • Motor current | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Missing two motor speed samples per day
Communications with the EPC (Motor Current) • Motor current | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | AIC disk space not enough to store 14 days of logs | N/A
Communications with the EPC (Motor Current) • Motor current | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | Motor speed shared memory data may out of sync in a packet by 4ms or more
OPM – Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Sensor System Failure
OPM – Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Sensor System Failure
PMD Hardware Compatibility Given the AIC console is powered on, when the PMD hardware revision is not 0, 1, 2, 256, 257 or 258, then the PMD software is not compatible with the hardware. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | MC Scaling Factor is not compatible the PMD version compatibility issue | Programming / Software Error
Advanced Pump Monitoring Metrics Given the Impella CP with SmartAssist advanced pump monitoring metrics display setting is enabled, when an initialized Impella CP optical pump is connected, then advanced pump monitoring metrics must be enabled with the following metrics: • Left Ventricular Pressure (LVP) • End Diastolic Pressure (EDP) • Advanced Cardiac Output (CO) • Advanced Cardiac Power Output (CPO) | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Software algorithm falsely detects condition related to Pump Monitoring Metrics | Programming / Software Error
Cardiac Output calibration control Cardiac Output calibration Pause time | Incorrect Algorithms | Decreased Impella speed below P-2 during pinging sequences per design | Incorrect Pump Performance Metrics
Final qualification motor current value for Cardiac Output calibration algorithm | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | EEPROM is corrupted, diastolic flow calculation is not accurate | Programming / Software Error
Alarms to be disabled during CO calibration | Incorrect Algorithms | Suction alarms are masked during CO calibration | Programming / Software Error
250 Hz Data synchronization | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Unable to handle/generate the samples at 250Hz | False Alarm/ Notification
Log Pinging Start | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | Log is not stored or retrieved correctly | N/A
Log Pinging Stop | Failure of Critical Functions (RTLog, Logging, Alarm Generation, etc.) | Log is not stored or retrieved correctly | N/A
Cardiac output calibration Prep screen - entry | Faulty or Incompatible I/O | Impella pump automatically goes into AUTO mode | Programming / Software Error
Trend Data Histories Given the same pump is reconnected, then the Metrics Trend time scale, Y-scale setting, and data histories shall be maintained. Given a different pump is reconnected, then the Metrics Trend Y-scale setting shall be restored to default, and the data histories shall be cleared. | Database Error (storing, retrieving data from a file) | Data is stored in a file which has read/write access errors | AIC Startup Failure
Metrics Trend History - after power cycle The Metrics Trend History for 15 min, 1 hr, 8 hr and 12 hr data must be retained through a manual power cycle and abnormal resets and shutdowns. | Database Error (storing, retrieving data from a file) | Data is stored in a file which has read/write access errors | AIC Startup Failure
Heartrate on Motor Current Channel | Incorrect Algorithms | Software algorithm falsely detects condition related to Pump Monitoring Metrics | Incorrect Pump Performance Metrics
Power-Up Sequence At Power-Up, the IMC Console SW shall start up console operation by initiating the software-based control. | Faulty or Incompatible I/O | Graphics driver crashes and is not restarted resulting in loss of User Interface | AIC Startup Failure | fail to initiate console | failure to reset
Pump Types & Sensor Types Supported The software must support the Impella 2.5. Corrupted message between EEPROM & EPM or between EPM &EPC -> EEPROM read fails | Faulty or Incompatible I/O | Corrupted message between EEPROM & EPM or between EPM &EPC -> EEPROM read fails | Programming / Software Error | message corrupted | loss of data
Pump Types & Sensor Types Supported The software must support the Impella 2.5. | Faulty or Incompatible I/O | Incorrect PMD SW installed which does not support pump type | Programming / Software Error | incompatible software version | application program version or upgrade problem
Pump Types Supported The software must support the Impella 2.5. | Faulty or Incompatible I/O | New pump type/sensor available at site without updated software | Programming / Software Error | fail to update software | application program version or upgrade problem
Step 4 Completion Criteria Given an Impella with an optical sensor is detected and the purge cassette is primed and the luers are detected as connected, when Step 3 is complete, then Step 4 is skipped and Step 6: Purge Fluid Entry is displayed and the purge system is set to standard configuration. | False Positive Condition (SW detects the condition when there is none) | Kink in the purge line causes false detection of luers connected and cassette primed and kink is removed before kink error screen is displayed | Inadequate Priming Of Purge Line | false positive condition detected | invalid sensing
Step 4 Completion Criteria Given an Impella with an optical sensor is detected and the purge cassette is primed and the luers are detected as connected, when Step 3 is complete, then Step 4 is skipped and Step 6: Purge Fluid Entry is displayed and the purge system is set to standard configuration. | Faulty or Incompatible I/O | Communication error between the EPC and AHP causes the purge system to not be set to standard configuration |  | purge cassette prime condition detected | communication or transmission issue
Placement Lumen Flushing Detected Given on Step 5: Flush Placement Lumen or Placement Lumen Not Flushed Error, when the Placement Signal is greater than 100 mmHg and less than 3000 mmHg, then the Placement Lumen is detected as being flushed. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal is not reliable or not available | Inadequate Priming Of Purge Line | flushing error | failure to prime
Placement Lumen Flushing Detected Given on Step 5: Flush Placement Lumen or Placement Lumen Not Flushed Error, when the Placement Signal is greater than 100 mmHg and less than 3000 mmHg, then the Placement Lumen is detected as being flushed. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal is not reliable or not available | Inadequate Priming Of Purge Line | lumen flushing condition detected | failure to flush
Start Impella - P-Level Selection Given the Flow Control Input is displayed on the Step 7 Start Impella screen and a non-zero P-level is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must start running at the selected Flow Control Option and the Placement Screen must be displayed. Given the Flow Control Input is displayed on the Step 7 Start Impella screen and P0 is highlighted, when the rotary knob or the START IMPELLA soft key is pressed, then the pump must not start running and the Flow Control Input must be removed. | Faulty or Incompatible I/O | Communication failure between EPC and PMD | Programming / Software Error | pump activated inappropriately | pumping issue
Motor Control Safety (Pump Stop - Motor Current) If the total motor current of the pump motor exceeds the maximum allowed current, then the motor is stopped and the PMD SW alerts the EPC SW. | Faulty or Incompatible I/O | PMD comm failure - Pump stop is not reported to EPC | Motor Vibrates during the usage of the device
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Sensor System Failure
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Rough Surfaces
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Components damaged during insertion/removal
Pump Pressure Out of Range Given the placement signal is not reliable due to the pump pressure out of range, when the pump pressure is within the range defined in the pump EEPROM (/E_20/ & /E_30/, /F_50/), then the placement signal is reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM or EPM | Misuse of pump metrics
Pump Flow Calculation Based on Placement Signal Given a pump is connected, when the Flow Calculation pump EEPROM value in /M_65/ is 1, then the Flow Calculation based on Placement Signal Algorithm is enabled. The algorithms used for pump flow calculation based on placement signal can be found in 0042-96256. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked Placement signal adjustment applied when placement signal is not reliable | Programming / Software Error
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Incorrect Algorithms | Changes in ventricular or arterial compliance. | Incorrect Pump Performance Metrics
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Advanced Cardiac Output Calculation The Advanced Cardiac Output shall be calculated every 3 seconds using the equation below: CO = NCO + Flow where NCO is the Native Cardiac Output (see GID-SRS-516273) and Flow is the most recent flow value | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Basic Cardiac Output Calculation The Basic Cardiac Output must be the user entered Cardiac Output | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Native Cardiac Output Calculation The Native Cardiac Output shall be calculated using the equation below: NCO = ( NCO0 / PP0 ) * PP where • NCO0 is the Initial Native Cardiac Output (see GID-SRS-516275) • PP0 is the Initial Pulse Pressure (see GID-SRS-516274) • PP = PPmax - PPmin | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Native Cardiac Output Calculation The Native Cardiac Output shall be calculated using the equation below: NCO = ( NCO0 / PP0 ) * PP where • NCO0 is the Initial Native Cardiac Output (see GID-SRS-516275) • PP0 is the Initial Pulse Pressure (see GID-SRS-516274) • PP = PPmax - PPmin | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Advanced Cardiac Power Output Calculation Advanced Cardiac Power Output shall be calculated every 3 seconds using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-516272) MAP: Mean Arterial Pressure | Incorrect Algorithms | Incorrect pump position cause the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
Advanced Cardiac Power Output Calculation Advanced Cardiac Power Output shall be calculated every 3 seconds using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-516272) MAP: Mean Arterial Pressure | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Basic Cardiac Power Output Calculation Basic Cardiac Power Output shall be calculated using the equation below: CPO = (CO * MAP ) / 451 CPO: Cardiac Power Output CO: Cardiac Output (see GID-SRS-688038) MAP: Mean Arterial Pressure at time of Cardiac Output entry | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Update CO Reminder - Triggering Criteria Given the Update CO Reminder Counter is incremented, when the Update CO Reminder Counter reaches 150, then the Update CO Reminder must be displayed. The value 150 was chosen to balance detecting true error while minimizing alarming for false error, see ES2020-074 for more information.The value of 150 was also selected because this is approximately 7.5 minutes, which is a round value for ease of training and communication. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Suction Impella position not ok | Incorrect Pump Performance Metrics
EDP Calculation Per Beat When the LVP signal is calculated, then the EDP calculated value shall be the inflection point of the LVP signal at the end of diastole. | Incorrect Algorithms | Suction causes the pump metrics to be inaccurate | Incorrect Pump Performance Metrics
EDP Calculation Per Beat When the LVP signal is calculated, then the EDP calculated value shall be the inflection point of the LVP signal at the end of diastole. | Incorrect Algorithms | Pump position in the aorta causes the pump metrics to be inaccurate. | Incorrect Pump Performance Metrics
LVP Calculation Given LVP signal is calculated, when AoP and motor current values are available, then the AIC shall calculate LVP using the equation below: LVP = AoP - dP AoP: Placement Signal dP: differential pressure see GID-SRS-662382 | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Pump Placement Calculation – Open Pressure Measurement The open pressure measurement algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked | Sensor System Failure
Placement Signal Lumen Blocked Detection The placement signal lumen is blocked if the placement signal is valid, the placement signal pulsatility is ≤ 10 mmHg and the placement signal minimum is > 130 mmHg for 30 consecutive seconds. Alarm criteria time measurement accuracy: ±3 seconds | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Optical Sensor The optical sensor placement algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal lumen is blocked Placement signal adjustment applied when placement signal is not reliable | Sensor System Failure
Pump Placement Calculation – Phase Shift Detection Algorithm The phase shift detection algorithm shall report when the pump position is correct. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation – Flat Line Detection Algorithm The SW shall enable flat line detection if the Impella Pump EEPROM has the flatline placement monitoring flag set (/A_180/ ENABLE_FLATLINE_DP_MIN). The monitoring range for flat line protection (performance min < flat line monitoring < performance max) shall be defined in the Impella pump EEPROM (/E_130/). The allowable signal limits for the flat line monitoring functions are pump performance dependent (10 limit values) and defined in pump EEPROM (/E_110/, /E_120/). The SW shall initiate an error message if a wrong position based on flat line is detected. Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Sensor System Failure
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation – Differential Frequency Detection Algorithm The SW shall perform differential frequency detection if the Impella Pump EEPROM DFD flag (/A_180/ ENABLE_DIFF_FREQ_P_I_PLACE) or phase shift placement flag (/A_180/ ENABLE_PHASE_SHIFT_PLACE) is set. The algorithm shall perform calculations at the same rate as the motor current algorithm (i.e. 25 samples/sec). If the algorithm result exceeds 160 ms (difference between motor current and placement signal period) then a counter is incremented otherwise the same counter is decremented. If the counter exceeds a limit then the PSD alarm is suspended. For pump EEPROM versions < 03, the limit shall be set to 150. For pump EEPROM versions ≥ 03, the limit shall be defined in the pump EEPROM (/I_66/). Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Pump Placement Calculation The Motor Current Transition Detection Algorithm shall detect if the aortic valve is closing on the pump outlet. If the motor current transition exceeds the threshold (/E_230/, /E_240/) more than once in 2 minutes then the algorithm shall report the valve is contacting the pump outlet area. (This flag is currently checked by the open pressure measurement pump position monitoring.) Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | Incorrect Placement
Pump Placement Calculation The Motor Current Transition Detection Algorithm shall detect if the aortic valve is closing on the pump outlet. If the motor current transition exceeds the threshold (/E_230/, /E_240/) more than once in 2 minutes then the algorithm shall report the valve is contacting the pump outlet area. (This flag is currently checked by the open pressure measurement pump position monitoring.) Please refer to the Pump Placement Algorithms in 0042-96256 for further details on the implementation of this algorithm. | False Positive Condition (SW detects the condition when there is none) | External factors, such as low patient pressures, could cause the position algorithms to incorrectly detect position alarms and display false alarms. | False Alarm/ Notification
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Rough Surfaces
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Components damaged during insertion/removal
Suction / Blocked Inlet – Actions If a suction / blocked inlet is detected the SW shall automatically reduce the pump motor speed if auto suction response is enabled or initiate a Suction alarm if auto suction response is disabled. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available Placement signal adjustment applied when placement signal is not reliable | Misuse of pump metrics
Retrograde Flow – Prevention If retrograde flow control is enabled, the SW shall adjust the minimum rotational speed allowed such that the minimum pump flow is within the range 0.1 to 0.2 L/min. (The minimum P1 rotational speed limit defined in the pump EEPROM /A_170/ shall still apply). | False Positive Condition (SW detects the condition when there is none) | Software underestimates Impella pump flow under high diastolic LV pressure conditions
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Optical Sensor Failure
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Optical bench reports an invalid signal or a non-physiological spike in the placement signal. | Sensor System Failure
OPM - Sensor measuring The OPM SW component shall support the following measuring requirements: • Measuring range: Defined in the pump EEPROM (/E_20/ and /E_30/) • Resolution: +/- 0.1 mmHg • Sample rate: 250 Hz | CPU/Memory Management (out of CPU or memory) | CPU/Memory Management (out of CPU or memory) | AIC Startup Failure
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Sensor System Failure
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Rough Surfaces
OPM - Signal to Noise Ratio Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Components damaged during insertion/removal
OPM - Signal to Noise Ratio (Case Start) Given an Impella pump with an optical sensor is connected, when the optical sensor average signal-to-noise ratio (SNR) is less than the limit contained in the pump EEPROM (/OS_70/), then the average SNR used by Case Start shall be bad. The average SNR used by Case Start must be greater than the limit contained in the pump EEPROM plus 50 before the SNR is good again. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Misuse of pump metrics
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | AIC Startup Failure
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Sensor System Failure
OPM - Ambient Pressure (No Pump Connected) Given an Impella pump with an optical sensor is not connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg) for 20 ±2 seconds, then a controller error alarm must be reported. The alarm is resolved, if the ambient pressure is in the allowed range for 20 ±2 seconds consecutively. The alarm is removed if an Impella pump with an optical sensor is connected and the placement and suction monitoring suspended alarm has been set. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Rough Surfaces
OPM - Ambient Pressure (Pump Connected) Given an Impella pump with an optical sensor is connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Components damaged during insertion/removal
OPM - Ambient Pressure (Pump Connected) Given an Impella pump with an optical sensor is connected, when the ambient pressure for optical pressure measurement is out of range (550-825 mmHg), then the placement signal is not reliable. | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Communication failures with OPM | Misuse of pump metrics
Data Streaming: Enabled Given the AIC Controller is powered on, when an Impella Connect Module is detected, then data streaming must be enabled. | Incorrect Algorithms | Cybersecurity Attack Over USB Bi-directional communication on USB connection allows entry for cybersecurity attack | Programming / Software Error
Data Streaming: Enabled Given the AIC Controller is powered on, when an Impella Connect Module is detected, then data streaming must be enabled. | CPU/Memory Management (out of CPU or memory) | Software Process Lockup Software process consumes too much CPU/Memory or blocks critical processes from running | AIC Startup Failure
Data Streaming: Enabled Given the AIC Controller is powered on, when an Impella Connect Module is detected, then data streaming must be enabled. | Faulty or Incompatible I/O | QNX USB driver fails to start or fails during runtime Failure in software of unknown providence (SOUP)
Data Streaming: AIC System Information Given an AIC is powered on, when data streaming is first enabled, then the following AIC information must be sent to the Impella Connect Module: • System revision string • AIC serial number • EPC version string • AHP version string • EPM version string • PBM1 version string • PBM2 version string • KBD version string • PMD/CPLD version string • OPM version string • OPM serial number • IP address • Netmask • Gateway Refer to 0042-97042 Impella Connect Communication Interface Protocol. | Faulty or Incompatible I/O | For AIC SN, System Revision & network configuration: Software fails to trigger sending the data or data is lost in the transmission to the Impella Connect Module (ICM). Data is only sent once and never reaches the Impella Connect cloud.
Data Streaming: AIC System Information Given an AIC is powered on, when data streaming is first enabled, then the following AIC information must be sent to the Impella Connect Module: • System revision string • AIC serial number • EPC version string • AHP version string • EPM version string • PBM1 version string • PBM2 version string • KBD version string • PMD/CPLD version string • OPM version string • OPM serial number • IP address • Netmask • Gateway Refer to 0042-97042 Impella Connect Communication Interface Protocol. | Faulty or Incompatible I/O | For AIC software versions, OPM SN: Software fails to trigger sending the data or data is lost in the transmission to the Impella Connect Module (ICM). Data is only sent once and never reaches the Impella Connect cloud.
Screen Information Shared Memory Space | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Screen Brightness Shared Memory Space | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
System Information Shared Memory Space | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Placement Signal Information Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Motor Current Information Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Impella Flow Information Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Waveform Simulator Shared Memory Contents | CPU/Memory Management (out of CPU or memory) | Checking for waveform simulation consumes too much CPU | Motor Vibrates during the usage of the device
Waveform Simulation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Simulated data overrides actual data while on patient support | Sensor System Failure
Time Scale Information Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Y-Scale Shared Information Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Active Alarm Info - Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Differential Pressure Shared Memory | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
CO/CPO Indicator Status | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
CO Value String | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
CPO Value String | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
CPO Value Color | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Basic CO/CPO Timestamp String | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
LV Adjustment Screen Message | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
LV Adjustment Screen Default Value | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
LV Adjustment Screen Current Value | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Home Screen Message | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
Home Screen Message Color | CPU/Memory Management (out of CPU or memory) | Writing to shared memory consumes too much CPU | Motor Vibrates during the usage of the device
eAutoMode LVP Signal Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | LVP adjustment is incorrect because adjustment is made when: - suction is detected, or - the Impella position is not ok, or - the placement signal is not relaible, or - user enters the wrong adjustment value | Incorrect description/specification for use of device
eAutoMode LVP Signal Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | LVP adjustment is incorrect because adjustment is made when: - suction is detected, or - the Impella position is not ok, or - the placement signal is not relaible, or - user enters the wrong adjustment value | Incorrect Pump Performance Metrics
eAutoMode LVP Signal Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction | Incorrect description/specification for use of device
eAutoMode LVP Signal Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction | Incorrect Pump Performance Metrics
eAutoMode LVP Signal Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Placement signal not reliable or not available | Incorrect Pump Performance Metrics
Decoupling Factor Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
Decoupling Factor Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
Decoupling Factor Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
Decoupling Factor Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
Decoupling Factor Calculation | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Advanced Starting Procedure - Initial P-Level | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Advanced Starting Procedure - Increase P-Level | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Advanced Starting Procedure - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Advanced Starting Procedure - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Advanced Starting Procedure - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Advanced Starting Procedure - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Advanced Starting Procedure - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Advanced Starting Procedure - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Incorrect Impella Pump Position Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Sensor System Failure
eAutoMode Advanced Starting Procedure - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Rough Surfaces
eAutoMode Advanced Starting Procedure - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Components damaged during insertion/removal
eAutoMode Advanced Starting Procedure - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Misuse of pump metrics
eAutoMode Start Ramp Up Process eAutoMode Ramp Up Process - Complete eAutoMode Start Ramp Down Process eAutoMode Fast Ramp Down Process eAutoMode Ramp Down Process - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Start Ramp Up Process eAutoMode Ramp Up Process - Complete eAutoMode Start Ramp Down Process eAutoMode Fast Ramp Down Process eAutoMode Ramp Down Process - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Start Ramp Up Process eAutoMode Ramp Up Process - Complete eAutoMode Start Ramp Down Process eAutoMode Fast Ramp Down Process eAutoMode Ramp Down Process - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Start Ramp Up Process eAutoMode Ramp Up Process - Complete eAutoMode Start Ramp Down Process eAutoMode Fast Ramp Down Process eAutoMode Ramp Down Process - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Start Ramp Up Process eAutoMode Ramp Up Process - Complete eAutoMode Start Ramp Down Process eAutoMode Fast Ramp Down Process eAutoMode Ramp Down Process - Complete | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Fast Ramp Up Process eAutoMode Slow Ramp Up Process eAutoMode Slow Ramp Down Process eAutoMode Extra Fast Ramp Down Process | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Ramp Up Process - Exit eAutoMode Ramp Down Process - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Incorrect Impella Pump Position Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Sensor System Failure
eAutoMode Ramp Up Process - Exit eAutoMode Ramp Down Process - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Rough Surfaces
eAutoMode Ramp Up Process - Exit eAutoMode Ramp Down Process - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Components damaged during insertion/removal
eAutoMode Ramp Up Process - Exit eAutoMode Ramp Down Process - Exit | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Misuse of pump metrics
eAutoMode Impella Flow Reduction - Alarm Set eAutoMode Impella Flow Reduction - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Impella Flow Reduction - Alarm Set eAutoMode Impella Flow Reduction - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Impella Flow Reduction - Alarm Set eAutoMode Impella Flow Reduction - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Impella Flow Reduction - Alarm Set eAutoMode Impella Flow Reduction - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Impella Flow Reduction - Alarm Set eAutoMode Impella Flow Reduction - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Decoupling Factor Suction Detection - Alarm Set eAutoMode Decoupling Factor Suction Detection - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because LVP is incorrect due to: - Suction, or - An incorrect LVP adjustment, or - Placement signal not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Decoupling Factor Suction Detection - Alarm Set eAutoMode Decoupling Factor Suction Detection - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because AoP is incorrect due to the placement signal being not reliable or not available. | Incorrect Pump Performance Metrics
eAutoMode Decoupling Factor Suction Detection - Alarm Set eAutoMode Decoupling Factor Suction Detection - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Rough Surfaces
eAutoMode Decoupling Factor Suction Detection - Alarm Set eAutoMode Decoupling Factor Suction Detection - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Components damaged during insertion/removal
eAutoMode Decoupling Factor Suction Detection - Alarm Set eAutoMode Decoupling Factor Suction Detection - Alarm Resolved | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Decoupling factor (k) is incorrect because pump is in the incorrect position, which causes k to always equal to 1. | Misuse of pump metrics
eAutoMode Start Auto Suction Response eAutoMode Stop Auto Suction Response | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Rough Surfaces
eAutoMode Start Auto Suction Response eAutoMode Stop Auto Suction Response | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Components damaged during insertion/removal
eAutoMode Start Auto Suction Response eAutoMode Stop Auto Suction Response | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Suction Not Detected due to: - Placement signal not reliable or not available - Placement signal adjustment applied when placement signal is not reliable | Misuse of pump metrics
Switch from eAutoMode to P-Level Mode | Faulty or Incompatible I/O | Communication failures between EPC & PMD (dropped or corrupted message) | Programming / Software Error
eAutoMode Pump Connected | Faulty or Incompatible I/O | EEPROM read fails on pump transfer due to corrupted message between EEPROM & EPM or between EPM & EPC | Programming / Software Error
eAutoMode Pump Connected | Faulty or Incompatible I/O | Pump running in eAutoMode is transferred to a controller with software that does not support eAutoMode.
Purge System Blocked -Impella 5.5 | Faulty or Incompatible I/O | Communication failures with AHP, occlusion or break in the purge sytem during use | Purge System Failure
Purge System Blocked -Impella 5.5 | Faulty or Incompatible I/O | Communication failures with AHP, purge system is compromised due to handling | Purge System Failure
Purge System Blocked -Impella 5.5 | Faulty or Incompatible I/O | Purge Pressure sensor failure, occlusion or break in the purge sytem during use | Purge System Failure
Purge System Blocked -Impella 5.5 | Faulty or Incompatible I/O | Purge Pressure sensor failure, purge system is compromised due to handling | Purge System Failure
Purge Pressure High -Impella 5.5 Given an Impella 5.5 has been detected as connected and the purge system blocked alarm is not active, when the average purge pressure is greater than or equal to Pmax and the purge flow is less than Fmin, then a purge pressure high alarm must be triggered. Where Pmax and Fmin are defined in the purge cassette RFID (0043-9600 Purge Cassette – RFID Parameters). | Faulty or Incompatible I/O | Communication failures with AHP, occlusion or break in the purge sytem during use | Purge System Failure
Purge Pressure High -Impella 5.5 Given an Impella 5.5 has been detected as connected and the purge system blocked alarm is not active, when the average purge pressure is greater than or equal to Pmax and the purge flow is less than Fmin, then a purge pressure high alarm must be triggered. Where Pmax and Fmin are defined in the purge cassette RFID (0043-9600 Purge Cassette – RFID Parameters). | Faulty or Incompatible I/O | Communication failures with AHP, purge system is compromised due to handling | Purge System Failure
Purge Pressure High -Impella 5.5 Given an Impella 5.5 has been detected as connected and the purge system blocked alarm is not active, when the average purge pressure is greater than or equal to Pmax and the purge flow is less than Fmin, then a purge pressure high alarm must be triggered. Where Pmax and Fmin are defined in the purge cassette RFID (0043-9600 Purge Cassette – RFID Parameters). | Faulty or Incompatible I/O | Purge Pressure sensor failure, occlusion or break in the purge sytem during use | Purge System Failure
Purge Pressure High -Impella 5.5 Given an Impella 5.5 has been detected as connected and the purge system blocked alarm is not active, when the average purge pressure is greater than or equal to Pmax and the purge flow is less than Fmin, then a purge pressure high alarm must be triggered. Where Pmax and Fmin are defined in the purge cassette RFID (0043-9600 Purge Cassette – RFID Parameters). | Faulty or Incompatible I/O | Purge Pressure sensor failure, purge system is compromised due to handling | Purge System Failure
Complete Priming - Completion | False Positive Condition (SW detects the condition when there is none) | Case Start Step completed priming does not exit because too high of a pressure was detected of an already primed pump. | Incorrect Programing/Insufficient Power
Pump Flow Calculation – Output | Faulty Data (data is corrupt, invalid, incomplete or incorrect) | Motor Current for ECP pumps drifts over time causes Impella flow to be incorrect | Incorrect Programing/Insufficient Power
| False Positive Condition (SW detects the condition when there is none) | Maximum motor current deviation (Max%) calculated over the most recent motor current pulsatility monitoring period switches to detect suction instead of incorrect position (true underlying condition) due to noise. A suction alarm is triggered instead of an incorrect position alarm. | Incorrect Algorithm
| False Positive Condition (SW detects the condition when there is none) | unexpected motor current noise | Incorrect Repositioning
| False Positive Condition (SW detects the condition when there is none) | Maximum motor current deviation (Max%) calculated over the most recent motor current pulsatility monitoring period switches to detect incorrect position instead of suction (true underlying condition) due to noise. An incorrect position alarm is triggered instead of a suction alarm. | Incorrect Algorithm
| Incorrect Algorithms | Bi-directional communication on USB connection allows entry for cybersecurity attack. | N/A
| CPU/Memory Management (out of CPU or memory) | Software process consumes too much CPU/Memory or blocks critical processes from running | Console malfunction
| Faulty or Incompatible I/O | Failure in software of unknown providence (SOUP). | N/A
| Faulty Error Handling (wrong response when error is detected) | Optical bench (OPM) Controller Error alarms (e.g., #1039 for bad optical lamp) can be triggered when an Impella ECP pump is connected. Alarms are false because ECP does not have a pressure sensor. | ECP-HA-067
| Faulty or Incompatible I/O | After power cycle, EPM fails to establish communication with EPC | ECP-HA-066
| EEPROM Read Fails | EEPROM read fails due to corrupted message between EEPROM & EPM or between EPM & EPC | ECP-HA-066
| Faulty or Incompatible I/O | EEPROM read fails due to corrupted message between EEPROM & EPM or between EPM & EPC | Outflow in ventricle
| Faulty or Incompatible I/O | EEPROM read fails due to corrupted message between EEPROM & EPM or between EPM & EPC | Outflow in ventricle
Predictive algorithm calculates output | Faulty Software | Software malfunction or algorithm error. | Unsecured Device
Predictive algorithm calculates output | Faulty Software | Software malfunction or algorithm error. | Incorrect Repositioning
Receive and process data for the predictive algorithm | Faulty infrastructure | Infrastructure is unavailable Slow response time Data is lost or corrupted | Unsecured Device
Receive and process data for the predictive algorithm | Faulty infrastructure | Infrastructure is unavailable Slow response time Data is lost or corrupted | Incorrect Repositioning
Receive and process data for the predictive algorithm | Data communication error | Data is lost or corrupted | Unsecured Device
Receive and process data for the predictive algorithm | Data communication error | Incoming SaMD data communication error | Incorrect Repositioning
Receive and process data for the predictive algorithm | Software Of Unknown Provenance | SOUP in SaMD SOUP in infrastructure | Unsecured Device
Receive and process data for the predictive algorithm | Software Of Unknown Provenance | SOUP in SaMD SOUP in infrastructure | Incorrect Repositioning
Invalid Window - Mixed Source | Not displaying correct Prediction | Software malfunction | Incorrect Repositioning
Review predictive algorithm output | SaMD output (Use) | User misinterpretates SaMD output | Unsecured Device
Review predictive algorithm output | SaMD output (Use) | User misinterpretates SaMD output | Incorrect Repositioning

## dFMEA_Pool
input_statement | fm | Failure Mode | Hazard
Impeller. Propels the flow of blood. Impeller geometry (radius). | hemolysis | hemolysis | unexpected therapeutic results
Impeller. Interfaces with bearings. OD of Shaft. | Rotor Dynamic Issues | rotor dynamic issue | unintended system motion
Impeller. Interfaces with bearings. OD of Shaft. | Over constrained rotor leading to high motor current and/or shaft wear | rotor constrained | mechanical issue
Impeller. Interfaces with bearings. OD of Shaft. | Over constrained rotor leading to high motor current and/or shaft wear | shaft worn | device abrasion from instrument or another object
Impeller. Interfaces with bearings. Runout of Shaft. | Rotor Dynamic Issues | unitended rotor motion | unintended system motion
Impeller. Interfaces with bearings. Runout of Shaft. | Over constrained rotor leading to high motor current and/or shaft wear | high motor current | electrical issue
Impeller. Interfaces with bearings. DLC coating thickness. | Shaft wear/Rotor Dynamic Issues | shaft interfered with bearing | device or device component damaged by another device
Impeller. Interfaces with bearings. DLC coating thickness. | Interference with bearing causing damage to coating leading to shaft wear | coating damaged | device abrasion from instrument or another object
Impeller. Interfaces with bearings. DLC coating thickness. | Over constrained rotor leading to high motor current and/or shaft wear | rotor interfered with bearing | device abrasion from instrument or another object
Impeller. Interfaces with bearings. DLC coating thickness. | Over constrained rotor leading to high motor current and/or shaft wear | coating thickness out of specification | physical property issue
Impeller. Interfaces with bearings. DLC Coating adhesion strength. | Particulate generation blocking purge | particulate generated | particulates
Impeller. Interfaces with bearings. DLC Coating adhesion strength. | Shaft wear and unstable rotor dynamics | low coating adhesion strength | failure to adhere or bond
Impeller. Interfaces with bearings. Cross-sectional area of notches at end of shaft. | Rotor Dynamic Issues/Shaft wear | rotor dynamic issue | unintended system motion
Impeller. Interfaces with bearings. Cross-sectional area of notches at end of shaft. | (Poor journal bearing lubrication ) -Rotor Dynamic Issues | poor bearing lubrication | physical resistance
Impeller. Interfaces with bearings. Length. | Partial engagement leading to premature shaft/bearing wear | rotor partial engagement | disconnection
Impeller. Interfaces with bearings. Length. | Rotor Dynamic Issues (Collides with single pin) | rotor collided | unintended collision
Impeller. Interfaces with bearings. Length. | Purge blockage (Collides with single pin) | purge blockage | blockage within device or device component
Impeller. Aid in purge delivery. OD of Shaft/DLC Coating Thickness. | Too little purge, high purge pressure | insufficient purge | insufficient flow or underinfusion
Impeller. Aid in purge delivery. Cross-sectional area of notches at end of shaft. | Too little purge, high purge pressure | high purge pressure | increase in pressure
Impeller. Hold magnet and washer . OD of Shaft/DLC Coating Thickness. | washer and bearing wear | washer worn | device abrasion from instrument or another object
Impeller. Hold magnet and washer . OD of Shaft/DLC Coating Thickness. | Rotor Dynamic Issues | hold magnet worn | device abrasion from instrument or another object
Impeller. Hold magnet and washer . OD of Shaft/DLC Coating Thickness. | Glue migration to washer surface | glue migrated | migration of device or device component
Impeller. Define central axis of rotation and defines rotor foundation. Runout of Shaft. | Rotor Dynamic Issues | rotor central axis runout | unintended system motion
Impeller. Define central axis of rotation and defines rotor foundation. OD of Shaft. | Shaft wear | shaft worn | mechanical issue
Impeller. Define central axis of rotation and defines rotor foundation. DLC coating thickness. | Coating/Shaft Wear | coating worn | device abrasion from instrument or another object
Sensor. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | high readings
Sensor. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | inappropriate material used | material integrity issue
Sensor. Determining position of pump within the body. Depth of the cavity. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | sensor malfunctioned | device inoperable
Sensor. Determining position of pump within the body. Silicone gel thickness. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to determine device position | failure to sense
Sensor. Determining position of pump within the body. Silicone gel material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to detect device position | malposition of device
Sensor. Determining position of pump within the body. Membrane dimensions. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | inappropriate material used | installation-related problem
Sensor. Determining position of pump within the body. Sensor cable material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | sensor disconnected | loose or intermittent connection
Sensor. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect | incorrect sensor reading | low readings
Sensor. Biocompatible with body. Silicone gel material. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor biocompatibility compromised | biocompatibility issue
Sensor. Determining position of pump within the body. Diaphragm dimensions. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor reading drifted | incorrect measurement
Sensor. Determining position of pump within the body. Silicone gel adhesion (shear) strength. | Silicone gel displaces from sensor and cirulates through the body | gel displaced | dislodged or dislocated
Sensor. Determining position of pump within the body. Silicone gel adhesion (shear) strength. | Silicone gel displaces from sensor and cirulates through the body | low adhesion strength | failure to adhere or bond
Sensor. Measure diagnostic pressure. Sensor Material. | Readings from sensor incorrect and/or failure of the sensor to provide pressure reading | incorrect pressure reading | incorrect measurement
Sensor. Measure diagnostic pressure. Depth of the cavity. | Readings from sensor incorrect and/or failure of the sensor to provide pressure reading | fail to provide pressure reading | unable to obtain readings
Sensor. Measure diagnostic pressure. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | inappropriate material used | defective component
Sensor. Measure diagnostic pressure. Silicone gel material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | inappropriate material used | installation-related problem
Sensor. Measure diagnostic pressure. Membrane dimensions. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | inappropriate membrane used | incorrect measurement
Sensor. Measure diagnostic pressure. Sensor cable material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | sensor disconnected | unable to obtain readings
Sensor. Measure diagnostic pressure. Silicone gel material. Sensor drifts beyond specification (less accurate, issues with pressure reading) | Sensor drifts beyond specification (less accurate, issues with pressure reading) | sensor reading drifted | incorrect measurement
Sensor. Measure diagnostic pressure. Silicone gel material. Sensor drifts beyond specification (less accurate, issues with pressure reading) | Sensor drifts beyond specification (less accurate, issues with pressure reading) | sensor drifted | dislodged or dislocated
Pump Housing. Structural support for the Impella. Material. | Buckling/interface of Impeller to pump housing (pump stop) | housing buckled | buckled material
Pump Housing. Structural support for the Impella. Strut design features. | Buckling/interface of Impeller to pump housing (pump stop) | housing structure deformed | structural problem
Pump Housing. Allows blood to exit pump safely. Surface finish. | Damage to red blood cells | red blood cell damaged | unexpected therapeutic results
Pump Housing. Allows blood to exit pump safely. Surface finish. | Thrombus formation | thrombus formation | structural problem
Pump Housing. Allows blood to exit pump safely. Window design features. | Damage to red blood cells | blood blocked | blockage within device or device component
Pump Housing. Allows blood to exit pump safely. Window design features. | Thrombus formation | thrombus formation | unexpected therapeutic results
Pump Housing. Allows blood to exit pump safely. Window design features. | Reduced flow | blood flow reduced | insufficient flow or underinfusion
Pump Housing. Provides protected location for sensor to mount on. Sensor deflection bump/sensor recessed feature. | Sensor damage/sensor dislodgement | sensor deflected | dislodged or dislocated
Pump Housing. Provides protected location for sensor to mount on. Sensor deflection bump/sensor recessed feature. | Sensor damage/sensor dislodgement | sensor recessed | pitted
Pump Housing. Provides protected location for sensor to mount on. Sensor deflection bump/sensor recessed feature. | Sensor damage | sensor broken or damaged | break
Pump Housing. Provides window for guidewire translation. Window design features. | unable to track pump along guidewire | unable to track pump | difficult to position
Pump Housing. Provides window for guidewire translation. Window design features. | Skiving of guidewire (debry into blood path) | debris in blood path | particulates
Impeller. Interfaces with bearings. OD of Shaft. | unable to assemble | shaft out of dimension | physical property issue
Impeller. Interfaces with bearings. Runout of Shaft. | Unable to assemble | shaft runout | unintended system motion
Impeller. Interfaces with bearings. DLC coating thickness. | Unable to Assemble | unable to assemble | installation-related problem
Impeller. Interfaces with bearings. Length. | Does not engage leading bearings not being functional | fail to engage | installation-related problem
Impeller. Hold magnet and washer . OD of Shaft/DLC Coating Thickness. | Unable to assemble | unable to assemble | defective component
Impeller. Define central axis of rotation and defines rotor foundation. Runout of Shaft. | Unable to assemble | shaft runout | unintended system motion
Impeller. Define central axis of rotation and defines rotor foundation. OD of Shaft. | Unable to assemble | unable to assemble | disassembly
Impeller. Define central axis of rotation and defines rotor foundation. DLC coating thickness. | Unable to assemble | unable to assemble | defective item
Yoke. Provides material for electronmagnetic fields for motor operation. Cross sectional area of yoke. | RPM instability and/or loss of efficiency (reduced flow/support) | loss of efficiency | output issue
Yoke. Provides heat sink. Material. | Loss of efficiency (reduced flow/support) | electronmagnetic field reduced | output energy incorrect
Yoke. Provides structural support for pump/motor. Mechanical Properties (i.e. hardness). | Component breaks under stress | component broken under stress | break
Yoke. Provides welding junction for endcap and pump housing. Material. | Weld failure | weld broken | solder joint failure
Yoke. Provides welding junction for endcap and pump housing. Geometery. | Weld failure | incomplete welding | solder joint failure
Yoke. Provides structural support for pump/motor. Material. | Not biocompatible Leachable compound on material | non-biocompatible compound leached | biocompatibility issue
Yoke. Provides structural support for pump/motor. Material. | Not biocompatible Leachable compound on material | fail to provide structural support | structural problem
Single Pin. Provides independent connection points for (Purge and Placement lumen and nitnol). Slot Location. | Purge lumen fracture leading to purge fluid leakage | purge lumen fractured | fracture
Single Pin. Provides independent connection points for (Purge and Placement lumen and nitnol). Slot Location. | Purge lumen fracture leading to purge fluid leakage | purge fluid leakage | fluid leak
Single Pin. Creates hermetically seal between purge lumen and motor. OD of the single pin head. | Damage the optical lumen and fiber resulting in signal issues | failure to read input signal | failure to read input signal
Single Pin. Creates hermetically seal between purge lumen and motor. Length of single pin stem. | Rotor dynamic issues due to increased engagement of the shaft unto bearing | rotor dynamic issue | unintended movement
Single Pin. Creates hermetically seal between purge lumen and motor. Length of single pin stem. | Unable to deliver purge fluid to motor | hermetic seal damaged | fluid leak
Single Pin. Creates hermetically seal between purge lumen and motor. Length of single pin stem. | Unable to deliver purge fluid to motor | unable to deliver purge fluid | no flow
Single Pin. Provides strain relief and torque support for all lumen connections. Overall length. | Purge leakage/Kink | fail to provide strain relief | mechanical issue
Single Pin. Provides strain relief and torque support for all lumen connections. Overall length. | Purge leakage/Kink | fail to provide torque support | mechanical issue
Single Pin. Provides strain relief and torque support for all lumen connections. Overall length. | Purge leakage/Kink | lumen connection leakage | fluid leak
End Cap. Connection point of the motor to catheter. Large OD of Endcap (motor side). Poor interaction between endcap and motor housing leading to weak bond (unable to weld) | Poor interaction between endcap and motor housing leading to weak bond (unable to weld) | poor interaction | solder joint failure
End Cap. Connection point of the motor to catheter. Large OD of Endcap (motor side). Poor interaction between endcap and motor housing leading to weak bond (unable to weld) | Poor interaction between endcap and motor housing leading to weak bond (unable to weld) | unable to weld | solder joint failure
End Cap. Connection point of the motor to catheter. Large OD of Endcap (motor side). High interaction forces with introducer which leads to being to unable to push through introducer tip | High interaction forces with introducer which leads to being to unable to push through introducer tip | high interaction force applied | manufacturing or shipping issue associated with device
End Cap. Connection point of the motor to catheter. Large OD of Endcap (motor side). High force pulling from valve causing damage to internal mechanisms (pump failure) | High force pulling from valve causing damage to internal mechanisms (pump failure) | high pull force applied | installation-related problem
End Cap. Connection point of the motor to catheter. Large OD of Endcap (motor side). High force pulling from valve causing damage to internal mechanisms (pump failure) | High force pulling from valve causing damage to internal mechanisms (pump failure) | internal mechanism damaged | solder joint failure
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Unable to complete reflow of catheter material which causing low joint strength between catheter and endcap. | Unable to complete reflow of catheter material which causing low joint strength between catheter and endcap | low joint strength | solder joint failure
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Unable to complete reflow of catheter material which causing low joint strength between catheter and endcap. | Unable to complete reflow of catheter material which causing low joint strength between catheter and endcap | incomplete material reflow | manufacturing or shipping issue associated with device
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Leaking between joint | Leaking between joint | joint leakage | fluid leak
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Endcap deforms during off axis insertion (loss of function, pushability) | Endcap deforms during off axis insertion (loss of function, pushability) | endcap deformed | material distortion
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Endcap deforms during off axis insertion (loss of function, pushability) | Endcap deforms during off axis insertion (loss of function, pushability) | loss of pushability | difficult to insert
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Endcap deforms during off axis insertion (block of purge) | Endcap deforms during off axis insertion (block of purge) | purge fluid blocked | no flow
End Cap. Connection point of the motor to catheter. Small OD of Endcap (catheter). Endcap deforms during off axis insertion (reduced bond joint strength) | Endcap deforms during off axis insertion (reduced bond joint strength) | joint strength reduced | solder joint failure
End Cap. Connection point of the motor to catheter. Surface Finish (Barbed Feature). | Reduced bond strength | bond strength reduced | failure to adhere or bond
End Cap. Connection point of the motor to catheter. Height (Barbed Feature). Unable to reflow (reduced strength of bond) | Unable to reflow (reduced strength of bond) | unable to reflow | solder joint failure
End Cap. Connection point of the motor to catheter. Height (Barbed Feature). | Reduced bond strength | bond strength reduced | mechanical issue
End Cap. Connection point of the motor to catheter. Spacing between barbs. Unable to reflow (reduced strength of bond) | Unable to reflow (reduced strength of bond) | unable to reflow | manufacturing or shipping issue associated with device
End Cap. Connection point of the motor to catheter. Number of barbs. | Reduced bond strength | bond strength reduced | solder joint failure
End Cap. Connection point of the motor to catheter. Anti-rotation slots. | Loss of toqueability (damage to functionality, can not insert pump) | loss of toqueability | mechanical issue
End Cap. Connection point of the motor to catheter. Anti-rotation slots. Loss of functionality (twisting of internal lumens causing reduced purge flow) | Loss of functionality (twisting of internal lumens causing reduced purge flow) | internal lumen twisted | material twisted
End Cap. Connection point of the motor to catheter. Anti-rotation slots. | Reduced bond strength | loss of anti-rotation mechanism | unintended system motion
End Cap. Provides recess for fiber exiting catheter. End Cap IDs. | Endcap deforms during off axis insertion (loss of function, pushability) | endcap deformed | material distortion
End Cap. Provides recess for fiber exiting catheter. End Cap IDs. Endcap deforms during off axis insertion (block of purge) | Endcap deforms during off axis insertion (block of purge) | endcap deformed | no flow
End Cap. Provides recess for fiber exiting catheter. End Cap IDs. | Endcap deforms during off axis insertion (reduced bond joint stregnth) | endcap deformed | loss of or failure to bond
End Cap. Provides recess for fiber exiting catheter. End Cap IDs. | Unable to assembly | unable to assemble | installation-related problem
Co-extruded catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. | reduced deliverability of device | device deliverability reduced | failure to deliver
Co-extruded catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. | Kink during device deliverability | device kinked | delivery system failure
Co-extruded catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material thickness. | Kink of catheter | catheter kinked | delivery system failure
Co-extruded catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material thickness. | Tear/Puncture | fail to provide wire protection | break
Co-extruded catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Kink of catheter | fail to provide lumen protection | kinked
Co-extruded catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Tear/Puncture | fail to provide wire protection | torn material
Co-extruded catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Intensity of marking. | Unable to read marking | device torqueability reduced | mechanical issue
Co-extruded catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Intensity of marking. | Unable to read marking | unable to read marking | device markings issue
Co-extruded catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Suface finish. | Thrombus formation | thrombus formation | adverse event without identified device or use problem; structural problem
Strain Relief. Relief strain of cable, catheter, and pressure storage set assembly. Material (Durometer). | kink leading to pump stop | catheter kinked | device inoperable
Strain Relief. Relief strain of cable, catheter, and pressure storage set assembly. Material (Durometer). | kink leading to placement signal failure | cable kinked | failure to read input signal
Strain Relief. Relief strain of cable, catheter, and pressure storage set assembly. Material (Durometer). | kink leading to deliverability issues | pressure storage set kinked | delivery system failure
Strain Relief. Relief strain of cable, catheter, and pressure storage set assembly. ID. | unable to assemble | unable to assemble | installation-related problem
Strain Relief. Relief strain of cable, catheter, and pressure storage set assembly. ID. | Insufficient engagement for bonding leading to weak bond | insufficient bond engagement | failure to adhere or bond
Strain Relief. Mechanically connect cable, catheter, and pressure storage set. Interlock geometry. | dislodgement of connection | interlock geometry failed | dislodged or dislocated
PCA Board. Creates electrical path from the patient cable through catheter wires to motor. Trace Width and Thickness. | Over current condition/excessive heating leading to board failure due to insufficient conductivity | insufficient conductivity | failure to conduct
PCA Board. Creates electrical path from the patient cable through catheter wires to motor. Solder pads. | Excessive heating during assembly leading to board failure (i.e. pad delamination) | excessive heating | overheating of device or device component
PCA Board. Creates electrical path from the patient cable through catheter wires to motor. Trace spacing. Interference leakage current which leads to motor phase short. | Interference leakage current which leads to motor phase short | over leakage current | electrical issue
PCA Board. Heat generated should not exceed 40 C. Temperature. | Component fails | board burned | overheating of device or device component
PCA Board. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Width. | Over current condition/excessive heating leading to board failure due to insufficient conductivity | over current condition | electrical overstress
PCA Board. Provides motor, EEPROM communication and sensor access paths. Solder Pads. Excessive heating during assembly leading to board failure (i.e. pad delamination) | Excessive heating during assembly leading to board failure (i.e. pad delamination) | pad delaminated | delamination
PCA Board. Provides path and communication for EEPROM. Component Ratings in Circuit (Capacitor Voltage Value and Rating Resistor Value and Power Dissipation Semi-Conductor Component Voltage Rating Semi-Conductor Component Power Dissipation). | Component fails and circuit not functioning as intended | faulty component | electrical issue
PCA Board. Provides path and communication for EEPROM. Power Consumption. Component fails and could lead to AIC failure and affect patient sensor accuracy | Component fails and could lead to AIC failure and affect patient sensor accuracy | component damaged | electrical issue
PCA Board. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Type. | Loss/ Corruption of Data during use | component damaged | loss of data
PCA Board. Creates electrical path from the patient cable to pressure sensor. Trace Width. | Over current condition/excessive heating leading to board failure due to insufficient conductivity | trace width short | electrical overstress
PCA Board. Creates electrical path from the patient cable to pressure sensor. Trace Clearance. Interference leakage current which leads to sensor wire getting damaged | Interference leakage current which leads to sensor wire getting damaged | over leakage current | grounding malfunction
PCA Board. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Clearance. | Interference which leads EEPROM communication failure | trace clearance short | electrical shorting
PCA Board. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Size. | Unable to provide operational support | fail to retain data | loss of data
PCA Board. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Clearance. | Interference which leads to EEPROM communication failure | loss of communication | communication or transmission issue
PCA Board. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Type. | Loss/ Corruption of Data during storage | data corrupted | loss of data
PCA Board. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Mousebite Location. | Board and/or trace damage | trace damaged | circuit failure
PCA Board. Houses electrical components (i.e. resistors, capacitors, ICs). PCA Board and EEPROM. | Unable to program EEPROM | unable to program EEPROM | programming issue
PCA Board. Houses electrical components (i.e. resistors, capacitors, ICs). PCA Board and EEPROM. Radiated or conducted emissions from microcontroller in pump plug. | Radiated or conducted emissions from microcontroller in pump plug. | microcontroller emission | radiofrequency interference (RFI)
PCA Board. Houses electrical components (i.e. resistors, capacitors, ICs). PCA Board and EEPROM. Loss of communication between AIC and EEPROM during support due to electromagnetic interference including RFID, electrocautery. | Loss of communication between AIC and EEPROM during support due to electromagnetic interference including RFID, electrocautery. | board interfered with electromagnetic | electromagnetic interference (EMI)
Single Pin. Provides independent connection points for (Purge and Placement lumen and nitnol). Slot Diameter. | Unable to assemble | unable to assemble pin | installation-related problem
Single Pin. Creates hermetically seal between purge lumen and motor. OD of the single pin head. | Unable to assemble | pin head out of specification | defective component
Single Pin. Creates hermetically seal between purge lumen and motor. Length of single pin stem. | Unable to assemble end cap | pin length out of specification | defective component
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Crack/damaged leading to pump stop | handle halves broken or damaged | break
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Poor bonding surface leading to internal components being exposed | poor bonding surface | failure to adhere or bond
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Poor bonding surface leading to internal components being exposed | internal component exposed | electrical shorting
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Warpage of material leading to unable to assemble | material warpage | material distortion
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Internal dimensional features. Warpage of material leading to unable to assemble | Warpage of material leading to unable to assemble | unable to assemble | material integrity issue
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Internal dimensional features. | Unable to assemble | insecure housing | unstable; loose or intermittent connection; protective measure issue
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Internal dimensional features. Damage to internal components leading to pump start issues or stoppage | Damage to internal components leading to pump start issues or stoppage | internal component damaged | break
Upper and Lower Handle halfs. Provides handle for user to manipulate device. External geometry. Unable to handle (difficulty with insertion of device) | Unable to handle (difficulty with insertion of device) | unable to handle | difficult to insert
Upper and Lower Handle halfs. Provides handle for user to manipulate device. External geometry. | Unable to place device effectively | unable to place device | device handling issue
Upper and Lower Handle halfs. Provides handle for user to manipulate device. External geometry. | Damage to gloves or clinicians hand | sharp surface | material puncture
Upper and Lower Handle halfs. Provides location for device identification (Product ID and Serialization). Material. | Unable to read identification | unable to read identification | device markings issue
Upper and Lower Handle halfs. Provides location for device identification (Product ID and Serialization). Material. | Unable to read identification | unable to read serialization | device markings issue
Upper and Lower Handle halfs. Provides location for device identification (Product ID and Serialization). Material. | Unable to print/laser information on pump | unable to print info on handle | device markings issue
Upper and Lower Handle halfs. Provides connection point for purge side arm, cable, catheter. Interlock geometry. | dislodgement of connection | connection dislodged | dislodged or dislocated
End Cap. Connection point of the motor to catheter. Number of barbs. | Reduced manufacturability | difficult to manufacture | manufacturing or shipping issue associated with device
Patient Cable. Provides optical connection from console to handle. DB loss < 0.5. Poor placement signal/Loss of placement signal | Poor placement signal/Loss of placement signal | loss of signal | failure to read input signal
Patient Cable. Provides optical connection from console to handle. DB loss < 0.5. Poor placement signal/Loss of placement signal | Poor placement signal/Loss of placement signal | fail to connect | connection issue
Patient Cable. Provides optical connection from console to handle. Alignment of ferrule in the pump plug. | Poor placement signa/Loss of placeent Signall | connection misaligned | failure to align
Patient Cable. Provides optical connection from console to handle. Height of ferrule in the pump plug. | Poor placement signa/Loss of placeent Signall | intermittent connection | loose or intermittent connection
Patient Cable. Provides optical connection from console to handle. Geometery of end face of ferrule. | Poor placement signa/Loss of placeent Signall | poor connection | intermittent continuity
Patient Cable. Provides optical connection from console to handle. Fiber configuration. | Poor placement signa/Loss of placeent Signall | incorrect fiber configuration | use of incorrect control settings
Patient Cable. Provides electrical signals from console to handle. Wire Gauge. | Intermittent or loss of motor power | intermittent connection | loss of power
Patient Cable. Provides electrical signals from console to handle. Wire Resistance. | Intermittent or loss of motor power | intermittent connection | failure to read input signal
Patient Cable. Provides electrical signals from console to handle. Wire Resistance. | Intermittent or loss of motor power | high wire resistance | physical resistance
Patient Cable. Provides transfer of data from console to handle. Circuit Schematic. | Pump unable to proceed through case start or performs poorly | fail to transfer data | failure to service
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Low process capability (low yeild) | fail to provide internal lumen | kinked
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Low process capability (low yeild) | fail to provide internal wire | break
Pump Tray. Properly house pump during transport. Tray Cavity (LxWxH). | Pump unable to fit | insufficient tray cavity | inadequate storage
Pump Tray. Properly house pump during transport. Tray Cavity (LxWxH). Unable to seal properly leading to unit not being sterilized | Unable to seal properly leading to unit not being sterilized | incorrect seal | delivered as unsterile product
Pump Tray. Properly house pump during transport. Tray Cavity (LxWxH). | Unable to package product | unable to package | packaging issue
Pump Tray. Properly house pump during transport. Tray Cavity (LxWxH). Unable to kit components properly | Unable to kit components properly | unable to kit component | packaging issue
Pump Tray. Properly house pump during transport. Pump handle cavity (features in tray). Handle won't be secured and could breach sterile barrier | Handle won't be secured and could breach sterile barrier | sterile barrier breached | delivered as unsterile product
Pump Tray. Properly house pump during transport. Pump handle cavity (features in tray). | Unable to assemble packaging | unable to package | device packaging compromised
Pump Tray. Properly house pump during transport. Pump handle cavity (features in tray). User inconvience when removing the pump | User inconvience when removing the pump | unable to remove package | difficult to open or remove packaging material
Pump Tray. Properly house pump during transport. Patient Cable cavity and wrapping feature. | Unable to assemble packaging | fail to package | packaging issue
Pump Tray. Properly house pump during transport. Patient Cable cavity and wrapping feature. Patient cable head won't fit (unable to assmebly packaging) | Patient cable head won't fit (unable to assmebly packaging) | inappropriate fit | fitting problem
Pump Tray. Properly house pump during transport. Patient Cable cavity and wrapping feature. Unsecure patient cable head causing breach in sterile barrier | Unsecure patient cable head causing breach in sterile barrier | unable to secure part | migration of device or device component
Pump Tray. Properly house pump during transport. Rail for purge side arm clip. | Uncable to secure purge side arm leading to kink in purge line. | unable to secure part | unstable
Pump Tray. Provides container for pump for sterilization process. Tray Cavity (LxWxH). | Impact ability to sterilize product completely | incomplete sterilization | delivered as unsterile product
Pump Tray. Provides container for pump for sterilization process. Seal parameter. | Low seal strength | low seal strength | unsealed device packaging
Cannula Inserts. Provides protection of motor and cannula during shipment. Inserts (LxWxH). | Insert will not fit tray | inappropriate fit | fitting problem
Cannula Inserts. Provides protection of motor and cannula during shipment. Inserts (LxWxH). | Unable to seal tyvek properly to tray | unable to seal properly | failure to adhere or bond
Cannula Inserts. Provides protection of motor and cannula during shipment. Inserts (LxWxH). | Leading damage to pigtail and/or cannula | part damaged | shipping damage or problem
Cannula Inserts. Provides protection of motor and cannula during shipment. Inserts (LxWxH). | damage cannula | part damaged | packaging issue
Cannula Inserts. Provides protection of motor and cannula during shipment. Cradle geometery. | Cannula damaged (crushed) | fail to provide protection | protective measure issue
Tyvek Lid. Provides sterilization path. Thickness. | Poor flow of sterilization gas into tray | sterilization path opened | delivered as unsterile product
Tyvek Lid. Provides sterilization path. Thickness. Poor flow of sterilization gas into tray | Poor flow of sterilization gas into tray | inadequate sterilization gas flowed into tray | failure to sterilize
Tyvek Lid. Provides sterilization path. Thickness. Easly punctured/damage causing breach in sterile barrier | Easly punctured/damage causing breach in sterile barrier | sterilization path breached | failure to sterilize
Tyvek Lid. Provides sterilization path. Material. | Poor flow of sterilization gas into tray | insufficient sterilization gas flowed into tray | failure to sterilize
Tyvek Lid. Provides sterile barrier. adhesive type. Unable to seal properly leading to unit not being sterilized | Unable to seal properly leading to unit not being sterilized | incorrect adhesive type | failure to sterilize
Tyvek Lid. Provides containment for entire device. Lid (LxW). Unable to seal properly leading to unit not being sterilized | Unable to seal properly leading to unit not being sterilized | unable to seal properly | manufacturing or shipping issue associated with device
Loctite. Bonding handles together. Viscosity. | Lower bond strength | low bond strength | failure to adhere or bond
Loctite. Bonding handles together. Viscosity. | Difficult to assemble | inappropriate viscosity | loss of or failure to bond
Loctite. Bonding handles together. Bond/Retention Strength. | Bond Failure | inappropriate loctite applied | installation-related problem
Loctite. Bonding handles together. Cure Time. | Bond Failure | inappropriate cure time | installation-related problem
Loctite. Bonding handles together. Cure Time. | Difficult to assemble | inappropriate cure time | failure to adhere or bond
PCA Board. Provides structural support surface ferral clip and fiber guides. Thickness of baord. | Difficult to assemble | difficult to assemble | installation-related problem
PCA Board. Provides structural support surface ferral clip and fiber guides. Thickness of baord. | Difficult to assemble | board thickness out of specification | physical property issue
PCA Board. Houses electrical components (i.e. resistors, capacitors, ICs). Board layout. | Board failure/temporary shut down and restart of board (error:115) | inappropriate board layout | electrical shorting
Single Pin Subassembly. Provides reinforcement of deliverability. Placement of nitonol wire. | Unable deliver | unable to deliver | failure to deliver
Single Pin Subassembly. Provides purge fluid to motor. Connection of the purge lumen. | Unable to deliver purge fluid to motor | unable to deliver purge fluid | failure to deliver
Single Pin Subassembly. Provides purge fluid to motor. Connection of the purge lumen. | Leaking between joint | joint leakage | fluid leak
Single Pin Subassembly. Provides fixation for the optical lumen. Location of the optical lumen. | Unable assemble | unable to fix | installation-related problem
Single Pin Subassembly. Protects both PCBs from high pressure purge fluid. Interference length of single assembly. | PCB board failure | board damaged | protective measure issue
Tipped Catheter. Provides connection for the pump to the pump handle. Overall Length. | Unable to assemble | unable to connect | component or accessory incompatibility
Tipped Catheter. Provides deliverability for the pump into patient and long term use. Material (Rigidity). | Unabe deliver unit | unable to deliver | no flow
Pump-Pump Handle AU. Finished pump unit which provides flow of blood and pressure sensing (without purge). Ferrule connection of the sensor. | Sensor Failure | sensor disconnected | unable to obtain readings
Pump-Pump Handle AU. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soder locations. | Failure to start pump | fail to start pump | fail to pump
Pump-Pump Handle AU. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. | Bad connections to soldering board leading to pump stop/failure to start pump | no delivery | pumping stopped
Pump-Pump Handle AU. Finished pump unit which provides flow of blood and pressure sensing (without purge). Specification of bin and flow (EPROM). | Incorrect flow, reduced support | reduced flow | insufficient flow or underinfusion
Pump-Pump Handle AU. Provide orientation to deliver pump into body. 90 orientation of port arm to cannula bend. | Unable to assmeble | unable to assemble | incompatibility problem
Pump-Pump Handle AU. Provide orientation to deliver pump into body. 90 orientation of port arm to cannula bend. | Increased suction/low flow | low flow | insufficient flow or underinfusion
Sensor to Motor Assembly. Provides pump placement information. Sensor location on pump housing. | Incorrect placement signal | incorrect sensor placement | installation-related problem
Sensor to Motor Assembly. Provides pump placement information. Delo location (application). Increased insertion force causing difficulty to insert | Increased insertion force causing difficulty to insert | excessive insert force | difficult to insert
Sensor to Motor Assembly. Provides pump placement information. Wash hole on pump housing. | Inaccurate placement signal readings | inaccurate sensor reading | incorrect measurement
Sensor to Motor Assembly. Measure diagnostic pressure. Sensor location on pump housing. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | incorrect measurement
Sensor to Motor Assembly. Measure diagnostic pressure. Delo location (application). | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect pressure reading | incorrect measurement
Sensor to Motor Assembly. Measure diagnostic pressure. Wash hole on pump housing. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to provide pressure reading | unable to obtain readings
Motor Housing Assembly. Provide power to the Impella. Coil location relative to the yoke. | Reduced power/flow | reduced power | power source issue
Motor Housing Assembly. Provides interface for Rotor Assembly. Overall Length. Unable to engage into ball bearing | Unable to engage into ball bearing | unable to engage | connection issue
Motor Housing Assembly. Provides interface for Rotor Assembly. Overall Length. Weak bond to ball bearing | Weak bond to ball bearing | weak bond | failure to adhere or bond
Motor Housing Assembly. Provides interface for pump housing. Overall Length. Unable to weld pump housing | Unable to weld pump housing | unable to weld | solder joint failure
Motor Housing Assembly. Provides interface for pump housing. Overall Length. Weak bond of pump housing | Weak bond of pump housing | unable to connect | component incompatible
Motor Housing Assembly. Providing interface for End cap. OD of Endcap Interface. | Unable to assemble | unable to assemble | component incompatible
Motor Housing Assembly. Provides interface for single pin. Small OD on Endcap interface. | Unable to assemble | unable to connect | component incompatible
Motor Housing Assembly. Provides interface for single pin. Length on Endcap interface. | Weak joint, leak | weak interface strength | loss of or failure to bond
Motor Housing Assembly. Provides interface for single pin. Slot alignment between single pin yoke. | Unable to assemble | part misaligned | installation-related problem
Pump Handle Assembly. Provides product identification for pump. location of serial number logo on pump shell. | unable to determine pump type | unable to determine pump type | device markings issue
Pump Handle Assembly. Connects the catheter to patient cable and pressure side arm. Strain relief arrangement. | Unable assemble | unable to assemble | installation-related problem
Pump Handle Assembly. Connects the catheter to patient cable and pressure side arm. Alignment of cannula to handle. | Unable to assemble (packaging) | unable to assemble | failure to align
Pump Handle Assembly. Connects the catheter to patient cable and pressure side arm. Placement of adhesive. | Bond Failure, lead to electrical failure | fail to bond | failure to adhere or bond
Pump Handle Assembly. Contains electrical, fiber optic, and fluid connections. routing of fiber (> 5mm radius). | loss of signal/light | inappropriate fiberoptic routing | installation-related problem
Pump Handle Assembly. Contains electrical, fiber optic, and fluid connections. Shell internal layout. | Damage the fiber, loss of signal/light | fiber broken or damaged | disconnection
Motor Pump Assembly. Provides flow for the pump . Gap bewteen Impeller and sleeve bearing. | High purge pressure (Purge blockage) | high purge pressure | increase in pressure
Motor Pump Assembly. Provides flow for the pump . Gap bewteen Impeller and sleeve bearing. | Poor engagement of rotor into bearing (bond failure) | poor engagement | loss of or failure to bond
Motor Pump Assembly. Provides flow for the pump . Delo location (application). Failure of shaft to rotate | Failure of shaft to rotate | unable to rotate shaft | device inoperable
Cannula Inserts. Provides protection of motor and cannula during shipment. Cradle geometery. | Cannula will not fit in package | inappropriate fit | fitting problem
Catheter to Motor Housing Subassembly. Assembling purge system to motor housing. Delo location (application). | Bond Failure, purge leakage | purge leakage | fluid leak
Catheter to Motor Housing Subassembly. Assembling purge system to motor housing. Fiber lumen alignment with motor housing. | unable to assemble | unable to assemble | failure to align
Cannula to Pump Assembly . Placement in the heart and valve. Length from pigtail to endcap. | Unable to sit appropriately in the heart, low flow | fail to place catheter correctly | malposition of device
Cannula to Pump Assembly . Placement in the heart and valve. Length from pigtail to endcap. | Unable to sit appropriately in the heart, hemolysis | incorrect pigtail length | physical property issue
Cannula to Pump Assembly . Placement in the heart and valve. Max OD of Assembly. | Difficulty inserting/withdrawing device | unable to place | difficult to insert
Cannula to Pump Assembly . Placement in the heart and valve. Delo locations/Transitions. | Bond Failures | bond failure | failure to adhere or bond
Cannula to Pump Assembly . Placement in the heart and valve. Alignment of Pigtail to catheter. | Unable to sit appropriately in the heart, low flow | incorrect placement | malposition of device
Cannula to Pump Assembly . Placement in the heart and valve. Alignment of Pigtail to catheter. | Unable to sit appropriately in the heart, hemolysis | implant placed at improper location | installation-related problem
Pump Housing to Motor Assembly. To provide flow. Positioning of Impeller to Pump housing. | hemolysis | rough surface | unexpected therapeutic results
Pump Housing to Motor Assembly. To provide flow. Positioning of Impeller to Pump housing. | Pump stop due to contact between rotor and pump housing | pumping stopped | pumping stopped
Pump Housing to Motor Assembly. To provide flow. Positioning of Impeller to Pump housing. | Low flow | low blood flow | no flow
Pump Housing to Motor Assembly. Creating geometry for the sensor. Alignment of the fiber channel through the assembly. | Unable to assembly | fail to create geometry | physical property issue
Pump Housing to Motor Assembly. Provide mechanical integrity for motor assembly. Weld locations. | Failure of weld at low loads | weld in wrong location | solder joint failure
Pump Housing to Motor Assembly. Provide mechanical integrity for motor assembly. Depth Integrity. | Failure of weld at low loads | fail to provide mechanical integrity | material integrity issue
Magnet. Provide power to the Impella. Magnet Coating. | Early degredation of magnet leading to loss of efficiency/pump stop | magnet degraded | pumping stopped
Magnet. Provide power to the Impella. Magnet Coating. | Swelling of magnet | magnet swelled | degraded
Magnet. Provide power to the Impella. Magnet Coating. | Early degredation of magnet leading to loss of efficiency/pump stop | incorrect magnetic coating | unintended magnet quench
Magnet. Provide power to the Impella. Magnetic moment. | Insufficient power supplied to rotate impeller | insufficient magnetic field generated | loss of power
Magnet. Provide power to the Impella. Magnetic moment. | Excessive power supplied to rotate Impella | excess magnetic field generated | power conditioning issue
Magnet. Provide power to the Impella. Magnet material. | Insufficient power supplied to rotate impeller | fail to provide power | pumping stopped
Magnet. Provide power to the Impella. Coating Material Magnet Material. | Not biocompatible | incorrect coating material | biocompatibility issue
Magnet. Provide power to the Impella. Coating Material. Contamination, Particle in Coating . | Contamination, Partikel in Coating | particulate in coating | device contamination with biological material
Magnet. Provide power to the Impella. Coating Material . | Early degredation of magnet leading to loss of efficiency/pump stop | inappropriate coating material | degraded
Magnet. Provide power to the Impella. Bonding with Impeller. Early degradation of AU Rotor leading to loss of efficiency/pump stop | Early degredation of AU Rotor leading to loss of efficiency/pump stop | rotor degraded | pumping stopped
Fiber lumen. Provides a protective jacket over fiber from handle to motor. OD. | Unable to assemble | fail to provide protection | protective measure issue
Butterfly. Secures catheter from translating within the tray. Channel for catheter. | catheter unsecure leading damage to pigtail and/or cannula | catheter unsecured | device or device component damaged by another device
Butterfly. Secures catheter from translating within the tray. Clipping mechanism features. catheter unsecure leading damage to pigtail and/or cannula | catheter unsecure leading damage to pigtail and/or cannula | catheter unsecured | unstable
Butterfly. Secures catheter from translating within the tray. Clipping mechanism features. catheter unsecure leading damage to pigtail and/or cannula | User inconvience when removing the pump | cannula broken or damaged | device or device component damaged by another device
Nitinol Wire. Provides reinforcement of catheter for deliverability. Material. | Unable to insert device successfully | fail to provide reinforcement | physical property issue
Nitinol Wire. Provides reinforcement of catheter for deliverability. Length. | Unable to insert device successfully | unable to insert device | difficult to insert
Nitinol Wire. Increased Kink resistance. Material. | Kinkage of purge line | fail to resist kink | kinked
Cannula. Transportation blood from inlet to outlet. OD . | Unable to insert device successfully | unable to insert cannula | difficult to insert
Cannula. Transportation blood from inlet to outlet. ID. | Restriction of flow (low flow condition) | low flow condition | insufficient flow or underinfusion
Cannula. Transportation blood from inlet to outlet. Surface Finish of ID. | Damage to blood cells | rough cannula surface | unexpected therapeutic results
Cannula. Transportation blood from inlet to outlet. Surface Finish of ID. | Damage to blood cells | blood cell damaged | unexpected therapeutic results
Cannula. Provides support for positioning for inlet and outlet of device. Length. | Inlet/outlet in incorrect position, no flow | improper device position | malposition of device
Cannula. Provides support for positioning for inlet and outlet of device. Length. | Damage ventricle (perforation) | ventricle perforated | use of device issue
Cannula. Provides support for positioning for inlet and outlet of device. Angle. | Increased suction/low flow | suction increased | increase in suction
Cannula. Provides support for positioning for inlet and outlet of device. Surface finish of OD. | Clot formation | clot formated | infusion or flow issue
Cannula. Provides support for positioning for inlet and outlet of device. Marker Band. | incorrect positioning of device (reduced support) | arotic valve perforated | device handling issue
Cannula. Provides support for positioning for inlet and outlet of device. Marker Band OD. | Damage to arotic valve/vasculature/ventricle wall | vasculature perforated | use of device issue
Cannula. Provides support for insertion. Marker Band OD. | Unable to see band under imaging causing incorrect placement issues | unable to see band | poor quality image
Cannula. Provides support for insertion. Marker Band. | Unable to see band under imaging causing incorrect placement issues | device misplaced | malposition of device
Cannula. Provides support for insertion. Marker Band OD. Unable to successfully deliver device through introducer. | Unable to successfully deliver device through introducer | unable to insert device through introducer | difficult to insert
Cannula. Provides support for insertion. ID. | Insufficient stiffeness for use with guidewire, poor insertion | cannula stiffed | material rigid or stiff
Cannula. Provides support for insertion. OD. | Insufficient stiffeness for use with guidewire, poor insertion | cannula stiffed | difficult to insert
Guidewire Repositioning Unit. Provides hemostatis of arteriotomy. OD of GWRU. | Excessive bleeding at site | excessive bleeding | bleed back
Guidewire Repositioning Unit. Allows for repositioning of device. ID of GWRU. Increased interaction forces between catheter causing difficulty for repositioning | Increased interaction forces between catheter causing difficulty for repositioning | excessive friction | device abrasion from instrument or another object
Guidewire Repositioning Unit. Allows for repositioning of device. Surface Finish. Increased interaction forces between catheter causing difficulty for repositioning | Increased interaction forces between catheter causing difficulty for repositioning | rough surface | difficult to insert
Guidewire Repositioning Unit. Allows for repositioning of device. Surface Finish. Increased interaction forces between the GWRU and tissue causing damage | Increased interaction forces between the GWRU and tissue causing damage | vasculature perforated | physical property issue
Guidewire Repositioning Unit. Allows for repositioning of device. Hub Internal Geometery. Increased pressure in valve leading to increased interaction forces between catheter causing difficulty for repositioning. | Increased pressure in valve leading to increased interaction forces between catheter causing difficulty for repositioning | valve pressure increased | difficult to insert
Guidewire Repositioning Unit. Maintains pump position. Butterfly. | Device movement impacting positioning | pump moved | unintended movement
Guidewire Repositioning Unit. Maintains pump position. Butterfly. | Unable to secure device into position | unable to secure position | unintended movement
Guidewire Repositioning Unit. Maintains pump position. Tuohy Valve. | Unable to secure device into position | fail to maintain pump position | positioning issue
Guidewire Repositioning Unit. Maintains pump position. Tuohy Valve. Unable to join components and secure device into position. | Unable to join components and secure device into position | fail to secure pump position | positioning issue
Guidewire Repositioning Unit. Maintains sterility of catheter. Material thickness. tear and/or break during use causing sterility to be compromised. | tear and/or break during use causing sterility to be compromised | fail to maintain sterility | disinfection or sterilization issue at user location
Guidewire Repositioning Unit. Maintains sterility of catheter. Valve sleeve Tuohy. Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | fail to secure sterile sleeve | disinfection or sterilization issue at user location
Guidewire Repositioning Unit. Provides hemostatis of arteriotomy. ID of GWRU. Damage to arteriotomy/tissue tract. | Damage to arteriotomy/tissue tract | tissue tract damaged | device handling issue
Guidewire Repositioning Unit. Provides hemostatis of arteriotomy. ID of GWRU. Damage to arteriotomy/tissue tract. | Damage to arteriotomy/tissue tract | arteriotomy damaged | device handling issue
Guidewire Repositioning Unit. Provides hemostatis of arteriotomy. Material Stiffness. | fail to provide homeostasis | fail to provide homeostasis | bleed back
Guidewire Repositioning Unit. Provides hemostatis of catheter. Valve ID. | excessive bleeding | excessive bleeding | bleed back
Guidewire Repositioning Unit. Provides hemostatis of catheter. Hub Internal Geometery. Excessive bleeding between the valve and catheter | Excessive bleeding between the valve and catheter | excessive bleeding | physical property issue
Guidewire Repositioning Unit. Allows distal profusion of leg. OD . | Leg Ischemia | excessive bleeding | unexpected therapeutic results
Backloading Lumen, PTFE Backloading Lumen. Allow for easy installation of device. Length. | unable to track pump along guidewire | unable to track pump along guidewire | installation-related problem
Backloading Lumen, PTFE Backloading Lumen. Allow for easy installation of device. Material. | Breakage of lumen during removal causing pump stop | lumen broken or damaged | break
Backloading Lumen, PTFE Backloading Lumen. Allow for easy installation of device. Material. | increased friction causing breakage during removal | excessive friction | physical resistance
Backloading Lumen, PTFE Backloading Lumen. Allow for easy installation of device. Thickness. | breakage of the lumen during removal | lumen broken or damaged | fracture
Ball Bearing. Provides radial and axial during rotation of Impeller. Material of races. Early wear of balls leading to reduce life | Early wear of balls leading to reduce life | bearing worn out | degraded
Ball Bearing. Provides radial and axial during rotation of Impeller. Material of balls. | Early wear of balls leading to reduce life | inapppropriate radial rotation | device operates differently than expected
Ball Bearing. Provides radial and axial during rotation of Impeller. OD of Outer Race. | Unable to assemble | unable to assemble | installation-related problem
Ball Bearing. Provides radial and axial during rotation of Impeller. OD of Outer Race. | Bond Failure | inapppropriate axial rotation | installation-related problem
Ball Bearing. Provides radial and axial during rotation of Impeller. OD of Outer Race. Radial misalignment leading too increased motor current. | Radial misalignment leading too increased motor current | radial misalignment | failure to align
Ball Bearing. Provides radial and axial during rotation of Impeller. OD of Outer Race. axial misalignment leading too increased motor current. | Radial misalignment leading too increased motor current | axial misalignment | failure to align
Ball Bearing. Provides radial and axial during rotation of Impeller. OD of Outer Race. | Radial misalignment leading too unable to run pump | OD out of specification | physical property issue
Ball Bearing. Provides radial and axial during rotation of Impeller. ID of Inner Race. | Radial misaalignment leading too increased motor current | motor current increased | electrical issue
Ball Bearing. Provides radial and axial during rotation of Impeller. ID of Inner Race. | Radial misaalignment leading too unable to run pump | ID out of specification | physical property issue
Ball Bearing. Provides radial and axial during rotation of Impeller. Contact surface between inner and outer and balls. | Early wear of components | contact surface worn out | degraded
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Material. | Damage ventricle | ventricle damaged | material rigid or stiff
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Material. Kinkage of pigtail causing damage to ventricle | Kinkage of pigtail causing damage to ventricle | piagtail kinked | kinked
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Material. | Unable to assemble | unable to assemble | incompatibility problem
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Material. Fracture backloading lumen inside pump | Fracture backloading lumen inside pump | pigtail fractured | fracture
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Material. Split during insertion causing damage to vasculature/ventricle | Split during insertion causing damage to vasculature/ventricle | pigtail split | split
Pigtail, KR pigtail. Provides support of device in ventricle. Base Material. incorrect positioning of device due buckling of base | incorrect positioning of device due buckling of base | pigtail buckled | buckled material
Pigtail, KR pigtail. Provides support of device in ventricle. Base Material. Kinkage of pigtail causing damage to ventricle | Kinkage of pigtail causing damage to ventricle | piagtail kinked | kinked
Pigtail, KR pigtail. Provides support of device in ventricle. Base Material. Fracture backloading lumen inside pump | Fracture backloading lumen inside pump | pigtail fractured | fracture
Pigtail, KR pigtail. Provides support of device in ventricle. Tip Profile. Damage ventricle due to sharper point/easily kink causing point | Damage ventricle due to sharper point/easily kink causing point | ventricle perforated | material rigid or stiff
Pigtail, KR pigtail. Provides support of device in ventricle. Material. | Infection | patient infection | biocompatibility issue
Pigtail, KR pigtail. Provides support of device in ventricle. OD of Base and Tip. | Clot formation | clot formated | unexpected therapeutic results
Pigtail, KR pigtail. Provides support of device in ventricle. OD of Base. | Clot formation | clot formated | physical property issue
Pigtail, KR pigtail. Provides support of device in ventricle. Length. | Incorrect positioning of device | pigtail malposition | malposition of device
Pigtail, KR pigtail. Provides support of device in ventricle. Length. | Damage to ventricle | ventricle damaged | physical property issue
Pigtail, KR pigtail. Provides support of device in ventricle. Lap joint length. Overmold reflow transition. | Bond failure | pigtail overmold | loss of or failure to bond
Pigtail, KR pigtail. Reduce suction issues. Profile. | Suction issues, reduced flow | suction reduced | suction issue
Pigtail, KR pigtail. Reduce suction issues. Length. | Suction issues, reduced flow | blood flow reduced | insufficient flow or underinfusion
Pigtail, KR pigtail. Provides support of device in ventricle. Curl diameter OD. | Entanglement with bio-structures within ventricle | pigtail entangled with ventricle | patient device interaction problem
Inflow Cage. Provides opening for blood flow. Window design features. | Restriction of flow (low flow condition) | low flow condition | insufficient flow or underinfusion
Inflow Cage. Provides opening for blood flow. Window design features. Clot formation due blood stagnation zones | Clot formation due blood stagnation zones | blood stagnation zone | patient device interaction problem
Inflow Cage. Provides opening for blood flow. Window design features. | Increased blood damage | red blood cell damaged | patient device interaction problem
Inflow Cage. Provides opening for blood flow. Window design features. | Damage to ventricle/vasculature | vasculature damaged | patient device interaction problem
Inflow Cage. Provides opening for blood flow. Window design features. | Clot formation | inappropriate window design | unexpected therapeutic results
Inflow Cage. Provides opening for blood flow. Junction between teardrop and struts. | Clot formation | clot formated | unexpected therapeutic results
Inflow Cage. Provide junction between pigtail and cannula. M2 thread. | Weak bond | weak bond | loss of or failure to bond
Inflow Cage. Provide junction between pigtail and cannula. M2 thread. | Unable to assemble | difficult to assemble | installation-related problem
Inflow Cage. Provide junction between pigtail and cannula. ID of Tear drop. | Fracture of backloading lumen causing pump stop | junction fractured | fracture
Inflow Cage. Provide junction between pigtail and cannula. ID. | Clot Formation | clot formated | patient device interaction problem
Inflow Cage. Provide junction between pigtail and cannula. ID. Reduced rigidity causing it flatten/ovalize leading to reduce flow | Reduced rigidity causing it flatten/ovalize leading to reduce flow | cage rigidity reduced | material rigid or stiff
Inflow Cage. Provide junction between pigtail and cannula. ID. Reduced rigidity causing it flatten/ovalize leading to reduce flow | Reduced rigidity causing it flatten/ovalize leading to reduce flow | cage structure deformed | structural problem
Delo. Utilized for bonding. Tensile strength. | bond Failure | low bond strength | failure to adhere or bond
Delo. Utilized for bonding. Tensile strength. | unable to assemble | unable to bond | installation-related problem
Delo. Utilized for bonding. Water absorption. | bond Failure | adhesive water absorption | failure to adhere or bond
Delo. Utilized for bonding. Transition temperature. | bond Failure | incorrect transition temperature | failure to adhere or bond
Delo. Utilized for bonding. Elongation at break. | bond Failure | elongation shortened | loss of or failure to bond
Delo. Utilized for bonding. Elongation at break. | unable to assemble | unable to bond | failure to adhere or bond
Delo. Providing smooth tranisition of surfaces. Surface quality . | Abrasion issues causing damage to vasculature | rough adhesive surface | patient device interaction problem
Delo. Providing smooth tranisition of surfaces. Surface quality . | Abrasion issues causing damage to vasculature | vasculature damaged | patient device interaction problem
Delo. Providing smooth tranisition of surfaces. Surface quality . | unable to assemble | unable to assemble | scratched material
Delo. Providing smooth tranisition of surfaces. Tensile strength. Material breaks off during insertion causing damage/stroke | Material breaks off during insertion causing damage/stroke | unable to insert | difficult to insert
Delo. Providing smooth tranisition of surfaces. Compatibility to substrates. | bond Failure | fail to compatible with substrate | incompatibility problem
Delo. Joint tightenss (sealant). Viscosity. | unable to assemble | inappropriate viscosity level | installation-related problem
Delo. Joint tightenss (sealant). Viscosity. weak seal leading to leakage of purge fluid | weak seal leading to leakage of purge fluid | inappropriate viscosity level | loss of or failure to bond
Delo. Joint tightenss (sealant). Elognation at break. | Component breaks under stress leading to purge loss performance | elongation shortened | loss of or failure to bond
Delo. Joint tightenss (sealant). Compatibility to substrates. | weak seal leading to leakage of purge fluid | fail to compatible with substrate | chemical issue
Epot-Tek. Providing smooth transition on critical areas. Tensile strength. Material breaks off during insertion causing damage/stroke | Material breaks off during insertion causing damage/stroke | bond fractured | loss of or failure to bond
Epot-Tek. Providing smooth transition on critical areas. Viscosity. Unable to apply | Unable to apply | inappropriate viscosity level | installation-related problem
Epot-Tek. Providing smooth transition on critical areas. Water absorbtion. | Material breaks off during insertion causing damage/stroke | adhesive water absorption | loss of or failure to bond
Epot-Tek. Provide rotational securement for thread on inlet cage to pigtail. Bond. | Pigtail not properly secured to inlet cage | part unsecured | unstable
Epot-Tek. Provides structural support . Tensile strength. | loss performance of pump | low bond strength | loss of or failure to bond
Epot-Tek. Provides structural support . Viscosity. | loss performance of pump | inappropriate viscosity level | failure to adhere or bond
Epot-Tek. Provides structural support . Homogeneously of mixture. reduced joint strength leading to loss of pump performance. | reduced joint strength leading to loss of pump performance | inhomogeneous mixture | installation-related problem
Epot-Tek. Provides structural support . Water absorbtion. reduced joint strength leading to loss of pump performance | reduced joint strength leading to loss of pump performance | adhesive water absorption | loss of or failure to bond
Epot-Tek. Provides structural support . Glass Tranistion (Tg). increase heat leading to shrink tube delamination | increase heat leading to shrink tube delamination | adhesive overheated | temperature issue
Epot-Tek. Provides structural support . Curing shrinkage ratio. | Eptok delamination leading to pump stop | adhesive curing shrinkage | loss of or failure to bond
Epot-Tek. Provides structural support . Compatibility to substrates. | Eptok delamination leading to pump stop | adhesive delaminated | loss of or failure to bond
Epot-Tek. Provides structural support . Compatibility to substrates. | Eptok delamination leading to pump stop | fail to compatible with substrate | failure to adhere or bond
Epot-Tek. . Heat dissipation rate. increase heat leading to shrink tube delamination | increase heat leading to shrink tube delamination | shrink tube delaminated | delamination
Epot-Tek. Providing smooth transition on critical areas. Viscosity. Rough surfaces causing hemolysis | Rough surfaces causing hemolysis | rough adhesive surface | patient device interaction problem
Epot-Tek. Providing smooth transition on critical areas. Viscosity. Rough surfaces causing hemolysis | Rough surfaces causing hemolysis | patient hemolysis | unexpected therapeutic results
Epot-Tek. Providing smooth transition on critical areas. Viscosity. | Rough surfaces causing thrombosis formation | patient thrombosis formation | unexpected therapeutic results
Epot-Tek. Providing smooth transition on critical areas. Viscosity. Unable to cover critical surfaces causing corrosion issues. | Unable to cover critical surfaces causing corrosion issues | surface corrosion | corrosion
Epot-Tek. Providing smooth transition on critical areas. Compatibility to substrates. flake off causing particles in blood leading to stroke | flake off causing particles in blood leading to stroke | patient stroke | unexpected therapeutic results
Epot-Tek. Providing smooth transition on critical areas. Compatibility to substrates. flake off causing particles in blood leading to stroke | flake off causing particles in blood leading to stroke | adhesive flaked off | particulates
Epot-Tek. Providing smooth transition on critical areas. Compatibility to substrates. | flake off leading to rough surfaces being uncovered causing hemolysis | adhesive peeled off | unexpected therapeutic results
Epot-Tek. Providing smooth transition on critical areas. Compatibility to substrates. unexpected therapeutic results | flake off leading to rough surfaces being uncovered thrombosis formation | fail to compatible with substrate | unexpected therapeutic results
Coil. Create electromagnetic field. Wire Material. High motor current due to high resistance | High motor current due to high resistance | high coil resistance | electrical issue
Coil. Create electromagnetic field. Wire Material. Inflexibility leading to fracture during forming (coil short) | Inflexibility leading to fracture during forming (coil short) | coil fractured | fracture
Coil. Create electromagnetic field. Wire Insulation material. | Coil short | coil short | electrical shorting
Coil. Create electromagnetic field. Wire Insulation material. Damage during forming leading to coil short | Damage during forming leading to coil short | insulation damage | electrical shorting
Coil. Create electromagnetic field. | Damage during forming leading to coil short | insufficient electromagnetic field created | output energy incorrect
Coil. Create electromagnetic field. Wire Diameter. | High motor current due to high resistance | coil diameter out of specification | physical property issue
Coil. Create electromagnetic field. Phase Angle. | Reduced coil efficiency leading to high motor current | coil efficiency reduced | output issue
Coil. Create electromagnetic field. Wire Formation (Wire location and alignment). | Coil short | fail to form coil | uncoiled
Coil. Create electromagnetic field. Overal Length. | unable to assemble | coil length out of specification | installation-related problem
Coil. Create electromagnetic field. Overal Length. | High motor current | unable to assemble | installation-related problem
Coil. Create electromagnetic field. Coil OD. | Unable to assemble | unable to assemble | installation-related problem
Coil. Create electromagnetic field. Coil OD. Damage due to contact with ID of yoke | Damage due to contact with ID of yoke | coil abrased | device abrasion from instrument or another object
Coil. Create electromagnetic field. Coil OD. Damage to insulation leading to coil shorts during forming process | Damage to insulation leading to coil shorts during forming process | coil insulation damaged | electrical shorting
Coil. Create electromagnetic field. Lead lengths. Incorrect connections made during soldering (scrap) | Incorrect connections made during soldering (scrap) | coil misconnected | solder joint failure
Tubing Polyester Shrink Clear. Protecting coil from purge solution. ID. Inner shrink tubing delamination during molding | Inner shrink tubing delamination during molding | tubing delaminated | delamination
Tubing Polyester Shrink Clear. Protecting coil from purge solution. ID.Coil short due to increased thickness of subassembly | Coil short due to increased thickness of subassembly | coil short | electrical shorting
Tubing Polyester Shrink Clear. Protecting coil from purge solution. ID. Unable to assemble or remove component after molding | Unable to assemble or remove component after molding | unable to remove | difficult to remove
Tubing Polyester Shrink Clear. Protecting coil from purge solution. Material. coil short due fluid contact coil | coil short due fluid contact coil | electrical shorting | electrical shorting
PUR Purge Lumen 0.55X0.75mm Frosted. Deliver purge fluid to motor. ID. | High purge rate | high purge rate | inaccurate flowrate
PUR Purge Lumen 0.55X0.75mm Frosted. Deliver purge fluid to motor. Material Durometer. | Material has high kinkability leading to incorrect purge rate issues/occlusions | lumen kinked | structural problem
PUR Purge Lumen 0.55X0.75mm Frosted. Deliver purge fluid to motor. Wall thickness. | Material has high kinkability leading to incorrect purge rate issues/occlusions | lumen kinked | physical property issue
PUR Purge Lumen 0.55X0.75mm Frosted. Deliver purge fluid to motor. Surface finish. | Unable to assemble | inappropriate lumen frosted | physical resistance
Pressure reservoir. Provides purge pressure when purge cassette/purge bag gets changed. Membrane Material. | Pressure reservoir does not store sufficient purge volume for change of purge cassette/purge fluid bag | unable to provide sufficient pressure | pressure issue
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Membrane Material. Pressure reservoir does not provide sufficient purge pressure for change of purge cassette/purge fluid bag | Pressure reservoir does not provide sufficient purge pressure for change of purge cassette/purge fluid bag | fail to provide sufficient pressure | installation-related problem
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Membrane Dimensions. | Pressure reservoir does not provide sufficient purge pressure for change of purge cassette/purge fluid bag | fail to provide pressure | installation-related problem
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Interface between Components. Leakage of purge fluid | Leakage of purge fluid | reservoir leakage | fluid leak
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Material. Infection | infection | reservoir infection | bacterial contamination of device
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Dimensions of Check Valve. Check valve allows flow in both directions leading to leak during change of purge cassette/purge fluid bag | Check valve allows flow in both directions leading to leak during change of purge cassette/purge fluid bag | valve opened in wrong direction | fluid leak
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Dimensions of Check Valve. Check valve blocks flow in both directions leading to no adequate purge pressure | Check valve blocks flow in both directions leading to no adequate purge pressure | valve blocked in wrong direction | no flow
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Locking dimensions Cap Nut. | Unable to assemble | unable to assemble | installation-related problem
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Locking dimensions Cap Nut. | Unable to assemble | locking dimension out of specification | physical property issue
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Locking dimensions check valve. | Unable to connect to purge cassette | unable to connect cassette | connection issue
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Material. Aging effects leading to embrittlement of membrane and damaged membrane | Aging effects leading to embrittlement of membrane and damaged membrane | membrane embrittled | degraded
Pressure reservoir. Provides purge pressure when purge cassette/purge fluid bag gets changed. Material. Moisture absorption leading to out of specification tolerance | Moisture absorption leading to out of specification tolerance | moisture absorbed | moisture damage
Pump 371, Unsterile. Biocompatible with body. Material. | Not biocompatible with body | material bio-incompatibility | biocompatibility issue
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES | Hemolysis | fail to measure cardiac output | incorrect measurement
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Incorrect pump position | Incorrect pump position | incorrect pump position measured | incorrect measurement
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Excessive pump movement | Excessive pump movement | incorrect movement calculated | programming issue
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Motor current drift due to pump wear | Motor current drift due to pump wear | motor current drifted | electrical issue
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Suction at high speed during calibration sequence | Suction at high speed during calibration sequence | suction at high speed | increase in suction
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Reduced support (lower mean flow/MAP) during calibration sequence. | Reduced support (lower mean flow/MAP) during calibration sequence | flow reduced | inadequate service
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Loss of diagnostic accuracy | Loss of diagnostic accuracy | loss of accuracy | incorrect software programming calculations
Algorithm. Measure Cardiac Ouput. Algorithm is described in ES. Pump speed deviation from set speed during calibration | Pump speed deviation from set speed during calibration | pump speed deviation | calibration issue
Algorithm. Measure left ventricular and diastolic pressure. Algorithm is described in ES. Loss of diagnostic accuracy | Loss of diagnostic accuracy | loss of diagnostic accuracy | calibration issue
Algorithm. Estimate Flow. Flow displayed is +/- 10% of actual flow. displayed flow is greater than actual flow | displayed flow is greater than actual flow | flow displayed high | incorrect display
Front Panel Optical Connection. Establishing optical connection between the optical bench and optical sensor on pump. Insertion loss (< 0.5). | No placement signal/Poor placement signal | loss of optical connection | disconnection
Front Panel Optical Connection. Establishing optical connection between the optical bench and optical sensor on pump. Alignment of ferrule in the recepticle. No placement signal/Poor placement signal | No placement signal/Poor placement signal | loss of signal | failure to read input signal
Front Panel Optical Connection. Establishing optical connection between the optical bench and optical sensor on pump. Geometry of end face of ferrule. | No placement signal/Poor placement signal | no placement signal | connection issue
Front Panel Optical Connection. Establishing electrical connections between pump motor and pump motor driver. Wiring arrangement/configuration. | pump will not run | incorrect wiring arrangement | device misassembled during manufacturing or shipping
Front Panel Optical Connection. Establishnig electrical connections between EEPROM and Impellatronics Board. Wiring arrangement/configuration. | unable recognize pump, pump will not run | unable to recognize pump | failure to sense
Front Panel Optical Connection. Establishing electrical connections between the pressure sensors and Impellatronics board. Wiring arrangement/configuration. | No placement signal/Poor placement signal | poor placement signal | disconnection
Optical Bench Opsens Assembly Test. Translating the light signal coming from the optical pump to a digital signal that the AIC can read. Contrast (%). | No placement signal/Poor placement signal | fail to translate light signal to digital signal | optical issue
Optical Bench Opsens Assembly Test. Translating the light signal coming from the optical pump to a digital signal that the AIC can read. Signal to Noise Ratio (SNR). | No placement signal/Poor placement signal | fail to translate light signal to digital signal | computer software issue
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Engagement surfaces. | Contamination on ferrule face leading to poor/no sensor performance | sensor contaminated | contamination during use
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Engagement surfaces. | Contamination on ferrule face leading to poor/no sensor performance | sensor surface contaminated | decreased sensitivity
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Spring force. | Unable to close cover leading to contamination | unable to close cover | difficult to open or close
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Material interface at push washer. | Wear out of spring force overtime leading to contamination | spring worn out | naturally worn
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Material interface at push washer. Particulate contamination leading to contamination on ferrule surface | Particulate contamination leading to contamination on ferrule surface | particulate contamination | contamination during use
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Opening features. | Difficult to open by user, delay in use | difficult to open cover | difficult to open or close
Optical Receptacle Cover. To protect optical ferrule face in plug from contamination/damage. Spring force. | Difficult to open by user, delay in use | spring force worn out | difficult to open or close
Container, Shipping, Impella Controller. Protect AIC console during transit. Foam cutout layout. | Damaged AIC unit, non-operable | device damaged in shipping | shipping damage or problem
Assembly Motor Purge AIC. Drives the rack and piston on the purge cassette in order to provide purge flow. Power source. | Failure of motor to run | fail to run | failure to power-up
Assembly Motor Purge AIC. Drives the rack and piston on the purge cassette in order to provide purge flow. Stepper motor driver. | Failure of motor to run | fail to run | defective component
Assembly Motor Purge AIC. Drives the rack and piston on the purge cassette in order to provide purge flow. Spline assembly. | failure to engage with purge cassette properly | fail to engage with motor | disconnection
Assembly Motor Purge AIC. Drives the rack and piston on the purge cassette in order to provide purge flow. Wiring arrangement/configuration. | Failure of motor to run | fail to run motor | defective item
Purge pressure sensor. Translates pressure of the feedback system to drive the stepper motor. pressure reservour. | Failure of motor to run | fail to translate pressure | failure to select signal
Purge pressure sensor. Translates pressure of the feedback system to drive the stepper motor. Pressure transducer present switch. | Prevents switch from acknowledging tranducer is present causing motor not to run | feel to feeback pressure | unable to obtain readings
RFID PCA. Sensor presence of purge cassette to alert system to run. Signal range. | fails to recognize purge cassette causing motor not to run | fail to recognize cassette | failure to sense
Purge Unit Drive. Translates pressure of the feedback system to drive the stepper motor. Wiring arrangement/ configuration. | fails to translate pressure adequately causing the motor not to run | fail to drive motor | failure to deploy
Purge Unit Drive. Translates pressure of the feedback system to drive the stepper motor. Board configuration. | fails to translate pressure adequately causing the motor not to run | fail to receive pressure | failure to read input signal
Rear Housing. Enclosure for the AIC unit and helps with mounting components. Conductive paint. failure to prevent interference/failures through outside RF. | failure to prevent interreference/failures through outside RF | fail to prevent RF interference | Radiofrequency Interference (RFI)
Rear Housing. Enclosure for the AIC unit and helps with mounting components. Conductive paint. | failure to prevent interreference/failures through inside RF | fail to prevent RF interference | Radiofrequency Interference (RFI)
Rear Housing. Enclosure for the AIC unit and helps with mounting components. Threaded inserts. | unable to mount components/unable to assemble | unable to mount | installation-related problem
Push Button Switch. Turns on the unit. Wiring arrangement/configuration. | Unable to turn on unit | unable to turn on | failure to deploy
Push Button Switch. Turns on the unit. Wiring arrangement/configuration. | Unable to properly shut down unit | unable to shut down | failure to shut off
Fuse Holder (Neutral). Shutdown system if there is a serious fault condition occurs (safety control). Type of fuse (Size, blow time). | System faults too early/faults too late | system shut down too late | protective measure issue
Fuse Holder (Line). Shutdown system if there is a serious fault condition occurs (safety control). Type of fuse (Size, blow time). | System faults too early/faults too late | system shut down too early | protective measure issue
Power Entry Module. Cleans up electrical noise coming from the wall. Wiring arrangement/configuration. | no power to the unit | unable to clean electrical noise | radiofrequency interference (RFI)
Power Entry Module. Cleans up electrical noise coming from the wall. Internal EMI filter. | unable to filter out EMI noise, leading insufficient protection against power surges | unable to filter out EMI noise | electromagnetic interference (EMI)
Power Entry Module. Charges battery of the AIC and powers the console. Internal component configuration. | Excess EMI noise, interferrance to other boards | excess EMI noise | application interface becomes nonfunctional or program exits abnormally
Power Entry Module. Charges battery of the AIC and powers the console. Internal component configuration. | unable to provide power to unit | power outage | loss of power
Power Supply. Charges battery of the AIC and powers the console. Internal component configuration. | Pull too much power and fuse blows | power pulled too much | power conditioning issue
Fan. Provide cooling for the AIC unit. Wiring arrangement/configuration. | Failure of fan motor | fail to provide cooling | overheating of device or device component
Speakers. Provide audio sounds/warnings to the user. Wiring arrangement/configuration. | Failure of speaker | fail to provide audio warning | inaudible voice prompts
Speakers. Provide audio sounds/warnings to the user. Coil Configuration. | Failure of speaker | fail to provide audio warning | device alarm system issue
Door Latch Assembly. Holds the door to the purge cassetter. Locking mechanism. | Unable to open door assembly | fail to unlatch | difficult to open or close
Door Latch Assembly. Holds the door to the purge cassetter. Locking mechanism. | Unable to keep door closed | unable to keep door closed | difficult to open or close
Door Latch Assembly. Holds the door to the purge cassetter. Spring. Locking mechanism. | Unable to keep door closed | locking mechanism failure | fail-safe mechanism issue
Battery Switch. To disengage batteries from unit to allow for shipping of units. Plastic rocker switch. | Unable to open unit to allow for shipping to patients | fail to disengage | failure to separate
Battery Switch. To disengage batteries from unit to allow for shipping of units. Plastic rocker switch. Unable to close causing power loss when unplugged during transportation | Unable to close causing power loss when unplugged during transportation | unable to disengage | loss of power
Rotary Selector. Navigate the screen on the console and select functions. Switch mechanism. | Fails to actuate | fail to navigate | human-device interface issue
Rotary Selector. Navigate the screen on the console and select functions. Switch mechanism. | Fails to actuate | fail to select | use of device issue
Rotary Selector. Navigate the screen on the console and select functions. Switch mechanism. | Fails to actuate | fail to actuate | invalid sensing
Rotary Selector. Navigate the screen on the console and select functions. Shaft. | does not register actuation | fail to register actuation | failure to sense
Rotary Selector. Navigate the screen on the console and select functions. Clockwise/counter clockwise detection. | Damage electronics, unable to detect rotation | fail to detect rotation | device operates differently than expected
Front Housing. Holds the scree, nob, PUD assembly, and other functional components. Conductive paint. | failure to prevent interreference/failures through outside RF | fail to prevent RF interference | Radiofrequency Interference (RFI)
Front Housing. Holds the scree, nob, PUD assembly, and other functional components. Threaded inserts. | unable to mount components/unable to assemble | unable to assemble | installation-related problem
AC Line Cord & Plug. Electrical Power Mains Connector. Electrical Cord and Connector. Open/Short Circuit | Open/Short Circuit | plug short circuit | electrical shorting
AC Line Cord & Plug. Electrical Power Mains Connector. Electrical Cord and Connector. Unseated plug from console AC Power socket. | Unseated plug from console AC Power socket | plug unseated | loose or intermittent connection
AC Line Cord & Plug. Electrical Power Mains Connector. Electrical Cord and Connector. Unseated plug from wall AC Power Outlet. | Unseated plug from wall AC Power Outlet | plug unseated | disconnection
LCD Display. Electronic component performance. component within specification. | electrical short / over load | blank screen | circuit failure
LCD Display. Visual User Interface. Visual Screen. | No or poor screen image | poor screen image | poor quality image
Smart Battery Pack. Backup Power. Nominal voltage, current and power capacity. | Fails to provide / sustain power / current | fail to sustain power | battery issue
Smart Battery Pack. Backup Power. Charge time to full capacity. | Slow or no charge rate | slow charge rate | charging issue
Smart Battery Pack. Backup Power. Discharge time with full capacity. Battery unable to power AIC operating under worst case conditions for a minimum of 1 hour. | Battery unable to power AIC operating under worst case conditions for a minimum of 1 hour. | reduced battery capacity | low battery
Smart Battery Pack. Backup Power. Smart Battery Data interface. | Fails to provide data | fail to provide battery data | data issue
Power and Battery Manager PCA. Electronic component performance. component within specification. | electrical short / over load | electronical component over load | electrical overstress
Power and Battery Manager PCA. Provides AC Power. . AC Connected, Power Console from AC/DC Power Supply. | Fails to Power Console | fail to power on AC | failure to power-up
Power and Battery Manager PCA. Battery Power. AC not connected, Power Console from battery. | Fails to Power Console | fail to power on battery | battery issue
Power and Battery Manager PCA. Charges the Battery. Recharges the battery under AC connected power. | Fails to Charge Battery | fail to recharge battery | failure to charge
Power and Battery Manager PCA. Activates Alarm. Upon power loss, Capacitor activate Alarms. | Fails to Activate power loss Alarms | fail to activate alarm | electrical power problem
Power and Battery Manager PCA. Activates Alarm. Upon power loss, Capacitor activate Alarms. | Fails to Activate power loss Alarms | low capacitor charged | charging issue
Power and Battery Manager PCA. Communication to 4-1 Assembly. Provides Battery Level and Charging level indicators to User. | Fails to Provide information to EPC | fail to provide battery level | battery issue
Power and Battery Manager PCA. Communication to 4-1 Assembly. Provides Battery Level and Charging level indicators to User. | Fails to Provide information to EPC | fail to provide charge level | battery issue
Power and Battery Manager PCA. Powers Fan. Provides Cooling. | Fails to Cool system | fail to cool system | overheating of device or device component
Power and Battery Manager PCA. Powers Indication LED. Provides power source and On/Off mode indication. | Fails to indicate power mode state | fail to indicate power mode | no device output
Power and Battery Manager PCA. Powers Indication LED. Provides power source and On/Off mode indication. | Fails to indicate power mode state | fail to indicate power state | no device output
Power and Battery Manager PCA. Patient Isolation / Power. Powers PMD Circuitry. | Loss of Power | loss of power | loss of power
Power and Battery Manager PCA. Patient Isolation / Power. Provide Isolation between patient and AC Mains. | Allows leakage current | over leakage current | electrical issue
Impellatronic Common Ground Interface. Electronic component performance. component within specification. | electrical short / over load | electronical component over load | electrical overstress
Impellatronic Common Ground Interface. Relays communication from EPM and DSP to EPC. EPM Communication. | Loss of Communication | loss of communication | communication or transmission issue
Impellatronic Common Ground Interface. Provides rotary push button interface to 4-in-1 Assembly. Direction Detector and Switch. | Fail to detect rotation or press | fail to detect rotation | failure to sense
Impellatronic Common Ground Interface. Provides LED Power Indication from PBM. LED color indication. | Fails to indicate proper console power state | fail to indicate power state | inadequate user interface
Impellatronics PMD. Electronic component performance. component within specification. | electrical short / over load | electronical component short | electrical shorting
Impellatronics PMD. Pump Driver. DSP Executes pump drive algorithm. | DSP fails to execute algorithm | fail to execution algorithm | programming issue
Impellatronics PMD. Pump Driver. Power Circuitry Switches pump motor phases on and off. | Fails to switch motor phases | fail to switch motor phase | pumping issue
Impellatronics PMD. EPC Communication /Patient Isolation. Communicates with EPC. | fails to communicate | fail to communicate | communication or transmission issue
Impellatronics PMD. EPC Communication /Patient Isolation. Provide Isolation between patient and AC Mains. | Allows leakage current | excess electrical leakage | electrical issue
Impellatronics PMD. Pump drive electrical parameters information . DSP ADC - Reads/Calculates pump current and speed. | fails to read/calculate | fail to read | failure to read input signal
Impellatronics PMD. Pump drive electrical parameters information . DSP ADC - Reads/Calculates pump current and speed. | fails to read/calculate | fail to calculate | parameter calculation error due to software problem
Impellatronics PMD. Pump drive electrical parameters information . Analog Current sensing circuitry - Conditions current signals. | fails to accurately condition signals | fail to condition signal | use of incorrect control settings
Impellatronics PMD. CPLD Backup system. Monitors pump current and speed range and activates alarm. | Fails to detect error and alarm | fail to detect error | device sensing issue
Impellatronics PMD. DSP Flash Memory. Sustain program and configuration data. | flash memory corruption. | flash memory corrupted | loss of data
Impellatronic EPM. Pump EEPROM Communication. Provides Communication to Pump EEPROM. | Failed to read pump from EEPROM | fail to read | failure to read input signal
Impellatronic EPM. EPC Communication CPU / RS232 Circuitry. Provides Communication to EPC. | Failed to communicate to EPC | fail to communicate | communication or transmission issue
Impellatronic EPM. Sensor Reading CPU / Sensor Circuitry. Obtain data from hydraulic pressure sensor. | Failed to read Hydraulic sensor | fail to read | unable to obtain readings
Impellatronic EPM. Control Algorithm Execution. Program Execution. | Fails to function | fail to execution algorithm | computer software issue
4-in-1 Assembly Carrier Board. Electronic component performance. component within specification. | electrical short / over load | board short circuit | electrical shorting
4-in-1 Assembly Carrier Board. Electronic component performance. component within specification. | electrical short / over load | component overloaded | electrical overstress
4-in-1 Assembly Carrier Board. CPU Functionality . Executes System Software. | Fails to start or run EPC | fail to run | electrical issue
Assembly. Carrier Board. Power to Display Circuitry (LCD Panel and Backlight). Visual User Interface. | No User Interface | no display | no display or display failure
Assembly. Carrier Board. Audio Circuitry. Amplifies EPC output into speaker. | No / poor quality audio from speaker | low quality audio | audible noise
Assembly. Carrier Board. Audio Circuitry. Amplifies EPC output into speaker. | Audio Amplifier excessive power load | poor quality audio | audible noise
Assembly. Carrier Board. External Video Output to Patient Monitors and Impella Connect. VGA Display output. | No or poor video Output | poor video output | poor quality image
4-in-1 Assembly Carrier Board. Real Time Clock functionality. Runs electronic calendar and system time. | Inaccurate time keeping | inaccurate time keeping | date-related software issue
4-in-1 Assembly Carrier Board. Real Time Clock functionality. Runs electronic calendar and system time. | resets time / calendar | incorrect electronic calendar | date-related software issue
4-in-1 Assembly Carrier Board. Interface and Power to Compact Flash Cards. System boot and mass storage device. | System fails to boot / start application | fail to boot up | reset issue
4-in-1 Assembly Carrier Board. USB Interface. Software Upgrade / Log download. | System fails to log data | fail to download log data | data back-up problem
4-in-1 Assembly Carrier Board. USB Interface. Software Upgrade / Log download. | Fails to boot from USB / install software | fail to upgrade software from USB | problem with software installation
4-in-1 Assembly Carrier Board. USB Interface. Power to Impella Connect. | Fails to power / insufficient power to impella connect | fail to power system | failure to power-up
Assembly. Carrier Board. Display Brightness Memory. Adjustable Display backlight controller. | No or poor display brightness | Insufficient brightness | optical issue
4-in-1 Assembly Carrier Board. Ethernet. Network connection. | System configuration | incorrect network configuration | use of incorrect control settings
4-in-1 Assembly Carrier Board. Keypad circuitry. Provides Power Path to Keypad. | Inoperable Console | fail to power keypad | electrical power problem
4-in-1 Assembly Carrier Board. Serial Port Circuitry. Provides power path. | Inoperable Console | fail to power port | power source issue
4-in-1 Assembly Carrier Board. Serial Port Circuitry. Provides signal communication path. | Inoperable Console | signal path open | circuit failure
4-in-1 Assembly Carrier Board. Serial Port Configuration. Provides communication configuration. | Inoperable Console | incorrect communication configuration | use of incorrect control settings
4-in-1 Assembly Carrier Board. Provide interface between computer module and peripheral device. Network connection. | Compromised SW | network failure | application network issue
4-in-1 Assembly Carrier Board. Provide interface between computer module and peripheral device . Provides power and signal paths. | Loss of Power or Communication | no power output | loss of power
4-in-1 Assembly Carrier Board. CPU Functionality . Executes System Software. | Inoperable Console | board malfunctioned | device inoperable
Back Light Cable. Carry Power / Signal. Provides Power / Signal Path. | Short or Open | cable short | electrical shorting
LVDS Cable. Carry Signal. Provides Signal Path. | Short or Open | cable open | failure to read input signal
Service Dongle. Carry Signal. Provides Signal Path. | Short or Open | service dongle open | electrical issue
4-in-1 to Optical Bench Cable. Carry Signal. Provides Signal Path. | Short or Open | cable opened | failure to read input signal
Optical Bench Power Cable. Carry Power. Provides Power Path. | Short or Open | cable short | electrical shorting
Optical Bench Power Cable. Carry Power . Provides Power Path. | Short or Open | cable short | failure to read input signal
4-in-1 to PUD Cable. Carry Power / Signal. Provides Power / Signal Path. | Short or Open | cable open | loss of power
AU Purge Lumen Pump 401. Provides purge fluid to the motor. Connection of purge lumen. | leaking between joint | lumen connection leakage | fluid leak
AU Motor Housing. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. | Bad connections to soldering board leading to pump stop/failure to start pump | inappropriate soldering | solder joint failure
AU Motor Housing. Provide power to the Impella. Coil location relative to the yoke. | Reduced power/flow | reduced power | power source issue
AU Motor Housing. Provides interface for Rotor Assembly. Overall Length. | Unable to engage into ball bearing | unable to engage into bearing | disconnection
AU Motor Housing. Provides interface for Rotor Assembly. Overall Length. Weak bond to ball bearing | Weak bond to ball bearing | weak bond | failure to adhere or bond
AU Motor Housing. Provides interface for pump housing. ID of insert. | Unable to assemble hub | unable to assemble | installation-related problem
Stator Pump. Provides heat sink. Material. Loss of efficiency (reduced flow/support). | Loss of efficiency (reduced flow/support) | stator overheated | overheating of device or device component
Stator Pump. Provides heat sink. Material. Loss of efficiency (reduced flow/support). | Loss of efficiency (reduced flow/support) | loss of efficiency | pumping issue
Stator Pump. Provides heat sink. Material. Loss of efficiency (reduced flow/support). | Loss of efficiency (reduced flow/support) | pump flow reduced | pumping issue
Stator Pump. Provides structural support for pump/motor. Mechanical Properties (i.e. hardness). | Component breaks under stress | stator broken or damaged | material integrity issue
Motor Stator. Provides material for electronmagnetic fields for motor operation. Cross sectional area of yoke. RPM instability and/or loss of efficiency (reduced flow/support) | RPM instability and/or loss of efficiency (reduced flow/support) | stator rotation instable | unstable
Ball Bearing. Provides axial and radial support during rotation of Impeller. Material of races. | Early wear of balls leading to reduced runtime | bearing life cycle reduced | degraded
Ball Bearing. Provides axial and radial support during rotation of Impeller. Material of balls. | Early wear of balls leading to reduced runtime | bearing runtime reduced | degraded
Ball Bearing. Provides axial and radial support during rotation of Impeller. OD of Outer Race. Radial misalignment leading too increased motor current | Radial misaalignment leading too increased motor current | radial misalignment | failure to align
Ball Bearing. Provides axial and radial support during rotation of Impeller. ID of Inner Race. | Radial misaalignment leading to increased motor current | motor current increased | electrical issue
Ball Bearing. Provides axial and radial support during rotation of Impeller. Contact surface between inner and outer and balls. | Early wear of components | contact surface worn out | degraded
Sleeve Bearing. Provides radial support during rotation of Impeller. OD. | Unable to assemble | fail to provide support | material integrity issue
Sleeve Bearing. Provides radial support during rotation of Impeller. ID. | Unable to assemble | unable to assemble | installation-related problem
Magnet. Provide power to the Impella. Magnet Dimensions. | Insufficient power supplied to rotate impeller | magnet out of dimension | physical property issue
Magnet. Provide power to the Impella. Magnet Dimensions. | Unable to assemble | magnet out of dimension | installation-related problem
No Hole Insert , Yoke to Hub. Provides interface for pump housing. end face perpendicularity. | Weak bond of pump housing | weak bond | failure to adhere or bond
AU Sensor-Sensorcable, Pump. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | false reading from device noncompliance
AU Sensor-Sensorcable, Pump. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect | incorrect reading | low readings
AU Sensor-Sensorcable, Pump. Biocompatible with body. Sensor Material. | Not biocompatible with body | sensor biocompatibility compromised | biocompatibility issue
AU Sensor-Sensorcable, Pump. Determining position of pump within the body. Diaphragm dimensions. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor drifted | migration of device or device component
BG Rotor Pump. Provides flow for the pump . Gap between Shaft and sleeve bearing. | Insufficient radial support of rotor shaft | excessive gap | misassembled
BG Rotor Pump. Provides flow for the pump . Gap between Shaft and sleeve bearing. | Unable Pass final qual | insufficient radial support | misassembled
BG Rotor Pump. Provides flow for the pump . Adhesive Epoxy Duopox 1895 locations/Transitions. | Bond failure | epoxy mislocated | failure to adhere or bond
BG Rotor Pump. Provides flow for the pump . Delo location (application). Failure of shaft to rotate | Failure of shaft to rotate | unable to rotate | device inoperable
BG Motor-Pumphouse Pump. Determining position of pump within the body. Silicone gel thickness. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect pressure reading | incorrect measurement
BG Motor-Pumphouse Pump. Determining position of pump within the body. Silicone gel thickness. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to provide reading | unable to obtain readings
BG Motor-Pumphouse Pump. Biocompatible with body. Silicone gel material. | Not biocompatible with body | non-biocompatible gel applied | biocompatibility issue
BG Motor-Pumphouse Pump. Biocompatible with body. Silicone gel material. Sensor drifts beyond specification (less accurate, issues with placement) | Sensor drifts beyond specification (less accurate, issues with placement) | sensor drifted | incorrect measurement
BG Motor-pump housing. Provides pump placement information. Sensor location on pump housing. | Incorrect placement signal | incorrect location information | positioning issue
BG Motor-pump housing. Provides pump placement information. Sensor hole on pump housing. | Inaccurate placement signal readings | incorrect placement signal | positioning issue
BG Motor-pump housing. Assembling purge system to motor housing. Sensor Cable alignment with motor housing. | unable to assemble | unable to assemble | installation-related problem
BG Motor-pump housing. To provide flow. Positioning of Impeller to Pump housing. | hemolysis | patient hemolysis | unexpected therapeutic results
BG Motor-pump housing. To provide flow. Positioning of Impeller to Pump housing. Pump stop due to contact between rotor and pump housing | Pump stop due to contact between rotor and pump housing | housing abrased | device abrasion from instrument or another object
BG Motor-pump housing. To provide flow. Positioning of Impeller to Pump housing. | Low flow | low flow | insufficient flow or underinfusion
BG Motor-Pump housing. Creating geometry for the sensor. Alignment of the sensor cable channel through the assembly. | Unable to assembly | assembly misaligned | misassembled
BG Motor-Pumphouse Pump. Provide mechanical integrity for motor assembly. Weld locations. | Failure of weld at low loads | loss of mechanical integrity | solder joint failure
BG Motor-Pumphouse Pump. Provide mechanical integrity for motor assembly. Depth Integrity. | Failure of weld at low loads | loss of mechanical integrity | installation-related problem
BG Motor-pump housing. To provide flow. Positioning of Impeller to Pump housing. Biomaterial gets caught between impeller and pump housing | Biomaterial gets caught between impeller and pump housing | biomaterial caught | insufficient flow or underinfusion
Impeller. Air in purge delivery. OD of Shaft/DLC Coating Thickness. | Low purge pressure | air in purge path | air leak
BG Pump Housing. Structural support for the Impella. Material. | Buckling/interface of Impeller to pump housing (pump stop) | housing surface buckled | structural problem
BG Pump Housing. Structural support for the Impella. Strut design features. | Buckling/interface of Impeller to pump housing (pump stop) | housing surface buckled | material distortion
BG Pump Housing. Allows blood to enter pump safely. Suface finish. | Damage to red blood cells | red blood cell damaged | unexpected therapeutic results
BG Pump Housing. Allows blood to enter pump safely. Suface finish. | Thrombus formation | fail to allow blood flow safely | unexpected therapeutic results
BG Pump Housing. Allows blood to enter pump safely. Window design features. | Damage to red blood cells | inappropriate hosing design | unexpected therapeutic results
BG Pump Housing. Allows blood to enter pump safely. Window design features. | Thrombus formation | thrombus formated | structural problem
BG Pump Housing. Allows blood to enter pump safely. Window design features. | Reduced flow | inappropriate hosing design | insufficient flow or underinfusion
BG Pump Housing. Provides protected location for sensor to mount on. Sensor recessed feature. | Sensor damage/sensor dislodgement | sensor dislodged | dislodged or dislocated
BG Pump Housing. Provides protected location for sensor to mount on. Sensor recessed feature. | Sensor damage | sensor recessed | pitted
BG Pump Housing. Provides protected location for sensor to mount on. Sensor deflection bump/sensor recessed feature. | Sensor damage | sensor deflected | dislodged or dislocated
BG Pump Housing. Provides interface for sensor to read pressure. Sensor hole. | Insufficient area to read pressure | sensor holed | incorrect measurement
BG Pump Housing. Provides interface for sensor to read pressure. Sensor hole. | Insufficient area to read pressure | insufficient sensor area | incorrect measurement
BG Pump Housing. Provides welding junction for pump housing. Material. | Weld failure | material weldability compromised | solder joint failure
AU Pump Catheter. Provides reinforcement of deliverability. Placement of nitonol wire. | Unable deliver | fail to reinforce deliverability | delivery system failure
AU Pump Catheter. Provides deliverability for the pump into patient and long term use. Material (Rigidity). | Unabe deliver unit | unable to deliver | failure to deliver
AU Pump Catheter. Provides deliverability for the pump into patient and long term use. Amount of flash. | Thrombus formation | thrombus formation | unexpected therapeutic results
AU Pump Catheter. Finished pump unit which provides flow of blood and pressure sensing (without purge). Connection of Sensor Cable to Pump Housing. | Sensor Failure | unsecured connection | installation-related problem
AU Pump Catheter. Provides deliverability for the pump into patient and long term use. Amount of flash. | Thrombus formation | thrombus formated | improper flow or infusion
Catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. | reduced deliverability of device | catheter deliverability reduced | inaccurate delivery
Catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. | Kink during device deliverability | fail to push | failure to advance
Catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. | Kink during device deliverability | catheter torqued | material torqued
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material thickness. | Kink of catheter | fail to recover kink | physical property issue
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material thickness. | Tear/Puncture | fail to resist kink | structural problem
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Kink of catheter | catheter kinked | kinked
Catheter. Provides protection for internal wires and lumens (kink recovery and kink resistance). Material selection. | Tear/Puncture | catheter torn | torn material
Stiffening Pin. Provides independent connection points for Purge lumen. Slot Size. Reduced tensile strength with bearing cap | Reduced tensile strength with bearing cap | tensile strength reduced | mechanical issue
Stiffening Pin. Creates seal between purge lumen and motor. ID of the pin. | Leak of purge fluid into catheter | lumen leakage | fluid leak
Stiffening Pin. Creates seal between purge lumen and motor. OD of the pin . | Leak of purge fluid into catheter | purge fluid leakage | fluid leak
Stiffening Pin. Provides independent connection points for Purge lumen. pin OD surface finish. | Reduced tensile strength with bearing cap | tensile strength reduced | mechanical issue
Stiffening Pin. Provides strain relief and torque support for all lumen connections. Overall length. | Purge leakage/Kink | fail to provide torque support | mechanical issue
Stiffening Pin. Provides strain relief and torque support for all lumen connections. Overall length. | Purge leakage/Kink | fail to provide strain relief | pumping stopped
Nitinol Wire. Increased Kink resistance. OD. | Unable to assemble | unable to assemble | installation-related problem
Nitinol Wire. Bond to motor housing. Surface finish of distal end. | Reduced bond strength - damage to purge or catheter | low bond strength | failure to adhere or bond
Nylon Catheter support Ring. Provides recess for sensor cable exiting catheter. Support ring ID. | Support ring deforms during off axis insertion (loss of function, pushability) | support ring deformed | installation-related problem
Nylon Catheter support Ring. Provides recess for sensor cable exiting catheter. Support ring ID. | Support ring deforms during off axis insertion (loss of function, pushability) | support ring deformed | material distortion
Nylon Catheter support Ring. Connection point of the motor to catheter. OD of rings. | rings deform during off axis insertion (loss of function, pushability) | ring deformed | material distortion
11F Transition Ring. Connection point of the motor to catheter. OD of ring or Barb height. Stretching of catheter reduces wall thickness and tensile strength. | Stretching of catheter reduces wall thickness and tensile strength | catheter stretched | stretched
11F Transition Ring. Connection point of the motor to catheter. Spacing between barbs. | Insufficient contact area with catheter | insufficient contact area | disconnection
Pump Plug Assembly. Provides product identification for pump. location of serial number logo on pump shell. | unable to determine pump type | fail to determine pump type | device markings issue
Pump Plug Assembly. Contains electrical and fluid connections. Shell internal layout. Damage the sensor cable | Damage the sensor cable, loss of signal | electrical connection damaged | failure to read input signal
Pump Plug Assembly. Contains electrical and fluid connections. Shell internal layout. | Damage the purge lumen | fluid connection damaged | fluid leak
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Crack/damaged leading to pump stop | housing broken or damaged | protective measure issue
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Not biocompatible with body (skin iteration) | non-biocompatible material used | biocompatibility issue
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Material. Poor bonding surface leading to internal components being exposed | Poor bonding surface leading to internal components being exposed | internal component exposed | failure to adhere or bond
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Warpage of material leading to unable to assemble | material warped | disassembly
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Internal dimensional features. | Warpage of material leading to unable to assemble | material warped | installation-related problem
Upper and Lower Plug Halves. Provides secure housing of the various components at the junction point from cable to catheter. Internal dimensional features. | Damage to internal components leading to pump start issues or stoppage | internal component damaged | device inoperable
Upper and Lower Plug Halves. Provides handle for user to manipulate device. External geometry. | Unable to handle (difficulty with insertion of device) | unable to handle | difficult to insert
Upper and Lower Plug Halves. Provides handle for user to manipulate device. External geometry. | Unable to place device effectively | difficult to place device | device handling issue
Upper and Lower Plug Halves. Provides handle for user to manipulate device. External geometry. | Damage to gloves or clinicians hand | user hand injured | human factors issue
Upper and Lower Plug Halves. Provides location for device identification (Product ID and Serialization). Material. | Unable to read identification | unable to read product serialization | device markings issue
Upper and Lower Plug Halves. Provides location for device identification (Product ID and Serialization). Material. | Unable to print information on pump | unable to print identification | device markings issue
Upper and Lower Plug Halves. Provides location for device identification (Product ID and Serialization). Material. | Unable to print information on pump | unable to read product ID | device markings issue
Upper and Lower Plug Halves. Provides connection point catheter strain relief. Interlock geometry. | dislodgement of connection | connection dislodged | dislodged or dislocated
Strain Relief. Relief strain of catheter and pressure storage set assembly. Material (Durometer). | kink leading to pump stop | fail to provide strain relief | kinked
Strain Relief. Relief strain of catheter and pressure storage set assembly. Material (Durometer). | kink leading to placement signal failure | catheter kinked | pumping stopped
Strain Relief. Relief strain of catheter and pressure storage set assembly. Material (Durometer). | kink leading to deliverability issues | catheter kinked | no flow
Connector with PCB. Creates electrical path from the patient cable through catheter to motor. Trace Width. | Over current condition/excessive heating leading to board failure | trace width short | electrical overstress
Connector with PCB. Creates electrical path from the patient cable through catheter to motor. Solder pads. | Excessive heating during assembly leading to board failure (i.e. pad delamination) | pad delaminated | overheating of device or device component
Connector with PCB. Creates electrical path from the patient cable through catheter to motor. Trace Thickness. | Over current condition/excessive heating leading to board failure | trace thickness thin | electrical overstress
Connector with PCB. Creates electrical path from the patient cable through catheter to motor. Trace spacing. | Interferance which leads EEPROM communication failure | trace spacing short | electrical shorting
Connector with PCB. Assembles onto pump plug shell. OD of connector and/or length/width of PCB. | Unable to assemble | unable to assemble | disassembly
Connector with PCB. Assembles onto pump plug shell. OD of connector and/or length/width of PCB. | Fluid Ingress | fluid ingress | fluid leak
Connector with PCB. Connection point to Patient Cable. OD of connector and/or key. | Unable to make connection | unable to make connection | installation-related problem
Connector with PCB. Connection point to Patient Cable. Pin layout of connector. | Unable to make connection | incorrect pin layout | physical property issue
Connector with PCB. Connection point to Patient Cable. Pin length. | Unable to make connection | pin length out of specification | installation-related problem
Connector with PCB. Creates electrical path from the patient cable through catheter wires to motor. Trace spacing. Interference leakage current which leads to motor phase short. | Interference leakage current which leads to motor phase short | motor phase short | electronic property issue
Connector with PCB. Creates electrical path from the patient cable through catheter wires to motor. Trace spacing. Interference leakage current which leads to motor phase short. | Interference leakage current which leads to motor phase short | current leakage | electrical issue
Connector with PCB. Heat generated should not exceed 40 C. Temperature. | Component fails | connector overheated | temperature issue
Connector with PCB. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Width. Over current condition/excessive heating leading to board failure due to insufficient conductivity | Over current condition/excessive heating leading to board failure due to insufficient conductivity | insufficient conductivity | failure to conduct
Connector with PCB. Provides motor, EEPROM communication and sensor access paths. Solder Pads. | Excessive heating during assembly leading to board failure (i.e. pad delamination) | pad delaminated | overheating of device or device component
Connector with PCB. Provides path and communication for EEPROM. Component Ratings in Circuit (Capacitor Voltage Value and Rating Resistor Value and Power Dissipation Semi-Conductor Component Voltage Rating Semi-Conductor Component Power Dissipation). | Component fails and circuit not functioning as intended | incorrect component rating | electrical issue
Connector with PCB. Provides path and communication for EEPROM. Power Consumption. | Component fails and could lead to AIC failure and affect patient sensor accuracy | incorrect power consumption | electrical power problem
Connector with PCB. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Type. | Loss/ Corruption of Data during use | data corrupted | data issue
Connector with PCB. Creates electrical path from the patient cable to pressure sensor. Trace Width. | Over current condition/excessive heating leading to board failure due to insufficient conductivity | trace width out of specification | failure to conduct
Connector with PCB. Creates electrical path from the patient cable to pressure sensor. Trace Clearance. | Interference leakage current which leads to sensor wire getting damaged | current leakage | electrical issue
Connector with PCB. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Clearance. | Interference which leads EEPROM communication failure | trace clearance out of specification | failure to conduct
Connector with PCB. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Size. | Unable to provide operational support | insufficient memory size | loss of data
Connector with PCB. Provides electrical path from the patient cable to the EEPROM microcontroller communication and power supply circuitry. Trace Clearance. | Interference which leads to EEPROM communication failure | electrical leakage path | electrical issue
Connector with PCB. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Microprocessor Memory Type. | Loss/ Corruption of Data during storage | fail to store data | loss of data
Connector with PCB. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Mousebite Location. | Board and/or trace damage | fail to retain data | loss of data
Connector with PCB. Houses electrical components (i.e. resistors, capacitors, ICs). Board layout. | Unable to program EEPROM | inappropriate board layout | electrical issue
Connector with PCB. Retains all the pump characteristics and operational data (i.e. gain, time, status of pump etc). Mouse bite Location. | Difficult to assemble | board damaged | installation-related problem
Connector with PCB. Houses electrical components (i.e. resistors, capacitors, ICs). PCA Board and EEPROM. | Radiated or conducted emissions from microcontroller in pump plug. | microcontroller radiated | radiofrequency interference (RFI)
Connector with PCB. Houses electrical components (i.e. resistors, capacitors, ICs). PCA Board and EEPROM. | Loss of communication between AIC and EEPROM during support due to electromagnetic interference including RFID, electrocautery. | loss of communication | electronic property issue
Repo. Provides hemostatis of arteriotomy. OD of Repo. | Excessive bleeding at site | exessive bleeding | fluid leak
Repo. Allows for repositioning of device. ID of Repo. Increased interaction forces between catheter causing difficulty for repositioning | Increased interaction forces between catheter causing difficulty for repositioning | difficult to reposition | positioning issue
Repo. Allows for repositioning of device. Surface Finish. Increased interaction forces between the repo and tissue causing damage | Increased interaction forces between the repo and tissue causing damage | interaction force increased | patient device interaction problem
Repo. Allows for repositioning of device. OD of repo. Unable to insert device successfully | Unable to insert device successfully | fail to insert | positioning issue
Repo. Allows for repositioning of device. Hub Internal Geometery. Increased pressure in valve leading to increased interaction forces between catheter causing difficulty for repositioning. | Increased pressure in valve leading to increased interaction forces between catheter causing difficulty for repositioning | interaction force increased | installation-related problem
Repo. Maintains pump position. Suture ring. | Device movement impacting positioning | unable to secure position | unintended movement
Repo. Maintains sterility of catheter. Material thickness. tear and/or break during use causing sterility to be compromised | tear and/or break during use causing sterility to be compromised | catheter broken or damaged | disinfection or sterilization issue at user location
Repo. Maintains sterility of catheter. Material. Sleeve leaks | Sleeve leaks | sleeve leakage | fluid leak
Repo. Maintains sterility of catheter. Valve sleeve Tuohy. Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | fail to secure sleeve | installation-related problem
Repo. Advance/retract along catheter. ID of repo. Increased interaction forces between catheter causing difficulty for repositioning. | Increased interaction forces between catheter causing difficulty for repositioning | difficult to advance | physical resistance
Repo. Provides hemostatis of catheter. Valve ID. Excessive bleeding between the valve and catheter. | Excessive bleeding between the valve and catheter | fail to provide hemostasis | fluid leak
Outlet Cage. Provides opening for blood flow. Window design features. | Restriction of flow (low flow condition) | inappropriate window design | restricted flowrate
Outlet Cage. Provides opening for blood flow. Window design features. Clot formation due blood stagnation zones. | Clot formation due blood stagnation zones | blood stagnated | insufficient flow or underinfusion
Outlet Cage. Provides opening for blood flow. Window design features.Increased blood damage. | Increased blood damage | blood damaged | unexpected therapeutic results
Outlet Cage. Provides opening for blood flow. Window design features. | Clot formation | clot formated | atient device interaction problem
Outlet Cage. Provides opening for blood flow. Junction between teardrop and struts. | Clot formation | clot formated | structural problem
Outlet Cage. Provide structural support for pigtail. Strut cross-section. | Kinking of struts | strut kinked | kinked
Outlet Cage. Provide junction between pigtail and cannula. M2 thread. | Weak bond | weak bond | failure to adhere or bond
Outlet Cage. Provide junction between pigtail and cannula. M2 thread. | Unable to assemble | unable to assemble | incompatibility problem
Outlet Cage. Provide junction between pigtail and cannula. ID. Reduced rigidity causing it flatten/ovalize leading to reduce flow. | Reduced rigidity causing it flatten/ovalize leading to reduce flow | cage flattened | material distortion
Outlet Cage. Provide structural support for pigtail. Material. Cross-section too small | Cross-section too small | cage out of specification | structural problem
Outlet Cage. Provide structural support for pigtail. Material. | Reduced rigidity causing it flatten/ovalize leading to reduce flow | cage flattened | structural problem
Pigtail 6F Pumpe 150. Provides support of device in PA. Tip Material. | Kinkage of pigtail causing damage to PA | piagtail kinked | kinked
Pigtail 6F Pumpe 150. Provides support of device in PA. Material. | Infection | patient infection | biocompatibility issue
Pigtail 6F Pumpe 150. Provides support of device in PA. OD. | Difficulty inserting/withdrawing device | difficult to withdraw pigtail | difficult to remove
AU Pump-Cannula. Finished pump unit which provides flow of blood and pressure sensing (without purge). Specification of bin and flow (EPROM). | Incorrect flow, reduced support | incorrect flow | improper flow or infusion
AU Pump-Cannula. Placement in the heart and valve. Length from pigtail to support ring. | Unable to sit appropriately in the heart, low flow | incorrect pigtail length | physical property issue
AU Pump-Cannula. Placement in the heart and valve. Length from pigtail to support ring. | Unable to sit appropriately in the heart, hemolysis | incorrect pigtail length | malposition of device
AU Pump-Cannula. Provides deliverability. Shape of cannula. | Unable to deliver | unable to deliver | difficult to insert
AU Pump-Cannula. Provides deliverability. Shape of cannula. Unable to sit appropriately in the heart, hemolysis | Unable to sit appropriately in the heart, hemolysis | cannula misplaced | positioning issue
AU Pump-Cannula. Placement in the heart and valve. Adhesive Epoxy Duopox locations/Transitions. | Bond Failures | adhesive applied to wrong place | failure to adhere or bond
AU Pump-Cannula. Placement in the heart and valve. Adhesive Epoxy Duopox locations/Transitions. Increase insertion/withdrawing force | Increase insertion/withdrawing force | excessive adhesive applied | installation-related problem
AU Pump-Cannula. Placement in the heart and valve. Alignment of Pigtail to catheter. | Unable to sit appropriately in the heart, low flow | pump misplaced | positioning issue
Cannula. Provides support for positioning for inlet and outlet of device. Length. | Damage heart muscle (perforation) | heart muscle perforated | device handling issue
Cannula. Provides support for positioning for inlet and outlet of device. Shape. | Damage heart muscle (perforation) | improper outlet position | malposition of device
Cannula. Provides support for positioning for inlet and outlet of device. Shape. | Inlet/outlet in incorrect position, no flow | improper inlet position | malposition of device
Sterile Connector Cable. Provides electrical signals from console to handle. Wire Gauge. | Intermittent or loss of motor power | incorrect wire gauge | overheating of device or device component
Sterile Connector Cable. Provides electrical signals from console to handle. Wire Resistance. | Intermittent or loss of motor power | incorrect wire resistance | overheating of device or device component
Sterile Connector Cable. Provides transfer of data from console to handle. Circuit Schematic. | Pump unable to proceed through case start or performs poorly | fail to transfer data | disconnection
Tyvek Lid. Provides sterile barrier. Material selection.Unable to seal properly leading to unit not being sterilized. | Unable to seal properly leading to unit not being sterilized | inadequate seal | device disinfection or sterilization issue
Epot-Tek. Providing smooth transition on critical areas. Viscosity. | Rough transition causing damage to vasculature or increase insertion force | vasculature damaged | patient device interaction problem
Epot-Tek. Providing smooth transition on critical areas. Viscosity. Unable to cover critical surfaces causing damage to vasculature or increase insertion force | Unable to cover critical surfaces causing damage to vasculature or increase insertion force | rough adhesive surface | unexpected therapeutic results
Contamination Sleeve. Maintains pump position. Tuohy Valve. | Unable to secure device into position | unable to secure device | installation-related problem
Contamination Sleeve. Maintains pump position. Tuohy Valve. | Unable to join components and secure device into position | fail to join part | installation-related problem
click adapter. Advance/retract along catheter. ID of adapter. Increased interaction forces between catheter causing difficulty for repositioning. | Increased interaction forces between catheter causing difficulty for repositioning | ID out of specification | difficult to insert
click adapter. click adapter will not rotate. Slot Width. Damge to repo sidearm. | Damge to repo sidearm | fail to rotate | device inoperable
Pressure Storage Set. filter purge fluid to remove air bubbles, and hold pressure when purge disconnected. purge reservoir. | Purge Reservoir Membrane damaged/tear | fail to hold pressure | decrease in pressure
Pressure Storage Set. filter purge fluid to remove air bubbles, and hold pressure when purge disconnected. assembly. | THF bond failure between filter and kink protector lumen | fail to hold pressure | loss of or failure to bond
Pressure Storage Set. filter purge fluid to remove air bubbles, and hold pressure when purge disconnected. assembly. | Yellow/Purple Luer Connector Broken | fail to hold pressure | loose or intermittent connection
Pressure Storage Set. filter purge fluid to remove air bubbles, and hold pressure when purge disconnected. purge filter. | Damaged/ cracked purge filter | fail to purge fluid | delivery system failure
Loctite. cover electrical connections on PCB. Viscosity. | delamination | loctite delaminated | loss of or failure to bond
Inner Blister. Properly house pump during transport. Tray Cavity (LxWxH). Unable to seal properly leading to unit not being sterilized | Unable to seal properly leading to unit not being sterilized | fail to seal properly | failure to sterilize
Inner Blister. Properly house pump during transport. Pump handle cavity (features in tray). Handle won't be secured and could breach sterile barrier | Handle won't be secured and could breach sterile barrier | fail to secure part in packaging | device or device component damaged by another device
Inner Blister. Properly house pump during transport. Pump handle cavity (features in tray). | User inconvience when removing the pump | difficult to unpackage | difficult to remove
Inner Blister. Properly house pump during transport. Patient Cable cavity and wrapping feature. | Unsecure patient cable head causing breach in sterile barrier | fail to secure pump position | manufacturing or shipping issue associated with device
Inner Blister. Properly house pump during transport. Rail for purge side arm clip. | Uncable to secure purge side arm leading to kink in purge line. | unable to secure part in packaging | packaging issue
Inner Blister. Provides container for pump for sterilization process. Tray Cavity (LxWxH). Impact ability to sterilize product completely. | Impact ability to sterilize product completely | blister packaging compromised | failure to sterilize
Inner Blister. Provides container for pump for sterilization process. Seal parameter. | Low seal strength | low seal strength | failure to sterilize
Plug Cover. Holds pump plug in place during transport. dimensions. | Does not stay in place - leading to pump damage | plug cover dislodged | dislodged or dislocated
Pump Cover. Holds pump in place during transport. dimensions. Does not stay in place - leading to pump damage | Does not stay in place - leading to pump damage | fail to secure part in packaging | device or device component damaged by another device
Packaging butterfly. Secures catheter from translating within the tray. Clipping mechanism features. | User inconvience when removing the pump | difficult to remove pump | difficult to remove
AU Pump Handle. Finished pump unit which provides flow of blood and pressure sensing (without purge). Ferrule connection of the sensor. | Sensor Failure | unable to sense | failure to sense
AU Pump Handle. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soder locations. | Failure to start pump | fail to start pump | failure to pump
AU Pump Handle. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. | Bad connections to soldering board leading to pump stop/failure to start pump | pumping stopped | pumping stopped
AU Pump Handle. Finished pump unit which provides flow of blood and pressure sensing (without purge). Specification of bin and flow (EPROM). | Incorrect flow, reduced support | pump flow reduced | pumping issue
AU Pump Handle. Provide orientation to deliver pump into body. 90 orientation of port arm to cannula bend. | Increased suction/low flow | pump misoriented | installation-related problem
AU Pump Handle. Connects the catheter to patient cable and pressure side arm. Placement of adhesive. | Bond Failure, lead to electrical failure | connection detached | disconnection
AU Pump Handle. Connects the catheter to patient cable and pressure side arm. Alignment of cannula to handle. | Not aligned (difficulty to deliver) | cannula misoriented | failure to align
AU Pump Handle. Contains electrical, fiber optic, and fluid connections. routing of fiber (> 5mm radius). | loss of signal/light | inappropriate fiberoptic routing | disassembly
AU Pump Handle. Contains electrical, fiber optic, and fluid connections. Shell internal layout. | Damage the fiber, loss of signal/light | fiber broken or damaged | break
Loctite. Bond kink protector to handle. Bond/Retention Strength. | Bond Failure | low retention strength | loss of or failure to bond
Ceramic Sleeve. separates coil from inside of motor housing. ID - D1. | Unable to assemble | unable to assemble | incompatibility problem
Ceramic Sleeve. separates coil from inside of motor housing. OD. | Coil does not fit/ coil insulation damaged | coil insulation damaged | device abrasion from instrument or another object
Ceramic Sleeve. separates coil from inside of motor housing. Material. | Unable to assemble | fail to assemble | material rigid or stiff
Ceramic Sleeve. molding boundary. ID - D1. Material. | Molding failure | molding cracked | crack
Ceramic Sleeve. interfaces proximal bearing. ID - D1. | Proximal bearing does not fit | fail to fit in sleeve | incompatibility problem
Ceramic Sleeve. interfaces proximal bearing. Length. | Wear on magnet - corrosion | sleeve worn | corrosion
Ceramic Sleeve. In contact with purge fluid. Material. | Not biocompatible | sleeve biocompatibility compromised | biocompatibility issue
Loctite. Bond kink protector to handle. Cure Time. | Bond Failure | short cure time | failure to adhere or bond
Coil. Create electromagnetic field. number of turns. Overheating of motor | Overheating of motor | incorrect coil turn | overheating of device or device component
Coil. Create electromagnetic field. number of turns. OD too large (unable to assemble) | OD too large (unable to assemble) | coil OD out of specification | difficult to insert
Coil. fits in Yoke. OD after forming. | does not fit in yoke | inappropriate fit | fitting problem
Coil. Interfaces with ceramic sleeve. ID before forming. | does not fit sleeve | inappropriate fit | fitting problem
Coil. Create electromagnetic field. Wire Material. Increase heat due to high resistance | Increase heat due to high resistance | high coil resistance | overheating of device or device component
Coil. Create electromagnetic field. Wire Diameter. oversized coil | oversized coil | coil diameter out of specification | defective component
Coil. Create electromagnetic field. Wire Diameter. increased heat | increased heat | coil temperature increased | overheating of device or device component
Coil. Create electromagnetic field. Overall Length. unable to assemble | unable to assemble | coil length out of specification | installation-related problem
Coil. Create electromagnetic field. Overall Length. | High motor current (increased heat) | motor current increased | electrical issue
Coil. Create electromagnetic field. Position on Ceramic Sleeve. Unable to assemble | Unable to assemble | coil malposition | installation-related problem
Coil. Create electromagnetic field. Coil OD. Damage due to contact with ID of yoke (leakage current) | Damage due to contact with ID of yoke (leakage current) | leakage current out of specification | electrical issue
Coil. Create electromagnetic field. Wire Bonding Material. unable to wind coil (size too large) | unable to wind coil (size too large) | unable to wind coil | failure to fold
Coil. Create electromagnetic field. Gap between lead wires at exit. high motor current - higher heat | high motor current - higher heat | coil temperature increased | electrical issue
Coil. Create electromagnetic field. Gap between lead wires at exit. | no effect | coil efficiency reduced | output issue
Yoke. Holds coil/sleeve subassembly for molding and defines molding boundary. ID - D3. | Coil does not fit | fail to hold part | installation-related problem
Yoke. Holds coil/sleeve subassembly for molding and defines molding boundary. ID - D3. | Pump stop/coil short (insulation damaged) | coil short | electrical shorting
Yoke. Interfaces with hub. Distal OD - D2. | unable to assemble | unable to assemble | installation-related problem
Yoke. Interfaces with hub. Length of overlap with hub - L2. Magnet wear against proximal bearing - purge lumen blocked | Magnet wear against proximal bearing - purge lumen blocked | magnet worn out | degraded
Yoke. Interfaces with hub. Length of overlap with hub - L2. | gap at bond | excess gap | failure to adhere or bond
Yoke. interfaces with catheter. OD of Bonding Surface - D5. | insufficient bond strength | insufficient bond strength | loss of or failure to bond
Yoke. interfaces with catheter. OD of Bonding Surface - D5. | blood ingress into catheter | interface leakage | fluid leak
Yoke. interfaces with catheter. texture of bonding surface. | insufficient tensile strength | insufficient tensile strength | failure to adhere or bond
Yoke. Blood contact. surface finish. | hemolysis | patient hemolysis | patient device interaction problem
Yoke. Blood contact. surface finish. | corrosion | yoke corroded | corrosion
Yoke. contains magnetic field. Material. | corrosion | magnet corroded | corrosion
Yoke. allows for routing of motor cable. inside geometry. | damage to motor cable insulation | insulation broken or damaged | electrical shorting
Magnet. Provide power to the Impella. OD of magnet. | Unable to assemble or fails motor current or purge flow test | unable to assemble | defective component
Magnet. Provide power to the Impella. Magnet Coating. | Corrosion of magnet leads to pump stop | magnet corrosion | corrosion
Magnet. Provide power to the Impella. Length. | Unable to assemble or fails motor current | unable to assemble | physical property issue
Magnet. Provide power to the Impella. ID. | Fails motor current test or unable to assemble | fail to pass current test | failure of device to self-test
Magnet. Provide power to the Impella. Runout between ID and OD. | Fails motor current test or unable to assemble | fail to pass current test | failure of device to self-test
Impeller. Interfaces with bearings. Length. Rotor Dynamic Issues (Collides with purge lumen) | Rotor Dynamic Issues (Collides with purge lumen) | rotor collided | decreased pump speed
Impeller. Hold magnet and washer . OD of Shaft. | washer and bearing wear | washer worn | device abrasion from instrument or another object
Impeller. Hold magnet and washer . OD of Shaft. Rotor Dynamic Issues | Rotor Dynamic Issues | unstable rotor dynamics | unintended movement
Impeller. Hold magnet and washer . OD of Shaft. Glue migration to washer surface. | Glue migration to washer surface | glue migrated | sticking
Impeller. Interfaces with bearings. OD of Shaft. | Shaft breaks | shaft broken or damaged | break
Handle, Bottom. contains PCB with sensor. dimensions. | unable to assemble | unable to assemble | physical property issue
AU Pump Sensor. Placement in the heart and valve. Max OD of Assembly (with sensor visor installed). | Difficulty inserting/withdrawing device | difficult to withdraw sensor | difficult to remove
AU Pump Cannula. Placement in the heart and valve. Delo locations/Transitions. | Bond Failures | bond strength compromised | loss of or failure to bond
AU Pump Cannula. Placement in the heart and valve. Delo locations/Transitions. | Increase insertion/withdrawing force | difficult to withdraw cannula | retraction problem
AU Pump Cannula. Placement in the heart and valve. Alignment of Cannula to sensor. | Sensor is pushed against wall of ventricle | sensor against ventricle wall | malposition of device
AU Pump Cannula. Placement in the heart and valve. Length from inflow cage to endcap. Unable to sit appropriately in the heart, low flow | Unable to sit appropriately in the heart, low flow | pump misplaced | malposition of device
AU Inflow Cage – Cannula. provides inlet for blood flow. Bond of cannula to inlet. | weak bond | weak bond | failure to adhere or bond
Cannula (inner layer). Transport blood from inlet to outlet. Wall thickness. Cannula is too stiff | Cannula is too stiff | cannula stiffed | material rigid or stiff
Cannula (inner layer). Transport blood from inlet to outlet. Wall thickness. OD is too large | OD is too large | OD out of specification | physical property issue
Cannula (inner layer). Transport blood from inlet to outlet. Wall thickness. | hole in canula | hole in cannula | hole in material
Cannula (inner layer). Transport blood from inlet to outlet. Wall thickness. | Cannula breaks | cannula broken or damaged | break
Cannula (inner layer). Transport blood from inlet to outlet. Material. | incorrect material chosen | incorrect material selected | material integrity issue
Cannula outer layer. Transport blood from inlet to outlet. Wall thickness. nitinol pokes through | nitinol pokes through | cannula poked | hole in material
Cannula outer layer. Transport blood from inlet to outlet. Wall thickness. | OD is too large | OD out of specification | physical property issue
Cannula Coil. Transport blood from inlet to outlet. rigidity. column strength. material. | Cannula Kinks | cannula kinked | kinked
Optical Sensor, Pump. Fit over sensor fiber. Nitinol Tube ID. | Does not fit over sensor fiber | inappropriate fit | fitting problem
Optical Sensor, Pump. fit in yoke slot. Nitinol Tube OD. | Does not fit in yoke slot | fail to fit in sensor | component or accessory incompatibility
Optical Sensor, Pump. protect sensor fiber. Nitinol Tube Material properties. | Does not protect sensor fiber | fail to protect fiber | break
Optical Sensor, Pump. protect sensor fiber. Nitinol Tube Length. | Fiber break - unable to pass final qual | fiber broken or damaged | break
Optical Sensor, Pump. protect sensor fiber. buffer length. | unable to assemble | unable to assemble | installation-related problem
Optical Sensor, Pump. fit in catheter. Nitinol Tube OD. | Does not fit in catheter | fail to fit in catheter | fitting problem
Optical Sensor, Pump. fit in loading lumen. Nitinol Tube OD. | does not fit in loading lumen | fail to fit in lumen | incompatibility problem
Optical Sensor, Pump. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | incorrect measurement
Optical Sensor, Pump. Determining position of pump within the body. Depth of the cavity. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to provide pressure reading | unable to obtain readings
Optical Sensor, Pump. Determining position of pump within the body. Sensor Material. | Readings from sensor incorrect | incorrect sensor reading | incorrect measurement
Optical Sensor, Pump. Biocompatible with body. Sensor Material. | Not biocompatible with body | sensor material biocompatibility compromised | biocompatibility issue
Optical Sensor, Pump. Biocompatible with body. Silicone gel material. | Not biocompatible with body | silicone gel biocompatibility compromised | biocompatibility issue
Optical Sensor, Pump. Biocompatible with body. Silicone gel material. Sensor drifts beyond specification (less accurate, issues with placement) | Sensor drifts beyond specification (less accurate, issues with placement) | sensor drifted | migration of device or device component
Optical Sensor, Pump. Determining position of pump within the body. Diaphragm dimensions. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor reading drifted | imprecision
Optical Sensor, Pump. Determining position of pump within the body. Silicone gel. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | imprecision
Distal Bearing. Holding of Impeller radially. ID of Bearing - D1. | Unstable rotor dynamics | unstable rotor dynamics | unintended system motion
Distal Bearing. Holding of Impeller radially. ID of Bearing - D1. | Shaft wear | shalft worn out | degraded
Distal Bearing. Holding of Impeller radially. ID of Bearing - D1. | Shaft or bearing wear | bearing worn out | naturally worn
Distal Bearing. Facillitate Purge Flow. Bearing thickness - L1. | Blood ingress into motor | blood ingress into motor | bleed back
Distal Bearing. Facillitate Purge Flow. Bearing thickness - L1. | High purge pressure | high purge pressure | increase in pressure
Distal Bearing. Defines purge gap. ID of Bearing - D1. | No and/or insufficient purge delivery | insufficient purge | insufficient flow or underinfusion
Distal Bearing. Defines purge gap. ID of Bearing - D1. | Increased purge flow to patient | purge flow increased | excess flow or over-infusion
Distal Bearing. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Bearing thickness - L1. Not enough lubrication allowed on distal bearing and thrust washer leading to wear on rotor, impeding Rotor Rotation. | Not enough lubrication allowed on distal bearing and thrust washer leading to wear on rotor, impeding Rotor Rotation | rotor rotation impeded | physical resistance
Distal Bearing. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Bearing thickness - L1. Excessive magnet preload leading to wear on washer and bearing | Excessive magnet preload leading to wear on washer and bearing | excessive magnet preloaded | output above specifications
Distal Bearing. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Bearing thickness - L1.Magnet preload reduced leading to impeller collapse on bearing | Magnet preload reduced leading to impeller collapse on bearing | pump collapsed | collapse
Distal Bearing. Seals purge solution at the hub interface. OD and adjacent surface. | Leakage of purge | purge leakage | loss of or failure to bond
Distal Bearing. Seals purge solution at the hub interface. OD and adjacent surface. Limiting amount of glue in area to hold unit together affecting bond strengh. | Limiting amount of glue in area to hold unit together affecting bond strengh | inadequate glue | loss of or failure to bond
Distal Bearing. Seals purge solution at the hub interface. Bearing thickness - L1. Rotor Assembly dislodges from motor | Rotor Assembly dislodges from motor | rotor dislodged | pumping stopped
Distal Bearing. Maintain axial load of rotor assembly under all axial loads. Bearing thickness - L1. Inadequate surface area for bond which affects bond strength | Inadequate surface area for bond which affects bond strength | inadequate bond surface | failure to adhere or bond
Distal Bearing. Maintain axial load of rotor assembly under all axial loads. OD of Bearing - D2. Inadequate space for epoxy to adhere affecting bond strength | Inadequate space for epoxy to adhere affecting bond strength | inadequate bond surface | loss of or failure to bond
Distal Bearing. Maintain axial load of rotor assembly under all axial loads. OD of Bearing - D2. | Excessive delo required to hold load affecting bond strength | fail to maintain axial load | structural problem
Distal Bearing. Interfaces with Magnet. Length - L1. Stiff rotor - motor current high | Stiff rotor - motor current high | rotor stiffed | material rigid or stiff
Distal Bearing. Interfaces with Magnet. Length - L1. | unable to assemble | fail to assemble | installation-related problem
Distal Bearing. Interfaces with Magnet. Length - L1. | Unstable rotor dynamics | unstable rotor dynamics | unintended system motion
Proximal Bearing. Holding of Impeller radially. ID of bearing - D1. | Rotor Dynamic Issues | unstable rotor dynamics | structural problem
Proximal Bearing. Holding of Impeller radially. ID of bearing - D1. | Wear on shaft or bearing leading to high motor current | shaft worn out | degraded
Proximal Bearing. Holding of Impeller radially. ID of bearing - D1. Wear on shaft or bearing leading to high motor current. | Unable to assemble | motor current increased | pumping issue
Proximal Bearing. Interface with purge lumen. ID of bearing - D1. | Unable to assemble | unable to assemble | installation-related problem
Proximal Bearing. Interface with purge lumen. ID of bearing - D1. Purge leak towards PCB | Purge leak towards PCB | lumen leakage | electrical shorting
Proximal Bearing. Interface with purge lumen. ID of bearing - D1. | Loss of purge pressure | loss of purge pressure | decrease in pressure
Proximal Bearing. Interface with ceramic sleeve. intermediate OD - D2. | Unable to assemble | difficult to assemble | installation-related problem
Proximal Bearing. Interface with magnet. L1. | Wear on magnet - cause corrosion | magnet worn out | corrosion
Proximal Bearing. Holding of Impeller radially. L1. | Unstable rotor dynamics | fail to provide radial support | structural problem
PCB. Connect coil leads to motor cable. Pad surface area. | Broken/lack electrical connection | electrical connection damaged | intermittent continuity
PCB. Connect coil leads to motor cable. Distance between pads. | Electrical short | pad clearance short | electrical shorting
PCB. Stick to proximal ceramic bearing. Glue Type. | Connection between leads and motor cable lost | insufficient glue | loss of or failure to bond
Inflow Cage. Allow blood to enter cannula. window radii. | thinner strut - bending during introducer interaction | thin strut | mechanical issue
Inflow Cage. Allow blood to enter cannula. window radii. | blood turbulence | blood turbulence | patient device interaction problem
Inflow Cage. Bond to Cannula. OD of Bonding Surface - D2. | insufficient tensile strength | low tensile strength | mechanical issue
Repositioning Unit. Allows for repositioning of device. ID of Valve. | Increased interaction forces between catheter causing difficulty for repositioning | difficult to reposition | physical resistance
Repositioning Unit. Allows for repositioning of device. Hub Internal Geometery. | Increased pressure in valve leading to increased interaction forces between catheter causing difficulty for repositioning | fail to reposition | physical property issue
Repositioning Unit. Provides hemostasis of catheter. Valve OD. | Excessive bleeding between the valve and catheter | exessive bleeding | free or unrestricted flow
Repositioning Unit. Maintains hemostasis with the butterfly. butterfly ID. Excessive bleeding between butterfly and sheath | Excessive bleeding between butterfly and sheath | excessive bleeding | improper or incorrect procedure or method
Repositioning Unit. Maintains hemostasis with the butterfly. Sheath OD. | Excessive bleeding between butterfly and sheath | fail to maintain hemostasis | defective component
Repositioning Unit. Transition catheter from graft to skin. butterfly material. | catheter exits at higher angle | device transitted inappropriately | device handling issue
Repositioning Unit. Transition catheter from graft to skin. sheath material. | catheter exits at higher angle | catheter transited inappropriately | installation-related problem
Repositioning Unit. Maintains pump position. Hub external Geometry. | Butterfly separates from hub | part separated | detachment of device component
Repositioning Unit. Maintains pump position. Butterfly internal geometry. | Butterfly separates from hub | part separated | material separation
Repositioning Unit. Maintains pump position. Sheath OD. | Butterfly separates from hub | fail to retain position | malposition of device
Purge Lumen. deliver purge solution through the catheter to the motor. length. | kinks in handle | lumen kinked | material twisted
Purge Lumen. deliver purge solution through the catheter to the motor. ID. | insufficient glue - bond strength comprimised | bond strength compromised | loss of or failure to bond
Motor Cable. deliver power to the motor. diameter. | unable to assemble | fail to receive power | pumping stopped
Motor Cable. deliver power to the motor. insulation. | pump stop | insulation broken or damaged | pumping stopped
Motor Cable. deliver power to the motor. stripped length. | unable to assemble | inappropriate cable stripped length | installation-related problem
Motor Cable. Interface with other components in the catheter. length of shielding cut vs. length of outer insulation stripping (Note: desire for no shielding exposed). | Damage to purge lumen | shielding exposed | installation-related problem
Motor Cable. Interface with other components in the catheter. length of shielding cut vs. length of outer insulation stripping (Note: desire for no shielding exposed). | Damage to purge lumen | inappropriate insulation stripped | electrical shorting
Motor Cable. Interface with other components in the catheter. length of shielding cut vs. length of outer insulation stripping (Note: desire for no shielding exposed). | damage to optical fiber | inappropriate shielding cut | installation-related problem
Hub. Connect to Yoke. Larger ID -D2. | unable to assemble | unable to assemble | incompatibility problem
Hub. Connect to Yoke. Larger ID -D2. insufficient glue - bond strength comprimised. | insufficient glue - bond strength comprimised | insufficient glue | loss of or failure to bond
Hub. weld to pump housing struts. material selection. | unable to weld | unable to weld | solder joint failure
Hub. Connect to pump housing struts. material selection. welding grooves (angle). | weak weld | weak weld | solder joint failure
Hub. weld to pump housing struts. material selection. | corrosion around weld | weld corroded | corrosion
Hub. Connect to pump housing struts. welding grooves (width) - L2. | Poor weld strength | inadequate weld strength | mechanical issue
Hub. Blood contact. surface finish. | hemolysis | patient hemolysis | patient device interaction problem
Hub. Blood contact. material/processing selection. | corrosion | hub corroded | corrosion
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). ID - D1. Angular Alignment of washer affected leading to washer wear (L/D) | Angular Alignment of washer affected leading to washer wear (L/D) | washer worn | mechanical issue
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). ID - D1. Glue could migrate to functional bearing surface (Rotors glue together, No purge fluid delivery, fluid bearing surface contaminated) | Glue could migrate to functional bearing surface (Rotors glue together, No purge fluid delivery, fluid bearing surface contaminated) | bearing surface contaminated | contamination during use
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Thickness of washer - L1. Wear of washer to the magnet surface/lose of fluid bearing surface. | Wear of washer to the magnet surface/lose of fluid bearing surface | loss of bearing surface | degraded
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Thickness of washer - L1. Excessive magnet preload leading to wear on washer and bearing | Excessive magnet preload leading to wear on washer and bearing | bearing overloaded | mechanical issue
Thrust Washer. Prevent delo migration through ID. ID - D1. Delo migrating into bearing ID | Delo migrating into bearing ID | adhesive migrated | migration of device or device component
Thrust Washer. Interface with Impeller Shaft. ID - D1. Unable to assemble | Unable to assemble | difficult to assemble | installation-related problem
Epotek 301. fill hub and yoke - combine coil and ceramic sleeve. mechanical properties. | Epotek does not adhere to internal components | fail to adhere component | failure to adhere or bond
Epotek 301. fill hub and yoke - provide thermal conductivity. material properties - viscosity. | high temperature of pump | inappropriate viscosity level | electrical issue
Epotek 301. fill hub and yoke - provide thermal conductivity. material properties - viscosity. | high temperature of pump | low thermal conductivity | overheating of device or device component
Coil Reinforced Catheter. interface with yoke. Distal OD of assembly. | insufficient tensile strength | insufficient tensile strength | mechanical issue
Coil Reinforced Catheter. interface with yoke. Distal OD of assembly. | unable to assemble | unable to assemble | installation-related problem
Coil Reinforced Catheter. Allow for lumens to pass through. Overall length. | unable to assemble | unable to pass through | installation-related problem
Coil Reinforced Catheter. Allow for lumens to pass through. Proximal ID of coil. | unable to assemble | unable to pass through | physical property issue
AU Rotor Pump Housing. Connect Motor to pump housing. Materials. | Weak weld | weld corroded | solder joint failure
AU Rotor Pump Housing. Connect Motor to pump housing. Materials. | Coroded weld | weld corroded | corrosion
AU Rotor Pump Housing. allow blood flow out of impella. area underneath strut. | thrombosis | thrombus formated | improper flow or infusion
Sheath. Provide concentricity control to purge lumen. ID/OD/material properties. | Unable to assemble | fail to provide concentricity control | failure to advance
Pump Housing. allow blood flow out of impella. Chamfer around window. | hemolysis | patient hemolysis | improper flow or infusion
Pump Housing. allow blood flow out of impella. Strut size. | bent pump housing | housing bent | bent
Pump Housing. allow blood flow out of impella. surface finish. | hemolysis | patient hemolysis | unexpected therapeutic results
Pump Housing. allow blood flow out of impella. ID. | loss of flow | loss of flow | no flow
Pump Housing. allow blood flow out of impella. Wall Thickness. | weak wall | housing wall too thin | physical property issue
Sensor Visor. Protect sensor face. ID. structural integrity of visor comprimised - sensor damaged | structural integrity of visor comprimised - sensor damaged | fail to protect sensor | mechanical issue
Sensor Visor. Protect sensor face. OD (D1). will not fit in pump slot | will not fit in pump slot | unable to assemble | incompatibility problem
Sensor Visor. Protect sensor face. surface finish. | thrombus formation | thrombus formated | unexpected therapeutic results
Sensor Visor. Protect sensor face. Length. gap to pump housing- thrombus formation | gap to pump housing- thrombus formation | excessive gap | unexpected therapeutic results
Sensor Visor. Protect sensor face. Length between visor "hat" start and bottom of visor (L1). will not protect sensor face - sensor damaged | will not protect sensor face - sensor damaged | sensor damaged | break
AU, Sleeve Proximal Bearing. Provide inner boundary for molding of motor housing. Seal of bearing to sleeve. | Stiff rotor | fail to provide inner boundary | structural problem
AU, Sleeve Proximal Bearing. Provide inner boundary for molding of motor housing. Seal of bearing to sleeve. | excess leakage current | excess leakage current | electrical issue
AU, Sleeve Proximal Bearing. provide alignment sleeve and magnet. alignment of sleeve to proximal bearing. | Stiff rotor | rotor stiffed | material rigid or stiff
AU, Sleeve Proximal Bearing. provide alignment sleeve and magnet. alignment of sleeve to proximal bearing. | excess vibration | excess vibration | vibration
AU, Sleeve Proximal Bearing. provide alignment sleeve and magnet. alignment of sleeve to proximal bearing. | high motor current | fail to align | installation-related problem
AU, Coil Sleeve. Secure coil to ceramic sleeve. Position on Ceramic Sleeve. | Reduce efficiency leading to increased motor current | fail to secure sleeve | installation-related problem
AU, Coil Sleeve. Secure coil to ceramic sleeve. OD of assembly. Damage due to contact with ID of yoke (leakage current). | Damage due to contact with ID of yoke (leakage current) | sleeve broken or damaged | device abrasion from instrument or another object
AU, Coil Sleeve. Secure coil to ceramic sleeve. OD of assembly. Damage to insulation leading to coil shorts during forming process | Damage to insulation leading to coil shorts during forming process | coil insulation damaged | electrical shorting
AU, Coil Sleeve. Secure coil to ceramic sleeve. Length of coil on sleeve. | High motor current (increased heat) | coil overheated | electrical issue
AU, Stator. connect Motor cable to coil via PCB. Soldering pattern. | spins in incorrect direction - hemolysis | stator spined in wrong direction | device operates differently than expected
AU, Stator. connect Motor cable to coil via PCB. tightness of purge lumen to proximal bearing. coil short or leakage current | coil short or leakage current | coil short | defective component
AU, Stator. connect Motor cable to coil via PCB. tightness of purge lumen to proximal bearing. coil short or leakage current | coil short or leakage current | leakage current out of specification | electrical issue
AU, Stator. connect Motor cable to coil via PCB. tightness of purge lumen to yoke. epotek leaks out of yoke during molding | epotek leaks out of yoke during molding | adhesive leakage | misassembled
AU Motor Housing. Allow gap between magnet and ceramic sleeve. concentricity. | comprimise magnet gap - pump stop | magnet gap compromised | unintended magnet quench
AU Motor Housing. Allow gap between magnet and ceramic sleeve. concentricity. | Stiff rotor or high motor current | rotor stiffed | material rigid or stiff
AU Motor Housing. Allow gap between magnet and ceramic sleeve. concentricity. | vibration | motor vibrated | vibration
Rotor Assembly. Pump blood. gap between purge lumen and shaft. | purge blocked - high purge pressure | purge pressure high | increase in pressure
Rotor Assembly. Pump blood. gap between magnet and proximal bearing. | high motor current | excessive gap | installation-related problem
Rotor Assembly. Pump blood. shaft and proximal bearing engagement. corrosion of proximal end of magnet. | corrosion of proximal end of magnet | magnet corrosion | corrosion
Rotor Assembly. Pump blood. gap between rotor and pump housing. | eccentricity | rotor eccentricity | malposition of device
AU, Pump Catheter Motor. Connect hub to motor housing. bond between hub and motor housing. | low tensile strength | inadequate bond tensile strength | failure to adhere or bond
AU, Pump Catheter Motor. Connect hub to motor housing. Concentricity. | eccentricity | motor eccentricity | failure to align
Contamination Sleeve. Maintains sterility of catheter. Material thickness of sterile sleeve. tear and/or break during use causing sterility to be compromised. | tear and/or break during use causing sterility to be compromised | sleeve torn | bacterial contamination of device
Contamination Sleeve. Maintains sterility of catheter. Valve sleeve Tuohy. Unable to secure sterile sleeve leading to exposed catheter and compromised sterility. | Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | unable to secure | bacterial contamination of device
Sterile Incision Template. Sterile tool. Sterility. | Does not sterilize | failure to sterilize | delivered as unsterile product
Sterile Incision Template. provide measurement for direct insertion incision. Length of ruler. | provides inaccurate measurement | inadeqaute measurement | measurement system incompatibility
Sterile Incision Template. Biocompatible with body. biocompatibility. | material not biocompatible | material biocompatibility compromised | biocompatibility issue
PSS Retainer Top, PSS Retainer Bottom. Provide protective housing for Pressure Storage Set. Internal dimensional features. | Unable to assemble | unable to assemble | defective component
PSS Retainer Top, PSS Retainer Bottom. Provide protective housing for Pressure Storage Set. Material. | Warpage of material leading to unable to assemble | fail to assemble | material distortion
PSS Retainer Top, PSS Retainer Bottom. Provide protective housing for Pressure Storage Set. Material. | Crack/damaged | set cracked | crack
PSS Retainer Top, PSS Retainer Bottom. Provide protective housing for Pressure Storage Set. External geometry. | Damage to gloves or clinicians hand | clinician hand injured | material puncture
Glass Cylinder Housing. Holding the glass cylinder. ID. | Unable to assemble | unable to assemble | installation-related problem
Glass Cylinder Housing. Holding the glass cylinder. ID. Off-axis assembly leading to high force on piston (purge system blockage) | Off-axis assembly leading to high force on piston (purge system blockage) | off-axis assembly | misassembled
Glass Cylinder Housing. Holding the glass cylinder. ID Ribs. Off-axis assembly leading to high force on piston (purge system blockage) | Off-axis assembly leading to high force on piston (purge system blockage) | off-axis assembly | installation-related problem
Glass Cylinder Housing. Connection of the suction line to housing. ID. | Weak bond leading to leak | connection leakage | loss of or failure to bond
Glass Cylinder Housing. Connection of the suction line to housing. ID. Tubing inserted to far blocking flow path (purge system blockage) | Tubing inserted to far blocking flow path (purge system blockage) | flow blocked | installation-related problem
Glass Cylinder Housing. Interface with cartidge cap. Alignment Hole ID. Shifting of cartridge cap leading to blockage of flow path (purge system blockage) | Shifting of cartridge cap leading to blockage of flow path (purge system blockage) | flow blocked | misassembled
Glass Cylinder Housing. Interface with cartidge cap. Alignment Location. | Unable to function in AIC | inappropriate interface design | accessory incompatible
Glass Cylinder Housing. Interface with cartidge cap. Alignment Location. | Shifting of cartridge cap leading to blockage of flow path (purge system blockage) | component shifted | installation-related problem
Glass Cylinder. Interface with Piston and houses fluid for pumping. ID. | Leakage of system | interface leakage | fluid leak
Glass Cylinder. Interface with Piston and houses fluid for pumping. ID. | High force leading to blockage of system | excessive force applied | installation-related problem
Glass Cylinder. Interface with Piston and houses fluid for pumping. Surface Quality. | Leakage of system | rough surface | scratched material
Glass Cylinder. Interface with Piston and houses fluid for pumping. Surface Quality. | High force leading to blockage of system | interface blocked | blocked connection
Glass Cylinder. Interface with Piston and houses fluid for pumping. Surface Quality. | Unable to seal properly leading to leakage of system | fail to seal properly | leak
Glass Cylinder. Interface with Piston and houses fluid for pumping. Length. | Piston comes out of assembly leading to leakage | interface separated | detachment of device component
Rack, Purge Cassette. Mates with drive shaft and piston, then supports the drive of the piston to move fluid through the purge cassette. Rack Teeth Features. | unable to drive piston properly leading to inadequate purge flow | fail to drive fluid | no flow
Rack, Purge Cassette. Mates with drive shaft and piston, then supports the drive of the piston to move fluid through the purge cassette. OD of Piston Mating Feature (D1). | Unable to assemble | unable to assemble | installation-related problem
Rack, Purge Cassette. Mates with drive shaft and piston, then supports the drive of the piston to move fluid through the purge cassette. OD of Piston Mating Feature (D1). | Piston falls off of assembly due improper engagement | piston disengaged | detachment of device component
Rack, Purge Cassette. Mates with drive shaft and piston, then supports the drive of the piston to move fluid through the purge cassette. Split pin hole ID (D3). | Piston falls off of assembly due improper engagement | piston falling off | component falling
Piston. Drive fluid and create seal with cylinder. Overall OD (D6). | Unable to assemble | unable to assemble | defective component
Piston. Drive fluid and create seal with cylinder. Overall OD (D6). Poor seal leading leakage of system. | Poor seal leading leakage of system | piston leakage | leak
Piston. Drive fluid and create seal with cylinder. Overall OD (D6). Poor seal leading leakage of system. | High force leading to unable to drive purge fluid properly causing inadequate purge flow | fail to create seal | fluid leak
Piston. Drive fluid and create seal with cylinder. Overall OD Cavity. particulates break free causing blockage/occulustion of system | particulates break free causing blockage/occulustion of system | piston broken or damaged | particulates
Piston. Drive fluid and create seal with cylinder. Material. | non-biocompatible | piston biocompatibility compromised | biocompatibility issue
Piston. Mates with Rack. ID of Piston. Piston falls off of assembly due improper engagement | Piston falls off of assembly due improper engagement | piston falling off | component falling
Piston. Mates with Rack. Split pin hole ID (D1). Piston falls off of assembly due improper engagement. | Piston falls off of assembly due improper engagement | piston disengaged | detachment of device component
Split Pin. Retains piston to rack. OD. | Unable to assemble | unable to assemble | incompatibility problem
Split Pin. Retains piston to rack. OD. | Pin becomes unseated and leads to deflection of piston causing leak | pin not seated properly | unstable
Split Pin. Retains piston to rack. Length. | interaction with cylinder/damage cylinder leading to leaking | pin broken or damaged | break
Split Pin. Retains piston to rack. Length. | Pin becomes unseated and leads to deflection of piston causing leak | pin unseated | installation-related problem
Split Pin. Retains piston to rack. Thickness. | Unable to assemble | fail to retain pin | installation-related problem
Gear Box. Houses drive shaft rack. Internal geometry (D1-D5, L3). | Unable to assemble | unable to assemble | installation-related problem
Gear Box. Facilitates interaction between drive shaft and rack and the rest of the assembly. Length (L1). | Misalignment with cylinder and rest of assemble causing high forces | drive shaft misaligned | failure to align
Gear Box. Facilitates interaction between drive shaft and rack and the rest of the assembly. Length (L2). insufficient clearance between drive shaft and rack (device binds up) | insufficient clearance between drive shaft and rack (device binds up) | insufficient clearance | installation-related problem
Gear Box. Facilitates interaction between drive shaft and rack and the rest of the assembly. Length (L2). | Gap between teeth leading to unable to actuate device | insufficient clearance | physical resistance
Drive Shaft. Interfaces with AIC stepper motor. Geometry of Teeth. Improper mating with the stepper motor causing binding/unable to actuate device. | Improper mating with the stepper motor causing binding/unable to actuate device | incompatible shaft | accessory incompatible
Drive Shaft. Interfaces with AIC stepper motor. Gear mate depth. | Unable seat properly in the AIC | shaft not seated properly | failure to align
Drive Shaft. Drives Rack and Piston. Geometry of the gear feature. | Binding with Rack/unable to actuate rack | unable to actuate rack | failure to deploy
Cylinder Pin. Retain gear box to purge cassette cartridge bottom. OD. | Unable to assemble | fail to retain | detachment of device component
Cylinder Pin. Retain gear box to purge cassette cartridge bottom. Length. | Unable to assemble | unable to assemble | installation-related problem
Cylinder Pin. Retain gear box to purge cassette cartridge bottom. Length. | Poor actuation of device | fail to actuate device | failure to deploy
RFID Transponder. Housing data for interaction between purge cassette and console. Integrity of Chip. Purge cassette and console unable to communicate with each other. | Purge cassette and console unable to communicate with each other | loss of communication | wireless communication issue
RFID Transponder. Housing data for interaction between purge cassette and console. Memory/Data Structure. | Unable to program chip properly leading to the device unable to communicate with console | unable to program | programming issue
RFID Transponder. Housing data for interaction between purge cassette and console. Length. | Unable to assemble | unable to assemble | incompatibility problem
Cartridge Top and Cartridge Bottom. Houses purge cassette components. Internal geometry. | Unable to assemble | unable to assemble | installation-related problem
Cartridge Top and Cartridge Bottom. Houses purge cassette components. Internal geometry. | Purge cassette does not provide support | fail to protect | protective measure issue
Cartridge Top and Cartridge Bottom. Houses purge cassette components. Internal geometry. | non-biocompatible | non-biocompatible cartridge applied | biocompatibility issue
Cartridge Top and Cartridge Bottom. Houses purge cassette components. Exterior grooves. | Unable to mate with AIC | incompatible cartridge | incompatibility problem
Cartridge Top and Cartridge Bottom. Houses purge cassette components. Exterior grooves. unable to utilize device. | unable to utilize device | fail to load | failure to service
Suction Line. Deliver purge fluid. Length. Kink causing occlusion of purge flow. | Kink causing occlusion of purge flow | line kinked | no flow
Suction Line. Deliver purge fluid. ID. inadquate flow of purge fluied | inadquate flow of purge fluied | fail to deliver fluid | failure to deliver
Suction Line. Deliver purge fluid. Material. insufficient pressure leading to low purge flow | insufficient pressure leading to low purge flow | insufficient pressure | decrease in pressure
Suction Line. Deliver purge fluid. Material. insufficient pressure leading to low purge flow | insufficient pressure leading to low purge flow | insufficient pressure | insufficient flow or underinfusion
Suction Line. Deliver purge fluid. Surface Quality. reduced purge flow due to obstruction | reduced purge flow due to obstruction | fluid flow reduced | obstruction within device
Suction Line. Deliver purge fluid. Surface Quality. occlusion/leakage of the system (reduced purge flow) | occlusion/leakage of the system (reduced purge flow) | line leakage | fluid leak
Luer Lock Connector. Allow connection of the suction line to the fluid bag. ID. | Bond failure leading to leaking of purge fluid | fluid leakage | fluid leak
Luer Lock Connector. Allow connection of the suction line to the fluid bag. Connection Geometry. | Unable to connect to fluid bag | unable to connect | component or accessory incompatibility
Inflow Valve. Control purge fluid flow on suction line. Valve Geometry. | Incorrect flow leading to purge flow issues | incorrect flow | improper flow or infusion
Inflow Valve. Control purge fluid flow on suction line. Valve Geometry. | retrograde flow causing backup of purge line | retrograde flow | improper flow or infusion
Inflow Valve. Control purge fluid flow on suction line. Valve Geometry. | air in purge system causing occlusions | air in purge system | occlusion within device
Inflow Valve. Control purge fluid flow on suction line. Burst pressure. | break of component leading to leak | valve broken or damaged | fluid leak
Inflow Valve. Control purge fluid flow on suction line. Inner geometry. | retrograde flow causing backup of purge line | back flow | improper flow or infusion
Outflow Valve. Control purge fluid flow on pressure line. Valve Geometry. | Unable to assemble | unable to assemble | incompatibility problem
Outflow Valve. Control purge fluid flow on pressure line. Valve Geometry. | Incorrect flow leading to purge flow issues | incorrect flow | improper flow or infusion
Outflow Valve. Control purge fluid flow on pressure line. Valve Geometry. | retrograde flow causing backup of purge line | retrograde flow | improper flow or infusion
Outflow Valve. Control purge fluid flow on pressure line. Back pressure . | break of component leading to leak | valve broken or damaged | fluid leak
Outflow Valve. Control purge fluid flow on pressure line. Inner geometry. | air in purge system causing occlusions | air in valve system | occlusion within device
Pressure Line. Deliver purge fluid. Length. | Kink causing occlusion of purge flow | line kinked | obstruction within device
Pressure Line. Deliver purge fluid. ID. | inadquate flow of purge fluied | inadequate flow | insufficient flow or underinfusion
Pressure Line. Deliver purge fluid. Material. | insufficient pressure leading to low purge flow | insufficient pressure | decrease in pressure
Pressure Line. Deliver purge fluid. Surface Quality. | reduced purge flow due to obstruction | fluid flow reduced | obstruction within device
Pressure Line. Deliver purge fluid. Surface Quality. | occlusion/leakage of the system (reduced purge flow) | line leakage | fluid leak
Pressure Line. Deliver purge fluid. Surface Quality. | Bond failure leading to leaking of purge fluid | fail to deliver fluid | failure to deliver
Spike Line. Deliver purge fluid. Length. | Kink causing occlusion of purge flow | line kinked | no flow
Spike Line. Deliver purge fluid. ID. | inadquate flow of purge fluied | fluid flow reduced | insufficient flow or underinfusion
Spike Line. Deliver purge fluid. Surface Quality. | Bond failure leading to leaking of purge fluid | insufficient bond engagement | failure to adhere or bond
Spike Line. Deliver purge fluid. Surface Quality. | occlusion/leakage of the system (reduced purge flow) | insufficient bond engagement | fluid leak
Pressure Transmitter. Transmit purge pressure to pressure sensor in console. membrane. | Leakage of system | fail to transmit signal | failure to transmit record
Pressure Transmitter. Transmit purge pressure to pressure sensor in console. Material. | non-biocompatible | non-biocompatible material applied | biocompatibility issue
Pressure Transmitter. Transmit purge pressure to pressure sensor in console. Outer geometry. unable to insert unit into console | unable to insert unit into console | unable to insert | difficult to insert
Pressure Transmitter. Transmit purge pressure to pressure sensor in console. internal geometry. | incorrect pressure reading leading incorrect purge flow | incorrect pressure reading | incorrect measurement
Outflow Valve. Mate with cylinder and catridge body. Center hole and channel. | inadquate flow of purge fluid | difficult to assemble | incompatibility problem
Outflow Valve. Mate with cylinder and catridge body. Material. | non-biocompatible | valve biocompatibility compromised | biocompatibility issue
Ventilaition Valve. Facilitate air flow and ventiliation of purge cassette. OD. | Unable to assemble | fail to facilitate airflow | infusion or flow issue
Ventilaition Valve. Facilitate air flow and ventiliation of purge cassette. Internal geometry. | unable to vent properly system seizes up leading to unable to provide purge flow | valve out of control | difficult to open or close
Ventilaition Valve. Facilitate air flow and ventiliation of purge cassette. Color. | Improper assembly leading to unable to vent properly system seizes up causing system to unable to provide purge flow | retrograde flow | improper flow or infusion
Ventilaition Valve. Facilitate air flow and ventiliation of purge cassette. Bacteria filter. | contamination of purge fluid | valve contaminated | contamination during use
Ventilaition Valve. Facilitate air flow and ventiliation of purge cassette. Bacteria filter. | unable to vent properly system seizes up leading to unable to provide purge flow | purge fluid contaminated | contamination during use
Loctite. Bond glass cylinder assembly together. Viscosity. | Lower bond strength leading to leaking | bond field leakage | fluid leak
Loctite. Bond glass cylinder assembly together. Bond/Retention Strength. | Bond failure leading to leaking of purge fluid | low retention strength | fluid leak
Stycast. Bond glass cylinder assembly together. Viscosity. | Lower bond strength leading to leaking | inappropriate viscosity level | failure to adhere or bond
Stycast. Bond glass cylinder assembly together. Bond/Retention Strength. | Bond failure leading to leaking of purge fluid | low bond strength | fluid leak
Stycast. Bond glass cylinder assembly together. Cure Time. | Bond failure leading to leaking of purge fluid | inappropriate curing time | failure to adhere or bond
Tetrahydrofuran. Bonding cartridge cap to the cylinder assembly. Viscosity. | Lower bond strength leading to leaking | inappropriate viscosity level | loss of or failure to bond
Tetrahydrofuran. Bonding cartridge cap to the cylinder assembly. Bond/Retention Strength. | Bond failure leading to leaking of purge fluid | low bond strength | fluid leak
Tetrahydrofuran. Bonding cartridge cap to the cylinder assembly. Cure Time. | Bond failure leading to leaking of purge fluid | inappropriate curing time | failure to adhere or bond
Tetrahydrafuran. Bonds outflow valve to pressure line and pressure transmitter to pressure line. Viscosity. | Lower bond strength leading to leaking | low bond strength | loss of or failure to bond
Tetrahydrafuran. Bonds outflow valve to pressure line and pressure transmitter to pressure line. Bond/Retention Strength. | Bond failure leading to leaking of purge fluid | low bond strength | fluid leak
Pump Assembly. Pump purge fluid through device. Matching of the piston to cylinder. | High force leading to unable to drive purge fluid properly causing inadequate purge flow | fail to drive motor | pumping issue
Pump Assembly. Pump purge fluid through device. Matching of the piston to cylinder. | bad seal leading to leakage | incorrect seal | leak
Pump Assembly. Pump purge fluid through device. Vertical alignment. | bad seal leading to leakage | fail to seal properly | pumping issue
Pressure Line Assembly. Deliver purge fluid. Engagement of the pressure line. | Poor bond leading to purge fluid leakage | weak bond | loss of or failure to bond
Loctite. Bonds cartridge halves together, cylinder pin, cartridge cap to cartridge, and ventilation valve. Viscosity. | Assembly false apart leading to unable to perfrom operation | bond field damaged | break
Suction Line Assembly. Deliver purge fluid. Engagement of the suction line. | Poor bond leading to purge fluid leakage | weak bond | fluid leak
Dilator Body. Assembled to dilator hub for attachment to sheath. Joint strength. | Dilator body separates from dilator hub during use | dilator body separated | material separation
Dilator Body. Predilates tissue track during delivery of sheath assembly. OD. | Sheath damage (hub or tip) | sheath broken or damaged | break
Dilator Body. Predilates tissue track during delivery of sheath assembly. OD. Step height between sheath and dilator exists during device use and insertion into vasculature. | Step height between sheath and dilator exists during device use and insertion into vasculature | fail to insert | difficult to insert
Dilator Body. Predilates tissue track during delivery of sheath assembly. OD. Unable to assemble dilator into sheath during preparation of device bedside. | Unable to assemble dilator into sheath during preparation of device bedside | unable to assemble | installation-related problem
Dilator Body. Predilates tissue track during delivery of sheath assembly. ID. | Unable to deliver sheath and dilator assembly over access guidewire (0.035") | fail to deliver | delivery system failure
Dilator Body. Provides visualization on fluorscopy during use. Material. | Unable to visualize dilator body on fluoroscopy | unable to visualize dilator | image display error
Dilator Body Coating. Provides lubricious surface to facilitate sheath insertion into vasculature. . Coating adherence to dilator surface. | Coating dislodges from dilator component (particulate) during use. | rough coating surface | scratched material
Dilator Body Coating. Provides lubricious surface to facilitate sheath insertion into vasculature. . Coating coefficient of friction. | Coating has unacceptable (too high) coefficient of friction. | rough coating surface | physical resistance
Dilator Guard Tube Sleeve. Protect the coating on the dilator and dilator tip. ID.Dilator sleeve rubs on outer coating of dilator leading to particulate generation; coating wear | Dilator sleeve rubs on outer coating of dilator leading to particulate generation; coating wear | dilator coating worn out | particulates
Dilator Hub. Allows for flushing and injection through inner lumen. Luer geometry. | Unable to flush/occlusion | fail to flush | device clogged
Dilator Hub. Describes the size of the sheath device. Pad Printing. | Pad printing of dilator size and wire compatibility is illegible | inappropriate dilator size | device markings issue
Dilator Locking Nut. Allows for connection fo the dilator to the sheath. Thread geometry. | Unable to fasten dilator to sheath during prep and use. | unable to fasten dilator | installation-related problem
Dilator Locking Nut. Allows for connection fo the dilator to the sheath. Material. Component cracks during use, dilator backs out of sheath | Component cracks during use, dilator backs out of sheath. | dilator cracked | crack
Dilator Tip. Predilates tissue track during delivery of sheath assembly. ID. | Tip of dilator splits during tracking over access guidewire (0.035") | dilator tip split | split
Dilator Tip. Provides an atraumatic dilation of vasculature over access guidewire. Tip shape. Tip splits during delivery over access wire | Tip splits during delivery over access wire | tip split | split
Dilator Tip. Provides an atraumatic dilation of vasculature over access guidewire. Taper angle. Blunt taper angle; inability to deliver sheath assembly into vasculature | Blunt taper angle; inability to deliver sheath assembly into vasculature | blunt tip | dull
Hemostasis Valve. Allow devices to pass through into the vessel. Material. | Non-biocompatible | valve biocompatibility compromised | biocompatibility issue
Hemostasis Valve. Maintain Hemostasis (prevents blood from leaving the introducer). Material. | Valve leaks with and without devices inserted | valve leakage | fluid leak
Hemostasis Valve. Allows user to peel away sheath from interventional device. . Material. | Valve leaks with and without devices inserted | valve broken or damaged | peeled
Hemostasis Valve. Allow devices to pass through into the vessel. Radial compression force. | Valve is damaged during device insertion through sheath. | unable to pass through | failure to advance
Hemostasis Valve. Allow devices to pass through into the vessel. Radial compression force. | High device insertion force through sheath | fail to pass through | physical resistance
Hemostasis Valve. Allows user to peel away sheath from interventional device. . Valve cut lines. | Valve unable to be peeled away | unable to peel | installation-related problem
Packaging - Tray. Sterile barrier that hold the inner contents of products. Inner shape. Inner geometry of tray unable to hold packaged contents in place during transit. Damage to components and assemblies. | Inner geometry of tray unable to hold packaged contents in place during transit. Damage to components and assemblies. | inappropirate tray cavity | packaging issue
Packaging - Tray. Sterile barrier that hold the inner contents of products. Wall Thickness. | Cracks in tray allow foreign body entrance into sterile components. | foreign matter entered | item contaminated during manufacturing or shipping
Packaging - Tray. Sterile barrier that hold the inner contents of products. Wall Thickness. | Cracks in tray allow foreign body entrance into sterile components. | tray cracked | crack
Packaging - Tyvek Lid. Provide a sterile barrier by sealing the inner contents in tray. Seal. | Seal integrity is compromised during shipping and transport | seal integrity compromised | delivered as unsterile product
Packaging - Tyvek Lid. Allow for EO gas to permeate into product while maintaining sterile barrier. Permeability. | Lid does not permeate EO gas during sterilization. Unable to sterilize product. | unable to permeate | failure to sterilize
Packaging - Tyvek Lid. Provide a sterile barrier by sealing the inner contents in tray. Seal. | Unable to open | unable to open | difficult to open or remove packaging material
Packaging Breather Bag. Breather Bag. Seal. Seal integrity is compromised during shipping and transport. | Seal integrity is compromised during shipping and transport | seal integrity compromised | shipping damage or problem
Packaging Breather Bag. Breather Bag. Seal. Seal integrity is compromised during aging through shelf life. | Seal integrity is compromised during aging through shelf life | seal integrity compromised | degraded
Packaging Breather Bag. Breather Bag. Size. Breather bag is damaged during shipping and transport. Foregin bodies enter product. | Breather bag is damaged during shipping and transport. Foregin bodies enter product. | foreign matter entered | contamination during use
Packaging Breather Bag. Breather Bag. Size. Breather bag is damaged during shipping and transport. Device is damaged | Breather bag is damaged during shipping and transport. Device is damaged | bag broken or damaged | shipping damage or problem
Packaging Breather Bag. Breather Bag. Seal. | Unable to open | unable to open | difficult to open or remove packaging material
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | Sheath kinks | sheath kinked | material twisted
Sheath Body - General. Provides visualization on fluorscopy during use. Material. | Unable to visualize on fluoroscopy during device insertion | unable to visualize sheath | image display error
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | High insertion forces of inteventional device | difficult to insert | physical resistance
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | Sheath delaminates | sheath delaminated | delamination
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | Unable to peel, sheath body too robust | unable to peel | difficult to remove
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. Tip shape. | Sheath separates unintendedly, tip splits | sheath separated | material separation
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. Tip shape. | Tip separates from dilator during use (fishmouths) | sheath tip separated | material separation
Sheath Hub. Allows for flushing. Sideport core in hub. | Unable to aspirate/flush sheath | unable to flush sheath | failure to flush
Sheath Hub. Allows user to peel away sheath from interventional device. . Sheath hub design. | Sheath unable to be peeled away | unable to peel | difficult to remove
Sheath Hub. Attaches to sheath body. Sheath to hub bond. | Sheath detaches from hub during use | sheath detached | detachment of device or device component
Sheath Hub. Attaches to sheath body. Sheath hub design. Sheath hub cracks during use | Sheath hub cracks during use | sheath cracked | crack
Sheath Hub. Attaches to sheath body. Sheath hub design. Sheath hub unable to maintain hemostasis, leaks | Sheath hub unable to maintain hemostasis, leaks | unable to maintain hemostasis | fluid leak
Sheath Hub. Attaches to sheath body. Inner geometry. | Damage to interventional device during insertion and removal | sheath broken or damaged | device damaged prior to use
Sheath Hub. Attaches to sheath body. Inner geometry. | Unable to pass interventional devices | unable to pass through | failure to advance
Sheath Hub. Connects to snap caps mechanically and with adhesive. Snap cap mechanical connection feature. | Snap cap detaches from sheath hub during use | cap detached | detachment of device or device component
Sheath Hub. Provides radial compression for hemostasis valve. Peg Height. Valve pulls out of sheath hub during peeling; remains on interventional device | Valve pulls out of sheath hub during peeling; remains on interventional device | sheath pulled out | material separation
Sideport. Provide conduit for aspiration and flushing of sheath . Material. | Non-biocompatible | non-biocompatible material applied | biocompatibility issue
Sideport. Provide conduit for aspiration and flushing of sheath . Stop cock. | Unable to rotate to close (bleeding) | unable to rotate sideport | difficult to open or close
Sideport. Provide conduit for aspiration and flushing of sheath . ID. | Unable to aspirate/flush sheath | unable to aspirate | aspiration issue
Sideport. Provide conduit for aspiration and flushing of sheath . Stop cock. | Unable to flush/occlusion | unable to flush | failure to flush
Sideport. Provide conduit for aspiration and flushing of sheath . Material. Collapse tubing during aspiration. | Collapse tubing during aspiration | tubing collapsed | collapse
Snap Cap. Threads to dilator hub. Thread geometry. Unable to fasten dilator to sheath during prep and use. | Unable to fasten dilator to sheath during prep and use. | unable to fasten dilator | loss of or failure to bond
Snap Cap. Threads to dilator hub. Thread geometry. Unable to fasten dilator to sheath during prep and use, threads crack. | Unable to fasten dilator to sheath during prep and use, threads crack. | thread cracked | crack
Snap Cap. Holds valve in place within the sheath hub. Applies axial compression. Wall Thickness. | Valve leaks due to inadequate compression (axial) | inadequate axial compression | pressure issue
Electric Supply. Provide electric energy to motor. Phase interruption (e.g. broken pin, cable breakage). | Phase interruption (e.g. broken pin, cable breakage) | motor phase interrupted | electrical power problem
Electric Supply. Provide pump specific data (EEPROM). Loss/corruption of data. | Loss/corruption of data | data corrupted | loss of data
Patient cable. Provide electric energy to motor. . | Cable overheating | cable overheated | overheating of device or device component
Motor. Provide torque to flexible shaft. | Motor stalling | motor stalled | device inoperable
Motor. Provide torque to flexible shaft. | Loss of mechanical coupling to shaft | loss of mechanical coupling | component missing
Motor. Electrical isolation of purge lumen. | Loss of eletrical isolation | loss of electrical isolation | electrical shorting
Motor. Connect to catheter (via flange). | Loss of connection (catheter - flange) | loss of connection | disconnection
Motor. High electro-mechanical efficiency. | Generation of high temperatures | high temperature generated | overheating of device or device component
Motor. High electro-mechanical efficiency. Generation of bubbles in purge due to heat. | Generation of bubbles in purge due to heat | bubble generated | occlusion within device
Heat Exchanger. Cooling of catheter. Cooling compromised, softening of catheter | Cooling compromised, softening of catheter | catheter softened | insufficient cooling
Heat Exchanger. Pre-Heating of purge fluid. Efficiency of motor compromised (highly viscous purge), heat concentration in motor. | Efficiency of motor compromised (highly viscous purge), heat concentration in motor | fail to pre-heating fluid | insufficient heating
Heat Exchanger. Cooling of motor. Efficiency of motor compromised (highly viscous purge), heat concentration in motor. | Efficiency of motor compromised (highly viscous purge), heat concentration in motor (ref. Line 16) | motor efficiency compromised | overheating of device or device component
Heat Exchanger. Contain purge fluid. | Puncture / rupture | exchanger ruptured | material rupture
Purge Supply. Supply purge fluid to pump. | No purge supply to motor | fail to supply purge fluid | no flow
Purge Supply. Connect to purge cassette. | Untight connection to purge cassette | connection untighted | fluid leak
Kink protection. Prevent catheter kinks. | Kink in drive shaft | kink protection compromised | kinked
Catheter. Protect spiral bearing and flexible shaft. . | Puncture / rupture | catheter ruptured | material rupture
Catheter. Prevent catheter kinks. softening of catheter enabling catheter kinking | softening of catheter enabling catheter kinking | catheter softened | kinked
Outflow Cannula. Crimp to 9Fr. . Air bubbles introduced during insertion | Air bubbles introduced during insertion | air bubble introduced | air leak
Pump Cage: Outflow Section. Structural integrity. | Broken strut | structural integrity compromised | structural problem
Pump Cage: Outflow Section. Structural integrity. Kink in drive shaft. | Kink in drive shaft | shaft kinked | kinked
Pump Cage: Outflow Section. Expand to effective diameter. Pump Housing not fully expanded. | Pump Housing not fully expanded | fail to fully expand housing | failure to expand
Pump Cage: Rotor Section. Structural integrity. Broken strut. | Broken strut | strut broken or damaged | break
Pump Cage: Rotor Section. Structural integrity. | Eccentric rotor position | rotor eccentricity | positioning issue
Pump Cage: Rotor Section. Expand to effective diameter. | Pump Housing not fully expanded | fail to expand | failure to expand
Pump Cage: Rotor Section. Provide constant head clearance. | Inconsistent inner surface | insufficient inner clearance | installation-related problem
Pump Cage: Rotor Section. Provide constant head clearance. | Rough inner surface | rough inner surface | scratched material
Pump Cage: Inflow Section. Expand to effective diameter. | Pump Housing not fully expanded | fail to fully expand housing | failure to expand
Bearing Section. Purge tightness. . | Purge leakage leading to blood ingress | purge leakage | fluid leak
Bearing Section. Axial/radial alignment of flexible shaft. Variation in head clearance leading to high shear stress | Variation in head clearance leading to high shear stress | high shear stress | mechanical issue
Bearing Section. Low friction. High temperatures due to high friction | High temperatures due to high friction | high bearing friction | physical resistance
Bearing Section. Low friction. High temperatures due to high friction | High temperatures due to high friction | bearing overheated | temperature issue
Bearing Section. Prevent blood ingress. Purge gap too large leading to blood ingress | Purge gap too large leading to blood ingress | purge gap too large | bleed back
Rotor / Flexible Shaft. Structural integrity. . | Rotor dislodges from shaft | structural integrity compromised | structural problem
Rotor / Flexible Shaft. Expand to effective diameter. | Rotor expanded asymmetrically | rotor expanded asymmetrically | failure to align
Rotor / Flexible Shaft. Expand to effective diameter. | Rotor overexpanded reducing head clearance | rotor overexpanded | physical resistance
Rotor / Flexible Shaft. Structural integrity. | Square bar dislodged | shaft dislodged | dislodged or dislocated
Rotor / Flexible Shaft. Smooth blood contacting surfaces. | Rough surface | rough surface | scratched material
Pump Cage: Rotor Section. Smooth blood contacting surfaces. | Rough inner surface | rough inner surface | unexpected therapeutic results
Drive Shaft Cover. Provide frontal bearing. Plastic deformation leading to eccentric rotor. | Plastic deformation leading to eccentric rotor | shaft deformed | mechanical issue
Drive Shaft Cover. Provide frontal bearing. Plastic deformation leading to increased wear. | Plastic deformation leading to increased wear | shaft worn | mechanical issue
Drive Shaft Cover. Prevent ingestion of cardiac structures. High temperatures due to high friction. | High temperatures due to high friction | shaft overheated | overheating of device or device component
Drive Shaft Cover. Prevent ingestion of cardiac structures. High temperatures due to high friction cause clots. | High temperatures due to high friction cause clots | clot formated | temperature issue
Inflow mesh. Prevent ingestion of cardiac structures. . | Single ended / loose structures still make it inside (see risks DE072-4) | fail to prevent ingestion | protective measure issue
Inflow mesh. Structural integrity. Mesh ripped from pump housing during removal. | Mesh ripped from pump housing during removal | mesh ripped | torn material
Inflow mesh. Allow blood intake. | Clot formation | clot formated | infusion or flow issue
Inflow mesh. Allow blood intake. High shear stress leads to blood activation and clotting. | High shear stress leads to blood activation and clotting | blood clot formated | mechanical issue
Pigtail. Centering of pump in ventricle. Pump too close to wall | Pump too close to wall | pigtail too close to wall | patient device interaction problem
Pigtail. Wireless insertion. Damage to AO valve leaflets | Damage to AO valve leaflets | valve leaflet damaged | device handling issue
Pigtail. Safe pump removal. Damage to mitral valve apparatus | Damage to mitral valve apparatus | mitral valve damaged | device handling issue
Peel Away Sheath. Receive crimped pump head. | Peel Away splits prematurely | sheath splitted prematurely | split
Peel Away Sheath. Receive crimped pump head. | Damage to pump head | pump head broken or damaged | break
Peel Away Sheath. Peel away from catheter. | Does not peel | fail to peel | peeled
Peel Away Sheath. Transfer pump into Introducer. Pigtail tip stuck in Introducer Main Body. | Pigtail tip stuck in Introducer Main Body | pigtail tip stuck | sticking
Dilator. Vascular access. Does not lock on to Introducer | Does not lock on to Introducer | fail to lock | disconnection
Dilator. Vascular access. | Deformed / split tip | dilator deformed | material distortion
Dilator. Vascular access. Diameter mismatch with introducer sheath. | Diameter mismatch with introducer sheath | dilator mismatched | component incompatible
Introducer. Vascular access. Deformed tip. | Deformed tip | introducer tip deformed | dull
Introducer. Hemostatic seal. Seal not tight during insertion. | Seal not tight during insertion | seal integrity compromised | unsealed device packaging
Introducer. Allow sufficient blood flow to leg. | Blockage of blood flow to leg | blood flow blocked | occlusion within device
Introducer. Maintain pump position. | Pump too deep in ventricle | fail to maintain position | positioning issue
Introducer. Maintain pump position. Pump slips from ventricle | Pump slips from ventricle | incorrect pump position | malposition of device
Introducer. Complete deaeration of introducer main-body. | Air bubbles introduced during insertion | air bubble introduced | obstruction within device
Introducer. Collapse pump from 10F to 9F head for insertion. | Pump transfer impossible | fail to collapse pump | retraction problem
Linear Crimp Tool. Collapse pump from 21F to 10F head for insertion. | Pump transfer in peel-away impossible | fail to collapse pump | retraction problem
Linear Crimp Tool. Safe pump transfer into peel-away for insertion. | Pump transfer in peel-away impossible | fail to transfer pump | retraction problem
Linear Crimp Tool . Lock to peel-away sheath. Peel-away sheath unable to be inserted into linear crimp tool | Peel-away unable to be inserted into linear crimp tool | unable to insert | difficult to insert
Linear Crimp Tool . Lock to peel-away sheath. Peel-away sheath does not engage with radial spring. | Peel-away does not engage with radial spring | unable to engage | disconnection
Linear Crimp Tool . Lock to peel-away sheath. Peel-away sheath does not release from radial spring. | Peel-away does not release from radial spring | fail to release | difficult to remove
Linear Crimp Tool . Lock to peel-away sheath. Seperation between hub and shaft. | Seperation between hub and shaft | sheath shaft separated | material separation
Linear Crimp Tool . Collapse pump from 21F to 10F head for insertion. Pump damaged during diameter reduction. | Pump damaged during diameter reduction | pump damaged during collapsing | break
Retainer. fit over PSS. dimensions. | slips off | retainer slipped | slippage of device or device component
Retainer. fit over PSS. material properties. traps air in filter. | traps air in filter | air trapped in filter | entrapment of device or device component
Retainer. provide support for PSS during handling. material properties. | would not protect pss | fail to protect | protective measure issue
Retainer. provide support for PSS during handling. material properties. negatively affects purge. | negatively affects purge | purge affected negatively | infusion or flow issue
Retainer. provide support for PSS during handling. material properties. Patient reactive to substance. | Patient reactive to substance | non-biocompatible retainer material applied | biocompatibility issue
Packaging. Provide sterile barrier. material properties. | Sterile barrier breaks | sterile barrier broken or damaged | delivered as unsterile product
Dilator Hub. 23Fr Axillary. 23Fr RP. 23Fr Flex. Allows for flushing and injection through inner lumen. Luer geometry. | Unable to flush/occlusion | unable to inject | occlusion within device
Dilator Hub. Describes the size of the sheath device. Pad Printing. | Pad printing of dilator size and wire compatibility is illegible | illegible printing | device markings issue
Dilator Locking Nut. Allows for connection fo the dilator to the sheath. Thread geometry. Unable to fasten dilator to sheath during prep and use. | Unable to fasten dilator to sheath during prep and use. | unable to connect dilator | device difficult to setup or prepare
Dilator Locking Nut. Allows for connection fo the dilator to the sheath. Material. Component cracks during use, dilator backs out of sheath. | Component cracks during use, dilator backs out of sheath | dilator moved back | installation-related problem
Dilator Tip. Provides an atraumatic dilation of vasculature over access guidewire. Subassembly Design. | Tip Split | dilator tip split | split
Hemostasis Valve. Maintain Hemostasis (prevents blood from leaving the introducer). Subassembly Design. | Damaged Valve | valve broken or damaged | break
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | Sheath delaminates | sheath delaminated | delamination
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | Unable to peel, sheath body too robust | unable to peel | structural problem
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. Tip shape. | Sheath separates unintendedly, tip splits | tip splitted | split
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. Tip shape. | High insertion forces of inteventional device through distal tip section | excessive insert force | difficult to insert
Sheath Hub. Allows user to peel away sheath from interventional device. Sheath hub design. Sheath unable to be peeled away. | Sheath unable to be peeled away | unable to peel away | difficult to remove
Sheath Hub. Attaches to sheath body. Sheath to hub bond. Sheath detaches from hub during use. | Sheath detaches from hub during use | sheath detached | material separation
Sheath Hub. Attaches to sheath body. Sheath hub design. Sheath hub cracks during use. | Sheath hub cracks during use | sheath cracked | crack
Sheath Hub. Attaches to sheath body. Sheath hub design. Sheath hub unable to maintain hemostasis, leaks | Sheath hub unable to maintain hemostasis, leaks | unable to maintain hemostasis | fluid leak
Sheath Hub. Attaches to sheath body. Inner geometry. Damage to interventional device during insertion and removal | Damage to interventional device during insertion and removal | device damaged during removal | break
Sheath Hub. Attaches to sheath body. Inner geometry. Unable to pass interventional devices | Unable to pass interventional devices | unable to pass through | difficult to insert
Sheath Hub. Connects to snap caps mechanically and with adhesive. Snap cap mechanical connection feature. | Snap cap detaches from sheath hub during use | cap detached | detachment of device or device component
Sheath Hub. Provides radial compression for hemostasis valve. Peg Height. | Valve pulls out of sheath hub during peeling; remains on interventional device | unable to provide compression | pressure issue
Sheath Hub. Allows for flushing. Sideport core in hub. | Unable to aspirate/flush sheath | unable to aspirate | aspiration issue
Sideport. Provide conduit for aspiration and flushing of sheath . Material. | Non-biocompatible | unable to flush | difficult to flush
Sideport. Provide conduit for aspiration and flushing of sheath . Stop cock. | Unable to rotate to close (bleeding) | unable to rotate sideport | difficult to open or close
Sideport. Provide conduit for aspiration and flushing of sheath . ID. | Unable to aspirate/flush sheath | unable to aspirate | aspiration issue
Sideport. Provide conduit for aspiration and flushing of sheath . Material. | Kinking of side arm, unable to flush | conduit kinked | failure to flush
Sideport. Provide conduit for aspiration and flushing of sheath . Material. | Collapse tubing during aspiration | tubing collapsed | collapse
Sideport. Provide conduit for aspiration and flushing of sheath . Stop cock. Bond failure of sideare tubing to stop cock. | Bond failure of sideare tubing to stop cock | fail to bond | failure to adhere or bond
Snap Cap, 23Fr Axillary, 23Fr RP, 23Fr Flex. Threads to dilator hub. Thread geometry. | Unable to fasten dilator to sheath during prep and use. | unable to fasten dilator | device damaged prior to use
Snap Cap, 23Fr Axillary, 23Fr RP, 23Fr Flex. Threads to dilator hub. Thread geometry. | Unable to fasten dilator to sheath during prep and use, threads crack. | thread cracked | device damaged prior to use
Snap Cap 23Fr Axillary, 23Fr RP, 23Fr Flex. Holds valve in place within the sheath hub. Applies axial compression. . Wall Thickness. | Valve leaks due to inadequate compression (axial) | valve leakage | fluid leak
Outer Pouch. Provides a sterile barrier for the sheath . Dimensions. | Too small | improper bag size | physical property issue
Outer Pouch. Allows passage of sterile gas (EO). Material. | Incorrect material | unable to permeate EO gas through bag | material integrity issue
Outer Shelf Box. Provides protection during handling. Material . | Incorrect material | fail to provide protection | device packaging compromised
Axillary Clamp/graft lock. Provides connection over graft/sheath assembly. Material. | Graft lock too abrasive or sharp | sharp graft lock | material puncture
Axillary Clamp/graft lock. Provides connection over graft/sheath assembly. Diameter. | Graft lock diameter too larger | abrasive graft lock | pitted
Introducer/graft assembly. Provides connection between graft and introducer. Diameter. | Draft diamater too small or too large | fail to connect | component incompatible
Introducer/graft assembly. Provides connection between graft and introducer. Material. | Graft coating degregration | graft coating degraded | degraded
Cylinder. Hold Purge Fluid. Shaft Wall Thickness. | Crack/Stress Fracture | cylinder fractured | fracture
Cylinder. Hold Purge Fluid. Surface Quality. | Leak | cylinder leakage | leak
Cylinder. Hold Purge Fluid. Surface Quality. | High Piston Glide Force | high piston glide force | physical resistance
Cylinder. Hold Purge Fluid. Surface Quality. | Piston Flange Damage leading to Leak | piston flange damaged | break
Cylinder. Hold Purge Fluid. Shaft Length. Leak due to distal piston flange entering port | Leak due to distal piston flange entering port | cylinder leakage | leak
Cylinder. Hold Purge Fluid. Shaft Length. Piston pulls out of Cylinder | Piston pulls out of Cylinder | piston pulled out | detachment of device component
Cylinder. Hold Purge Fluid. Shaft Length. | Cannot assemble | fail to assemble | installation-related problem
Cylinder. Hold Purge Fluid. Shaft Chamfer. | Damage to Piston leading to leak | piston broken or damaged | leak
Cylinder. Hold Cylinder within Cartiridge Pair. Cap OD. | Excessive Stress on Pressure Line | excessive stress | mechanical issue
Cylinder. Low Break and Glide Force. Cylinder ID Coating. | High Piston Break Force | high piston break force | mechanical issue
Cylinder. Low Break and Glide Force. Cylinder ID Coating. | High Piston Glide Force | high piston glide force | mechanical issue
Cylinder. Connect Cylinder to Pressure Line Assembly. Luer Connector ID. | Insufficient Tensile Strength (Leak) | insufficient tensile strength | mechanical issue
Cylinder. Connect Cylinder to Pressure Line Assembly. Luer Connector ID. | Leak | connector leakage | fluid leak
Cylinder. Connect Cylinder to Pressure Line Assembly. Luer Connector Depth. | Insufficient Tensile Strength (Leak) | low tensile strength | mechanical issue
Cylinder. Connect Cylinder to Pressure Line Assembly. Luer Connector Wall Thickness. | Crack/Stress Fracture | luer fractured | fracture
Cylinder. Compatible with Body. Biocompatibility. | Patient Adverse Reaction | cylinder material biocompatibility compromised | biocompatibility issue
Cylinder. Low Break and Glide Force. Cylinder ID Coating. | Large particles obstruct fluid path | particle presented | obstruction within device
Cylinder. Sterile fluid path. Sterility Assurance Level. | Material incompatible with EtO sterilization | material incompatible with EtO sterilization | failure to sterilize
Cylinder. Sterile fluid path. Sterility Assurance Level. | Inadequate SAL achieved | inadequate SAL achieved | failure to sterilize
Rack. Mates with Driveshaft. Rack Teeth Profile. | Low Purge Flow Rate | rack mismatched | improper flow or infusion
Rack. Mates with Driveshaft. Rack Teeth Profile. | Deformed/Damaged Teeth | rack teeth damaged | break
Rack. Mates with Driveshaft. Rack Teeth Pitch. | Unable to assemble | unable to assemble | installation-related problem
Rack. Mates with Driveshaft. Rack Teeth Pitch. | Damaged Driveshaft or Rack Teeth | rack deformed | mechanical issue
Rack. Mate with Piston. Head Diameter (D1). | Loss of Seal (Leak) | rack mismatched | failure to adhere or bond
Rack. Mate with Gearbox. Width (Flat-to-Flat). | Unable to assemble | width out of specification | installation-related problem
Rack. Mate with Gearbox. Width (Flat-to-Flat). | Excessive Torque | excessive torque | mechanical issue
Rack. Mate with Gearbox. Width (Flat-to-Flat). | Crack/Stress Fracture | rack cracked | crack
Rack. Mate with Gearbox. Body Diameter (D2). | Excessive Friction (AIC Motor Stalls) | excessive friction | physical resistance
Rack. Mate with Cylinder Pin. Main Groove Width (L1). | Excessive Torque | excessive torque | mechanical issue
Rack. Mate with Cylinder Pin. Main Groove Width (L1). Insufficient Wall Thickness, Leading to Rack Break | Insufficient Wall Thickness, Leading to Rack Break | rack broken or damaged | break
Rack. Retains piston on rack. Neck Diameter. | Piston Disassembly from Rack | diameter out of specification | disassembly
Rack. Retains piston on rack. Neck Diameter. | Leak | piston disengaged | leak
Rack. Retains piston on rack. Neck Length. | Inaccurate Flow Calculation | incorrect flow calculation | improper flow or infusion
Rack. Control Stroke Distance. Main Groove Length (L2). | Leak | fail to control stroke distance | pumping issue
Rack. Control Stroke Centering. Main Groove Location (L3). | Leak | fail to control stroke center | pumping issue
Rack. Align Rack with Piston. Concentricity of Head relative to Body. | Excessive Force | excessive force | mechanical issue
Rack. Align Rack with Piston. Concentricity of Head relative to Body. | Disassembly of Piston from Rack | rack misaligned | failure to align
Rack. Align Rack with Piston. Concentricity of Head relative to Body. | Leak | rack misaligned | leak
Rack. Stroke. Rack Straightness. | Warped | rack warped | bent
Piston. Maintain Seal. Material. EQT Failure due to tightness test | EQT Failure due to tightness test | fail to pass tightness test | leak
Piston. Durability. Material. | Leak | piston durability compromised | degraded
Piston. Sealing Surface. Diameter (D1). | Leak | seal surface leakage | fluid leak
Piston. Sealing Surface & Stabilizer (proximal flange). Diameter (D1 & proximal flange). EQT Failure due to mechanical integrity test. | EQT Failure due to mechanical integrity test | fail to pass mechanical integrity test | material integrity issue
Piston. Sealing Surface & Stabilizer (proximal flange). Material Hardness. | Particulate Generation | particulate generated | particulates
Piston. Sealing Surface & Stabilizer (proximal flange). Material Hardness. Removal of Cylinder Coating. | Removal of Cylinder Coating | coating flaked off | peeled
Piston. Hold Piston onto Rack. Neck Length. Disassembly of Piston from Rack - Tear in Piston. | Disassembly of Piston from Rack - Tear in Piston | piston torn | torn material
Piston. Hold Piston onto Rack. Neck Length. Disassembly of Piston from Rack - Inadequate Contact | Disassembly of Piston from Rack - Inadequate Contact | piston disassembled | disassembly
Piston. Hold Piston onto Rack. Neck ID. | Premature Fatigue Failure | piston fatigue failure | degraded
Piston. Hold Piston onto Rack. Neck ID. | Disassembly of Piston from Rack | piston disassembled | disassembly
Piston. Hold Piston onto Rack. Neck ID. | Unable to assemble | unable to assemble | installation-related problem
Piston. Hold Piston onto Rack. Body ID. | Material tear | piston torn | torn material
Piston. EO Penetration. Angle of Notch. | Insufficient EO penetration | insufficient EO penetration | failure to sterilize
Piston. EO Penetration. Depth of Notch. | Insufficient EO penetration | insufficient EO penetrated | delivered as unsterile product
Inflow Valve. Prevent backflow into Purge Bag. ID, Inflow Port. | Too small | fail to prevent backflow | improper flow or infusion
Inflow Valve. Prevent backflow into Purge Bag. ID, Outflow Port. | Too small | low flow | no flow
Inflow Valve. Prevent backflow into Purge Bag. Max Operating Pressure Rating. | Broken Valve Membrane | valve membrane damaged | break
Inflow Valve. Prevent backflow into Purge Bag. Crack pressure. | Valve does not open | low pressure | pressure issue
Inflow Valve. Prevent backflow into Purge Bag. Particulate. | Material contamination | particulate presented | contamination during use
Inflow Valve. Prevent backflow into Purge Bag. Biocompatibility. | Patient Adverse Reaction | valve biocompatibility compromised | biocompatibility issue
Outflow Valve. Prevent backflow into Cylinder. OD, Inflow Port. | Blocked Port | valve port blocked | occlusion within device
Outflow Valve. Prevent backflow into Cylinder. ID, Outflow Port. | Too small | fail to prevent backflow | improper flow or infusion
Outflow Valve. Prevent backflow into Cylinder. Crack pressure. | Valve does not open | fail to open valve | difficult to open or close
Loctite. Bond Suction Line and Outflow Valve to Cylinder. Bond/Retention Strength. | Bond failure leading to leaking of purge fluid | bond field damaged | fluid leak
Loctite. Bond Suction Line and Outflow Valve to Cylinder. Cure Time. | Bond failure leading to leaking of purge fluid | bond field leakage | loss of or failure to bond
Pressure Line Assembly. Deliver purge fluid. Engagement of the pressure line. | Poor luer connection leading to purge fluid leakage | insufficient bond engagement | loss of or failure to bond
Purge Cassette Gen 02 Final Assembly. Pump purge fluid through device. Assembly functions. | Inadequate flow of purge fluid | inadequate purge flow | pumping issue
Purge Cassette Gen 02 Non-Sterile. Biocompatible with body. Material. | non-biocompatible | cassette material biocompatibility compromised | biocompatibility issue
Purge Cassette Gen 02 Sterile. Sterile Material. Material. | Not Sterilized | non-sterile cassette | delivered as unsterile product
Purge Cassette Gen 02 Sterile Non-QSK. Sterile Material. Material. | Not Sterilized | non-sterile cassette | failure to sterilize
Quick Set-Up Kit. Ablility to be assembled to Impella during case start. Mating features. | Leak | kit mismatched | fluid leak
Short Pressure Line Assembly. Deliver purge fluid. Engagement of the pressure line. | Poor bond leading to purge fluid leakage | insufficient bond engagement | fluid leak
AU Sensor-Sensor cable. Provide reading for flow calculation. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | fail to transfer reading | failure to read input signal
AU Sensor-Sensor cable. Provide reading for flow calculation. Depth of the cavity. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | loss of signal | unable to obtain readings
AU Sensor-Sensor cable. Provide reading for flow calculation. Sensor Material. | Readings from sensor incorrect | incorrect sensor reading | false reading from device noncompliance
AU Sensor-Sensor cable. Determining position of pump within the body. Diaphragm dimensions. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor drifted | incorrect measurement
AU Sensor-Sensor cable. Provide reading for flow calculation. Diaphragm dimensions. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | incorrect measurement
AU Purge Lumen. Provides purge fluid to the motor. Connection of purge lumen. | unable to deliver purge fluid to motor | unable to deliver purge fluid | failure to deliver
AU Purge Lumen. Provides purge fluid to the motor. Connection of purge lumen. | leaking between joint | lumen joint leakage | fluid leak
Stator. Finished pump unit which provides flow of blood and pressure sensing (without purge). Solder locations. | Failure to start pump | fail to rotate | failure to pump
Stator. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. | Bad connections to soldering board leading to pump stop/failure to start pump | fail to provide blood flow | failure to service
Stator. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. | Bad connections to soldering board leading to pump stop/failure to start pump | fail to provide pressure sensing | failure to sense
Stator. Provides structural support for pump/motor. Mechanical Properties (i.e. hardness). | Component breaks under stress | fail to provide structural support | structural problem
Ball Bearing. Provides axial and radial support during rotation of Impeller. Material of retainer. | Early wear of components | part worn out | degraded
No Hole Insert, Yoke to Hub. Provides interface for pump housing. end face perpendicularity. | Weak bond of pump housing | insufficient bond strength | installation-related problem
AU Sensor-Sensor cable. Biocompatible with body. Sensor Material. | Not biocompatible with body | cable biocompatibility compromised | biocompatibility issue
AU Rotor. Provides flow for the pump . Gap between Shaft and sleeve bearing. High purge pressure (Purge blockage) | High purge pressure (Purge blockage) | purge fluid blocked | obstruction within device
AU Rotor. Provides flow for the pump . Gap between Shaft and sleeve bearing. Insufficient radial support of rotor shaft | Insufficient radial support of rotor shaft | insufficient radial support | structural problem
AU Rotor. Provides flow for the pump . Adhesive Epoxy Duopox 1895 locations/Transitions. | Bond failure | loss of bond | loss of or failure to bond
AU Rotor. Provides flow for the pump . Delo location (application).corrosion of magnet | corrosion of magnet | magnet corrosion | corrosion
AU Motor Housing. Biocompatible with body. Silicone gel material. | Not biocompatible with body | non-biocompatible gel applied | biocompatibility issue
AU Motor Housing. Biocompatible with body. Silicone gel material. Sensor drifts beyond specification (less accurate, issues with placement) | Sensor drifts beyond specification (less accurate, issues with placement) | sensor drifted out of specification | unintended movement
AU Motor Housing. Provides pump placement information. Sensor hole location on pump housing. | Inaccurate placement signal readings | incorrect sensor placement | false reading from device noncompliance
AU Motor Housing. Assembling purge system to motor housing. Sensor Cable alignment with motor housing. | unable to assemble | housing misaligned | installation-related problem
AU Motor Housing. To provide flow. Positioning of Impeller to Pump housing. Pump stop due to contact between rotor and pump housing. | Pump stop due to contact between rotor and pump housing | rotor jammed | pumping stopped
AU Motor Housing. Provide mechanical integrity for motor assembly. Depth Integrity. | Failure of weld at low loads | loss of mechanical integrity | structural problem
AU Motor Housing. To provide flow. Positioning of Impeller to Pump housing. Biomaterial gets caught between impeller and pump housing. | Biomaterial gets caught between impeller and pump housing | biomaterial caught | entrapment of device or device component
AU Pump Housing. Structural support for the Impella. Material. Buckling/interface of Impeller to pump housing (pump stop). | Buckling/interface of Impeller to pump housing (pump stop) | housing surface buckled | material distortion
AU Pump Housing. Structural support for the Impella. Strut design features.Buckling/interface of Impeller to pump housing (pump stop). | Buckling/interface of Impeller to pump housing (pump stop) | housing structure deformed | pumping stopped
AU Pump Housing. Allows blood to enter pump safely. Suface finish. | Damage to red blood cells | blood entered unsafely | unexpected therapeutic results
AU Pump Housing. Allows blood to enter pump safely. Suface finish. Abrasion issues causing damage to vasculature. | Abrasion issues causing damage to vasculature | patient vasculature damaged | unexpected therapeutic results
AU Pump Housing. Allows blood to enter pump safely. Suface finish. | Thrombus formation | thrombus formated | unexpected therapeutic results
AU Pump Housing. Allows blood to enter pump safely. Window design features. | Damage to red blood cells | patient red blood cell damaged | unexpected therapeutic results
AU Pump Housing. Provides protected location for sensor to mount on. Sensor recessed feature. | Sensor damage/sensor dislodgement | sensor dislodged | dislodged or dislocated
AU Pump Housing. Provides protected location for sensor to mount on. Sensor deflection bump/sensor recessed feature. | Sensor damage | sensor damaged | break
AU Pump Housing. Provides protected location for sensor to mount on. Sensor recessed feature. | Sensor damage | sensor recessed | false reading from device noncompliance
AU Pump Housing. Provides interface for sensor to read pressure. Sensor hole. | Insufficient area to read pressure | insufficient sensor area | low readings
AU Pump Housing. Providing measurement for CVP, PAP, and PAPi calculations. Silicone gel thickness. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect sensor reading | false reading from device noncompliance
AU Motor Housing. Providing measurement for CVP, PAP, and PAPi calculations. Position of optical sensor. | inaccurate CVP | inaccurate signal calculation | electronic property issue
AU Pump Catheter. Assembling purge system to motor housing. Adhesive Epoxy Duopox AD840 locations/Transitions. | Bond Failure, purge leakage | purge leakage | fluid leak
Catheter. Provides deliverability (which includes the kink resistance, torqueability, pushability) of device. Material Selection. unable to secure in repo | unable to secure in repo | unable to secure | installation-related problem
AU Pump-Plug. Contains electrical, optical, and fluid connections. Shell internal layout. | Damage the sensor cable, loss of signal | optical connection damaged | failure to read input signal
Pigtail 6F Pumpe 150. Provides support of device in PA. OD. | Abrasion issues causing damage to vasculature | vasculature damaged | patient device interaction problem
AU Pump-Cannula. Placement in the heart and valve. Adhesive Epoxy Duopox AD840 locations/Transitions. | Increase insertion/withdrawing force | excessive adhesive applied | difficult to insert
Cannula. Provides support for positioning for inlet and outlet of device. Surface finish of OD. Abrasion issues causing damage to vasculature | Abrasion issues causing damage to vasculature | vasculature scratched | unexpected therapeutic results
Cannula. Provides support for positioning for inlet and outlet of device. Surface finish of OD. Abrasion issues causing damage to vasculature | Abrasion issues causing damage to vasculature | vasculature perforated | physical property issue
BG Pump Housing. Allows blood to enter pump safely. Suface finish. | Abrasion issues causing damage to vasculature | rough housing surface | material frayed
BG Pump Housing. Provides protected location for sensor to mount on. Optical Sensor offset. | offset altered | optical sensor migrated | unintended movement
Optical Sensor, Pump. Providing measurement for CVP, PAP, and PAPi calculations. Sensor Material. | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect measurement reading | incorrect measurement
Optical Sensor, Pump. Providing measurement for CVP, PAP, and PAPi calculations. Depth of the cavity. Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | Readings from sensor incorrect and/or Failure of the sensor to provide pressure reading | incorrect pressure reading | incorrect measurement
Optical Sensor, Pump. Providing measurement for CVP, PAP, and PAPi calculations. Diaphragm dimensions. | Sensor drifts beyond specification (less accurate, issues with placement) | sensor reading drifted | imprecision
Upper and Lower Handle halfs. Provides secure housing of the various components at the junction point from cable to catheter. Material. | Not biocompatible with body (skin iteration) | unable to secure housing | device or device component damaged by another device
Backloading Lumen. Allow for easy installation of guidewire. Length. guidewire damages sensor. | guidewire damages sensor | component damaged | installation-related problem
Backloading Lumen. Allow for easy installation of guidewire. Thickness. | breakage of the lumen during guidewire insertion | unable to insert guidewire | difficult to insert
Backloading Lumen. Allow for easy installation of guidewire. OD. | unable to assemble | lumen damaged | difficult to insert
Backloading Lumen. Allow for easy installation of guidewire. ID. | unable load pump on guidewire | unable to load device | device handling issue
Backloading Lumen. Allow for easy installation of guidewire. ID. | does not push out with guidewire insertion | unable to push out | device handling issue
Unsterile Pump. Allow for easy installation of guidewire. Assembly of red lumen into cannula. | unable load pump on guidewire | difficult to install guidewire | installation-related problem
AU Motor. Provide 14 days of support. Duration of use. | motor fails prior to 14 days | reduced motor life cycle | pumping stopped
AU Motor. Provide 14 days of support. Duration of use. | blood ingress into motor | blood ingressed into motor | fluid leak
AU Motor. Provide up to 4 L/min of flow. Flow requirement. | motor fails to achieve 4.0 L/min | motor capacity compromised | improper device output
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Flow requirement. | motor fails to achieve 4.0 L/min | fail to provide sufficient flow rate | insufficient flow or underinfusion
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | biomaterial | pump material biocompatibility compromised | biocompatibility issue
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | blood ingress | use duration reduced | degraded
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | ball bearing wear | usage duration reduced | device unsafe to use in environment
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | components separate | component separated | material separation
Pump. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | cannula loses shape | cannula misshapped | material torqued
Pump. Safely pump blood. Duration of use. | cannula loses shape | use duration reduced | material distortion
Pump. Safely pump blood. Suface finish. | Abrasion issues causing damage to vasculature | rough pump surface | unexpected therapeutic results
Hub. Contain motor. L1. interfere with shrink tubing - leakage current | interfere with shrink tubing - leakage current | motor current leakage | electrical issue
Hub. Contain motor. L1. blocks purge flow | blocks purge flow | purge flow blocked | no flow
Hub. Contain motor. L2. less overlap with bearing - instability | less overlap with bearing - instability | motor unstabled | pumping issue
Hub. Contain motor. L3. the wall thickness for the sensor groove is thin and break | the wall thickness for the sensor groove is thin and break | wall thickness out of specification | physical property issue
Hub. Contain motor. L3. can't fit to insert (resulting a gap between the insert and hub) | can't fit to insert (resulting a gap between the insert and hub) | unable to insert | difficult to insert
Hub. Contain motor. L4. the pump housing can't fit into the hub | the pump housing can't fit into the hub | inappropriate fit | fitting problem
Hub. Contain motor. D1. thin wall and easy to break | thin wall and easy to break | thin wall | break
Hub. Contain motor. D1. sleeve bearing glued tilt | sleeve bearing glued tilt | bearing glued | sticking
Hub. Contain motor. D2. rotor misalignment | rotor misalignment | rotor misaligned | failure to align
Hub. Contain motor. D3. Thrombus formation | Thrombus formation | thrombus formated | improper flow or infusion
Hub. Contain motor. Material selection. Corrosion | corrosion | hub corrosion | corrosion
Hub. Contain motor. Surface finish. Hemolysis | Hemolysis | patient hemolysis | patient device interaction problem
Hub. Contain motor. Surface finish. Abrasion issues causing damage to vasculature | Abrasion issues causing damage to vasculature | patient vasculature damaged | patient device interaction problem
Double Optical Bench Carrier Board. Translating the light signal coming from the second sensor to a digital signal that the AIC can read. Provide pressure signals. | Reduced software performance, delayed alarms and/or user interface | fail to translate to digital signal | electronic property issue
Double Optical Bench Carrier Board. Translating the light signal coming from the second sensor to a digital signal that the AIC can read. Provide pressure signals. | Power budget not sufficient for two optical benches | fail to translate to digital signal | failure to read input signal
Double Optical Bench Carrier Board. Translating the light signal coming from the second sensor to a digital signal that the AIC can read. Provide pressure signals. | Overcurrent in optical benches | over current in optical bench | electrical overstress
Double Optical Bench Carrier Board. Translating the light signal coming from the second sensor to a digital signal that the AIC can read. Provide pressure signals. | Over temperature in console due to additional module | board temperature high | overheating of device or device component
Second Optical Sensor connector. Establishing optical connections between the second sensors and Optical bench. Wiring arrangement/configuration. | Optical Sensor socket damage | sensor socket damaged | break
Second Optical Sensor connector. Establishing optical connections between the second sensors and Optical bench. Wiring arrangement/configuration. | Liquid or dust contamination in the socket | dust contamination | item contaminated during manufacturing or shipping
Second Optical Sensor connector. Establishing optical connections between the second sensors and Optical bench. Wiring arrangement/configuration. | ESD | liquid contamination | item contaminated during manufacturing or shipping
Second Optical Sensor connector. Establishing optical connections between the second sensors and Optical bench. Wiring arrangement/configuration. | Electric Shock | electrical shock | inappropriate shock
Double Optical Bench Carrier Board. Carry Signal. Provides Signal Path. | Double optical Carrier boards breaks, loose parts in console | board broken or damaged | unstable
PMD Impellatronic Board. Drive Motor. Provide power to the motor. Electrical over load leading to overheating of pump housing. | Electrical over load leading to overheating of pump housing | motor drive overloaded | electrical power problem
PMD Impellatronic Board. Drive Motor. Provide power to the motor. Electrical over load leading to overheating of pump housing. | Electrical over load leading to overheating of pump housing | pump housing overheated | overheating of device or device component
AU Motor Housing. Provide mechanical integrity for motor assembly. Depth Integrity. | Damage the epotek | loss of mechanical integrity | material integrity issue
AU Pump Housing. Providing measurement for CVP, PAP, and PAPi calculations. Position of optical sensor. | inaccurate CVP | optical sensor malposition | malposition of device
Nylon Catheter support Ring and 11F Transition Ring. Connection point of the motor to catheter. OD of rings. rings deform during off axis insertion (loss of function, pushability) | rings deform during off axis insertion (loss of function, pushability) | support ring deformed | installation-related problem
AU Pump-Cannula. Placement in the heart and valve. Adhesive locations/Transitions. | Bond Failures | adhesive applied to wrong place | failure to adhere or bond
AU Pump-Cannula. Placement in the heart and valve. Adhesive locations/Transitions. | Increase insertion/withdrawing force | adhesive applied to wrong place | difficult to remove
Cannula. Transportation blood from inlet to outlet. length. | Restriction of flow (low flow condition) | fail to transport blood | no flow
Cannula. Transportation blood from inlet to outlet. Surface finish of OD. | Clot formation | clot formated | physical property issue
Cannula. Provides support for positioning for inlet and outlet of device. Length. | incorrect position | incorrect cannula position | positioning issue
Cannula. Provides support for positioning for inlet and outlet of device. stiffness. | incorrect position | cannula stiffed | difficult to position
Cannula. Provides support for insertion. ID. Insufficient stiffness for use with guidewire, poor insertion | Insufficient stiffness for use with guidewire, poor insertion | insufficient cannula stiffness | mechanical issue
Cannula. Provides support for insertion. OD. Insufficient stiffness for use with guidewire, poor insertion | Insufficient stiffness for use with guidewire, poor insertion | insufficient cannula stiffness | physical property issue
Cannula. Provides support for insertion. OD . | Unable to insert device successfully | OD out of specification | difficult to insert
Cannula. Provides support for insertion. coil wire width. | Unable to insert device successfully | coil wire width out of specification | difficult to insert
Cannula. Provides support for insertion. coil wire thickness. | Unable to insert device successfully | coil wire thickness out of specification | difficult to insert
Cannula. Provides support for insertion. coil pitch. | Unable to insert device successfully | coil pitch out of specification | difficult to insert
Cannula. Provides support for insertion. extrusion durometer. | Unable to insert device successfully | extrusion durometer out of specification | difficult to insert
Cannula. Provides support for insertion. stiffness. | Unable to insert device successfully | cannula stiffness out of specification | difficult to insert
Pump Tray. Properly house pump during transport. Rail for purge side arm clip. Unable to secure purge side arm leading to kink in purge line. | Unable to secure purge side arm leading to kink in purge line. | unable to secure part | shipping damage or problem
Cannula retainer. Provides protection of motor and cannula during shipment. Inserts (LxWxH). | Insert will not fit tray | inappropriate fit | fitting problem
Cannula retainer. Provides protection of motor and cannula during shipment. Inserts (LxWxH).Unable to seal tyvek properly to tray | Unable to seal tyvek properly to tray | fail to seal properly | bacterial contamination of device
Cannula retainer. Provides protection of motor and cannula during shipment. Inserts (LxWxH). Leading damage to pigtail and/or cannula | Leading damage to pigtail and/or cannula | pigtail damaged | shipping damage or problem
Cannula retainer. Provides protection of motor and cannula during shipment. Inserts (LxWxH). Leading damage to pigtail and/or cannula | damage cannula | cannula damaged | shipping damage or problem
RPIJ Unsterile. Provides lumen for guidewire to pass through device. Assembly of red lumen into cannula. | unable load pump on guidewire | difficult to install | installation-related problem
RPIJ Unsterile. Provide up to 4 L/min of flow for a duration of 14 days. provide adequate flow. | low flow | pump duration reduced | pumping stopped
RPIJ Unsterile. Provide up to 4 L/min of flow for a duration of 14 days. Duration of use. | biomaterial | inadequate flow provided | pumping issue
RPIJ Unsterile. Be inserted through Abiomed 23Fr introducer. OD of pump. | unable to insert through introducer | pump incompatible with introducer | incompatibility problem
RPIJ Unsterile. Be inserted through Abiomed 23Fr introducer. material properties. | unable to insert through introducer | pump incompatible with introducer | component or accessory incompatibility
RPIJ Unsterile. Be delivered through interior jugular or femoral veins to pulmonary artery. trackability/pushability. | unable to deliver | unable to deliver pump | failure to deliver
RPIJ Unsterile. Provide accurate flow metric. EEPROM programming. | inaccurate flow | inaccurate flow provided | inaccurate flowrate
RPIJ Unsterile. Provide CVP, PAP, and PAPi metrics. optical sensor. | no CVP, PAP, PAPi metrics | incorrect wave metrics | electronic property issue
RPIJ Unsterile. Compatible with Abiomed guidewire. skiving. | incompatible with Abiomed guidewire | pump incompatible with guidewire | incompatibility problem
Sterile Pump. withstand shipping and storage conditions. pump design. | unable to withstand shipping and storage | pump damaged in shipping | shipping damage or problem
Sterile Pump. withstand shipping and storage conditions. pump packaging. | unable to withstand shipping and storage | pump damaged in storage | environmental control or utility issue
RPIJ Unsterile. Compatible with Abiomed 23Fr Introducer. outer diameter. | damage to hemostasis valve | incompatible with introducer | break
Cannula. Transportation blood from inlet to outlet. ID. | Cannula kinks during support | ID out of specification | kinked
Cannula. Transportation blood from inlet to outlet. OD. | Cannula kinks during support | OD out of specification | kinked
Cannula. Transportation blood from inlet to outlet. coil wire thickness. | Cannula kinks during support | coil wire thickness out of specification | kinked
Cannula. Transportation blood from inlet to outlet. coil pitch. | Cannula kinks during support | coil pitch out of specification | kinked
Cannula. Transportation blood from inlet to outlet. extrusion durometer. | Cannula kinks during support | extrusion durometer out of specification | kinked
Cannula. Transportation blood from inlet to outlet. stiffness. | Cannula kinks during support | cannula stiffness out of specification | kinked
Butterfly Suture Pad. Maintains pump position. Suture Pad Material. | Device movement impacting positioning | device moved | unintended movement
Butterfly Suture Pad. Maintains pump position. Suture Pad Suture Holes. | Unable to secure device into position | unable to secure position | unintended movement
Butterfly Suture Pad. Maintains pump position. Suture Pad Internal Geometry. | Butterfly moves along sheath | unable to secure position | device or device component damaged by another device
Dilator Body. Provide structure for dilator body. ID . | Dilator Buckle | dilator buckled | buckled material
Dilator Hub. Describes the size of the sheath device. Embossing. Embossing of dilator size is illegible. | Embossing of dilator size is illegible | illegible embossing | device markings issue
Dilator Hub. Allows for connection of the dilator to the sheath. Dilator cap to dilator hub tensile. Dilator hub disconnects from dilator cap during use, dilator backs out of sheath | Dilator hub disconnects from dilator cap during use, dilator backs out of sheath | dilator moved back | device difficult to setup or prepare
Dilator Tip. Provides an atraumatic dilation of vasculature over access guidewire. Taper angle. | Inability to deliver sheath assembly into vasculature | taper angle out of specification | difficult to insert
Dilator Tip. Provides an atraumatic dilation of vasculature over access guidewire. Tip shape . | Tip splits during delivery over access wire | incorrect tip shape | difficult to insert
Packaging - Outer. Contains and Identifies Inner Contents. Label Material. | Label damaged during Shipping | labeling damaged in shipping | shipping damage or problem
Packaging - Tray. Hold the inner contents of products. Inner shape. Inner geometry of tray unable to hold packaged contents in place during transit. Damage to components and assemblies. | Inner geometry of tray unable to hold packaged contents in place during transit. Damage to components and assemblies. | unable to secure part | shipping damage or problem
Packaging - Tray. Hold the inner contents of products. Inner shape . | Unable to fit contents of products into tray | unable to package | packaging issue
Packaging Breather Bag. Allow for EO gas to permeate into product while maintaining sterile barrier. Permeability. | Breather Bag does not permeate EO gas during sterilization. Unable to sterilize product. | inadequate sterilization gas flowed into | failure to sterilize
Packaging Breather Bag. Sterile barrier that holds the inner contents of product. Seal. | Seal integrity is compromised during shipping and transport | sterile barrier breached | shipping damage or problem
Packaging Breather Bag. Contain Inner Tray. Size. | Unable to fit tray and product in breather bag | fail to fit in bag | packaging issue
Packaging Breather Bag. Contain Inner Tray. Size. Breather bag is damaged during shipping and transport. Foregin bodies enter product. | Breather bag is damaged during shipping and transport. Foregin bodies enter product. | foreign matter entered | foreign material present in device
Sheath Body - Coil. Middle layer of sheath body that provides sheath column strength, radial strength, kink resistance. Coil Design. | Insufficient column strength | insufficient column strength | mechanical issue
Sheath Body - Coil. Middle layer of sheath body that provides sheath column strength, radial strength, kink resistance. Coil Design. | Radial stregth too low | insufficient radial strength | mechanical issue
Sheath Body - Coil. Middle layer of sheath body that provides sheath column strength, radial strength, kink resistance. Coil Design. | (Single access only) stiffness too low | low stiffness | material twisted
Sheath Body - Coil. Middle layer of sheath body that provides sheath column strength, radial strength, kink resistance. Coil Design. | (Single access only) stiffness too high | sheath stiffed | material rigid or stiff
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. | High insertion forces of inteventional device | excessive insert force | difficult to insert
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Wall Thickness. | Surface roughness too large | rough sheath surface | physical resistance
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Material. | Radial Strength too low | insufficient radial strength | mechanical issue
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Wall thickness. | (Single access only) sheath wall punctured by access needle | sheath punctured | material puncture
Sheath Body - General. Facilitates insertion of interventional devices into vasculature. Subassembly Design. (single access only) High removal forces of interventional device | (single access only) High removal forces of interventional device | excessive peel force | difficult to remove
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Material. High or unacceptable coefficient of friction between sheath ID and interventional device. | High or unacceptable coefficient of friction between sheath ID and interventional device | coil friction high | device abrasion from instrument or another object
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Material. Reinforcement layer (coil) exposed through liner | Reinforcement layer (coil) exposed through liner | coil exposed | installation-related problem
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Wall Thickness. | Reinforcement layer (coil) exposed through liner | coil exposed | physical property issue
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Material. Material not biocompatible for prolonged use | Material not biocompatible for prolonged use | biocompatibility compromised | biocompatibility issue
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Material. Material not biocompatible for prolonged use | Material not biocompatible for prolonged use | inappropriate prolonged use | biocompatibility issue
Sheath Body - Liner. Inner liner for laminated sheath body. Allows devices to pass through sheath ID and prevents exposure/contact with coil. Surface geometry. | Surface roughness too high | rough surface | scratched material
Sheath Body - Outer Extrusion. Outer Extrusion for laminated sheath body. Provides smooth and atraumatic outer sheath surface and prevents exposure/contact with reinforcement layer (coil). Material. | Reinforcement layer (coil) exposed through extrusion | coil exposed | protective measure issue
Sheath Body - Outer Extrusion. Outer Extrusion for laminated sheath body. Provides structure for sheath body. Material. | Material not biocompatible for prolonged use | sheath biocompatibility compromised | biocompatibility issue
Sheath Body - Outer Extrusion Document #10002130 (Long) Document #10002098 (Short). Outer Extrusion for laminated sheath body. Provides structure for sheath body. Surface geometry. | Surface roughness too high | rough sheath surface | scratched material
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. . Tip shape. | Tip separates from dilator during use (fishmouths) | tip separated | material separation
Sheath Body - Tip. Provides atraumatic leading edge of sheath (onto dilator). Facilitates insertion of sheath into anatomy. . Tip Design. | Sheath Tip Splitting | tip splitted | split
Sheath Body to Hub Sub-assembly. Sheath body attached to Sheath Hub. Sheath to hub bond. | Sheath detaches from hub during use | sheath detached | detachment of device or device component
Sheath Body to Hub Sub-assembly. Overmold to proximal sheath body. Overmold thickness. | Sheath detaches from hub during use | material separated | detachment of device component
Sheath foam. Retains silicone oil. Porosity. | High device insertion force through sheath | high insertion force | physical resistance
Sheath foam. Retains silicone oil. Material. | Degradation of material | foam degraded | device abrasion from instrument or another object
Sheath foam. Provides the target for the single access needle. Thickness. | (Single access only) Unable to insert access needle through foam | unable to access through | difficult to insert
Sheath foam. Allows devices to pass through into the vessel. Thickness. | (single Access only) High insertion forces of interventional device | unable to pass through | incompatibility problem
Sheath foam. Allows devices to pass through into the vessel. Thickness. | (single access only) Pump moves during SA procedure | device moved | incompatibility problem
Sheath hemostasis valve. Allow devices to pass through into the vessel. Material. | High device insertion force through sheath | high insertion force | physical resistance
Sheath hemostasis valve. Allow devices to pass through into the vessel. Radial compression force. | High device insertion force through sheath | high insertion force | mechanical issue
Sheath hemostasis valve. Allow devices to pass through into the vessel. Valve slits. | High device insertion force through sheath | valve slitted | split
Sheath hemostasis valve. Allow devices to pass through into the vessel. Radial compression force. | Valve damage | valve broken or damaged | break
Sheath hemostasis valve. Maintain Hemostasis (prevents blood from leaving the introducer). Material. Valve leaks with and without devices inserted. | Valve leaks with and without devices inserted | valve leakage | fluid leak
Sheath hemostasis valve. Allows devices to pass through into the vessel. Material. | (Single access only) unable to insert access needle through valve | unable to pass through | physical resistance
Sheath hemostasis valve. Allows devices to pass through into the vessel. Thickness. | (Single access only) unable to insert access needle through valve | unable to insert access | difficult to insert
Sheath hemostasis valve. Allow devices to pass through into the vessel. Material. | (single access only) High removal forces of interventional device | high removal force | physical resistance
Sheath hemostasis valve. Allow devices to pass through into the vessel. Material. | (single access only) Pump moves during SA procedure | pump moved | incompatibility problem
Sheath hub cap. Secures Foam Disc and Valve. Cap mechanical connection feature. | Cap detaches from sheath hub during use | cap detached | detachment of device or device component
Sheath hub cap. Secures Foam Disc. Four Ribs Thickness. Blockage of the valve entry due to incorrect foam alignment | Blockage of the valve entry due to incorrect foam alignment | foam misaligned | failure to align
Sheath hub cap. Transmits sealing pressure through compression ring on hemostasis valve. Four Ribs Height. | Reduced compression on valve | fail to transmit pressure | mechanical issue
Sheath hub cap. Secures Foam Disc. Four Ribs Height. Foam dislodging into patient | Foam dislodging into patient | foam dislodged | dislodged or dislocated
Sheath hub cap. Transmits sealing pressure through compression ring to hemostasis valve. Cap ID. Reduced compression on valve. | Reduced compression on valve | compression reduced | fluid leak
Sheath hub cap. Transmits sealing pressure through compression ring to hemostasis valve. Cap ID. | Compression force too high | compression increased | mechanical issue
Sheath Hub. Facilitates insertion of interventional devices into vasculature. Inner geometry. | Damage to interventional device during insertion and removal | device damaged during insertion | break
Sheath Hub. Housing the hemostasis valve. Depth of Valve Pocket. | Insufficient valve compression | insufficient valve compression | pressure issue
Sheath Hub. Housing the hemostasis valve. Depth of Valve Pocket. Inability to connect with dilator hub. | Inability to connect with dilator hub | unable to connect | installation-related problem
Sheath Hub. Housing the hemostasis valve. Valve Shoulder Height. | Insufficient valve compression | insufficient valve compression | fluid leak
Sheath Hub. Housing the hemostasis valve. Valve Shoulder ID. Valve unable to displace (to allow for pump insertion). | Valve unable to displace (to allow for pump insertion) | unable to assemble | installation-related problem
Sheath Hub. Attachment of suture pad. Barb Step Height. Insufficent tensile strength | Insufficent tensile strength | insufficient tensile strength | mechanical issue
Sheath Hub. Facilitates insertion of interventional devices into vasculature. Inner geometry. | (single access only) High removal forces of interventional device | fail to facilitate insertion | difficult to insert
Sheath Hub. Facilitates insertion of interventional devices into vasculature. Inner geometry. (single access only) High removal forces of interventional device. | (Single access only) Impella CP catheter punctured by access needle | high removal force | physical resistance
Sheath Hub. Facilitates insertion of interventional devices into vasculature. Inner geometry. (single Access only) High insertion forces of interventional device. | (single Access only) High insertion forces of interventional device | high insertion force | difficult to insert
Sheath Hub. Houses hemostasis valve. ID. (single access only) valve leaks with and without devices inserted. | (single access only) valve leaks with and without devices inserted | valve leakage | fluid leak
Sheath Hub. Facilitates insertion of interventional devices into vasculature. Inner geometry. (single access only) sheath hub and adjacent 7Fr sheath compress catheter. | (single access only) sheath hub and adjacent 7Fr sheath compress catheter | difficult to insert | mechanical issue
Sheath sideport. Join the sideport to the stopcock. Sideport ID. | Gaps between are too large or too small | excessive gap | installation-related problem
Sheath system. Provides ventricular access for devices to enter vasculature while maintaining hemostasis. Material. | Non-biocompatible | fail to maintain hemostasis | fluid leak
Sheath system. Provides ventricular access for devices to enter vasculature while maintaining hemostasis. Devices in the kit. | Required devices not included in the kit | fail to provide ventricular access | difficult to insert
Sheath system. Provides ventricular access for devices to enter vasculature while maintaining hemostasis. IFU. | Unable to assembly/prepare for use | difficult to prepare | device difficult to setup or prepare
Sideport Loctite 4311. Join the sideport to the sheath hub. Low viscosity for proper flow into the joint. | Gaps between are too large or too small | too small bond gap | installation-related problem
Sideport Loctite 4311. Join the sideport to the sheath hub. Low viscosity for proper flow into the joint. | Gaps between are too large or too small | too large bond gap | failure to adhere or bond
Supplemental Dilator Body. Assembled to dilator hub. Joint strength. | Dilator body separates from dilator hub during use | dialtor separated | detachment of device or device component
Supplemental Dilator Body. Predilates tissue track . ID. Unable to deliver dilator assembly over access guidewire. | Unable to deliver dilator assembly over access guidewire (0.035") | unable to deliver | failure to deliver
Supplemental Dilator Body. Provides visualization on fluorscopy during use. Material. | Unable to visualize dilator body on fluoroscopy | unable to visualize dilator | image display error
Supplemental Dilator Body. Provide structure for dilator body. ID . | Dilator Buckle | dilator buckled | structural problem
Supplemental Dilator Body. Provide structure for dilator body. Material. | Dilator Buckle | dilator buckled | material integrity issue
Supplemental Dilator Hub. Allows for flushing and injection through inner lumen. Luer geometry. | Unable to flush/occlusion | unable to inject | defective component
Supplemental Dilator Tip, 8Fr, 10Fr, 12Fr. Predilates tissue track. ID. | Tip of dilator splits during tracking over access guidewire (0.035") | tip ID out of specification | defective item
Supplemental Dilator Tip, 8Fr, 10Fr, 12Fr. Provides an atraumatic dilation of vasculature over access guidewire. Taper angle. | Inability to deliver dilator into vasculature | taper angle out of specification | difficult to insert
Supplemental Dilator Tip, 8Fr, 10Fr, 12Fr. Provides an atraumatic dilation of vasculature over access guidewire. Tip shape . | Tip splits during delivery over access wire | tip split | difficult to insert
Supplemental Dilator Tip, 8Fr, 10Fr, 12Fr. Provides an atraumatic dilation of vasculature over access guidewire. Tip shape. | Tip splits during delivery over access wire | incorrect tip shape | difficult to insert
Suture Pad. Fixates hub to the skin during use. Material. | Holes for suturing tears leading to displacement of sheath | pad displaced | unintended movement
Suture Pad. Fixates hub to the skin during use. Holes. | Unable to suture down unit leading to displacement of sheath | unable to suture down | use of device issue
Suture Pad. Fixates hub to the skin during use. ID. | Unable to rotate 180° | unable to rotate | difficult to remove
Suture Pad. Fixates hub to the skin during use. Material. | Unable to rotate 180° | pad detached | detachment of device component
AU Inflow Cage – Cannula. provides inlet for blood flow. Bond of cannula to inlet. | weak bond | weak bond | mechanical issue
AU Motor Housing. Allow gap between magnet and ceramic sleeve. concentricity. | compromise magnet gap - pump stop | magnet gap compromised | pumping stopped
AU Pump Handle. Finished pump unit which provides flow of blood and pressure sensing (without purge). Soldering quality. Bad connections to soldering board leading to pump stop/failure to start pump. | Bad connections to soldering board leading to pump stop/failure to start pump | connection loosened | pumping stopped
AU Pump Handle. Provide orientation to deliver pump into body. 90 orientation of port arm to cannula bend. | Increased suction/low flow | fail to provide orientation | failure to align
AU Pump Handle. Connects the catheter to patient cable. Placement of adhesive. Bond Failure, lead to electrical failure. | Bond Failure, lead to electrical failure | insufficient adhesive applied | electrical issue
AU Pump Handle. Connects the catheter to patient cable. Alignment of cannula to handle. | Not aligned (difficulty to deliver) | fail to align | installation-related problem
AU Rotor Pump Housing. Connect Motor to pump housing. Materials. | Corroded weld | weld corroded | corrosion
Catheter Anchor. secure the catheter. tubing ID. | does not secure | fail to secure catheter | unstable
Catheter Anchor. secure the catheter. through hole in button. | does not secure | fail to secure catheter | installation-related problem
Catheter Anchor. 2 year shelf life. housing pin hole ID. | deforms catheter | catheter deformed | material distortion
Catheter Anchor. allow for positioning/repositioning. button size. | unable to reposition | unable to position | difficult to position
Catheter Anchor. allow for positioning/repositioning. button size. | difficult to reposition | unable to reposition | positioning issue
Catheter Anchor. EO sterilizable. material selection. | does not sterilize | insufficient EO penetration | failure to sterilize
Catheter Anchor. EO sterilizable. gap sizes. | does not sterilize | insufficient EO penetration | failure to sterilize
Catheter Anchor. Biocompatible. material selection. | not biocompatible | non-biocompatible material applied | biocompatibility issue
Catheter Anchor. remain disengaged until user interacts. pin size and material. | deforms catheter | fail to disengage | detachment of device component
Catheter Anchor. 2 year shelf life. material selection. | degrades over 2 years | shelf life reduced | degraded
Catheter Anchor. 30 day indication. spring load. | deform catheter | intended use compromised | device unsafe to use in environment
Catheter Anchor. 30 day indication. tubing size and material. | deform tube or catheter | intended use compromised | degraded
Catheter Anchor. 2 year shelf life. button pin hole ID. | deforms catheter | catheter deformed | device damaged prior to use
Catheter Anchor. Maintains sterility of catheter. Material thickness of sterile sleeve.tear and/or break during use causing sterility to be compromised. | tear and/or break during use causing sterility to be compromised | catheter torn | disinfection or sterilization issue at user location
Catheter Anchor. Maintains sterility of catheter. Valve sleeve Tuohy. Unable to secure sterile sleeve leading to exposed catheter and compromised sterility. | Unable to secure sterile sleeve leading to exposed catheter and compromised sterility | catheter exposed | disinfection or sterilization issue at user location
Delo. Providing smooth transition of surfaces. Surface quality. Abrasion issues causing damage to vasculature | Abrasion issues causing damage to vasculature | rough adhesive surface | patient device interaction problem
Delo. Providing smooth transition of surfaces. Compatibility to substrates. | bond Failure | unable to bond | installation-related problem
Delo. Utilized for bonding. Elongation at break. | unable to assemble | elongation shortened | failure to adhere or bond
Delo. Providing smooth transition of surfaces. Surface quality . | unable to assemble | rough adhesive surface | scratched material
Distal Bearing. Holding of Impeller radially and axially. surface finish. | lower tensile strength, loose bond, unstable rotor dynamics | fail to maintain radial load | structural problem
Distal Bearing. Holding of Impeller radially and axially. Material. Leachable compound in purge fluid path | Leachable compound in purge fluid path | leachable compound in purge path | chemical issue
Impella Pump. Provide Cardiac Support. Material. | Not biocompatible with body | fail to provide cardiac support | pumping stopped
Impella Pump. Provide Cardiac Support. Material. | infection | patient infection | biocompatibility issue
Impella Pump. Provide Cardiac Support. Material. | Impact ability to sterilize product completely | pump sterility compromised | delivered as unsterile product
Impella Pump. Provide Cardiac Support. provide adequate flow. | low flow | inadequate flow provided | pumping issue
Impella Pump. Provide Cardiac Support. pump durability. | pump falls apart | pump duration reduced | degraded
Impella Pump. Provide Cardiac Support. pump durability. | pump stops | pump durability reduced | pumping stopped
Impella Pump. Be inserted through Abiomed 23Fr introducer. OD of pump. | unable to insert through introducer | difficult to insert pump | difficult to insert
Impella Pump. Be inserted through Abiomed 23Fr introducer. pushability of pump. | unable to insert through introducer | fail to insert pump | difficult to insert
Impella Pump. Be delivered through axillary artery to LV. trackability/pushability. | unable to deliver | unable to deliver pump | failure to deliver
Impella Pump. Compatible with Abiomed guidewire. drag of guidewire through pump. | incompatible with Abiomed guidewire | pump incompatible with guidewire | incompatibility problem
Impella Pump. Provide Cardiac Support. Pump System. | Long set up time | setup time long | device difficult to setup or prepare
Inner Blister. Properly house pump during transport. Pump handle cavity (features in tray). User inconvenience when removing the pump | User inconvenience when removing the pump | difficult to open or remove part | difficult to open or remove packaging material
Inner Blister. Properly house pump during transport. PSS Housing cavity (features in tray). PSS Housing won't be secured and could breach sterile barrier. | PSS Housing won't be secured and could breach sterile barrier | sterile barrier breached | shipping damage or problem
Inner Blister. Properly house pump during transport. PSS Housing cavity (features in tray).Unable to assemble packaging | Unable to assemble packaging | fail to package | packaging issue
Inner Blister. Properly house pump during transport. PSS Housing cavity (features in tray). | Damage to PSS Housing | housing broken or damaged | shipping damage or problem
Packaging butterfly. Secures catheter from translating within the tray. Channel for catheter. | catheter unsecure leading damage to cannula | part unsecured | device or device component damaged by another device
Packaging butterfly. Secures catheter from translating within the tray. Clipping mechanism features. | catheter unsecure leading damage to cannula | clipping mechanism failed | protective measure issue
Packaging butterfly. Secures catheter from translating within the tray. Clipping mechanism features. | User inconvenience when removing the pump | difficult to remove part | difficult to remove
Patient Cable. Relieve strain of patient cable at PSS Housing. Material (Durometer). | kink leading to pump stop | cable kinked | pumping stopped
Patient Cable. Relieve strain of patient cable at PSS Housing . Material (Durometer). | kink leading to placement signal failure | fail to release cable strain | installation-related problem
Patient Cable. Provides optical connection from console to handle. Fiber configuration. | Poor placement signal/Loss of placement Signal | incorrect fiber configuration | failure to conduct
Patient Cable. Provides optical connection from console to handle. Geometry of end face of ferrule. | Poor placement signal/Loss of placement Signal | incorrect fiber gauge | installation-related problem
Patient Cable. Provides optical connection from console to handle. Height of ferrule in the pump plug. | Poor placement signal/Loss of placement Signal | unable to assemble | physical property issue
Patient Cable. Provides optical connection from console to handle. Alignment of ferrule in the pump plug. | Poor placement signal/Loss of placement Signal | fail to assemble | failure to align
Pressure Storage Set. filter purge fluid to remove air bubbles, and hold pressure when purge disconnected. assembly. Yellow Luer Connector Broken | Yellow Luer Connector Broken | luer connector broken or damaged | break
PSS Housing. Provide protective housing for Pressure Storage Set. Material. Warpage of material leading to unable to assemble. | Warpage of material leading to unable to assemble | housing warped | installation-related problem
PSS Housing. Provide protective housing for Pressure Storage Set. Material. | Crack/damaged | housing cracked | crack
PSS Housing. Provide protective housing for Pressure Storage Set. External geometry. rough housing surface. | Damage to gloves or clinicians hand | clinician hand injured | patient device interaction problem
PSS Housing. provides mechanical connection for patient cable with overmold. Interlock geometry. | dislodgement of connection | connection dislodged | dislodged or dislocated
PSS Housing. Provide protective housing for Pressure Storage Set. Interlock geometry. Twisting of assembly on patient cable. | Twisting of assembly on patient cable | cable twisted | kinked
PSS Housing. Provide protective housing for Pressure Storage Set. Interlock geometry. | Fluid Ingress (leading to damage to PSS) | fluid ingress | fluid leak
PSS Housing. Informs user on basic instructions/warnings for use. Printed Text / Visual Indications. | Text and/or visual indications partially or fully missing | indication missing | device alarm system issue
Purge Lumen. deliver purge solution through the catheter to the motor. length. | unable to assemble | difficult to assemble | installation-related problem
Purge Lumen. deliver purge solution through the catheter to the motor. ID. | insufficient glue - bond strength compromised | insufficient glue | failure to adhere or bond
Purge Lumen. Deliver purge fluid to motor. Material Durometer. Material has high kinkability leading to incorrect purge rate issues/occlusions | Material has high kinkability leading to incorrect purge rate issues/occlusions | inappropriate material durometer | physical property issue
Purge Lumen. Deliver purge fluid to motor. Wall thickness. | Material has high kinkability leading to incorrect purge rate issues/occlusions | incorrect lumen thickness | physical property issue
Purge Lumen. Deliver purge fluid to motor. Surface finish. | Unable to assemble | inappropriate surface finish | physical resistance
Rotor Assembly. Pump blood. gap between purge lumen and shaft. | purge blocked - high purge pressure | purge blocked | no flow
Rotor Assembly. Pump blood. gap between magnet and proximal bearing. | high motor current | motor current increased | electrical issue
Rotor Assembly. Pump blood. gap between rotor and pump housing. | eccentricity | rotor shifted | positioning issue
Stainless Steel Connector. Facilitates connection of catheter purge line to cable purge line within Pump Handle AU. OD. | Unable to assemble purge tubing onto connector | unable to make connection | installation-related problem
Stainless Steel Connector. Facilitates connection of catheter purge line to cable purge line within Pump Handle AU. ID. | Purge fluid flow restriction | purge flow restricted | restricted flowrate
Strain Relief. Relief strain of cable and catheter. Material (Durometer). | kink leading to placement signal failure | cable kinked | disconnection
Strain Relief. Relief strain of cable and catheter. ID. | unable to assemble | unable to assemble | installation-related problem
Strain Relief. Relief strain of cable and catheter. ID. | Insufficient engagement for bonding leading to weak bond | insufficient engagement | loss of or failure to bond
Strain Relief. Relief strain of cable and catheter. Interlock geometry. | dislodgement of connection | connection dislodged | dislodged or dislocated
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Length. Larger Gap between magnet and thrust washer, lower tensile strength, loose bond, unstable rotor dynamics | Larger Gap between magnet and thrust washer, lower tensile strength, loose bond, unstable rotor dynamics | excessive gap | misassembled
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). surface finish. lower tensile strength, loose bond, unstable rotor dynamics | lower tensile strength, loose bond, unstable rotor dynamics | unstable rotor dynamics | installation-related problem
Thrust Washer. Functions as part of fluid bearing, carriers thrust load of the Impeller (Holds Impeller Axially). Material. Leachable compound in purge fluid path | Leachable compound in purge fluid path | leachable compound | improper chemical reaction
Tyvek Lid. Provides sterilization path. Thickness. Easily punctured/damage causing breach in sterile barrier | Easily punctured/damage causing breach in sterile barrier | lid punctured | delivered as unsterile product
Upper and Lower Handle halfs. Provides connection point cable and catheter. Interlock geometry. | dislodgement of connection | connection dislodged | dislodged or dislocated
Impellatronic PMD. Board calibrated by PAIC Fixture. Communicates with PAIC fixture. | PMD communication check failure | fail to calibrate board | communication or transmission issue
Impellatronic PMD. Board calibrated by PAIC Fixture. Communicates with PAIC fixture. | Correct hardware revision update fails | incorrect board revision | installation-related problem
PAIC Fixture. Checks and calibrates the ImpellaTronic board motor current measurement. The PAIC Fixture consists of 1 PAIC controller, 1 PAIC Application (PAIC App), 1 Computer, 1 Resistive Load, and a Digital Multimeter (DMM). . | System fails to complete calibration | fail to calibrate current measurement | failure to calibrate
PMD firmware. Firmware to be upload to Impellatronic Board. Communicates with PAIC fixture and DMM. | PMD communication check failure | fail to communicate | communication or transmission issue
Resistor. Improves motor current precision. 0.1Ohm, 1.0%, 3W | Resistor overheats | resistor overheated | overheating of device or device component
Digital multimeter (DMM). Measures motor current. Measures motor current for PAIC fixture. | PAIC fixture does not communicate with DMM | incorrect reading | incorrect measurement
Resistive Load. PAIC Fixture measures motor current against resistive load . 15 ohm, 50 W, 1%. | PAIC Fixture does not communicate with Resistive Load | fail to communicate with load | communication or transmission issue
Resistive Load. PAIC Fixture measures motor current against resistive load . 15 ohm, 50 W, 1%. | Incorrect resistive load | incorrect resistive load | electrical issue
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. ID - repositioning sheath nest. | Unable to reposition pump | unable to reposition | difficult to position
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. ID - repositioning sheath nest. | Repo unit not properly secured | unable to secure sheath | detachment of device or device component
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. Latch/lever geometry. | Lever not able to engage - ICU Sheath Lock unable to be secured | lever disengaged | disconnection
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. Latch/lever geometry. | Lever engaged - gap between body components | lever engaged | difficult to open or close
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. Latch/lever geometry. | ICU Sheath Lock unable to be detached | unable to detach | difficult to open or close
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. Sheath hub nest mating geometry. | ICU Sheath Lock unable to be secured (Sheath hub not able to fit in ID) | unable to secure sheath | installation-related problem
ICU Sheath Lock. Secure Impella CP to 14Fr Low Profile Sheath. Hinge mating geometry. | Damage to components due to assembly | hinge damaged | break
ICU Sheath Lock. EO sterilizable. gap sizes. | Does not sterilize | inadequate gap for EO penetration | failure to sterilize
ICU Sheath Lock. 3 year shelf life. material selection. | Degrades over 3 years | shelf life reduced | degraded
ICU Sheath Lock, Sterile. Parts sterile. Sterile pouch. | Not sterile | inappropriate pouch material for EO penetration | failure to sterilize
Console, Gas Supply, Air Intake / Filter. Filter airflow to/from system . Filter geometry. | Reduced airflow to oxygen concentrator | airflow reduced | medical gas supply problem
Console, Gas Supply, Air Intake / Filter. (DF04-07). Filter geometry. | Reduced airflow from blower | air intake reduced | medical gas supply problem
Console, gas supply, Blower. Provide airflow to system. Blower internal components. | Cannot set/change blower output. | fail to provide airflow | medical gas supply problem
Console, gas supply, Blower. Provide airflow to system. Blower internal components. | Loss of blower control range and resolution | loss of blower control | device operates differently than expected
Console, gas supply, flow sensor. Provide gasflow reading to console/user. Flow sensor internal components. | No gas flow reading | fail to provide gasflow reading | unable to obtain readings
Console, gas supply, flow sensor. Provide gasflow reading to console/user. Flow sensor internal components. | Airflow not properly sensed or controlled | fail to measure airflow | unable to obtain readings
Console, Gas Supply, Oxygen Concentrator. Provide Oxygen to Patient. Oxygen Concentrator internal components. | Loss of Oxygen Concentrator function | fail to provide oxygen | medical gas supply problem
Console, Gas Supply, Oxygen Concentrator. Provide Oxygen to Patient. Oxygen Concentrator internal components. | Oxygen concentration low | low oxygen concentration | medication error
Console, Gas Supply, Tubing. Provide Oxygen to Patient. Tubing Integrity. | Oxygen leak | oxygen leakage | gas leak
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Oxygen leak | valve leakage | gas leak
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Isolating Solenoid opens, fails to close (internal gas sources) | fail to close | difficult to open or close
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Isolating solenoid remains closed (external gas sources) | fail to open | medical gas supply problem
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Safety relief valve leaks | solenoid leakage | air leak
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Excessive gas pressure | excess gas pressure | medical gas supply problem
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. | Inlet Check Valve stuck open | valve stuck open | mechanical jam
Console, Gas Supply, Valves. Control Oxygen Supply to Patient. Valve Specifications. External oxygen flow stops and pressure read by inlet pressure sensor is greater than switchover point for internal gas mode. | External oxygen flow stops and pressure read by inlet pressure sensor is greater than switchover point for internal gas mode | oxygen flow stopped | medical gas supply problem
Console, Gas Supply,Pressure Sensor. Control Oxygen Supply to Patient. Valve Specifications. | Oxygen leak | oxygen leakage | medical gas supply problem
Console, Gas Supply,Pressure Sensor. Control Oxygen Supply to Patient. Valve Specifications. | No pressure reading | no pressure reading | data issue
Console, Power Supply, Batteries / Chargers. Provide power to console. Battery specifications. | Failure I2C Battery Controller Communication | battery communication failure | battery issue
Console, Power Supply, Batteries / Chargers. Provide power to console. Battery specifications. | Battery charger failure | charger malfunctioned | failure to service
Console, Power Supply, Batteries / Chargers. Provide power to console. Battery specifications. | Battery over temperature | battery over temperature | temperature issue
Console, Power Supply, Batteries / Chargers. Indicator Light. Component Specification(s). LED does not illuminate when Console is plugged in. | LED does not illuminate when Console is plugged in | fail to illuminate indicator | incorrect display
Console, Power Supply, Cabling. Provide power to console. Cable Integrity. | AC power failure | loss of AC power | loss of power
Console, Power Supply, Cabling. Provide power to console. Cable Integrity. | DC power supply cable to charger failure | loss of DC power | loss of power
Console, Power Supply, Power Control Board. Provide power to console. Power Supply specifications. | DC power failure | fail to provide power | electrical power problem
Console, Power Supply, Switching Power Supply. Provide power to console. Power Supply specifications. | DC power failure | fail to switch power supply | power source issue
Console, Power Supply, Switching Power Supply. Provide power to console. Power Supply specifications. | DC power failure | DC power failed | electrical power problem
Console, Structure, Cooling Fans / Ducts. Regulate Console Temperature. Cooling System Geometry. | Air filter clogged/blocked | air filter blocked | blockage within device or device component
Console, Structure, Cooling Fans / Ducts. Regulate Console Temperature. Cooling System Geometry. | DC power failure | air filter clogged | blockage within device or device component
Console, Structure, internal brackets. House inner componenents. Brackets specifications. | Internal bracket mechanical failure | bracket damaged | break
Console, Structure, internal brackets. House inner componenents. Components connect from ground to signal ground | Components not grounded | fail to ground | spark
Console, Structure, Outer Shell. House inner componenents. Shell Specifications. | Housing crack/fracture | housing fractured | fracture
Console, structure, sound insulation. House inner componenents. Insulation specifications. | Sound insulation detached | insulation detached | audible noise
Console, Structure, Tether / O2 Sat Interfaces. Connect Tubing to Console. Tether Geometry. | Failure of gas tubing connection | fail to connect | disconnection
Console, Structure, Wheels. Allow for transport of console. Wheel Geometry. | Wheel lock failure | wheel lock open | protective measure issue
Console, Structure, Wheels. Floor casters shall connect from ground to signal ground.. Wheel Grounding. | Wheels not grounded | fail to ground | grounding malfunction
Console, Structure, Wheels. Allow for transport of console. Wheel Geometry. | Wheel failure | wheel broken or damaged | device inoperable
Console, system control, alarm mgmt. Notify customer of alarms. System Specifications. | Fails to detect error and alarm | fail to detect alarm | failure to sense
Console, System Control, Cabling. Provide electrical connection for touch screen. Cable Geometry. | Display failure in use | fail to display | no display or display failure
Console, System control, Computer / Carrier. Provide Controls to System. System Specifications. | Carrier board failure | board failure | electronic property issue
Console, System control, Computer / Carrier. Monitor System Voltage. RIB Board specifications. | System Shutdown with Watchdog Alarm | console shutdown | failure to service
Console, System control, Computer / Carrier. Monitor System Voltage. RIB Board specifications. | System Shutdown without Watchdog Alarm | loss of watchdog alarm | protective measure issue
Console, System control, Data mgmt. Stores data. Software specifications. | Data export unavaliable | fail to export data | data back-up problem
Console, system control, SW Processes. Controls System Software. Software specifications. | Software systemic failure from hardware or software cause | loss of software application | computer software issue
Console, system control, SW Processes. Controls System Software. Software specifications. | Lose software function or control of hardware | loss of software function | application program issue
Console, User interface, Graphical User Interface. Provides interface for customer. Screen properties. Screen detects touches not performed by user | Screen detects touches not performed by user | false detection | over-sensing
Console, User interface, Graphical User Interface. Provides interface for customer. Screen properties. | Display failure in use | inadequate message display | incorrect display
Console, User interface, Graphical User Interface. Provides interface for customer. Screen properties. | Touch-screen failure in use | touchscreen malfunctioned | device inoperable
Console, User interface, Graphical User Interface. Displays gas and power remaining when disconnected from external sources. Component Specification(s). | Battery charger failure | display remained on screen | incorrect display
Console, User interface, Graphical User Interface. Provides interface for customer. Screen properties. | Cosmetic defect on display | no display | no display or display failure
Console, User Interface, Handle. Provides ability to transport. Handle Geometry. | Handle mechanical failure | handle broken or damaged | shipping damage or problem
Console, User Interface, Handle. Handle shall connect from ground to signal ground.. Handle Grounding. | Handle not grounded | fail to ground | grounding malfunction
Console, User Interface, Speaker. Provides audio feedback . Speaker specifications. | Speaker failure | audio malfunctioned | not audible alarm
Lung, Blood circuit/access, Barb-to-Tube Joints. Provides tubing connection point. Barb Geometry. | Joint leaks | joint leakage | fluid leak
Lung, Blood circuit/access, Luer Ports. Provides tubing connection point. Barb Geometry. | Blood Port leaks | port leakage | fluid leak
Lung, Blood circuit/access, Tubing. Transports blood. Tubing Mechanical properties. | Obstructed blood path | blood path obstructed | obstruction within device
Lung, Blood movement motor controller. controls pump. Controller specifications. | Pump control lost (cannot be adjusted) | loss of pump control | pumping issue
Lung, Blood movement motor controller. controls pump. Controller specifications. | Controller Over Heats | controller overheated | overheating of device or device component
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Reported Bubbles less than actual and bubble alarm enabled | bubble generated | occlusion within device
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Reported Flow less than actual | false bubble alarm | false alarm
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Reported Flow more than actual | incorrect flow reading | incorrect measurement
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Reported Bubbles more than actual | incorrect bubble reading | incorrect measurement
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Loss of 5V power to flow sensor | loss of DC power | power source issue
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Flow Sensor specifications. | Loss of RS485 Bus signals to/from flow sensor | loss of bus signal | failure to read input signal
Lung, Blood movement, flow/bubble sensor. Provides flow/bubble readings. Sensor specifications. | Reported Flow more than actual | flow reading high | high readings
Lung, Blood movement, Motor Encoder Cable. Provides electrical connection to Motor. Cable Specifications. | Loss of one or more encoder signals | loss of encoder signal | electronic property issue
Lung, Blood movement, Motor Hall Cable. Provides electrical connection to Motor. Cable Specifications. | Loss of one or more hall signals | loss of signal | unable to obtain readings
Lung, Blood movement, Motor Phase Cable. Provides electrical connection to Motor. Cable Specifications. | Loss of one or more phase connections | loss of motor phase connection | disconnection
Lung, Blood Movement, Pump. Pumps blood. Pump geometry. | Pump impeller contacts housing while rotating | pump jammed | pumping stopped
Lung, Blood Movement, Pump. Pumps blood. Pump geometry. | Pump impeller stops rotating | pump rotation stopped | pumping stopped
Lung, Blood Movement, Pump. Pumps blood. Pump geometry. | Pump impeller rotates less than the desired speed | pump rotation decreased | decreased pump speed
Lung, Blood Movement, Pump. Pumps blood. Pump geometry. | Pump impeller rotates the wrong direction | pump rotated in wrong direction | pumping issue
Lung, Blood Movement, Pump. Pumps blood. Pump geometry. | Pump impeller rotates greater than the desired speed | pump rotation increased | increased pump speed
Lung, Gas Exchange, Condensation Collection. Collects Condensation. Condensation Tray Specifications. Condensate leak to outside. | Condensate leak to outside | condensate leakage | moisture or humidity problem
Lung, Gas Exchange, Condensation Collection. Collects Condensation. Condensation Tray Specifications. | Condensation chamber exhaust obstructed | chamber exhaust obstructed | obstruction within device
Lung, Gas Exchange, Condensation Collection. Collects Condensation. Condensation Tray Specifications. Condensate leak onto floor. | Condensate leak onto floor | condensate leakage | fluid leak
Lung, Gas Exchange, Oxygenator. Oxygenates blood. Oxygenator geometry. | Blood path leak to outside | fail to oxygenate blood | medical gas supply problem
Lung, Gas Exchange, Oxygenator. Oxygenates blood. Oxygenator geometry. | Blood path obstructed | blood path obstructed | obstruction within device
Lung, Gas Exchange, Oxygenator. Oxygenates blood. Oxygenator geometry. | Air/O2 chamber leak to blood path | oxygenator leakage | gas leak
Lung, Gas Exchange, Oxygenator. Oxygenates blood. Oxygenator geometry. | Blood clots in oxygenator | blood clotted in oxygenator | infusion or flow issue
Lung, IV Pole. Holds IV Bag. Pole Geomtery. | Connection to IV pole mechanically fails | pole falling down | slippage of device or device component
Lung, Structure, Accessory connection interface. Provides connection to Lung Structure. Interface Geometry. | Connection breaks/fails | connection broken or damaged | break
Lung, Structure, Handle. Provides Handle for Lung Structure. Handle Geometry. | Handle Fails / Breaks | handle broken or damaged | structural problem
Lung, Structure, Protective housing. Provides housing for Lung Structure. Structure Geometry. | Excessive surface temperature | housing overheated | temperature issue
Lung, Tether, Air Tubing. Provides air flow to oxygenator. Tubing Mechanical properties. | Loss of Air flow to oxygenator | air flow obstructed | obstruction within device
Lung, Tether, Air Tubing. Provides air flow to oxygenator. Tubing Mechanical properties. | Reduction of Air flow to oxygenator | air flow reduced | obstruction within device
Lung, Tether, Electrical Cable. Provides electrical connection. Cable Integrity. | Loss of power signals | loss of power signal | loss of power
Lung, Tether, Electrical Cable. Provides electrical connection. Cable Integrity. Loss of CAN Bus signals | Loss of CAN Bus signals | loss of bus signal | connection issue
Lung, Tether, Electrical Cable. Provides electrical connection. Cable Integrity. Loss of RS485 Bus signals | Loss of RS485 Bus signals | loss of bus signal | failure to conduct
Lung, Tether, Oxygen Tubing Console to Oxygenator. Provides oxygen to oxygenator. Tubing Mechanical properties. | Loss of Oxygen flow | loss of oxygen flow | no flow
Lung, Blood Movement, Pump Indicator LED. Indicator Light. Cable Specifications. | LED does not illuminate when motor is running | fail to illuminate indicator | no display or display failure
Lung, Blood Movement, Pump Indicator LED. Indicator Light. Cable Specifications. | LED illuminates when motor is not running | fail to deactivate indicator | device remains activated
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Gear box. | E-drive unable to rotate | unable to rotate | failure to pump
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Gear box. | Pump control lost | loss of pump control | pumping issue
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Gear box. | Handle unable to stop rotating | unable to stop rotation | pumping issue
Lung, Blood Movement, E-Drive. Secure PLU in E-drive. Latch Assembly. | Unable to latch | unable to latch | difficult to open or close
Lung, Blood Movement, E-Drive. Secure PLU in E-drive. Latch Assembly. | unable to unlatch | unable to unlatch | difficult to open or close
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Fly Wheel. | Gear box rotating but not connected to fly wheel | wheel disengaged | disconnection
Lung, Blood Movement, E-Drive. Secure PLU in E-drive. Cradle Subassembly. | Unable to retain PLU | unable to retain | unstable
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Rotation Speed LED indicator. | Rotational speed indicator not illuminated | fail to illuminate indicator | no display or display failure
Lung, Blood Movement, E-Drive. Pumps blood in case of pump driver failure. Rotation Speed LED indicator. | LED does not illuminate to indicate rotational speed | fail to indicate rotation speed | no display or display failure
Lung, IV Pole Mount. Secure Pump driver to IV pole mount. Interface Geometry. | Connection to IV pole mount mechanically fails | mount falling off | component falling
Lung, IV Pole Mount. Rotate Pump driver on IV pole mount. Interface Geometry. | Unable to rotate pump driver on IV pole mount | unable to mount pump driver | installation-related problem
Lung, Recirculation Bag. Recirculate priming fluid during priming. Recirculation bag specification. | Recirculation bag leaks | bag leakage | fluid leak
Lung, Recirculation Bag. Recirculate priming fluid during priming. Recirculation bag specification. | Recirculation bag pathway blockage | bag pathway blocked | blockage within device or device component
Lung, Tether, Oxygen Tubing External Oxygen to Console. Provides External Oxygen to console. Tubing Mechanical properties. | Loss of Oxygen flow | loss of oxygen flow | medical gas supply problem
Lung, Tether, Oxygen Tubing External Oxygen to Console. Provides External Oxygen to console. Tubing Mechanical properties. | Reduction of Oxygen flow to oxygenator | oxygen flow reduced | medical gas supply problem
Lung, Tether, Oxygen Tubing External Oxygen to Console. Filter external oxygen to console. Filter geometry. | Lack of oxygen filtration to console | lack of oxygen filtration | filtration issue
AU Pressure Reservoir. To deliver filtered purge solution and provide purge pressure during purge events (purge cassette and/or purge fluid bag change). assembly. | bond failure (reservoir to filter) | fail to deliver purge solution | failure to deliver
AU Pressure Reservoir. To deliver filtered purge solution and provide purge pressure during purge events (purge cassette and/or purge fluid bag change). assembly. | bond failure (reservoir to filter) | weak bond | failure to adhere or bond
AU MABS Check Valve to Smiths Reservoir. To provide purge pressure during purge events (purge cassette and/or purge fluid bag change). assembly. | bond failure (check valve to reservoir) resulting in purge fluid leak | cassette leakage | fluid leak
Smiths PSS Reservoir, No Valve. To provide purge pressure during purge events (purge cassette and/or purge fluid bag change). reservoir housing. | damaged / cracked resulting in purge fluid leak | cassette cracked | crack
Smiths PSS Reservoir. To provide purge pressure during purge events (purge cassette and/or purge fluid bag change). membrane material. | damaged / cracked (tear) resulting in purge fluid leak | cassette torn | torn material
Smiths PSS Reservoir. To provide purge pressure during purge events (purge cassette and/or purge fluid bag change). membrane material. | unable to store sufficient purge volume | insufficient purge volume | short fill
Smiths PSS Reservoir. To provide purge pressure during purge events (purge cassette and/or purge fluid bag change). membrane material. | unable to store sufficent pressure | unable to store sufficient pressure | pressure issue
Logica Check Valve with Yellow Overlay. To prevent backflow of fluid. check valve. | damaged / cracked resulting in purge fluid leak | fail to prevent backflow | improper flow or infusion
Protective Cap. To prevent over expansion of membrane. radius of inner dome. unable to achieve adequate purge pressure | unable to achieve adequate purge pressure | inadequate purge pressure | pressure issue
Protective Cap. To prevent over expansion of membrane. radius of inner dome. | over expansion of membrane resulting in purge fluid leak | membrane expanded | material distortion
Protective Cap. Integrated Sidearm. To be assembled onto the pressure reservoir. thickness. | unable to secure onto pressure reservoir | unable to secure onto pressure reservoir | unstable
Protective Cap. Integrated Sidearm. To be assembled onto the pressure reservoir. OD. | unable to secure onto pressure reservoir | unable to assemble onto pressure reservoir | installation-related problem
Filter. To remove air bubbles and particulates from the purge solution. purge filter. | damaged / cracked resulting in purge fluid leak | unable to remove air bubble | filtration issue
Filter. To remove air bubbles and particulates from the purge solution. purge filter. | damaged / cracked "neck" resulting in purge fluid leak | unable to remove particulate | particulates
Filter. To remove air bubbles and particulates from the purge solution. purge filter material. | unable to filter out particles | unable to remove particulate | filtration issue
Filter. To remove air bubbles and particulates from the purge solution. purge filter. | damaged / cracked housing resulting in purge fluid leak | filter damaged | fluid leak

## uFMEA_Pool
Input_Statement | Use fm | Failure Mode | Hazard
Before starting eAUTO Mode, make sure that the adjustment value of the LV placement signal is 0 mmHg.. User checks LV adjustment is 0 mmHg prior to entering 'eAUTO' Mode. | User forgets to check LV adjustment prior to entering 'eAUTO' Mode
Before starting eAUTO Mode, make sure that the adjustment value of the LV placement signal is 0 mmHg.. User checks LV adjustment is 0 mmHg prior to entering 'eAUTO' Mode. | User forgets to check LV adjustment prior to entering 'eAUTO' Mode
Press the soft button FLOW CONTROL to open the FLOW CONTROL menu. User presses soft button to open the FLOW CONTROL menu. | User does not know how to operate soft button, assumes that the AIC is a touch screen
Rotate the selection button to select eAUTO. User rotates the selection button to select eAUTO. | User does not know how to operate soft button, assumes that the AIC is a touch screen
Press the selection button to start eAUTO.. User selects eAUTO Mode when Impella and ECMO are operating simultaneously. | User incorrectly selects AUTO Mode instead of eAUTO Mode
Press the selection button to start eAUTO.. User selects AUTO Mode when there is ECMO operating simultaneously to the Impella. | User incorrectly selects AUTO Mode instead of eAUTO Mode
Press the selection button to start eAUTO.. User selects eAUTO Mode when there is no ECMO operating simultaneously to the Impella. | User incorrectly selects eAUTO Mode instead of AUTO Mode
User gets an 'Impella Flow Reduction (eAUTO)' Alarm. User ignores Advisory Alarm and takes no action. | No failure mode; there is no Impella action
User gets an 'Maximum Flow Limited (eAUTO)' Alarm. User ignores Advisory Alarm and takes no action. | No failure mode; there is no Impella action
Remove device and accessories from package. Device and accessories are available for use. | Purge lumen ruptures during removal
Remove device and accessories from package. Device and accessories are available for use. | Unable to insert catheter
Start the Automated Impella Controller. AIC turns on and performs system test. | Unit fails to start
Spike Dextrose. Dextrose bag is prepared for use. | Damages fluid bag
Insert Purge Cassette. System primes purge cassette. | Unable to insert purge cassette
Insert Purge Cassette. System primes purge cassette. | Damage to purge cassette and/or purge disc
Connect Impella to controller via Cable . Controller detects Impella device. | Intermediate connectivity for sensor
Disconnect and discard the Y connector with the red and yellow luers from the purge tubing. Y connector discarded. | Unable to complete case start/confusion
Connect the yellow luer on the end of the purge tubing to the yellow luer on the clear sidearm of the Impella RP. Yellow luers connected. | Unable to complete case start/confusion
Connect the yellow luer on the end of the purge tubing to the yellow luer on the clear sidearm of the Impella RP. Yellow luers connected. | damage purge side arm causing leak
Enter Purge Fluid information into AIC. Purge fluid information is displayed on system. | Incorrect purge entered
Determine access site for femoral vein. Optimal access point is identified. | Potential for bleeding, hematoma, dissection
Determine access site for femoral vein. Optimal access point is identified. | Arterial access site is choosen instead of venous
Pre-dilate and place introducer. Introducer is placed in patient. | Potential for bleeding, hematoma, dissection
Pre-dilate and place introducer. Introducer is placed in patient. | Unable to place introducer/kink introducer
Advance into vein. Introducer gains access to vein. | Potential for bleeding, hematoma, dissection
Achieve ACT of 250 seconds or higher. Patient ACT is at acceptable levels to proceed with procedure. | Potential for clot development within introducer prior to device insertion
Remove dilator. Dilator is removed from patient. | Damage to introducer
Insert diagnostic catheter over a diagnostic guidewire to cross the aortic valve. diagnostic catheter gains access to LPA. | Potential for dissection and/or damage to the tricuspid valve
Insert diagnostic catheter over a diagnostic guidewire. diagnostic catheter gains access to LPA. | Potential for dissection and/or damage to the tricuspid valve
Exchange diagnostic guidewire for placement guidewire. Guidewire gains access to LPA via diagnostic catheter. | Potential for arrhythmias or damage to vasculature
Exchange diagnostic guidewire for placement guidewire then remove diagnostic catheter. Guidewire gains access to LPA via diagnostic catheter. | Potential for arrhythmias or damage to vasculature
Backload placement guidewire through pigtail, cannula, and inlet. Guidewire enters pigtail an exits inlet. | Unable to continue
Backload placement guidewire through easy guide lumen on pump, then pull white flag in line with the catheter shaft. Pump is properly loaded onto guidewire and easy guide lumen is removed. | Potential for damage to easy guide lumen (piece left in device)
Backload placement guidewire through easy guide lumen on pump, then pull white flag in line with the catheter shaft. Pump is properly loaded onto guidewire and easy guide lumen is removed. | User removes red lumen before loading
Advance Impella through the sheath and confirm position within heart using fluoroscopy. Device is positioned within heart. | Potential for arrhythmias, incorrect positioning and lack of support from the device
Advance Impella through the sheath and confirm position within heart using fluoroscopy. Device is positioned within heart. | damage pump/unable to place pump
Advance Impella through the sheath and confirm position within heart using fluoroscopy. Device is positioned within heart. | Kinking of catheter leading to bleeding
Advance Impella through the sheath and confirm position within heart using fluoroscopy. Device is positioned within heart. | Excessive bleeding
While the entire pump is in the abdominal IVC, calibrate the sensor by pressing the Zero Sensor soft button.. Pressure sensor is zeroed. | Sensor reading incorrect
Remove placement guidewire. Guidewire is removed. | Lack of support from the device
Start device via controller. Device starts and begins pumping blood. | Lack of support from the device
Start device via controller. Device starts and begins pumping blood. | biomaterial gets pulled into pump
Start device via controller. Device starts and begins pumping blood. | biomaterial gets pulled into pump leading to bleeding issues
Start device via controller. Device starts and begins pumping blood. | biomaterial gets pulled into purge gap causing high purge pressure
Start device via controller. Device starts and begins pumping blood. | biomaterial pulled onto sensor/biomaterial build-up causing failure
Verify and monitor placement of device. AIC provides information on pump placement together with fluoroscopy. | Lack of support from the device
Exchange Peel-Away Introducer for repositioning sheath. Introducer is peeled away and repositioning sheath is installed. | Damage to femoral vein; blood loss
Lock the anchoring ring in place by turning it clockwise . Anticontamination sleeve is connected to repositioning sheath. | User does not hold both sides of the anchoring rings and klickadapter rotates
Reference "CVP", "PAP" or PAPi" signal to monitor RV recovery for weaning decision. Supplement information from other diagnostic devices. | Make diagnostic decision without consulting diagnostic tools: Wean too early
Reference "CVP", "PAP" or PAPi" signal to monitor patient volume. Supplement information from other diagnostic devices. | Make diagnostic decision without consulting diagnostic tools: Make wrong decision about adding / take away volume
Reference "CVP", "PAP" or PAPi" signal to monitor patient status (trending for recovery / decompensation). Supplement information from other diagnostic devices. | Make diagnostic decision without consulting diagnostic tools: Make wrong diagnostic decision
Perform "Fast Wean": lower P level and watch "CVP", "PAP" and "PAPi" response. Help inform decision to wean vs. stay on support. | Make diagnostic decision without consulting diagnostic tools: Wean too early, make wrong diagnostic decision
Reference Trending Screen to monitor RV recovery for weaning decision. Help inform decision to wean vs. stay on support. | Make diagnostic decision without consulting diagnostic tools: Wean too early
Reference Trending Screen to monitor patient volume. Help inform decision on IV fluid management and diuretic drug management. | Make diagnostic decision without consulting diagnostic tools, Make wrong decision about adding / take away volume
Reference Trending Screen to monitor patient status (trending for recovery / decompensation). Help inform pump performance over time within cardiac system. | Make diagnostic decision without consulting diagnostic tools
Zeoring the differential pressure sensor (either initially or post-case start). Zeroed sensor signal, "CVP" algorithm has accurate input. | Wrong patient treatment decision
Spike Dextrose. Dextrose bag is prepared for use . | high purge pressure due biomaterial buildup
Connect Impella to controller via Cable . Controller detects Impella device. | No sensor signal/cable not properly seated leading to sensor signal issues
Flush Placement Lumen . The placement lumen will be de-aired. | Air introduced in body
Connect yellow and red Luers to Impella device. Controller detects Impella device. | Sensors connected incorrectly leading to pump setup failure
Assess axillary insertion site. Optimal access point is identified. | Potential for bleeding, hematoma, neuro vascular compromise
Isolate and expose the axillary artery and obtain control via proximal and distal vessel loops. Axillary artery is exposed and vessel loops are in place. | Potential for bleeding, hematoma, neuro vascular compromise
Isolate and expose the axillary artery and obtain control via proximal and distal vessel loops. Axillary artery is exposed and vessel loops are in place. | Dislodgement of pump or pump unstable do to position
Isolate and expose the axillary artery and obtain control via proximal and distal vessel loops. Axillary artery is exposed and vessel loops are in place. | damage to vasculature/infection
Attach a 10 mm diameter x 20 cm long vascular graft to the axillary artery using a standard end-to-side anastomosis. Graft secured to artery. | Potential for bleeding, hematoma, neuro vascular compromise
Attach a 10 mm diameter x 20 cm long vascular graft to the axillary artery using a standard end-to-side anastomosis. Graft secured to artery. | Unable to insert device, difficulty inserting device
Clamp the graft with a vascular clamp just above the anastomosis and loosen the vessel loops to allow blood to flow into the graft to assess for hemostasis at the anastomosis.. Clamp prevents blood from flowing through graft. | Potential for bleeding, hematoma, neuro vascular compromise
Insert the introducer into the graft . Introducer fits inside of graft. | Potential for bleeding, hematoma, neuro vascular compromise
Secure it with one (1) provided graft lock. Graft lock is secure around introducer. | Potential for bleeding, hematoma, neuro vascular compromise
Secure it with one (1) provided graft lock. Graft lock is secure around introducer. | Damage to graft lock
Insert diagnostic catheter over a diagnostic guidewire to cross the aortic valve. diagnostic catheter gains access to heart and is placed over the aortic arch. | Potential for dissection and/or damage to the aortic valve
Exchange diagnostic guidewire for placement guidewire. Guidewire gains access to heart and is placed over the aortic arch via diagnostic catheter. | Potential for arrhythmias or damage to vasculature
Exchange diagnostic guidewire for placement guidewire then remove diagnostic catheter. Guidewire gains access to heart and is placed over the aortic arch via diagnostic catheter. | Potential for arrhythmias or damage to vasculature
Remove the protective sleeve on the provided silicone-coated lubrication dilator and insert the dilator into the introducer over the placement guidewire. Dilator removed from packaging and inserted into sheath. | Difficulty inserting the dilator
Once fully inserted, remove the dilator, keeping the placement guidewire in place.. Dilator is successfully removed from sheath. | Difficulty removing the dilator
Once fully inserted, remove the dilator, keeping the placement guidewire in place.. Dilator is successfully removed from sheath. | Difficulty inserting the pump through valve
Once fully inserted, remove the dilator, keeping the placement guidewire in place.. Dilator is successfully removed from sheath. | Components damaged during insertion
Clamp the graft with a vascular clamp just above the anastomosis . Graft is properly clamped. | Potential for bleeding, hematoma, neuro vascular compromise
Backload the Impella onto the placement guidewire and advance catheter over the guidewire through the intorudcuer. Pump is properly loaded and motor housing resides in the graft, only the catheter shaft is seen exiting the valve. | Potential for bleeding, hematoma, neuro vascular compromise
Remove the vascular clamp and continue inserting the Impella pump into the aorta. Device is positioned within heart. | Potential for arrhythmias, incorrect positioning and lack of support from the device
Remove the vascular clamp and continue inserting the Impella pump into the aorta. Device is positioned within heart. | Dislodgement of pump or pump unstable do to position
Re-zero the pressure sensor (5.0 only). Pressure sensor is re-zeroed for accurate placement signal. | Difficultly positioning pump
Continue advancing across the aortic valve using fluroscopic imaging to properly position . Device is positioned within heart. | Potential for arrhythmias, incorrect positioning and lack of support from the device
Continue advancing across the aortic valve using fluroscopic imaging to properly position . Device is positioned within heart. | Dislodgement of pump or pump unstable do to position
Continue advancing across the aortic valve using fluroscopic imaging to properly position . Device is positioned within heart. | pump goes into posterior position or contacts mitral subannular structures causing placement issues
Remove the placement wire and initiate Impella catheter support by starting device via controller. Device starts and begins pumping blood. | Lack of support from the device
Clamp the graft adjacent to the axillary artery with a soft jawed vascular clamp or apply pressure. Graft is properly clamped. | Potential for bleeding, hematoma, neuro vascular compromise
Release the graft lock and remove it from the graft. Graft lock is removed. | Potential for bleeding, hematoma, neuro vascular compromise
Slide the introducer fully out of the graft then peel away the introducer . Introducer is fully removed from graft then peeled away. | Potential for bleeding, hematoma, neuro vascular compromise
Trim excess graft and slide the repositioning sheath into the graft. Repositioning sheath is fully in graft. | Potential for bleeding, hematoma, neuro vascular compromise
Using a heavy silk suture, secure the graft around the hub of the repositioning sheath, suture bluie hub to skin, engage Tuohy-Borst, secure catheter near rep unit and red handle on patient cable side. Repositioning sheath is fully in graft and pump secured in correct position within heart having final TEE position being 5 cm below AV and with flows which are acceptable (i.e. no position alarms or frequent suction alarms). | Potential for bleeding, hematoma, neuro vascular compromise
Using a heavy silk suture, secure the graft around the hub of the repositioning sheath, suture bluie hub to skin, engage Tuohy-Borst, secure catheter near rep unit and red handle on patient cable side. Repositioning sheath is fully in graft and pump secured in correct position within heart having final TEE position being 5 cm below AV and with flows which are acceptable (i.e. no position alarms or frequent suction alarms). | Dislodgement of pump or pump unstable do to position
Using a heavy silk suture, secure the graft around the hub of the repositioning sheath, suture bluie hub to skin, engage Tuohy-Borst, secure catheter near rep unit and red handle on patient cable side. Repositioning sheath is fully in graft and pump secured in correct position within heart having final TEE position being 5 cm below AV and with flows which are acceptable (i.e. no position alarms or frequent suction alarms). | high amount of position alarms
Using a heavy silk suture, secure the graft around the hub of the repositioning sheath, suture bluie hub to skin, engage Tuohy-Borst, secure catheter near rep unit and red handle on patient cable side. Repositioning sheath is fully in graft and pump secured in correct position within heart having final TEE position being 5 cm below AV and with flows which are acceptable (i.e. no position alarms or frequent suction alarms). | Pump not providing support
Remove the placement wire and initiate Impella catheter support by starting device via controller. Device starts and begins pumping blood. | Ingested biomaterial
Case Preparation. All required material at hand. | no or wrong on incomplete seldinger kit
Case Preparation. All required material at hand. | no or wrong super stiff guidewire
Case Preparation. All required material at hand. | wrong purge fluid (saline instead of glucose)
Case Preparation. All required material at hand. | purge fluid without heparine
Case Preparation. AIC setup for ECP. | AIC unable to run ECP
AIC & Purge System. purge system setup für support. | sterilty of purge system lost
AIC & Purge System. purge system setup für support. | inadequately extended purge-line
AIC & Purge System. purge system setup für support. | kink in purge-line
AIC & Purge System. purge system setup für support. | handle not held upright
AIC & Purge System. pump connection to AIC. | breakage of connectors or pins
Insertion Site & Introducer. Introducer deaired and lubricated with heparinized saline. | wrong flushing fluid without heparine
Insertion Site & Introducer. Introducer deaired and lubricated with heparinized saline. | insufficient flushing
Insertion Site & Introducer. Introducer deaired and lubricated with heparinized saline. | insufficient lubrication
Insertion Site & Introducer. Introducer deaired and lubricated with heparinized saline. | ACT too low at insertion of introducer
Insertion Site & Introducer. fully deaired introducer system. | backbleed stopped too early/bubble-free
Insertion & Case Supervision. ECP safely removed out of the tray. | pulling too hard without guiding of catheter
Insertion & Case Supervision. wetted & safe partial pump insertion into peel-away. | pulling the catheter repetively hard in contact with peel-away causing rupture of outflow cannula
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | peel-away positioned wrongly in crimp-tool
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | surplus forces on crimp tool
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | insufficient forces on crimp tool
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | misaligned pump in crimp tool
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | pulling the pump too deep into peel-away
Insertion & Case Supervision. safely crimped pump pulled back into peel-away. | pulling the pump insufficiently deep into peel-away
Insertion & Case Supervision. peel-away inserted and locked in introducer main body. | pushing peel-away insufficiently deep into main-body (no click)
Insertion & Case Supervision. peel-away inserted and locked in introducer main body. | pushing peel-away too deep into main-body (more than click)
Insertion & Case Supervision. pump transferred over sheath into vasculature. | kinking of catheter
Insertion & Case Supervision. pump transferred over sheath into vasculature. | aborting tranfer procedure
Insertion & Case Supervision. Introducer sheath successfully deaired . | insufficient deairing or introducer system
Insertion & Case Supervision. peel-away unlocked from main-body and pulled back. | introducer position lost
Insertion & Case Supervision. peel-away peeled and removed from catheter. | pump position changed
Insertion & Case Supervision. peel-away peeled and removed from catheter. | force applied to insertion site
Insertion & Case Supervision. deaired introducer main body. | air entrapped in introducer main body
Insertion & Case Supervision. washed and heparinized introducer system . | saline without heparine used
Insertion & Case Supervision. central positioning of pump system over aortic arch and valve in ventricle . | pump transferred into vessel (renal, carotid, subclavia, brachiocephalic)
Insertion & Case Supervision. central positioning of pump system over aortic arch and valve in ventricle . | pump entangling with cardic structures
Insertion & Case Supervision. central positioning of pump system over aortic arch and valve in ventricle . | pump positioned in pockets of vascular wall
Insertion & Case Supervision. central positioning of pump system over aortic arch and valve in ventricle . | pump positioned too deep into apex
Insertion & Case Supervision. central positioning of pump system over aortic arch and valve in ventricle . | pump positioned in contact with septal/ventricular conduction system
Insertion & Case Supervision. position fixed/secured in optimal position. | pump position changed uncontrolled
Insertion & Case Supervision. position fixed/secured in optimal position. | pump damage
Insertion & Case Supervision. position fixed/secured in optimal position. | tuohy damaged
Insertion & Case Supervision. stable ACT level throughout the case. | flow rate too high
Insertion & Case Supervision. stable ACT level throughout the case. | flow rate too low
Insertion & Case Supervision. optimal support selected. | P-Level too high
Insertion & Case Supervision. optimal support selected. | P-Level too low
Insertion & Case Supervision. stable ACT level throughout the case. | ACT level too high
Insertion & Case Supervision. stable ACT level throughout the case. | ACT level too low
Case Maintenance. purge cassette securely exchanged. | procedure consumes too much time
Case Maintenance. purge bag securely exchanged. | air entrapped in purge system
Case Maintenance. purge bag securely exchanged. | wrong purge fluid (saline instead of glucose)
Case Maintenance. purge bag securely exchanged. | purge fluid without heparine
Case Maintenance. purge pressure low occasion removed. | procedure consumes too much time
Case Maintenance. purge pressure high occasion removed. | procedure consumes too much time
Removal. patient securely weaned of the pump. | support unsuitable for patient disease
Removal. pump securely retracted out of ventricle. | pump interacting unfavorably with cardiac structures
Case Maintenance. Pump stop condition resolved. | pump stop condition worsened
Case Maintenance. Deair alarm condition resolved. | air entrapped in purge system comes with noise, vibration, wear
Case Maintenance. Motor current high condition resolved. | motor current high condition worsened
Case Maintenance. Motor current high condition resolved. | pump exchange too early
Case Maintenance. audio off used too often. | ignorance
Case Maintenance. prolongued usage of surgical mode. | ignorance
Removal. Access site closed after pump removal. | Unnoticed delayed restart of bleeding
Insertion Site & Introducer. Guidewire inserted in vasculature for additional intervention. | Re-use of Introducer system after pump removal
Determine access site for femoral artery. Optimal access point is identified. | Potential for bleeding, hematoma, dissection
Advance into artery. Introducer gains access to artery. | Potential for bleeding, hematoma, dissection
Advance into artery. Introducer gains access to artery. | Unable to insert device
Advance into artery. Introducer gains access to artery. | Damage to artery/vasculature
Advance into artery. Introducer gains access to artery. | Damage causing bleeding
Insert access needle into valve at 10 or 2 o'clock position. Needle inserts through valve. | Needle stick is too close to helix or valve breakline
Insert access needle into valve at 10 or 2 o'clock position. Needle inserts through valve. | Damage to introducer
Insert 0.035" guidewire through access needle. 0.035" guidewire inserts through needle. | Unable to insert device
Insert 0.035" guidewire through access needle. 0.035" guidewire inserts through needle. | Potential for damage to vasculature
Remove access needle. Access needle removes from valve. | Unable to remove device
Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | Unable to insert device
Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | Damage to introducer
Remove 0.035" guidewire. 0.035" guidewire removed. | Unable to remove device
Insert PCI equipment through PCI sheath. PCI equipment inserts through PCI sheath. | Unable to insert device
Remove PCI equipment through PCI sheath. PCI equipment removes through PCI sheath. | Unable to remove device
Remove PCI sheath. PCI sheath removes. | Damage to introducer
Remove PCI sheath. PCI sheath removes. | Unable to remove device
Press and hold the power switch on the right side of the AIC for 3 seconds. AIC turns on and performs system test. | AIC fails to start
Press the START NEW CASE soft button from the startup screen. New case has mode has been selected. | New case has not been selected
Remove the Impella BTR Catheter/Pump from its package using sterile technique and inspect, including its connector, for damage. Impella catheter is removed from packaging and ready for use. | Impella catheter/ pump or second optical sensor is damaged
Inspect the cables and for damage, including damage to the connector pins at the controller end. Connectors with damaged pins are scrapped. | Damaged pins are not recognized and user continues with procedure
Inspect the optical fibers in the first and second optical plug. Plugs with damage fibers are scrapped. | Damgaged fibers are not recognized and users continues with procedure
Open the cover on the blue catheter plug by rotating clockwise. Line up the notch on the connector cable with the notch in the blue catheter plug on the front of the Automated Impella Controller and plug the cable into the controller.. Controller detects Impella device/ First Optical Sensor. | Sensor signal issues, Not plugged in
Open the cover on the blue catheter plug by rotating clockwise. Line up the notch on the connector cable with the notch in the blue catheter plug on the front of the Automated Impella Controller and plug the cable into the controller.. Controller detects Impella device/ First Optical Sensor. | No sensor signal/cable not properly seated leading to sensor signal issues
Remove small black cap from the second sensor socket. Insert the second optical sensor plug into the optical sensor socket. The red line on the second optical sensor and the red dot on the socket must be aligned. Controller detects Impella Second Optical Sensor. | User cannot line up second sensor
Remove small black cap from the second sensor socket. Insert the second optical sensor plug into the optical sensor socket. The red line on the second optical sensor and the red dot on the socket must be aligned. Controller detects Impella Second Optical Sensor. | No sensor signal/cable not properly seated leading to sensor signal issues
Press the NEXT soft button to continue. Soft button has been pressed. | NEXT has not been selected
Leave impella catheter on the table while the sensors are calibrated. Do not touch the sensors or move the Impella Catheter during this time.. Controller Calibrates Sensors. | First, second, or both sensors are not calibrated properly
Press START IMPELLA soft button to open the P-level menu. Turnthe selector knob to increase the P-level from P-0 to P-2.. Impella catheter starts and begins pumping blood. | Ingested biomaterial
Press START IMPELLA soft button to open the P-level menu. Turn the selector knob to increase the P-level from P-0 to P-2.. Impella catheter starts and begins pumping blood. | Ingested biomaterial
Press the selector knob to select the new P-level. Increase the P-level to P-9 to confirm correct and stable placement.. N/A, no patient harm is associated with this use step. | N/A, no patient harm is associated with this use step
Interpretation of Placement Signal and Motor Current (by default the Display Speed Pulse is not displayed). The user recognizes that the signal blanking 'flat line' in the Placement signal and motor current every 5s is to filter out the S- pulse spikes from the signal. | User does not correctly interpret the Placement Signal and Motor Current Waveforms
Interpretation of Placement Signal an Motor Current (user uses softkey to select DISPLAY-> DISPLAY SPEED PULSE). The user enables the S-Pulse and is confused by the placement signal and motor current. . | User does not correctly interpret the Placement Signal and Motor Current Waveforms
Controller Error (2131, 2035, 2039, 2040, 2041, 2043, 2044, 2045) Switch to back-up Controller (Yellow, Serious Alarm). User switches to back up Contoller. | Ignores alarm; placement and suction monitoring algorithms are not reliable
Inlet Signal Not Reliable* (2130, 2036, 2048) Flow, Placement and Suction Monitoring are Suspended (Yellow, Serious Alarm) 1. Monitor patient hemodynamics and Impella position with imaging. 2. Check the orange optical connector for good fit in the controller.. User monitors the patient hemodynamic and Impella positioning with imaging. User check the orange optical connector for good fit in the controller. . | User is confused by alarm and looks at the wrong optical sensor; colorblind; does not understand that Pin is out of range
Inlet Signal Not Reliable* (2130, 2036, 2048) Flow, Placement and Suction Monitoring are Suspended (Yellow, Serious Alarm) 1. Monitor patient hemodynamics and Impella position with imaging. 2. Check the orange optical connector for good fit in the controller.. User monitors the patient hemodynamic and Impella positioning with imaging. User check the orange optical connector for good fit in the controller. . | Alarm fatigue; alarm ignored; Audio for alarm is disabled
Remove device from packaging. Contents successfully removed from packaging. | Device damaged during removal
Flush the dilator through luer. Dilator succcessfully flushed with reduction of air bubbles. | User omits step
Flush the sheath through sideport. Sheath succcessfully flushed with reduction of air bubbles. | User omits step
Flush the sheath through sideport. Sheath succcessfully flushed with reduction of air bubbles. | Transparent sheath leads user to think bubbles still remain in sheath
Flush the sheath through sideport. Sheath succcessfully flushed with reduction of air bubbles. | Sheath unable to be flushed
Close the valve on sideport. Sideport successfully closed. | Unable to close valve
Load the dilator tip through sheath hub. Dilator successfully inserted through sheath. | Unable to insert dilator tip through sheath
Connect sheath hub to dilator hub. Sheath hub successfully connected to dilator hub. | Unable to connect the hubs
Connect sheath hub to dilator hub. Sheath hub successfully connected to dilator hub. | Hubs are not fully connected
Connect sheath hub to dilator hub. Sheath hub successfully connected to dilator hub. | Hub mechanism(s) break
Hydrate hydrophilic Coating. Coating hydrated. | User misses coating hydration or performs incorrectly
Hydrate hydrophilic Coating. Coating hydrated. | User hydrates coating with fluids not recommended for use
Hydrate hydrophilic Coating. Coating hydrated. | User hydrates coating with tools/medical supplies not recommended for use
Hydrate hydrophilic Coating. Coating hydrated. | User hydrates coating with tools/medical supplies not recommended for use
Hydrate hydrophilic Coating. Coating hydrated. | User hydrates coating with fluids not recommended for use (any fluid besides saline solution)
Hydrate hydrophilic Coating. Coating hydrated. | User hydrates coating with fluids not recommended for use (any fluid beside saline solution)
Follow accepted clinical practice for vessel puncture or incision and guidewire insertion. Optimal access point has been punctured. | Potential for bleeding, hematoma, dissection
Insert soft tip of a 0.035” guidewire through the introducer needle into the vessel . Guidewire is successfully in place. | Guidewire tip damaged
Advance sheath into artery over 0.035" guidewire. Guidewire successfully inserted through dilator tip. | Dilator tip damaged
Advance sheath into artery over 0.035" guidewire. Guidewire successfully inserted through dilator tip and sheath is fully in place. | Dilator tip damaged
Advance sheath into artery over 0.035" guidewire. Guidewire successfully inserted through dilator tip and sheath is fully in place. | Sheath not fully in place
Remove the dilator, leaving the sheath in position. Remove the 0.035” guidewire.. Dilator removed with sheath successfully in position. | Sheath and dilator are fully removed from artery
Open the valve on the sheath sidearm, aspirate and flush the sheath with a syringe connected to the flush port with saline. Sheath successfully flushed to cleanse the sheath of blood. | Thrombus formation
Close the valve on the sidearm. Sideport successfully closed. | Air bubbles enter bloodstream
Introduce device(s) up to 14Fr through the hemostasis valve/sheath and advance into position. Devices gain access into ventricle. | Potential for dissection and/or damage to the aortic valve
Introduce device(s) up to 14Fr through the hemostasis valve/sheath and advance into position. Devices gain access into ventricle. | Potential for arrhythmias or damage to vasculature
Introduce device(s) up to 14Fr through the hemostasis valve/sheath and advance into position. Devices gain access into ventricle. | Device greater than 14Fr tried to be inserted through sheath
Remove all devices from sheath. Devices removed from sheath. | Devices are unable to be removed through sheath
Aspirate and flush. Sheath succssfully flushed to cleanse the sheath of blood. | Thrombus formation
Hold pressure distal to sheath and remove sheath while applying manual pressure until hemostasis is achieved. Sheath successfully removed and hemostasis is achieved . | Excessive bleeding
Connect sideport to manometer. Sideport female luer connects to manometer male luer. | Luers do not connect
(Single access only) Insert access needle into valve. Needle inserts through valve. | Needle inserted too close to the center of the valve
(Single access only) Insert 0.035" guidewire through access needle. 0.035" guidewire inserts through needle. | User does not place 0.035" wire far enough into the sheath/vessel
(Single access only) Remove access needle. Access needle removes from valve. | N/A, no patient harm is associated with this use step
(Single access only) Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | User attempts to place too large of a PCI sheath
(Single access only) Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | User attempts to place PCI sheath without hydrophilic coating
(Single access only) Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | User attempts to place PCI sheath with limited hydrophilic coating and guide catheter does not reach ostium of the vessel
(Single access only) Insert PCI sheath over 0.035" guidewire through valve. PCI sheath inserts over 0.035" guidewire, through valve, and through 14 Fr sheath. | User does not secure Impella during PCI
(Single access only) Remove 0.035" guidewire. 0.035" guidewire removed. | N/A, no patient harm is associated with this use step
(Single access only) Insert PCI equipment through PCI sheath. PCI equipment inserts through PCI sheath. | User does not secure Impella during PCI
(Single access only) Remove PCI equipment through PCI sheath. PCI equipment removes through PCI sheath. | User does not secure Impella during PCI
Press and hold the power switch on the right side of the AIC for 3 seconds. AIC turns on and performs system test. | User does not start AIC
Press the START NEW CASE soft button from the startup screen. New case has been selected. | New case has not been selected
Open the purge cassette package. Purge cassette has been removed from package. | Purge cassette packaging unable to be opened
Secure the red and yellow luers to the sterile field. Red and yellow luers are in the sterile field. | User drops luers
Pass the purge cassette and spike off the sterile field. Purge Cassette and spike are passed off. | User passes off the entire purge cassette including the yellow luer
Spike Dextrose Bag. Dextrose bag is prepared for use. | User does not adequately puncture the dextrose bag
Open door using button on left side. Purge cassette door is open. | User is unable to locate button on left side of console
Open door using button on left side. Purge cassette door is open. | User attempts to push door to disconnect
Insert the purge cassette into the compartment on the front of the controller. System primes purge cassette. | User does not fully engage purge cassette into compartment, system not primed
Slide the purge disc into the slot to the right of the purge cassette until it snaps into place. System primes purge cassette. | User unable to connect pressure transducer into AIC compartment
Slide the purge disc into the slot to the right of the purge cassette until it snaps into place. System primes purge cassette. | User does not slide pressure transducer all the way to the left to lock in, disc does not click into place
Extend the purge tubing. Purge tubing is extended. | User omits step
Close the purge cassette door. Purge cassette door is closed. | User omits step
Remove the Impella Catheter from its package using sterile technique and inspect, including its connector, for damage. Impella catheter is removed from packaging and ready for use. | User applies force to catheter and parts are damaged
Snap the purge clip (located on the pressure reservoir of the clear sidearm) to the connector cable. Purge clip is connected to the connector cable. | User omits step
If included, discard the Y-connector with red and yellow luers from the purge tubing. Y-connector is discarded and luers and secure to purge tubing. | User omits step
Connect Impella Catheter using luers: yellow to yellow. Controller detects purge pressure and priming status of Impella device. | User damages purge sidarm
Connect Impella Catheter using luers: yellow to yellow. Controller detects purge pressure and priming status of Impella device. | User does not recognize air in line, air is introduced in body
Connect Impella Catheter using luers: yellow to yellow. Controller detects purge pressure and priming status of Impella device. | User does not know how to connect luers
To set the default purge values displayed on the screen press the ACCEPT soft button. User sets purge values. | User omits step
Enter Purge Fluid information into AIC. Press the EDIT soft button, scroll to the appropriate item and push the selector knob. . Purge fluid information is displayed on system. | Incorrect purge entered
Scroll through the values and push the selector knob or press SELECT to make a new selection. Press DONE button to finish editing.. User sets purge values. | User omits step
Confirm Purge Fluid Information. Purge fluid information is displayed on system. | User omits step
To complete the set-up, connect the purge tubing to the white connector cable by pushing the purge tubing into the clips attached to the white connector. Purge tubing is connected. | User omits step
Prep 5-8Fr and 23Fr sheaths and dilators by flushing the sidearms. Sheath and dilators are flushed. | User omits step
Insert dilators and lock into sheath. Sheath and dilators are connected. | User omits step
Flush dilators from 23Fr kit. Supplemental dilators are flushed. | User omits step
Insert 23Fr dilator into 23Fr sheath. 23Fr system is connected. | User omits step
Obtain access to the jugular or femoral vein.. Access is obtained. | User omits step
Insert needle into vessel. The needle position should be verified by blood return.. Needle stick is in the correct angle. | User sticks too low or too high in vessel
Insert needle into vessel. The needle position should be verified by blood return.. Needle stick is in the correct angle. | User does not utilize recommended stick angle
Remove the syringe and insert soft tip of guidewire through the introducer needle into the vessel. Guidewire is inserted into vessel. | Incorrect guidewire used (too soft)
Advance the guidewire to the required depth. Leave an appropriate amount of guidewire exposed.. Guidewire is inserted into vessel. | User does not leave enough guidewire for the needle exchange
Hold guidewire in place and remove introducer needle.. Needle is removed while wire access is maintained. | User doesn't hold pressure over venotomy
If necessary, pre-dilate the vessel using supplemental dilators by threading the dilator over the guidewire and advancing into the vessel. User pre-dilates the vessel. | User does not pre-dilate
Insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. User connects the dilator hub to sheath hub. | User does not connect the dilator to hub
Insert vessel dilator into sheath and rotate the dilator cap over valve housing to secure the dilator onto sheath assembly. User connects the dilator hub to sheath hub. | User wipes silicone coating off dilator prior to prepping
Thread the dilator/sheath assembly over the guidewire. Sheath assembly is inserted into vessel over guidewire. | User does not put the sheath completely in the vessel
Thread the dilator/sheath assembly over the guidewire. Sheath assembly is inserted into vessel over guidewire. | User selects wrong 23 French sheat length for access site
Achieve ACT of 250 seconds or higher. Patient ACT is at acceptable levels to proceed with procedure. | User removes dilator prior to ACT being 250 or higher
Once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the cap off the hub . Dilator cap is removed from sheath cap. | User attempts to dislodge dilator from sheath hub by cramming vs. twisting and the hub cracks
Once assembly is fully introduced into the vascular system, separate the dilator cap from the sheath valve housing by rotating the cap off the hub . Dilator cap is removed from sheath cap. | User is unable to disconnect dilator
Slowly retract and remove the dilator, leaving the sheath and guidewire in position. (diff wording from IFU). Dilator is removed from patient. | User does not recognize that sheath backed out from vein
Aspirate all air from the sheath valve assembly by using syringe connected to the sideport.. Air is removed from sheath . | User omits step
Flush the introducer through the sideport. Sheath is flushed. | User omits step
Insert diagnostic catheter over a 0.035" diagnostic guidewire into the pulmonary artery. Diagnostic catheter gains access to heart and is placed over the aortic arch. | Incorrect guidewire used
Exchange 0.035" diagnostic guidewire for 0.018" placement guidewire and advance the placement guidewire deep into either PA branch . Remove diagnostic catheter.. Guidewire gains access to heart and is placed over the aortic arch via diagnostic catheter. | Incorrect guidewire used
Aspirate and flush the sheath using a syringe connected to the sideport. Sheath is flushed. | User omits step
Insert the placement guidewire into the red EasyGuide lumen at the tip of the pigtail. Guidewire is inserted into EasyGuide lumen. | User removes EasyGuide lumen before insertion of guidewire
Insert the placement guidewire into the red EasyGuide lumen at the tip of the pigtail. Guidewire is inserted into EasyGuide lumen. | User kinks EasyGuide lumen and the wire will not advance through it
Advance the guidewire until it exits the red lumen near the label. Guidewire is inserted into EasyGuide lumen. | User removes EasyGuide lumen before guidewire exits pump housing
Remove the EasyGuide lumen by gently pulling the label in line with the catheter shaft while holding the Impella catheter. Guidewire is inserted into EasyGuide lumen. | User pulls EasyGuide lumen away from catheter shaft and fractures lumen
(Backloading without EasyGuide lumen) Wet the cannula with sterile water and backload the catheter onto the placement guidewire. Catheter backloaded over placement guidewire. | User does not know how to backload placement guidewire
(Backloading without EasyGuide lumen) Wet the cannula with sterile water and backload the catheter onto the placement guidewire. Catheter backloaded over placement guidewire. | User does not know how to backload placement guidewire, wire damages outflow cage
(Backloading without EasyGuide lumen, one-person) Advance the guidewire into the Impella Catheter to stabilize the cannula between fingers. The guidewire must exit theoutlet area and align with the straight black line on the catheter. Catheter backloaded over placement guidewire. | User does not know how to backload placement guidewire, wire damages outflow cage
(Backloading without EasyGuide lumen, two-person) The scrub assistant can help stabilize the catheter and the physician can focus on advancing the guidewire to follow the straight black line on catheter. Catheter backloaded over placement guidewire. | User does not know how to backload placement guidewire, wire damages outflow cage
Advance the catheter through the hemostatic valve into the jugular or femoral vein and along the placement guidewire. Impella catheter advances through sheath into vein. | User does not secure placement guidewire and it advances into heart
Advance the catheter through the hemostatic valve into the jugular or femoral vein and along the placement guidewire. Impella catheter advances through sheath into vein. | User pushes too hard and inflow cannula is bent
Advance the catheter through the hemostatic valve into the jugular or femoral vein and along the placement guidewire. Impella catheter advances through sheath into vein. | User does not use fluoro and wire advances too far in ventricle
Advance the catheter through the hemostatic valve into the jugular or femoral vein and along the placement guidewire. Impella catheter advances through sheath into vein. | User pulls on guidewire while advancing the pump and loses access
Follow the catheter under fluroscopy as it is advanced into proper position. Impella catheter is positioned within heart. | User does not have the repo sheath at the blue plug
Remove the placement guidewire. Placement guidewire is removed. | User does not remove placement guidewire before starting pump
Confirm position with fluroscopy . Position is confirmed. | User omits step
Confirm Impella is inserted in patient and guidewire is removed. Impella is inserted and guidewire is removed. | User omits step
Start the Impella Catheter? Yes. Impella catheter starts and begins pumping blood. | User does not assess patient pre-existing conditions
Start the Impella Catheter? Yes. Impella catheter starts and begins pumping blood. | User does not utilize fluoroscopy guidance
Wait 30 seconds for flow to reach its maximum value. Impella catheter starts and begins pumping blood. | User does not monitor catheter placement and it advances too deep while ramping up
Confirm correct and stable placement. AIC provides information on pump placement. | User does not utilize waveforms and fluoro to verify position and pump is malpositioned
Slide the respositioning sheath back to the blue Impella plug . Repositioning sheath is pulled back. | User does not hold Impella catheter in position while removing repositioning sheath from vasculature
Stitch 1st purse string suture around introducer sheath with circular shape while being careful not to puncture introducer sheath – follow contour of introducer about 1-2 mm from its edge, DO NOT TIE at this point. suture is around introducer and not tied. | sheath splits
Stitch 1st purse string suture around introducer sheath with circular shape while being careful not to puncture introducer sheath – follow contour of introducer about 1-2 mm from its edge, DO NOT TIE at this point. suture is around introducer and not tied. | user must remove tied suture and re-suture
Place 2nd purse string suture concentric to 1st suture but about 1-2 mm away from 1st suture line. Leave tying tails of 2nd suture 180 degrees from the tails of the 1st suture, DO NOT TIE yet. suture is around introducer and not tied. | sheath splits
Place 2nd purse string suture concentric to 1st suture but about 1-2 mm away from 1st suture line. Leave tying tails of 2nd suture 180 degrees from the tails of the 1st suture, DO NOT TIE yet. suture is around introducer and not tied. | user must remove tied suture and re-suture
Place 2nd purse string suture concentric to 1st suture but about 1-2 mm away from 1st suture line. Leave tying tails of 2nd suture 180 degrees from the tails of the 1st suture, DO NOT TIE yet. suture is around introducer and not tied. | suture interferes with tieing
Remove peel-away introducer completely from the vein. Peel-away introducer is removed and peeled away from vein. | User peels away sheath through skin in vessel
Remove peel-away introducer completely from the vein. Peel-away introducer is removed and peeled away from vein. | User does not utilize fluoroscopy guidance
Once peel-away is completely removed, hold pressure over the venotomy. Manual compression is applied. | User does not hold pressure
Grasp the two wings and bend until the valve assembly comes apart and until the introducer is completely separated from the catheter shaft. Peel-away introducer is removed and peeled away from vein. | User does not realize that the entire valve does not break and pulls on the catheter
Slide the respositioning sheath over the catheter shaft and advance it into the vein to the suture pad. Repositioning sheath is in place. | User does not hold Impella catheter in position while advancing the repositioning sheath up the shaft
Slide the respositioning sheath over the catheter shaft and advance it into the vein to the suture pad. Repositioning sheath is in place. | User advances the hub to the venotomy
Slide the respositioning sheath over the catheter shaft and advance it into the vein to the suture pad. Repositioning sheath is in place. | User does not use fluoroscopy and the catheter is advanced too far
Tie both purse string sutures so that tissue closes around repo sheath. Sutures are tied around the repo sheath. | sutures interfere with repo advancement
Anchor repositioning sheath to skin via suture pad with suturing that applies forward tension (pulling suture pad/hub towards wound). Repositioning sheath is sutured to patient. | Repositioning unit not secured tightly
Anchor repositioning sheath to skin via suture pad with suturing that applies forward tension (pulling suture pad/hub towards wound). Repositioning sheath is sutured to patient. | User does not maintain correct angle of entry and tenting of the venotomy causes bleeding
Anchor repositioning sheath to skin via suture pad with suturing that applies forward tension (pulling suture pad/hub towards wound). Repositioning sheath is sutured to patient. | Forward suture technique was not used and there is oozing in the unit
Attach the anticontamination sleeve to the klickadapter on the repositioning sheath. Anticontamination sleeve is connected to repositioning sheath. | User does not know how to attach anticontamination sleeve
Lock the anchoring ring in place by turning it clockwise . Anticontamination sleeve is connected to repositioning sheath. | User does not lock the anchoring ring in place
Carefully extend the anticontamination sleeve to maximum length and secure the end closest to the blue Impella plug by tightening the anchoring ring. Anticontamination sleeve is connected to repositioning sheath. | User does not lock the anchoring ring in place
Attach catheter to patient in 3 places using stat locks. cathether remains afixed to patient. | pump moves when patients moves
Zero differential pressure sensor. Differential pressure sensor is zeroed. | Zeroing differential pressure sensor reduces reported flow
Zero differential pressure sensor. Differential pressure sensor is zeroed. | Zeroing differential pressure sensor increases reported flow
Assess right ventricular function. Small changes in right ventricular systolic function as measured by echocardiography may be accompanied by significant improvement in right side forward flow; therefore, it is important to evaluate both echocardiographic evidence of improvement as well as CVP from an appropriate diagnostic tool, flow rate, and overall perfusion.. diagnostic tool is used for CVP measurement. | SmartAssist metric used as diagnostic tool
Initiate the weaning process by temporarily reducing the Impella RP with SmartAssist System Catheter flow to about 2 L/min.. FLOW CONTROL has been seleted and P-level has been reduced. | User does not know how to reduce P-level
Assess right ventricular function. Small changes in right ventricular systolic function as measured by echocardiography may be accompanied by significant improvement in right side forward flow; therefore, it is important to evaluate both echocardiographic evidence of improvement as well as CVP from an appropriate diagnostic tool, flow rate, and overall perfusion.. diagnostic tool is used for CVP measurement. | User does not use diagnostic tool for CVP
Record available information regarding flow rate, CVP from an appropriate diagnostic tool, echo parameters, and systemic hemodynamics.. diagnostic tool is used for CVP measurement. | User does not use diagnostic tool for CVP
After 15–20 minutes at the reduced flow rate, if there are signs of right ventricular recovery and no adverse effects from reduction in flow rate, continue the weaning process by reducing flow rate as tolerated to 0.5 L/min (P-1). At this flow rate there will no longer be any forward flow across the right heart.. 15-20 minutes is waited. | not enough time is waited
If the patient is maintained at a low flow rate (<1.5 L/min) for a prolonged period, increase ACT to at least 250 seconds.. ACT is increased to 250 seconds. | ACT is too low
When ACT is below 150 seconds and patient hemodynamics remain stable, remove initial purse string suture and place new purse string suture, but do not tie off. new suture is placed and not tied. | user must remove tied suture and re-suture
Reduce the P-level to P-0. P-level is reduced to 0. | User does not know how to reduce P-level
Reduce the P-level to P-0. P-level is reduced to 0. | User omits step
Remove the Impella RP IJ with SmartAssist system immediately after stopping the pump.. Impella Catheter is removed. | User applies force to catheter on removal and parts are damaged
Tie off purse string suture. Apply manual compression for 40 minutes or per hospital protocol. Manual compression is applied. | User does not apply enough pressure
Disconnect the connector cable from the AIC and turn the controller off by pressing the power switch on the side of the controller for 3 seconds. Cable is disconnected and controller is turned off. | User does not turn off AIC
Advance repositioning sheath to hemostatic valve and properly align accessory. Repositioning sheath is advanced to hemostatic valve. | User does not maintain catheter position during repositioning advancement, leading to pump migration
Open the ICU Sheath Lock and place it around the underside of the 14Fr Low Profile sheath hub. ICU Sheath Lock placed properly on underside of 14Fr Low Profile Sheath. | Low Profile sheath not properly mated to ICU Sheath Lock
Secure the Impella catheter and align the repositioning sheath in the channel and advance tip of the repositioning sheath to the indicated green region. Repositioning sheath is advanced to the green region on ICU Sheath Lock. | User does not maintain catheter position during repositioning advancement, leading to pump migration
Secure the Impella catheter and align the repositioning sheath in the channel and advance tip of the repositioning sheath to the indicated green region. Repositioning sheath is advanced to the green region on ICU Sheath Lock. | User does not advanced repositioning sheath far enough into accessory
Secure the Impella catheter and align the repositioning sheath in the channel and advance tip of the repositioning sheath to the indicated green region. Repositioning sheath is advanced to the green region on ICU Sheath Lock. | User advances repositioning sheath too far, cannot attach accessory
Secure the Impella catheter and align the repositioning sheath in the channel and advance tip of the repositioning sheath to the indicated green region. Repositioning sheath is advanced to the green region on ICU Sheath Lock. | User advances repositioning sheath too far, cannot reposition
Rotate the orange lever counterclockwise around the repositioning sheath into the locked position. Repositioning sheath secured in proper place with accessory. | User does not fully engage lever - too far out of accessory
Rotate the orange lever counterclockwise around the repositioning sheath into the locked position. Repositioning sheath secured in proper place with accessory. | User rotates lever in clockwise direction - pump not secured
To remove ICU Sheath Lock, rotate the orange lever clockwise, then remove ICU Sheath Lock from 14Fr Low Profile Sheath hub. Accessory removed from sheath hub. | User pulls ICU Sheath Lock away without fully disengaging from sheath hub
To remove ICU Sheath Lock, rotate the orange lever clockwise, then remove ICU Sheath Lock from 14Fr Low Profile Sheath hub. Accessory removed from sheath hub. | User removes accessory prior to support being ended
Setup: Power On (Mains), Connect Pump Driver sheath cable to Console. Pump Driver sheath cable is connected to Console. | No PLU function
Setup: Power On (Mains), Connect Pump Driver sheath cable to Console. Pump Driver is not connected to the Console before the Console is powered on. | Pump Driver connected to the Console after the Console is powered on
Setup: Power On (Mains), Connect line cord to AC wall outlet. Line cord is connected to AC wall outlet.. | Loss of AC mains power
Setup: Power On (Mains), Connect line cord to AC wall outlet. Line cord is connected to AC wall outlet. . | Loss of AC mains power
Setup: Power On (Mains), Turn on the Power button (Standby switch). System is powered on by turning on the Power Button (Standby switch). . | System remains off
Setup: Power On (Mains), Turn on the Power button (Standby switch). System is powered on by turning on the Power Button (Standby switch). . | User does not lift Power button (standby switch) cover
Setup: Power On (Battery), Connect Pump Driver sheath cable to Console. Pump Driver sheath cable is connected to Console. | No PLU Function
Setup: Power On (Battery), Turn on the Power button (Standby switch). System is powered on by turning on the Power Button (Standby switch). . | System remains off
Setup: Power On (Battery), Turn on the Power button (Standby switch). System is powered on by turning on the Power Button (Standby switch). . | User does not lift Power button (standby switch) cover
Setup: Power Off console, Set Pump RPM to 0.. Pump RPM is set to 0. . | No focus is set on "DOWN" element on UI, creating a delay
Setup: Power Off console, Set Pump RPM to 0.. Pump RPM is set to 0. . | Focus on "DOWN" element on UI is cleared, creating a delay
Setup: Power Off console, Set Pump RPM to 0.. Pump RPM is set to 0. . | RPM changes in opposite direction as intended
Setup: Power Off console, Set Pump RPM to 0.. Pump RPM is set to 0. . | Focus on "Pump" element on UI is not cleared, creating a delay
Setup: Power Off console, Set Pump RPM to 0.. Pump RPM is set to 0. . | User cycles the power button (Standby switch) when the pump driver motor is spinning
Setup: Power Off Console, Turn off the power button (Standby switch) for greater than 8 seconds. Console is powered off by turning off the power button (Standby switch) for greater than 8 seconds.. | System remains on because user does not press the power button (standby switch)
Setup: Power Off Console, Turn off the power button (Standby switch) for greater than 8 seconds. Console is powered off by turning off the power button (Standby switch) for greater than 8 seconds.. | User does not lift Power button (standby switch) cover
Transition from Mains Power to Battery Power, Check battery fuel gauge level, Disconnect mains cord. Mains cord is disconnected and system switches to Battery Power. . | system remains on and alarms
Transition from Mains Power to Battery Power, Check battery fuel gauge level, Disconnect mains cord. Mains cord is disconnected and system switches to Battery Power. . | loss of power
Transition from Mains Power to Battery Power, Acknowledge "Running on Battery" alarm. "Running on Battery" alarm acknowledged.. | Alarm continues
External Gas Mode, Connect 1/4" ID tubing from regulator of oxygen source (wall or tank) to Console oxygen inlet. 1/4" tubing from regulator of oxygen source (wall or tank) is connected to Console oxygen inlet.. | reduced O2 flow
External Gas Mode, Connect 1/4" ID tubing from regulator of oxygen source (wall or tank) to Console oxygen inlet. 1/4" tubing from regulator of oxygen source (wall or tank) is connected to Console oxygen inlet.. | blocked O2 flow
External Gas Mode, Connect 1/4" ID tubing from regulator of oxygen source (wall or tank) to Console oxygen inlet. 1/4" tubing from regulator of oxygen source (wall or tank) is connected to Console oxygen inlet.. | toxic gas passed through membrane and transferred to blood
External Gas Mode, Open (wall or tank) valve. Gas valve (wall or tank) is opened. . | no O2 flow
External Gas Mode, Open (wall or tank) valve. Gas valve (wall or tank) is opened. . | reduced O2 flow
External Gas Mode, Set Flow Regulator to desired LPM. Flow Regulator is set to desired LPM.. | reduced O2 flow
External Gas Mode, Set Flow Regulator to desired LPM. Flow Regulator is set to desired LPM.. | excessive O2 flow
External Gas Mode, Set Flow Regulator to desired LPM. Flow Regulator is set to desired LPM.. | Excessive positive pressure applied to membrane.
External Gas Mode, Confirm mode UI - External Gas element highlighted, Keep gas flow exhaust path clear and unrestricted. Gas flow exhaust path is clear and unrestricted.. | Excessive positive pressure applied to membrane.
Priming: OXY-1 System Setup, Securely attach the PLU Mounting Bracket to the IV Pole at the approximate height of 18" (0.45m) from the floor. . PLU Mounting Bracket is securely attached to the IV Pole approximately 18" (0.45m) from the floor. . | Cannot attach PLU
Priming: OXY-1 System Setup, Securely attach the PLU Mounting Bracket to the IV Pole at the approximate height of 18" (0.45m) from the floor. . PLU Mounting Bracket is securely attached to the IV Pole approximately 18" (0.45m) from the floor. . | PLU Disposable (oxygenator) located above patient's head
Priming: OXY-1 System Setup, Securely attach the PLU Mounting Bracket to the IV Pole at the approximate height of 18" (0.45m) from the floor. . PLU Mounting Bracket is securely attached to the IV Pole approximately 18" (0.45m) from the floor. . | Ergonomically difficult to install/remove PLU Disposable
Priming: OXY-1 System Setup, Securely attach the PLU Mounting Bracket to the IV Pole at the approximate height of 18" (0.45m) from the floor. . PLU Mounting Bracket is securely attached to the IV Pole approximately 18" (0.45m) from the floor. . | PLU Mounting bracket slips down during use.
Priming: OXY-1 System Setup, Adjust the PLU Mounting Bracket so the latch faces up (horizontal mounting position).. PLU Mounting Bracket latch is facing up (horizontal mounting position). . | PLU Durable in wrong orientation during PLU Disposable attachment
Priming: OXY-1 System Setup, Attach the Pump Driver to the PLU Mounting Bracket. Leave lid in open position.. Pump Driver is attached to the PLU Mounting Bracket and lid is in open position. . | Pump Driver not properly secured.
Priming: OXY-1 System Setup, Attach the Pump Driver to the PLU Mounting Bracket. Leave lid in open position.. Pump Driver is attached to the PLU Mounting Bracket and lid is in open position. . | Pump Driver lid may not stay open.
Priming: OXY-1 System Setup, Connect Pump Driver sheath cable to Console.. Pump Driver sheath cable is connected to Console. | No PLU Function
Priming: OXY-1 System Setup, Connect line cord to AC wall outlet.. Line cord is connected to AC wall outlet. . | Loss of AC mains power
Priming: OXY-1 System Setup, Turn on the power button (Standby switch). . Power button (Standby switch) is turned on. . | System remains off
Priming: OXY-1 System Setup, Turn on the power button (Standby switch). . Power button (Standby switch) is turned on. . | User does not lift Power button (standby switch) cover
Priming: Priming Kit Setup, Follow aseptic technique. Aseptic technique followed.. | Blood contacting surfaces exposed.
Priming: Priming Kit Setup, Peel off the top cover of the priming kit tub.. Top cover of priming kit tub is peeled off. . | Priming kit tub remains covered.
Priming: Priming Kit Setup, Remove all the pouched accessories and set aside. Do not open pouches.. All pouched accessories are removed and set aside. Pouches remain closed. . | Blood contacting surfaces exposed.
Priming: Priming Kit Setup, Remove the recirculation reservoir and hang it on the IV Pole. Remove the sterile tubing container, priming lines, and PLU Disposable assembly.. Recirculation reservoir is removed from Priming Kit and hung on the IV Pole. Sterile tubing container, priming lines, and PLU Disposable assembly are removed from Priming Kit. . | Recirculation reservoir, sterile tubing container, priming lines, and PLU Disposable assembly remain in Priming Kit .
Priming: Priming Kit Setup, Place the PLU Disposable on the Pump Driver. Lid should remain open.. PLU Disposable is placed on the Pump Driver and lid remains open. . | Pump Driver not properly secured.
Priming: Priming Kit Setup, Attach an included male luer cap to the arterial in-line access port.. Included male luer cap is attached to the arterial in-line access port. . | Arterial in-line access port remains open.
Priming: Priming Kit Setup, If needed, add spare in-line access port to arterial (high pressure side) blood circuit.. Spare in-line access port to arterial (high pressure side) blood circuit are added as needed. . | Connection leak draws in air.
Priming: Priming Kit Setup, On each of the two PLU Disposable access ports, replace the male luer cap with an included pigtail stopcock. (Ensure that the pigtail stopcocks are in the closed position.) Close all 3 clamps on the priming circuit. The two gas lines connected to the PLU Disposable are not used during priming and should be left coiled and secured. On each of the two PLU Disposable access ports, the male luer cap is replaced with an included pigtail stopcock. (Ensuring that the pigtail stopcocks are in the closed position.) All 3 clamps on the priming circuit are closed. The two gas lines connected to the PLU Disposable are not used during priming and are left coiled and secured.. | Disorder of setup creates inconvenience for the user.
Priming: Priming Fluid Prep, Hang the two IV Bags on the IV Pole hooks.. Two IV Bags are hung on the IV Pole hooks. . | Disorder of setup creates inconvenience for the user.
Priming: Priming Fluid Prep, Uncap and spike the line from the top of the recirculation reservoir into IV bag #2. (Ensure that all the priming fluid in IV bag #2 drains into the recirculation reservoir.). IV bag #2 is uncapped and line from the top of the recirculation reservoir is spiked into IV bag #2 (ensuring that all the priming fluid in IV bag #2 drains into the recirculation reservoir). . | Blood contacting surfaces exposed.
Priming: Priming Fluid Prep, Uncap and spike IV bag #1 (venous reservoir) with venous line IV spike. Unclamp #1. IV bag #1 is uncapped and spiked with the venous line IV spike. #1 is unclamped. . | Blood contacting surfaces exposed.
Priming: Priming Fluid Prep, Allow the recirculation bag to fill with fluid. Periodically squeeze bag to displace bubbles.. The recirculation bag fills with fluid. The bag is periodically squeezed to displace bubbles. . | Blood contacting surfaces exposed.
Priming: Priming Fluid Prep, Move the sterile tubing container onto a flat horizontal surface.. Sterile tubing container is moved onto a flat horizontal surface. . | Disorder of setup creates inconvenience for the user.
Priming: Venous, PLU, and Arterial, Unclamp the pinch clamps on the blue venous priming lines. (Ensure that priming fluid is flowing from each reservoir.). The pinch clamps on the blue venous priming lines are unclamped (ensuring that priming fluid is flowing from each reservoir). . | Disorder of setup creates delay for the user.
Priming: Venous, PLU, and Arterial, Unscrew the luer cap on the arterial priming line access port.. The luer cap on the arterial priming line access port is unscrewed. . | Blood contacting surfaces exposed.
Priming: Venous, PLU, and Arterial, Use a container or the priming kit tub to catch the expelled priming fluid from the arterial in-line access port.. Using a container or the priming tube kit, expelled priming fluid from the arterial in-line access port is caught. . | Disorder of setup creates inconvenience for the user.
Priming: Venous, PLU, and Arterial, Keep the arterial in-line access port open until most of the air bubbles have escaped and a constant flow of priming fluid is achieved.. Arterial in-line access port is kept open until most of the air bubbles have escaped and a constant flow of priming fluid is achieved.. | Disorder of setup creates delay for the user.
Priming: Venous, PLU, and Arterial, Unclamp the pinch clamp upstream of the arterial in-line access port.. The pinch clamp upstream of the arterial in-line access port is unclamped. . | Disorder of setup creates delay for the user.
Priming: Disposable De-Airing, Lift the PLU Disposable from the Pump Driver. Inspect the PLU Disposable for any trapped air bubbles.. PLU Disposable is inspected for any trapped air bubbles.. | Disorder of setup creates delay for the user.
Priming: Disposable De-Airing, Shake, tap, and manipulate the PLU Disposable with the ports facing up to uncover any air bubbles.. PLU Disposable is shaked, tapped, and manipulated with the ports facing up to uncover any air bubbles.. | Damage to PLU Disposable
Priming: Disposable De-Airing, Open the stopcock on one pigtail to release the bubbles. Use the tub to catch priming fluid. Close the stopcock. Repeat this process for the other pigtail.. Stopcock on one pigtail is opened to release bubbles. Tub is used to catch priming fluid. Stopcock is closed. Process is repeated for the other pigtail.. | Disorder of setup creates delay for the user.
Priming: Disposable De-Airing, Open the stopcock on one pigtail to release the bubbles. Use the tub to catch priming fluid. Close the stopcock. Repeat this process for the other pigtail.. Stopcock on one pigtail is opened to release bubbles. Tub is used to catch priming fluid. Stopcock is closed. Process is repeated for the other pigtail.. | User breaks pigtail stopcock
Priming: Disposable De-Airing, Repeat steps 9 and 10 until no air bubbles are visible. Place the PLU Disposable back onto the Pump Driver.. All visible air bubbles are removed and the PLU Disposable is placed back onto the Pump Driver. . | Pump Driver not properly secured.
Priming: Pump Circulation De-Airing, Squeeze the priming fluid from the venous reservoir bag into the recirculation reservoir until it is fully drained. Clamp the venous reservoir priming line so that it is removed from the circuit.. Priming fluid from venous reservoir bag is fully drained into recirculation reservoir. The venous reservoir priming line is clamped. . | Disorder of setup creates delay for the user.
Priming: Pump Circulation De-Airing, Repeat the inspection of the PLU Disposable to uncover any remaining air bubbles and release them through the PLU pigtail stopcocks.. Remaining air bubbles are released from PLU Disposable through PLU pigtail stopcocks. . | Arterial bloodstream bubbles
Priming: Pump Circulation De-Airing, Repeat the inspection of the PLU Disposable to uncover any remaining air bubbles and release them through the PLU pigtail stopcocks.. Remaining air bubbles are released from PLU Disposable through PLU pigtail stopcocks. . | Pump Bearing Wear
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to Pump Driver.
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driverand arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Lid closes before PLU Disposable is placed back onto Pump Driver
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Decreased or no blood flow
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Latch lever cannot close. Flow Sensor reading error
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Disorder of setup creates delay for the user.
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Top Latch plate cannot close far enough for latch arm to engage
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | Disposable not secure.
Priming: Pump Circulation De-Airing, Place the PLU Disposable back onto the Pump Driver. Route the arterial and venous lines through the flow sensor channels. Close the lid and engage the latch.. PLU Disposable is placed back on Pump Driver and arterial and venous lines are routed through flow sensor channels. Lid is closed and latch is engaged. . | FLow Sensor reading error.
Priming: Pump Circulation De-Airing, Hang the sterile tubing container on the IV pole hooks.. Sterile tubing contains is hung on IV pole hooks. | Disorder of setup creates inconvenience for the user.
Priming: Pump Circulation De-Airing, Using the Console user interface, set the Pump RPM to create a Blood Flow rate of approximately 2.0 LPM.. Pump RPM is set to create a Blood Flow rate of approximately 2.0 LPM.. | Disorder of setup creates delay for the user.
Priming: Pump Circulation De-Airing, Allow the priming fluid to circulate for several minutes. Engage the red arterial priming line clamp slightly to create resistance in the system. Adjust the Pump RPM as needed to maintain the desired Blood Flow rate.. Priming fluid circulates for several minutes. Red arterial priming line is clamped slightly to create resistance in the system. The Pump RPM is adjusted as needed to maintain the desired Blood Flow rate. . | Cavitation
Priming: Pump Circulation De-Airing, Continue to inspect the circuit for bubbles and take the appropriate measures to release them either through the pigtail stopcocks or the recirculation reservoir.. Circuit is inspected for bubbles and bubbles are removed using appropriate measures to release the bubbles through either the pigtail stopcocks or the recirculation reservoir. . | Arterial bloodstream bubbles
Priming: Pump Circulation De-Airing, Continue to inspect the circuit for bubbles and take the appropriate measures to release them either through the pigtail stopcocks or the recirculation reservoir.. Circuit is inspected for bubbles and bubbles are removed using appropriate measures to release the bubbles through either the pigtail stopcocks or the recirculation reservoir. . | Venous bloodstream bubbles
Priming: Pump Circulation De-Airing, Continue to inspect the circuit for bubbles and take the appropriate measures to release them either through the pigtail stopcocks or the recirculation reservoir.. Circuit is inspected for bubbles and bubbles are removed using appropriate measures to release the bubbles through either the pigtail stopcocks or the recirculation reservoir. . | User breaks pigtail stopcock
Priming: Pump Circulation De-Airing, Inspect the circuit for leaks. Tighten loose fittings and connections as needed.. Circuit is inspected for leaks. Loose fittings and connections are tightened as needed. . | leaking priming fluid and/or blood
Priming: Pump Circulation De-Airing, Set the Pump RPM to 0. Turn off the Console power button (standby switch).. The Pump RPM is set to 0 and Console power button (standby switch) is turned off. . | Continuous circulation of priming fluid
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Disorder of setup creates inconvenience for the user.
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Weak connection on O2 line
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Air feeding inner chamber when on internal gas
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Weak connection on Air line
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | O2 feeding outer chamber when on internal gas.
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Zipper does not open.
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Insufficient strain relief on Air and O2 lines
Priming: Pump Circulation De-Airing, If the Pump Driver and Console are to be put into use with the newly primed PLU Disposable, connect the gas lines to the Console, place the sheath over the gas lines and electrical cable, and attach the sheath tabs at each end to the Pump Driver and the Console.. If the Pump Driver and Console are being put into use with the newly primed PLU Disposable, the gas lines are connected to the Console, the sheath is placed over the gas lines and electrical cable, and the sheath tabs are attached at each end to the Pump Driver and the Console. . | Insufficient strain relief on Air and CO2 lines
Operator: Attach PLU Disposable to Pump Driver, Raise latch lever. Release latch arm from latch plate. Open top latch plate. . Latch lever is raised. Latch arm is released from latch plate. Top latch plate is opened. . | Disorder of setup creates inconvenience for the user.
Operator: Attach PLU Disposable to Pump Driver, Open lid far enough to engage hinge detents.. Lid is opened far enough to engage hinge detents. . | lid closes
Operator: Attach PLU Disposable to Pump Driver, While aligning pump motor cup to pump and condensation tray to oxygenator, place PLU disposable onto Pump Driver. . While aligning pump motor cup to pump and condensation tray to oxygenator, PLU Disposable is placed onto Pump Driver. . | Decreased or no blood flow
Operator: Attach PLU Disposable to Pump Driver, Route Air and O2 lines below latch lever.. Air and O2 lines are routed below latch lever. . | latch lever cannot close. Flow Sensor reading error
Operator: Attach PLU Disposable to Pump Driver, Press red Arterial Blood tube into sensor channel.. Red Arterial Blood tube is pressed into sensor channel. . | Disorder of setup creates delay for the user.
Operator: Attach PLU Disposable to Pump Driver, Press red Arterial Blood tube into sensor channel.. Red Arterial Blood tube is pressed into sensor channel. . | Top Latch plate cannot close far enough for latch arm to engage
Operator: Attach PLU Disposable to Pump Driver, Press blue Venous Blood tube into channel.. Blue Venous Blood tube is pressed into channel. . | Disorder of setup creates delay for the user.
Operator: Attach PLU Disposable to Pump Driver, Press blue Venous Blood tube into channel.. Blue Venous Blood tube is pressed into channel. . | Top Latch plate cannot close far enough for latch arm to engage
Operator: Attach PLU Disposable to Pump Driver, Close lid.. Lid is closed. . | Disposable not secure.
Operator: Attach PLU Disposable to Pump Driver, Close top latch plate. Engage latch arm to top latch plate. . Top latch plate is closed. Latch arm is engaged to top latch plate. . | Disorder of setup creates delay for the user.
Operator: Attach PLU Disposable to Pump Driver, Lower latch lever.. Latch level is lowered. . | FLow Sensor reading error.
Operator: Remove PLU Disposable from Pump Driver, Raise latch lever. Release latch arm from latch plate. Open top latch plate. Open lid far enough to engage hinge detents. Pull blue Venous Blood tube from channel. Pull red Arterial Blood tube from sensor channel. Route Air and O2 lines above latch lever. Remove PLU disposable from Pump Driver. Close lid. Close top latch plate. Engage latch arm to top latch plate. Lower latch lever. . PLU Disposable is removed from Pump Driver. . | Disorder of setup creates delay for the user.
Operator: Attach PLU sheath/Tubes to Console, Route 1/4" ID Air and 3/16" ID O2 tubes along same path as the electrical cable from PLU to Console.. 1/4" ID Air and 3/16" ID O2 tubes are routed along the same path as the electrical cable from PLU to Console.. | Disorder of setup creates inconvenience for the user.
Operator: Attach PLU sheath/Tubes to Console, At Console, connect 3/16" ID O2 tube to 3/16" barbed fitting.. 3/16" ID O2 tube is connected to 3/16" barbed fitting. . | Weak connection on O2 line
Operator: Attach PLU sheath/Tubes to Console, At Console, connect 3/16" ID O2 tube to 3/16" barbed fitting.. At Console, 3/16" ID O2 tube is connected to 3/16" barbed fitting. . | Air feeding inner chamber when on internal gas
Operator: Attach PLU sheath/Tubes to Console, At Console, connect 1/4" ID Air tube to 1/4" barbed fitting.. At Console, 1/4" ID Air tube is connected to 1/4" barbed fitting. . | Weak connection on Air line
Operator: Attach PLU sheath/Tubes to Console, At Console, connect 1/4" ID Air tube to 1/4" barbed fitting.. At Console, 1/4" ID Air tube is connected to 1/4" barbed fitting. . | O2 feeding outer chamber when on internal gas.
Operator: Attach PLU sheath/Tubes to Console, Open zipper on sheath. Zipper on sheath is opened. . | Zipper does not open.
Operator: Attach PLU sheath/Tubes to Console, Attached sheath tab to Console button.. sheath tab is attached to Console button. . | Insufficient strain relief on Air and O2 lines
Operator: Attach PLU sheath/Tubes to Console, Attached sheath tab to Console button. Wrap sheath around air tubes and electrical cable and start zipper at Console end. Zip sheath zipper along the entire length of cable and tubing. . sheath tab is attached to Console button. sheath is wrapped around air tubes and electrical cable and zipped along entire length of cable and tubing. . | Insufficient strain relief on Air and O2 lines
Operator: Attach PLU sheath/Tubes to Console, Attach sheath tab to PLU button.. sheath tab is attached to PLU button. . | Insufficient strain relief on Air and O2 lines
Operator: Attach PLU sheath/Tubes to Console, At Console, attach hook/loop strap to sheath. Hook/loop strap is attached to sheath. . | Insufficient strain relief on Air and CO2 lines
Operator: Remove PLU sheath/Tubes from Console, At Console, open hook/loop strap from sheath. Remove sheath tab from PLU button. Starting from PLU end, unzip the sheath zipper along the entire length of cable and tubing. Unwrap sheath from air tubes and electrical cable. Remove sheath tab from Console button. At Console, disconnect 1/4" ID Air tube from 1/4" barbed fitting. At Console, disconnect 3/16" ID O2 tube from 3/16" barbed fitting. Remove 1/4" ID Air and 3/16" ID O2 tubes. . At Console, hook/loop strap from sheath is opened. sheath tab is removed from PLU button. Starting from the PLU end, the sheath zipper is unzipped along the entire length of cable and tubing. The sheath is unwrapped from air tubes and electrical cable. sheath tab is removed from Console button. At Console, 1/4" ID Air tube is disconnected from 1/4" barbed fitting. At Console, 3/16" ID O2 tubing is disconnected from 3/16" barbed fitting. 1/4" ID Air and 3/16" ID O2 tubes are removed. . | Disorder of setup creates delay for the user.
Setup: Pneumatics Self-Check, Remove external gas source. Set Internal Gas, Air Flow to 15 LPM. Set Internal Gas, O2 Flow to 3 LPM. . External gas source is removed. Internal Gas, Air Flow is set to 15 LPM. Internal Gas, O2 Flow is set to 3 LPM. . | Disorder of setup creates delay for the user.
Setup: Pneumatics Self-Check, Clamp or pinch both Air and O2 lines. Both Air and O2 lines are clamped or pinched. . | Flow is not affected enough to generate error
Setup: Pneumatics Self-Check, Wait for system to detect flow errors. User waits for system to detect flow errors.. | Disorder of setup creates inconvenience for the user.
Setup: Pneumatics Self-Check, Observe UI Alarm "Air Flow Failure". UI Alarm "Air Flow Failure" is observed. . | Disorder of setup creates inconvenience for the user.
Setup: Pneumatics Self-Check, Observe UI Alarm "Concentrator Flow Failure". UI Alarm "Concentrator Flow Failure" is observed. . | Disorder of setup creates inconvenience for the user.
Setup: Pneumatics Self-Check, Release clamp or pinch on both Air and O2 lines. Clamp or pinch on both Air and O2 lines are released. . | Low gas flow
Setup: Pneumatics Self-Check, Observe UI Alarm condition is cleared. Cleared UI Alarm condition is observed.. | Disorder of setup creates inconvenience for the user.
Setup: Attach E-Drive to IV Pole, Place E-Drive bracket around IV Pole. E-Drive bracket is placed around IV Pole. . | Disorder of setup creates delay for the user.
Setup: Attach E-Drive to IV Pole, Set E-Drive bracket height to 18" from floor.. E-Drive bracket height is set to 18" from floor. . | Ergonomically difficult to crank E-Drive during use.
Setup: Attach E-Drive to IV Pole, Set E-Drive bracket height to 18" from floor.. E-Drive is attached to IV Pole, E-Drive bracket height is set to 18" from floor. . | During emergency, PLU Disposable located above patient's heart.
Setup: Attach E-Drive to IV Pole, Set E-Drive bracket height to 18" from floor.. E-Drive bracket height is set to 18" from floor. . | During emergency, PLU Disposable located above patient's heart.
Setup: Attach E-Drive to IV Pole, Tighten E-Drive bracket to IV Pole. E-Drive bracket is tightened to IV Pole. . | E-Drive bracket slips down during use.
Setup: Attach E-Drive to IV Pole, Place E-Drive into E-Drive mounting bracket.. E-Drive is placed into E-Drive mounting bracket. . | Disorder of setup creates delay for the user.
Setup: Attach E-Drive to Console, Place E-Drive into E-Drive mounting bracket.. E-Drive is placed into E-Drive mounting bracket. . | Disorder of setup creates delay for the user.
Setup: present primed system to surgeon, Clamp the patient side Arterial and Venous tubing near the sterile container. Patient side Arterial and Venous tubing is clamped near the sterile container. . | Disorder of setup creates delay for the user.
Setup: present primed system to surgeon, Remove the sterile tubing container from the IV pole hooks.. Sterile tubing container is removed from IV pole hooks. . | Disorder of setup creates delay for the user.
Setup: present primed system to surgeon, Open sterile tubing container.. Sterile tubing container is opened. . | Contamination of sterile surfaces
Setup: present primed system to surgeon, Present sterile tubing to surgeon. Surgeon clamps and cuts tubing inside container. Surgeon removes tubing from container. Bring tubing ends of the primed circuit near the ends of the cannulas. . Tubing inside container is clamped and cut by surgeon. Tubing is removed from container by surgeon. Tubing ends of primed circuit was brought near the ends of the cannulas. . | Disorder of setup creates delay for the user.
Setup: present primed system to surgeon, Have a helper push a saline flush into both ends to fill any air left in either tube.. Helper pushes a saline flush into both ends and any air left in either tube is filled. . | Arterial/Venous bloodstream bubbles
Setup: present primed system to surgeon, Mate the PLU arterial tubing to the patient arterial cannula tubing connector. PLU arterial tubing is mated to the patient arterial cannula tubing connector. . | leaking priming fluid and/or blood
Setup: present primed system to surgeon, Mate the PLU arterial tubing to the patient arterial cannula tubing connector. PLU arterial tubing is mated to the patient arterial cannula tubing connector. . | Reduced blood flow and reduced gas transfer into the blood
Setup: present primed system to surgeon, Mate the PLU venous tubing to the patient venous cannula tubing connector. PLU venous tubing is mated to the patient venous cannula tubing connector. . | Connection leak draws in air.
Setup: present primed system to surgeon, Mate the PLU venous tubing to the patient venous cannula tubing connector. PLU venous tubing is mated to the patient venous cannula tubing connector. . | Reduced blood flow and reduced gas transfer into the blood
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | reduced O2 flow
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | blocked O2 flow
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | toxic gas passed through membrane and transferred to blood
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | no O2 flow
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | excessive O2 flow
Operator: Initiate Therapy, Configure Console for External Gas Mode. Console is configured for External Gas Mode. . | Excessive positive pressure applied to membrane.
Operator: Initiate Therapy, Confirm bubble detection settings . Bubble detection settings are confirmed. . | Bubble alarms disabled
Operator: Initiate Therapy, Gradually increase Blood Pump RPM. Blood Pump RPM is gradually increased. . | Retrograde flow
Operator: Initiate Therapy, Gradually increase Blood Pump RPM. Blood Pump RPM is gradually increased. . | Pump cavitation
Operator: Initiate Therapy, Remove clamps (simultaneous with previous step). Clamps are removed (as Blood Pump RPM is simultaneously gradually increased). . | Pump Cavitation
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | No focus is set on "Pump" element on UI, creating a delay
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | Focus on "UP" or "DOWN" element on UI is cleared, creating a delay
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | RPM changes in opposite direction as intended
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | RPM changes faster than intended
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | Excessive positive/negative pressure applied to patient's blood stream.
Operator: Initiate Therapy, Adjust Blood Pump RPM to achieve desired Blood flow. Blood Pump RPM is adjusted to achieve desired Blood Flow. | Focus on "Pump" element on UI is not cleared, creating a delay
Operator: Initiate Therapy, Adjust upper blood flow alarm limit, as needed. . Upper blood flow alarm limit is adjusted as needed. . | Focus on element on UI is not set, creating a delay
Operator: Initiate Therapy, Adjust upper blood flow alarm limit, as needed. . Upper blood flow alarm limit is adjusted as needed. . | Focus on element on UI is cleared, creating a delay
Operator: Initiate Therapy, Adjust upper blood flow alarm limit, as needed. . Upper blood flow alarm limit is adjusted as needed. . | value changes in opposite direction as intended
Operator: Initiate Therapy, Adjust upper blood flow alarm limit, as needed. . Upper blood flow alarm limit is adjusted as needed. . | value changes faster than intended
Operator: Initiate Therapy, Adjust upper blood flow alarm limit, as needed. . Upper blood flow alarm limit is adjusted as needed. . | focus on element in UI is not cleared, creating a delay
Operator: Initiate Therapy, Adjust lower blood flow alarm limit, as needed. Lower blood flow alarm limit is adjusted as needed. . | Focus on element on UI is not set, creating a delay
Operator: Initiate Therapy, Adjust lower blood flow alarm limit, as needed. Lower blood flow alarm limit is adjusted as needed. . | focus on element in UI is cleared, creating a delay
Operator: Initiate Therapy, Adjust lower blood flow alarm limit, as needed. Lower blood flow alarm limit is adjusted as needed. . | value changes in opposite direction as intended
Operator: Initiate Therapy, Adjust lower blood flow alarm limit, as needed. Lower blood flow alarm limit is adjusted as needed. . | value changes faster than intended
Operator: Initiate Therapy, Adjust lower blood flow alarm limit, as needed. Lower blood flow alarm limit is adjusted as needed. . | focus on element in UI is not cleared, creating a delay
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Operator activates manual sequence for Periodic Sighing. Manual sequence for Periodic Sighing is correctly activated by operator. . | gas flow limited by excessive condensation
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Operator activates manual sequence for Periodic Sighing. Manual sequence for Periodic Sighing is correctly activated by operator. . | excessive CO2 removal
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Operator activates manual sequence for Periodic Sighing. Manual sequence for Periodic Sighing is correctly activated by operator. . | reduced O2 flow
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Remove Condensate Tray from Pump Driver to check for saturation. Condensate Tray is removed from Pump Driver and checked for saturation. . | Saturated Condensate Tray. Water drips on floor
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), If not saturated, re-install condensate tray. If not saturated, the Condensate Tray is re-installed. . | User does not re-install Condensate Tray. Condensate Tray missing. Water drips on floor.
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), If saturated, install new Condensate Tray. If saturated, a new Condensate Tray is installed. . | Condensate Tray missing. Water drips on floor.
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Keep Condensate Tray in protective packaging and store in a dry place prior to patient use. Install just prior to patient use.. Condensate Tray is kept in protective packaging and store in a dry place prior to patient use.. | User stores Condensate Tray. In the pump driver. Condensate tray may become damp and/or contaminated if stored in pump driver for extended periods prior to application to the patient.
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Adjust Internal Gas - Air Flow setting to "match" external setting. Internal Gas - Air Flow setting is adjusted to "match" external setting. . | If external gas lost, wrong sweep gas setting.
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Adjust Internal Gas - O2 Flow setting to "match" external setting . Internal Gas - O2 Flow setting is adjusted to "match" external setting. . | If external gas lost, no oxygenation gas.
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Maintain Air and O2 connections from PLU Disposable to Console. Air and O2 connections from PLU Disposable to Console are maintained. . | No sweep gas to outer chamber of oxygenator
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Maintain Air and O2 connections from PLU Disposable to Console. Air and O2 connections from PLU Disposable to Console are maintained. . | No oxygen gas to inner chamber of oxygenator
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Keep power button (standby switch) ON. Power button (Standby switch) remains ON. . | User turns Power button (Standby switch) off for greater than 8 seconds
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Access High Pressure Side Port. High pressure side port is opened.. | User breaks pigtail stopcock
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Access High Pressure Side Port. High pressure side port is opened.. | blood leaks from port
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Avoid spillage of liquids on device.. Spillage of liquids on device is avoided. . | Device failure
Operator: Maintain Therapy (Stable patient, AC Power, External Gas Mode), Check Air Inlet Filter. Air Inlet Filter is visually checked and replaced if there is any significant amount of lint and debris covering filter surface . | Air filter is not visually checked and not replaced
Operator: Terminate Therapy, Double clamp Arterial and Venous blood lines near patient.. Arterial and Venous blood lines near patient are double clamped. . | leaking blood
Operator: Terminate Therapy, Set Pump RPM to 0. Pump RPM is set to 0.. | Pump running with tubing clamped
Operator: Terminate Therapy, Cut the tubing between each pair of clamps. Tubing between each pair of clamps is cut. . | Disorder of procedure creates delay for user
Operator: Terminate Therapy, If external gas mode, set the Wall/Tank Flow Meter to zero. If internal gas mode, set Air/O2 flow settings to zero.. If external gas mode, the Wall/Tank Flow Meter is set to zero. If internal gas mode, the Air/O2 flow settings are set to zero. . | Disorder of procedure creates delay for user
Operator: Terminate Therapy, Turn off the power button (Standby switch) for greater than 8 seconds. power button (Standby switch) is turned off for greater than 8 seconds.. | System remains on because user does not press the power button (standby switch)
Operator: Terminate Therapy, Remove PLU Sheath/Tubes from Console. PLU Sheath/Tubes are removed from Console. . | Disorder of setup creates delay for the user.
Operator: Terminate Therapy, Remove PLU Disposable from Pump Driver. PLU Disposable is removed from Pump Driver. . | Disorder of setup creates delay for the user.
Operator Settings: Enter Settings Screen, Touch "Settings" icon. "Settings" icon is touched. . | Settings screen not opened
Operator Settings: Enter Settings Screen, Touch "Settings" icon. "Settings" icon is touched. . | "Lock" prompt is displayed
Operator Settings: Exit Settings Screen, Touch "Home" icon. "Home" icon is touched.. | Settings screen persists
Operator Setting: Alarm Volume, Touch "Settings" icon. "Settings" icon is touched. . | Settings screen not opened
Operator Setting: Alarm Volume, Touch "Settings" icon. "Settings" icon is touched. . | "Lock" prompt is displayed
Operator Setting: Alarm Volume, Touch "Alarm Volume" on the settings screen. "Alarm Volume" on the settings screen is touched. . | No focus set
Operator Setting: Alarm Volume, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | focus cleared
Operator Setting: Alarm Volume, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes in opposite direction as intended
Operator Setting: Alarm Volume, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes faster than intended
Operator Setting: Alarm Volume, Touch any other area to remove focus from "Alarm Volume". Any other area is touched to remove focus from "Alarm Volume". . | focus not cleared
Operator Setting: Alarm Volume, Touch "Home" icon. "Home" icon is touched.. | Settings screen persists
Operator Settings: Bubble Detection Size/Off, Touch "Settings" icon. "Settings" icon is touched.. | Settings screen not opened
Operator Settings: Bubble Detection Size/Off, Touch "Settings" icon. "Settings" icon is touched.. | "Lock" prompt is displayed
Operator Settings: Bubble Detection Size/Off, Touch "Bubble Size" on the settings screen. "Bubble Size" touched on settings screen.. | No focus set
Operator Settings: Bubble Detection Size/Off, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | focus cleared
Operator Settings: Bubble Detection Size/Off, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | Setting changes in opposite direction as intended
Operator Settings: Bubble Detection Size/Off, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | Bubble alarms disabled
Operator Settings: Bubble Detection Size/Off, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | Setting changes faster than intended
Operator Settings: Bubble Detection Size/Off, Touch any other area to remove focus from "Bubble Size". Any other area is touched to remove focus from "Bubble Size".. | focus not cleared
Operator Settings: Bubble Detection Size/Off, Touch "Home" icon. "Home" icon is touched.. | Settings screen persists
Operator Settings: Time/Date, Touch "Settings" icon (twice). "Settings" icon is touched twice.. | Settings screen not opened
Operator Settings: Time/Date, Touch "Settings" icon (twice). "Settings" icon is touched twice.. | "Lock" prompt is displayed
Operator Settings: Time/Date, Touch the desired "Time/Date" element on the settings screen. Desired "Time/Date" element is touched on the settings screen.. | No focus set
Operator Settings: Time/Date, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | focus cleared
Operator Settings: Time/Date, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | Setting changes in opposite direction as intended
Operator Settings: Time/Date, Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | Setting changes faster than intended
Operator Settings: Time/Date, Touch "Set" button to commit changes.. "Set" button is touched to commit changes.. | Date/Time not changed
Operator Settings: Time/Date, Touch "Home" icon. "Home" icon is touched.. | Settings screen persists
Operator Setting: Blood Pump Speed (RPM), Touch "Pump" element on UI to set focus for adjustment arrows. "Pump" element is touched on UI to set focus for adjustment arrows.. | No focus set
Operator Setting: Blood Pump Speed (RPM), Touch "UP" or "DOWN" controls on UI to change RPM value. "UP" or "DOWN" controls are touched on UI to change RPM value.. | focus cleared
Operator Setting: Blood Pump Speed (RPM), Touch "UP" or "DOWN" controls on UI to change RPM value. "UP" or "DOWN" controls are touched on UI to change RPM value.. | RPM changes in opposite direction as intended
Operator Setting: Blood Pump Speed (RPM), Touch "UP" or "DOWN" controls on UI to change RPM value. "UP" or "DOWN" controls are touched on UI to change RPM value.. | RPM changes faster than intended
Operator Setting: Blood Pump Speed (RPM), Stop RPM adjustments when desired blood flow is achieved.. RPM adjustments are stopped when desired blood fllow is achieved.. | Excessive positive/negative pressure applied to patient's blood stream.
Operator Setting: Blood Pump Speed (RPM), Touch any other area to remove focus from "Pump". Any other area is touched to remove focus from "Pump".. | focus not cleared
Operator Setting: Blood Pump Stop (0 RPM), Touch "Pump" element on UI to set focus for adjustment arrows. "Pump" element is touched on UI to set focus for adjustment arrows.. | No focus set
Operator Setting: Blood Pump Stop (0 RPM), Touch and hold "DOWN" control on UI to change RPM value. "DOWN" control is touched and held on UI to change RPM value.. | focus cleared
Operator Setting: Blood Pump Stop (0 RPM), Touch and hold "DOWN" control on UI to change RPM value. "DOWN" control is touched and held on UI to change RPM value.. | RPM changes in opposite direction as intended
Operator Setting: Blood Pump Stop (0 RPM), Touch any other area to remove focus from "Pump". Any other area is touched to remove focus from "Pump".. | focus not cleared
Operator Setting: Blood Flow Alarm Limit (Upper), Touch "Upper Alarm Limit" element on UI to set focus for adjustment arrows. "Upper Alarm Limit" element is touched on UI to set focus for adjustment arrows.. | No focus set
Operator Setting: Blood Flow Alarm Limit (Upper), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | focus cleared
Operator Setting: Blood Flow Alarm Limit (Upper), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes in opposite direction as intended
Operator Setting: Blood Flow Alarm Limit (Upper), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes faster than intended
Operator Setting: Blood Flow Alarm Limit (Upper), Touch any other area to remove focus from "Upper Alarm Limit". Any other area is touched to remove focus from "Upper Alarm Limit".. | focus not cleared
Operator Setting: Blood Flow Alarm Limit (Lower), Touch "Lower Alarm Limit" element on UI to set focus for adjustment arrows. "Lower Alarm Limit" element is touched on UI to set focus for adjustment arrows.. | No focus set
Operator Setting: Blood Flow Alarm Limit (Lower), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | focus cleared
Operator Setting: Blood Flow Alarm Limit (Lower), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes in opposite direction as intended
Operator Setting: Blood Flow Alarm Limit (Lower), Touch "UP" or "DOWN" controls on UI to change value. "UP" or "DOWN" controls are touched on UI to change value.. | value changes faster than intended
Operator Setting: Blood Flow Alarm Limit (Lower), Touch any other area to remove focus from "Lower Alarm Limit". Any other area is touched to remove focus from "Lower Alarm Limit".. | focus not cleared
Operator Setting: Internal Gas, Air Flow, Touch "Air" element on UI to set focus for adjustment arrows. "Air" element on UI touched to set focus for adjustment arrows.. | No focus set
Operator Setting: Internal Gas, Air Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. "UP" or "DOWN" controls are touched on UI to change LPM value.. | focus cleared
Operator Setting: Internal Gas, Air Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. "UP" or "DOWN" controls are touched on UI to change LPM value.. | LPM changes in opposite direction as intended
Operator Setting: Internal Gas, Air Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. Any other area is touched to remove focus from "Air".. | LPM changes faster than intended
Operator Setting: Internal Gas, Air Flow, Touch any other area to remove focus from "Air". Any other area is touched to remove focus from "Air".. | focus not cleared
Operator Setting: Internal Gas, O2 Flow, Touch "O2" element on UI to set focus for adjustment arrows. "O2" element is touched on UI to set focus for adjustment arrows.. | No focus set
Operator Setting: Internal Gas, O2 Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. "UP" or "DOWN" controls are touched on UI to change LPM value.. | focus cleared
Operator Setting: Internal Gas, O2 Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. "UP" or "DOWN" controls are touched on UI to change LPM value.. | LPM changes in opposite direction as intended
Operator Setting: Internal Gas, O2 Flow, Touch "UP" or "DOWN" controls on UI to change LPM value. "UP" or "DOWN" controls are touched on UI to change LPM value.. | LPM changes faster than intended
Operator Setting: Internal Gas, O2 Flow, Touch any other area to remove focus from "O2". Any other area is touched to remove focus from "O2".. | focus not cleared
Operator Setting: Initiate O2 Conc. Warmup, Touch "O2 Conc" icon. "O2 Conc" icon is touched.. | Cannot initiate O2 Conc warmup
Operator Setting: Initiate O2 Conc. Warmup, Touch "O2 Conc" icon. "O2 Conc" icon is touched.. | O2 Conc control clears
Operator Setting: Initiate O2 Conc. Warmup, Touch "On/Off" control to enable warmup. "On/Off" control is touched to enable warmup.. | Cannot initiate O2 Conc warmup
Operator Setting: Initiate O2 Conc. Warmup, Touch "On/Off" control to enable warmup. "On/Off" control is touched to enable warmup.. | Warmup aborted
Operator Setting: Initiate O2 Conc. Warmup, Wait for warmup complete indicator. Warmup complete indicator is displayed.. | warm-up not complete
Operator Setting: Initiate O2 Conc. Warmup, Wait for warmup complete indicator. Warmup complete indicator is displayed.. | O2 Concentrator runs indefinitely at 0.5 LPM
Operator Setting: Abort O2 Conc. Warmup, Touch "O2 Conc" icon. "O2 Conc" icon is touched.. | Cannot abort warm-up
Operator Setting: Abort O2 Conc. Warmup, Touch "O2 Conc" icon. "O2 Conc" icon is touched.. | Warm-up control clears
Operator Setting: Abort O2 Conc. Warmup, Touch "On/Off" control to abort warm-up. "On/Off" control is touched to abort warm-up.. | Warm-up continues
Operator Setting: Abort O2 Conc. Warmup, Touch "On/Off" control to abort warm-up. "On/Off" control is touched to abort warm-up.. | Warm-up re-starts
Operator: Transition from External Gas to Internal Gas Mode, Set Internal Gas, Air Flow - See UF032. Internal Gas Air Flow is set.. | CO2 transfer is insufficient.
Operator: Transition from External Gas to Internal Gas Mode, Set Internal Gas, O2 Flow - See UF033. Internal Gas O2 Flow is set.. | Oxygen transfer is insufficient.
Operator: Transition from External Gas to Internal Gas Mode, Initiate O2 Conc. Warmup - See UF034. O2 conc. Warmup is initiated.. | Delay in O2 Concentration output
Operator: Transition from External Gas to Internal Gas Mode, Remove external gas source. External gas source is removed.. | If external gas lost, wrong sweep gas setting.
Operator: Transition from External Gas to Internal Gas Mode, Remove external gas source. External gas source is removed.. | If external gas lost, no oxygenation gas.
Operator: Transition from External Gas to Internal Gas Mode, Remove external gas source. External gas source is removed.. | Pressure bleeds down slowly causing delay in switchover to internal gas mode.
Operator: Transition from External Gas to Internal Gas Mode, Acknowledge "Lost External Gas –No Internal Oxygen Set" alarm. "Lost External Gas –No Internal Oxygen Set" alarm is acknowledged.. | Disorder of procedure creates annoyance for user
Operator: Transition from External Gas to Internal Gas Mode, Acknowledge "Lost External Gas–Confirm Internal Settings" alarm. "Lost External Gas–Confirm Internal Settings" alarm is acknowledged.. | Disorder of procedure creates annoyance for user
Operator: Transition from External Gas to Internal Gas Mode, Use air with oxygen to supply sweep gas. Air with oxygen is used to supply sweep gas. | Air without oxygen is used to supply sweep gas
Operator: Transition from External Gas to Internal Gas Mode, Use air without oxygen to supply sweep gas. Air without oxygen is used to supply sweep gas. | Air with oxygen is used to supply sweep gas
Operator: Alarm Silence/Pause, Touch "Alarm Silence" icon. "Alarm Silence" icon is touched.. | Alarm not silenced/paused
Operator: Alarm Silence/Pause, Touch "Alarm Silence" icon. "Alarm Silence" icon is touched.. | Disorder of procedure creates delay for user
Operator: Alarm Silence/Pause, Touch "Alarm Silence" icon. "Alarm Silence" icon is touched.. | New alarm condition cleared too soon.
Operator: Lock Screen, Touch small "Lock" icon.. Small "Lock" icon is touched.. | Cannot lock screen
Operator: Lock Screen, Touch small "Lock" icon.. Small "Lock" icon is touched.. | Settings screen
Operator: Lock Screen, Touch small "Lock" icon.. Small "Lock" icon is touched.. | Locked "settings" screen
Operator: Lock Screen, Touch and Hold big "Lock" icon. Big "Lock" icon is touched and held.. | Big lock icon cleared
Operator: Lock Screen, Touch and Hold big "Lock" icon. Big "Lock" icon is touched and held.. | lock timer decrements and clears icon
Operator: Un-Lock Screen, Touch anywhere on screen. Anywhere on the screen is touched.. | Cannot unlock screen
Operator: Un-Lock Screen, Touch and Hold big "Lock" icon. Big "Lock" icon is touched and held.. | Big lock icon cleared
Operator: Un-Lock Screen, Touch and Hold big "Lock" icon. Big "Lock" icon is touched and held.. | lock timer decrements and clears icon
Manual sighing (Ext Gas Mode), Manually increase the Wall/Tank Flow Meter to the desired flow.. Wall/Tank Flow Meter is manually increased to the desired flow.. | gas flow limited by excessive condensation
Manual sighing (Ext Gas Mode), Manually increase the Wall/Tank Flow Meter to the desired flow.. Wall/Tank Flow Meter is manually increased to the desired flow.. | excessive CO2 removal
Manual sighing (Ext Gas Mode), Maintain high gas flow for duration of time required (according to patient status) to manually sigh. High gas flow (sweep rate) is maintained for long enough to complete manual sighing (according to patient status). . | gas flow limited by excessive condensation
Manual sighing (Ext Gas Mode), Maintain high gas flow for duration of time required (according to patient status) to manually sigh. High gas flow (sweep rate) is maintained for long enough to complete manual sighing (according to patient status). . | excessive CO2 removal
Manual sighing (Ext Gas Mode), Return flow to previous setting. Flow is returned to previous setting.. | reduced O2 flow
Manual sighing (Ext Gas Mode), Return flow to previous setting. Flow is returned to previous setting.. | excessive CO2 removal
Operator: Initiate Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | gas flow limited by excessive condensation
Operator: Initiate Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | excessive CO2 removal
Operator: Initiate Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | Cannot initiate Sigh
Operator: Initiate Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | Sigh control clears
Operator: Initiate Sigh Cycle, Touch "On/Off" control to enable sigh. "On/Off control touched to enable sigh.. | Cannot initiate Sigh
Operator: Initiate Sigh Cycle, Touch "On/Off" control to enable sigh. "On/Off" control is touched to enable sigh.. | Sigh aborted
Operator: Initiate Sigh Cycle, Wait 2 minutes until Sigh process completes. Sigh process completes after 2 minutes.. | Sigh aborted
Operator: Abort Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | Cannot abort Sigh
Operator: Abort Sigh Cycle, Touch "Sigh" icon. "Sigh" icon is touched.. | Sigh control clears
Operator: Abort Sigh Cycle, Touch "On/Off" control to abort sigh. "On/Off" control is touched to abort sigh.. | Sigh continues
Operator: Abort Sigh Cycle, Touch "On/Off" control to abort sigh. "On/Off" control is touched to abort sigh.. | Sigh re-starts
Operator: Point-of-Use Cleaning - Touch Screen, Place device in "Lock Screen" - See UF038. Device is placed in "Lock Screen".. | During cleaning, settings are changed
Operator: Point-of-Use Cleaning - Touch Screen, obtain recommended cleaning solution. Recommended cleaning solution is obtained.. | damage to screen results in surface etching and/or clouding
Operator: Point-of-Use Cleaning - Touch Screen, obtain recommended cleaning solution. Recommended cleaning solution is obtained.. | insufficient cleaning of screen
Operator: Point-of-Use Cleaning - Touch Screen, wipe with a soft, lint free cloth dampened with a cleaning solution. Screen is wiped with a soft, lint free cloth dampedned with a cleaning solution.. | damage to screen results in surface etching and/or clouding
Operator: Point-of-Use Cleaning - Touch Screen, wipe with a soft, lint free cloth dampened with a cleaning solution. Screen is wiped with a soft, lint free cloth dampedned with a cleaning solution.. | debris left on screen
Operator: Point-of-Use Cleaning - Touch Screen, allow cleaning solution to completely dry or dry with a lint free cloth. Cleaning solution is allowed to completely dry or it is dried with a lint free cloth.. | wet cleaning solution on screen
Operator: Point-of-Use Cleaning - Touch Screen, "Un-Lock Screen" - see UF039. Screen is unlocked.. | Screen remains locked.
Operator: Point-of-Use Cleaning, Remove contamination or soil from device using soft cloth. Contamination or soil is removed from device using soft cloth.. | Difficult to clean and disinfect later
Operator: Point-of-Use Cleaning, Remove contamination or soil from device using soft cloth. Contamination or soil is removed from device using soft cloth.. | damage to surface finish
Staff: Cleaning, Obtain recommended cleaning solution. Recommended cleaning solution is obtained.. | damage to surface finish
Staff: Cleaning, Spray soft cloth with cleaning solution. Soft cloth is sprayed with cleaning solution.. | Excess solution on device
Staff: Cleaning, Clean by wiping all device surfaces with moistened cloth. All device surfaces are cleaned by wiping with moistened cloth.. | damage to surface finish
Staff: Cleaning, Use a small brush or similar device to clean small areas or crevices. . Small areas or crevices are cleaned using a small brush or similar device.. | insufficient cleaning of device
Staff: Cleaning, Rinse by wiping all device surfaces with water moistened cloth. All device surfaces are rinsed by wiping with a water moistened cloth.. | Device has strong smell of cleaning solution
Staff: Cleaning, Dry all device surfaces with dry cloth. All device surfaces are dried with a dry cloth.. | Device damp from rinsing
Staff: Disinfecting, Obtain recommended disinfecting solution. Recommended disinfecting solution is obtained.. | damage to surface finish
Staff: Disinfecting, Spray soft cloth with disinfecting solution. Soft cloth is sprayed with disinfecting solution.. | Excess solution on device
Staff: Disinfecting, Disinfect by wiping all device surfaces with moistened cloth. All device surfaces are disinfected by wiping with a moistened cloth.. | damage to surface finish
Staff: Disinfecting, Use a small brush or similar device to disinfect small areas or crevices. . Small areas or crevices are disinfected using a small brush or similar device.. | insufficient disinfection of device
Staff: Disinfecting, Rinse by wiping all device surfaces with water moistened cloth. All device surfaces are rinsed by wiping with a water moistened cloth.. | Device has strong smell of disinfecting solution
Staff: Disinfecting, Dry all device surfaces with dry cloth. All device surfaces are dried with a dry cloth.. | Device damp from rinsing
Staff: PLU Lid Window Cleaning/Disinfecting, Remove lid window.. Lid window is removed.. | insufficient cleaning/disinfection of device
Staff: PLU Lid Window Cleaning/Disinfecting, Remove phillips head screw from underside of lid window.. Phillips head screw is removed from underside of lid window.. | insufficient cleaning/disinfection of device
Staff: PLU Lid Window Cleaning/Disinfecting, Pry window away from lid to release side snaps.. Window is pried away from lid to release side snaps.. | Mechanical retention of lid window compromised
Staff: PLU Lid Window Cleaning/Disinfecting, Rotate lid to disengage tab.. Lid is rotated to disengage tab.. | Mechanical retention of lid window compromised
Staff: PLU Lid Window Cleaning/Disinfecting, Clean lid window and lid mating surface.. Lid window and lid mating surface are opened.. | damage to surface finish
Staff: PLU Lid Window Cleaning/Disinfecting, Clean lid window and lid mating surface.. Lid window and lid mating surface are opened.. | Excess solution on device
Staff: PLU Lid Window Cleaning/Disinfecting, Clean lid window and lid mating surface.. Lid window and lid mating surface are opened.. | insufficient cleaning of device
Staff: PLU Lid Window Cleaning/Disinfecting, Clean lid window and lid mating surface.. Lid window and lid mating surface are opened.. | Device has strong smell of cleaning solution
Staff: PLU Lid Window Cleaning/Disinfecting, Clean lid window and lid mating surface.. Lid window and lid mating surface are opened.. | Device damp from rinsing
Staff: PLU Lid Window Cleaning/Disinfecting, Disinfect lid window and lid mating surface.. Lid window and lid mating surface are disinfected.. | damage to surface finish
Staff: PLU Lid Window Cleaning/Disinfecting, Disinfect lid window and lid mating surface.. Lid window and lid mating surface are disinfected.. | Excess solution on device
Staff: PLU Lid Window Cleaning/Disinfecting, Disinfect lid window and lid mating surface.. Lid window and lid mating surface are disinfected.. | insufficient disinfection of device
Staff: PLU Lid Window Cleaning/Disinfecting, Disinfect lid window and lid mating surface.. Lid window and lid mating surface are disinfected.. | Device has strong smell of disinfecting solution
Staff: PLU Lid Window Cleaning/Disinfecting, Disinfect lid window and lid mating surface.. Lid window and lid mating surface are disinfected.. | Device damp from rinsing
Staff: PLU Lid Window Cleaning/Disinfecting, Replace lid window. Lid window is replaced.. | Top of PLU Disposable exposed.
Staff: PLU Lid Window Cleaning/Disinfecting, Engage tab and rock window onto lid.. Tab and rock window is engaged onto lid.. | Mechanical retention of lid window compromised
Staff: PLU Lid Window Cleaning/Disinfecting, Press down on lid window to engage side snaps.. Lid window is pressed down on to engage side snaps.. | Mechanical retention of lid window compromised
Staff: PLU Lid Window Cleaning/Disinfecting, Install phillips head screw on underside of lid window.. Phillips head screw is installed on underside of lid window.. | Mechanical retention of lid window compromised
Operator: Log File Export, Insert an Oxy-1 data key into the export data port of the Console. OXY-1 data key is inserted into the export data port of the Console.. | data key not detected
Operator: Log File Export, Wait for the message that indicates that the export data has completed.. Message that indicates that the export data has completed is displayed.. | data file corrupted on data key.
Operator: Log File Export, Remove the data key. Data key is removed.. | data exported at each power on.
Operator: Log File Export, Connect the Oxy-1 data key to a PC using the USB adapter.. OXY-1 data key is connected to a PC using the USB adapter.. | data not available
Operator: Log File Export, Open the most recent .CSV file on the data key indicated by the file name format.. Most recent .CSV file on data key is opened.. | wrong data file accessed
Operator: Log File Delete, Touch "Settings" icon. "Settings" icon is touched.. | Settings screen not opened
Operator: Log File Delete, Touch "Settings" icon. "Settings" icon is touched.. | "Lock" prompt is displayed
Operator: Log File Delete, Touch the "Trash Can" icon. "Trash Can" icon is touched.. | Cannot delete log file
Operator: Log File Delete, Touch and Hold big "Trash Can" icon. Big "Trash Can" icon is touched and held. . | Big Trash Can icon cleared
Operator: Log File Delete, Touch and Hold big "Trash Can" icon. Big "Trash Can" icon is touched and held. . | delete timer decrements and clears icon
Operator: Log File Delete, Touch "Home" icon. "Home" icon is touched.. | Settings screen persists
Operator: PLU Disposable Change Out, Double clamp Arterial and Venous blood lines.. Arterial and Venous blood lines are double clamped.. | leaking priming fluid and/or blood
Operator: PLU Disposable Change Out, Set Pump RPM to 0. - See UF029. Pump RPM is set to 0.. | Pump running with tubing clamped
Operator: PLU Disposable Change Out, Cut the tubing between each pair of clamps. Tubing between each pair of clamps is cut.. | Disorder of procedure creates delay for user
Operator: PLU Disposable Change Out, Remove old PLU sheath/Tubes from Console. Old PLU sheath/Tubes are removed from Console.. | Disorder of setup creates delay for the user.
Operator: PLU Disposable Change Out, Remove old PLU Disposable from Pump Driver. Old PLU Disposable is removed from Pump Driver.. | Disorder of setup creates delay for the user.
Operator: PLU Disposable Change Out, Add a 3/8" barbed male-male connector to each patient blood line.. 3/8" barbed male-male connector is added to each patient blood line.. | Disorder of procedure creates delay for user
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to Pump Driver.
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Lid closes before PLU Disposable is placed back onto Pump Driver
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Decreased or no blood flow
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Latch lever cannot close. Flow Sensor reading error
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Disorder of setup creates delay for the user.
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Top Latch plate cannot close far enough for latch arm to engage
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | Disposable not secure.
Operator: PLU Disposable Change Out, Received a new primed PLU Disposable and attach new PLU Disposable to Pump Driver. A new primed PLU Disposable is received and new PLU Disposable is attached to Pump Driver.. | FLow Sensor reading error.
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Disorder of setup creates inconvenience for the user.
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Weak connection on O2 line
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Air feeding inner chamber when on internal gas
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Weak connection on Air line
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | O2 feeding outer chamber when on internal gas.
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Zipper does not open.
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Insufficient strain relief on Air and O2 lines
Operator: PLU Disposable Change Out, Attach new PLU sheath/Tubes to Console. New PLU sheath/Tubes are attached to Console.. | Insufficient strain relief on Air and CO2 lines
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | No focus is set on "Pump" element on UI, creating a delay
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | Focus on "UP" or "DOWN" element on UI is cleared, creating a delay
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | RPM changes in opposite direction as intended
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | RPM changes faster than intended
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | Excessive positive/negative pressure applied to patient's blood stream.
Operator: PLU Disposable Change Out, Set Pump RPM to desired setting.. Pump RPM is set to desired setting.. | Focus on "Pump" element on UI is not cleared, creating a delay
Operator: PLU Disposable Change Out, Remove clamps. Clamps are removed.. | Pump running with tubing clamped
Emergency: Transfer PLU Disposable to E-Drive, Un-zip sheath zipper.. sheath zipper is unzipped.. | No blood flow
Emergency: Transfer PLU Disposable to E-Drive, Clamp Arterial and Venous blood lines near patient. (to prevent retrograde flow). Arterial and Venous blood lines are clamped near patient (to prevent retrograde flow).. | Retrograde flow
Emergency: Transfer PLU Disposable to E-Drive, If pump not already stopped, set Pump RPM to 0. If pump not already stopped, Pump RPM set to 0.. | Pump RPM is not set to 0.
Emergency: Transfer PLU Disposable to E-Drive, If pump not already stopped, set Pump RPM to 0. If pump not already stopped, Pump RPM set to 0.. | No focus is set on "DOWN" element on UI, creating a delay
Emergency: Transfer PLU Disposable to E-Drive, If pump not already stopped, set Pump RPM to 0. If pump not already stopped, Pump RPM set to 0.. | Focus on "DOWN" element on UI is cleared, creating a delay
Emergency: Transfer PLU Disposable to E-Drive, If pump not already stopped, set Pump RPM to 0. If pump not already stopped, Pump RPM set to 0.. | RPM changes in opposite direction as intended
Emergency: Transfer PLU Disposable to E-Drive, If pump not already stopped, set Pump RPM to 0. If pump not already stopped, Pump RPM set to 0.. | Focus on "Pump" element on UI is not cleared, creating a delay
Emergency: Transfer PLU Disposable to E-Drive, Remove PLU Disposable from Pump Driver. PLU Disposable is removed from Pump Driver.. | Disorder of setup creates delay for the user.
Emergency: Transfer PLU Disposable to E-Drive, Align and latch PLU Disposable to E-Drive. PLU Disposable is aligned and latched to E-Drive.. | Pump and Pump Motor mis-aligned
Emergency: Transfer PLU Disposable to E-Drive, Align and latch PLU Disposable to E-Drive. PLU Disposable is aligned and latched to E-Drive.. | Damage to PLU Disposable
Emergency: Transfer PLU Disposable to E-Drive, Maintain Air and O2 connections from PLU Disposable to Console. Air and O2 connections from PLU Disposable to Console are maintained.. | No sweep gas to outer chamber of oxygenator
Emergency: Transfer PLU Disposable to E-Drive, Maintain Air and O2 connections from PLU Disposable to Console. Air and O2 connections from PLU Disposable to Console are maintained.. | No oxygen gas to inner chamber of oxygenator
Emergency: Transfer PLU Disposable to E-Drive, Maintain external Oxygen connection to Console. External Oxygen connection to Console is maintained.. | reduced O2 flow
Emergency: Hand cranking E-Drive, Unfold E-Drive crank handle. E-Drive crank handle is unfolded.. | Unable to turn blood pump
Emergency: Hand cranking E-Drive, Crank handle clockwise. Handle is cranked clockwise.. | blood pump stopped
Emergency: Hand cranking E-Drive, Crank handle clockwise. Handle is cranked clockwise.. | Pump within PLU Disposable is de-coupled. The pump is not fully engaged and can reduced forwared flow.
Emergency: Hand cranking E-Drive, Increase/Decrease cranking rate to achieve desired blood pump RPM.. Cranking rate is increased/decreased to achieve desired blood pump RPM.. | Low pump RPM
Emergency: Hand cranking E-Drive, Increase/Decrease cranking rate to achieve desired blood pump RPM.. Cranking rate is increased/decreased to achieve desired blood pump RPM.. | High pump RPM
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Turn-on new OXY-1 System. New OXY-1 System is turned on.. | No PLU function
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Turn-on new OXY-1 System. New OXY-1 System is turned on.. | Loss of AC mains power
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Turn-on new OXY-1 System. New OXY-1 System is turned on.. | System remains off
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Move 1/4" ID tubing from regulator of oxygen source (wall or tank) to new Console oxygen inlet.. 1/4" ID tubing is moved from regulator of oxygen source (wall or tank) to new Console oxygen inlet.. | reduced O2 flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Move 1/4" ID Air tube from old Console to new Console 1/4" barbed fitting.. 1/4" ID Air tube is moved from old Console to new Console 1/4" barbed fitting.. | reduced Air flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Move 3/16" ID O2 tube from old Console to new Console 3/16" barbed fitting.. 3/16" ID O2 tube is moved from old Console to new Console 3/16" barbed fitting.. | reduced O2 flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Clamp Arterial and Venous blood lines near patient. (to prevent retrograde flow). Arterial and Venous blood lines are clamped near patient (to prevent retrograde flow).. | Retrograde flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Stop cranking E-Drive handle. E-Drive handle is not cranked.. | Disorder of procedure creates delay for user
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Wait for E-Drive to stop.. E-Drive stops.. | Disorder of procedure creates inconvenience for user
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Press & hold E-Drive release button. E-Drive release button is pressed and held. . | Hard to unlatch PLU disposable
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Unlatch PLU Disposable from E-Drive. PLU Disposable is unlatched from E-Drive.. | Damage to PLU Disposable
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to Pump Driver.
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Lid closes before PLU Disposable is placed back onto Pump Driver
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Decreased or no blood flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Latch lever cannot close. Flow Sensor reading error
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disorder of setup creates delay for the user.
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Top Latch plate cannot close far enough for latch arm to engage
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disposable not secure.
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Attach PLU Disposable to new Pump Driver. PLU Disposable is attached to new Pump Driver.. | FLow Sensor reading error.
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Gradually increase Blood Pump RPM. Blood Pump RPM is gradually increased.. | Retrograde flow
Emergency: Transfer PLU Disposable from E-Drive to (new) OXY-1 System, Gradually increase Blood Pump RPM. Blood Pump RPM is gradually increased.. | Pump cavitation
Emergency: Leak in arterial blood tubing, Clamp tubing near leak. Tubing near leak is clamped.. | leaking priming fluid and/or blood
Emergency: Leak in arterial blood tubing, Set Pump RPM to 0. Pump RPM is set to 0.. | Pump running with tubing clamped
Emergency: Leak in arterial blood tubing, Repair leak. Leak is repaired.. | Arterial bloodstream bubbles
Emergency: Leak in arterial blood tubing, Repair leak. Leak is repaired.. | Blood stream is stagnant
Emergency: Leak in arterial blood tubing, Remove clamps. Clamps are removed.. | Excessive positive pressure applied to patient's blood stream.
Emergency: Leak in arterial blood tubing, Slowly increase Pump RPM to achieve desired flow.. Pump RPM is slowly increased to achieve desired flow.. | Retrograde flow
Emergency: Leak in venous blood tubing, Clamp tubing near leak. Tubing near leak is clamped.. | Leak draws in air
Emergency: Leak in venous blood tubing, Set Pump RPM to 0. Pump RPM is set to 0.. | Pump cavitation
Emergency: Leak in venous blood tubing, Repair leak. Leak is repaired.. | Venous bloodstream bubbles
Emergency: Leak in venous blood tubing, Repair leak. Leak is repaired.. | Blood stream is stagnant
Emergency: Leak in venous blood tubing, Remove clamps. Clamps are removed.. | Excessive negative pressure applied to patient's blood stream.
Emergency: Leak in venous blood tubing, Slowly increase Pump RPM to achieve desired flow.. Pump RPM is slowly increased to achieve desired flow.. | Retrograde flow
Alarms: Blood Flow < 0 LPM, If priming, ignore alarm. Alarm is ignored.. | Disorder of procedure creates annoyance for user
Alarms: Blood Flow < 0 LPM, If during patient care, adjust blood pump RPM > 0. Blood pump RPM is adjusted to > 0.. | Blood stream is stagnant
Alarms: High Blood Flow, If priming, ignore alarm. Alarm is ignored.. | Disorder of procedure creates annoyance for user
Alarms: High Blood Flow, If during patient care, Inspect the patient blood circuit for leaks. Patient blood circuit is inspected for leaks.. | Blood loss
Alarms: High Blood Flow, If during patient care, Check patient vital signs, blood gas. Patient vital signs and blood gas are checked.. | Blood loss
Alarms: High Blood Flow, If during patient care, Check patient vital signs, blood gas. Patient vital signs and blood gas are checked.. | Insufficient oxygen transfer & Hypoperfusion
Alarms: High Blood Flow, If during patient care, Adjust blood pump RPM. Blood pump RPM is adjusted.. | Pump cavitation
Alarms: High Blood Flow, If during patient care, Adjust Alarm Limit. Alarm Limit is adjusted.. | Disorder of procedure creates delay for user
Alarms: Low Blood Flow, If priming, ignore alarm. Alarm is ignored.. | Disorder of procedure creates annoyance for user
Alarms: Low Blood Flow, If during patient care, Inspect the patient blood circuit for restrictions. Patient blood circuit is inspected for restrictions.. | Incorrect blood flow
Alarms: Low Blood Flow, If during patient care, Adjust blood pump RPM. Blood pump RPM is adjusted.. | Incorrect blood flow
Alarms: Low Blood Flow, If during patient care, Adjust Alarm Limit. Alarm Limit is adjusted.. | Disorder of procedure creates delay for user
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Disorder of setup creates delay for the user.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Arterial bloodstream bubbles
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Pump Bearing Wear
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to PLU Durable.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Lid closes before PLU Disposable is placed back onto PLU Durable
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Decreased or no blood flow
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Latch lever cannot close. Flow Sensor reading error
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Top Latch plate cannot close far enough for latch arm to engage
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Disposable not secure.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | FLow Sensor reading error.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Disorder of setup creates inconvenience for the user.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Cavitation
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Venous bloodstream bubbles
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | leaking priming fluid and/or blood
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Continuous circulation of priming fluid
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Weak connection on O2 line
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Air feeding inner chamber when on internal gas
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Weak connection on Air line
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | O2 feeding outer chamber when on internal gas.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Zipper does not open.
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Insufficient strain relief on Air and O2 lines
Alarms: Bubbles Detected, If priming, Follow priming procedure to remove bubbles. Priming procedure is followed to remove bubbles.. | Insufficient strain relief on Air and CO2 lines
Alarms: Bubbles Detected, If during patient care, find and correct the cause of bubbles: Inspect the patient blood circuit for leaks, check that height of PLU is below patient's heart, check for clamps/restrictions on low pressure side causing cavitation, check for excess positive gas pressure on membrane, check for blocked/restricted gas flow exhaust path. Remove trapped air bubbles.. Patient blood circuit is inspected for leaks, the height of the PLU is confirmed to be below the patient's heart, low pressure side is inspected for clamps/restrictions which can cause cavitation, positive gas pressure on memrane is checked for, blocked/restricted gas flow exhaust path is checked for. Trapped air bubbles are removed.. | Arterial bloodstream bubbles
Alarms: Blood Flow Sensor Failure, If alarm persists, use 3rd party flow sensor to monitor blood flow until a new OXY-1 System is available. 3rd party flow sensor is used to monitor blood flow until a new OXY-1 System is available.. | There is no way to monitor flow. Any changes in flow will not be detected.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Remove PLU sheath/Tubes from (old) Console. PLU sheath/Tubes are removed from (old) Console.. | Disorder of setup creates delay for the user.
Alarms: Blood Flow Sensor Failure, Check Pump Driver’s electrical cable for connection. Pump Driver’s electrical cable is checked for connection. | Pump Driver’s electrical cable is not checked for connection
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Clamp Arterial and Venous blood lines near patient. (to prevent retrograde flow). Arterial and Venous blood lines are clamped near patient (to prevent retrograde flow).. | Retrograde flow
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Stop blood pump. Blood pump is stopped.. | No focus is set on "DOWN" element on UI, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Stop blood pump. Blood pump is stopped.. | Focus on "DOWN" element on UI is cleared, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Stop blood pump. Blood pump is stopped.. | RPM changes in opposite direction as intended
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Stop blood pump. Blood pump is stopped.. | Focus on "Pump" element on UI is not cleared, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Remove PLU Disposable from (old) Pump Driver. PLU Disposable is removed from (old) Pump Driver.. | Disorder of setup creates delay for the user.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to Pump Driver.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Lid closes before PLU Disposable is placed back onto Pump Driver
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Decreased or no blood flow
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Latch lever cannot close. Flow Sensor reading error
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disorder of setup creates delay for the user.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Top Latch plate cannot close far enough for latch arm to engage
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | Disposable not secure.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU Disposable to (new) Pump Driver. PLU Disposable is attached to new Pump Driver.. | FLow Sensor reading error.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | No focus is set on "Pump" element on UI, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | Focus on "UP" or "DOWN" element on UI is cleared, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | RPM changes in opposite direction as intended
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | RPM changes faster than intended
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | Excessive positive/negative pressure applied to patient's blood stream.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Restart blood pump. Blood pump is restarted.. | Focus on "Pump" element on UI is not cleared, creating a delay
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Disorder of setup creates inconvenience for the user.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Weak connection on O2 line
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Air feeding inner chamber when on internal gas
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Weak connection on Air line
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | O2 feeding outer chamber when on internal gas.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Zipper does not open.
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Insufficient strain relief on Air and O2 lines
Alarms: Blood Flow Sensor Failure, Transfer PLU Disposable to new OXY-1 System: Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Insufficient strain relief on Air and CO2 lines
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | No blood flow
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Retrograde flow
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Pump RPM is not set to 0.
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | No focus is set on "DOWN" element on UI, creating a delay
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Focus on "DOWN" element on UI is cleared, creating a delay
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | RPM changes in opposite direction as intended
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Focus on "Pump" element on UI is not cleared, creating a delay
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Disorder of setup creates delay for the user.
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Pump and Pump Motor mis-aligned
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | Damage to PLU Disposable
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | No sweep gas to outer chamber of oxygenator
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | No oxygen gas to inner chamber of oxygenator
If any of the following alarms trigger: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to E-Drive. PLU DIsposable is transferred to E-Drive.. | reduced O2 flow
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Hand crank E-Drive. E-Drive is hand cranked.. | Unable to turn blood pump
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Hand crank E-Drive. E-Drive is hand cranked.. | blood pump stopped
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Hand crank E-Drive. E-Drive is hand cranked.. | Low pump RPM
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Hand crank E-Drive. E-Drive is hand cranked.. | High pump RPM
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | No PLU function
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Loss of AC mains power
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | System remains off
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | reduced O2 flow
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | reduced Air flow
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Retrograde flow
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Disorder of procedure creates delay for user
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Disorder of procedure creates inconvenience for user
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Hard to unlatch PLU disposable
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Damage to PLU Disposable
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Disorder of setup creates inconvenience for the user. As a result, PLU Disposable is not attached to PLU Durable.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Lid closes before PLU Disposable is placed back onto PLU Durable
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Decreased or no blood flow
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Latch lever cannot close. Flow Sensor reading error
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Disorder of setup creates delay for the user.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Top Latch plate cannot close far enough for latch arm to engage
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Disposable not secure.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | FLow Sensor reading error.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Transfer PLU Disposable to new OXY-1 System. PLU Disposable is transferrred to new OXY-1 System.. | Pump cavitation
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Remove PLU sheath/Tubes from (old) Console. PLU sheath/Tubes are removed from (old) Console.. | PLU sheath/Tubes are not removed from (old) console.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Disorder of setup creates inconvenience for the user.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Weak connection on O2 line
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Air feeding inner chamber when on internal gas
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Weak connection on Air line
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | O2 feeding outer chamber when on internal gas.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Zipper does not open.
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Insufficient strain relief on Air and O2 lines
Alarms: Pump Drive Failure, Pump Control Failure 1, Pump Control Failure 2, Internal Power Failure 1, Internal Power Failure 2, Alarm System Disabled, Attach PLU sheath/Tubes to (new) Console. PLU sheath/Tubes are attached to (new) Console.. | Insufficient strain relief on Air and CO2 lines
Alarms: No External Power. Remove and replace console. | User does not replace the console and continues to use the console
Alarms: Low Battery Power, Battery Failure 1, Battery Failure 2, Connect line cord to AC wall outlet.. Line cord is connected to AC wall outlet.. | Loss of AC mains power
Alarms: Low Battery Power, Battery Failure 1, Battery Failure 2, Service Call to replace batteries. Service Call to replace batteries is placed.. | Batteries not replaced
Alarms: Battery Over Temperature. User checks console inlet and exhaust ports for blockages and corrects blockages if present.. | User does not correct blockage
Alarms: Battery Over Temperature. User checks console inlet and exhaust for blockages and corrects blockages if present.. | User does not correct blockage
Alarms: Battery Over Temperature. Replace console. | User does not replace the console and disconnects AC power and continues to use the console
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If "Air Flow Failure", check for blocked Air tubing. Blocked Air tubing is checked for.. | No sweep gas to outer chamber of oxygenator
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If "Oxygen Concentrator Flow", check for blocked O2 tubing. Blocked O2 tubing is checked for.. | No oxygen gas to inner chamber of oxygenator
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | reduced O2 flow
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | blocked O2 flow
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | toxic gas passed through membrane and transferred to blood
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | no O2 flow
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | excessive O2 flow
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, If any alarms persist, switch to External Gas mode. External Gas mode is activated.. | Excessive positive pressure applied to membrane.
Alarms: Air Flow Failure, Oxygen Concentrator Flow, Oxygen Concentrator FiO2 < 85%, Oxygen Concentrator FiO2 < 72%, Oxygen Concentrator Failure 1, Oxygen Concentrator Failure 2, Sigh Function Failure, Replace Console at next opportunity. Console is replaced at next opportunity.. | Disorder of procedure creates inconvenience for user
Alarms: Pneumatics Failure, Remove external gas source. External gas source is removed.. | Excessive positive pressure applied to membrane.
Alarms: Pneumatics Failure, If internal Gas Mode is automatically enabled, ensure Internal Gas settings are correct. Internal Gas settings are confirmed.. | Air Blower off
Alarms: Pneumatics Failure, If internal Gas Mode is automatically enabled, ensure Internal Gas settings are correct. Internal Gas settings are confirmed.. | O2 Concentrator off
Alarms: Pneumatics Failure, If internal Gas Mode is NOT automatically enabled, re-connect external gas source. External gas source is re-connected.. | External gas source not connected
Alarms: Pneumatics Failure, At Console, disconnect 1/4" ID Air tube from 1/4" barbed fitting.. 1/4" ID Air tube is disconnected from 1/4" barbed fitting.. | Can't check for air flow
Alarms: Pneumatics Failure, Check for gas flow at Console Air outlet. Gas flow at Console Air Outlet is checked.. | No Air flow
Alarms: Pneumatics Failure, At Console, disconnect 3/16" ID O2 tube from 3/16" barbed fitting.. 3/16" ID O2 tube is disconnected from 3/16" barbed fitting.. | Can't check for air flow
Alarms: Pneumatics Failure, Check for gas flow at Console O2 outlet. Gas flow at Console O2 outlet is checked.. | No O2 flow
Alarms: Pneumatics Failure, If no flow at either Air or O2 outlet, replace Console at next opportunity. Console is replaced at next opportunity.. | No flow at either Air or O2 outlet
Pre-Use Checks Console: Confirm Installation Checkout procedures are completed. Installation Checkout procedures are completed. | User does not complete installation checkout procedures
Pre-Use Checks Console: Confirm OXY-1 system is positioned near the patient in a well-ventilated area free of pollutants and high humidity.. OXY-1 system is positioned near the patient in a well-ventilated area free of pollutants and high humidity.. | User operates console outside of operational environmental limits
Pre-Use Checks Console: Confirm Console air intake and exhaust ports are not obstructed. Console air intake and exhaust ports are not obstructed. | User does not confirm Gas flow exhaust or intake ports are not obstructed
Pre-Use Checks Console: Confirm Console castors and IV pole castors (if equipped) are locked.. Console castors and IV pole castors (if equipped) are locked.. | Disorder of setup creates inconvenience for the user.
Pre-Use Checks Console: When the console is plugged in, confirm the power cord is connected directly to an outlet. the blue battery charging LED is on.. Power cord is connected directly to an outlet and the blue battery charging LED is on.. | User does not connect power cord to an outlet or the user does not confirm that blue battery charging LED is on.
Pre-Use Checks Console: Confirm the Power button (Standby Switch) to the ON position. Power button (Standby Switch) is switched to the ON position. | User does not lift Power button (standby switch) cover
Pre-Use Checks Console: Confirm after boot completes and the home page is displayed, confirm no alarms exist. After boot completes and the home page is displayed, confirm no alarms exist. If alarms exist, refer Alarm Troubleshooting. | User does not confirm that the home page is displayed and no alarms exist
Pre-Use Checks Pump Driver: Confirm the Pump Driver is securely latched onto the Pole Mount.. Pump Driver is securely latched onto the Pole Mount.. | User does not properly tighten PLU Mounting Bracket or uses the wrong diameter IV Pole
Pre-Use Checks Pump Driver: Verify that the primed PLU Disposable is properly seated in the Pump Driver.. The primed PLU Disposable is properly seated in the Pump Driver.. | User fails to raise latch lever, release latch arm from latch plate, and/or open top latch plate
Pre-Use Checks Pump Driver: Check that the blood tubing is properly seated into the channels of the Pump Driver flow sensor.. The blood tubing is properly seated into the channels of the Pump Driver flow sensor.. | User cannot close Top Latch plate far enough for latch arm to engage
Pre-Use Checks Pump Driver: Check that the lid is closed beneath the latch and the latch is fully engaged on the Pump Driver.. The lid is closed beneath the latch and the latch is fully engaged on the Pump Driver.. | User leaves the Disposable not secure.
Pre-Use Checks Pump Driver: Check the blood circuit access ports for leaks. Tighten caps or adjust stopcocks as needed.. No leaks in blood circuit access ports.. | User allows leaking of priming fluid
Pre-Use Checks Pump Driver: Check the blood circuit access ports for leaks. Tighten caps or adjust stopcocks as needed.. No leaks in blood circuit access ports.. | User allows air drawn into the blood circuit
Pre-Use Checks Pump Driver: The gas tubing between the PLU Disposable and Console is properly connected.. The gas tubing between the PLU Disposable and Console is properly connected.. | User leaves Air tubing blocked, restricted, or disconnected
Pre-Use Checks Pump Driver: The Cable Guard is zipped over the gas tubing and electrical cable and buttoned at both ends.. The Cable Guard is zipped over the gas tubing and electrical cable and buttoned at both ends.. | Operator does not zip the cable guard over the tubing and electrical cables
Pre-Use Checks Pump Driver: Check that the Condensate Tray is Installated in to the Pump Driver.. The Condensate Tray is installed.. | The condensate tray is not installed.
Pre-Use Checks Blood Circuit: Increase the Pump RPM until the Blood Flow is 3 LPM.. Blood Flow is 3 LPM.. | User does not set focus
Pre-Use Checks Blood Circuit: Increase the Pump RPM until the Blood Flow is 3 LPM.. Blood Flow is 3 LPM.. | User does not clear focus
Pre-Use Checks Blood Circuit: Inspect the entire blood circuit and blood pump for bubbles.. Blood circuit and blood pump are inspected for bubbles. . | User does not inspect the blood circuit and bubbles are trapped in blood lines
Pre-Use Checks Blood Circuit: If bubbles are present, stop the blood pump and eliminate bubbles. If bubbles are present, bubbles are eliminated.. | User does not eliminate bubbles trapped in blood lines
Pre-Use Checks Blood Circuit: Confirm the green indicator light comes on when the motor is rotating. Green indicator comes on when the motor is rotating.. | User does not recognize the green indicator comes on
Pre-Use Checks Blood Circuit: Set the Pump RPM to 0.. Pump RPM is 0.. | User does not set focus
Pre-Use Checks Blood Circuit: Set the Pump RPM to 0.. Pump RPM is 0.. | User does not clear focus
Pre-Use Checks Gas Circuit Integrity Test: Remove external gas source (if connected).. External Gas source is removed.. | User does not remove external gas source
Pre-Use Checks Gas Circuit Integrity Test: Set Internal Gas, Air Flow to 15 LPM.. Internal Gas, Air Flow is 15 LPM.. | User does not set Internal Gas, Air Flow to 15 LPM
Pre-Use Checks Gas Circuit Integrity Test: Set Internal Gas, O2 Flow to 3 LPM.. Internal Gas, O2 Flow is 3 LPM.. | User does not set Internal Gas, O2 Flow to 3 LPM
Pre-Use Checks Gas Circuit Integrity Test: Wait 60 seconds for internal gas sources to reach max flow settings.. Internal gas sources reach max flow settings.. | User does not wait 60 seconds for the internal gas sources reached max flow settings
Pre-Use Checks Gas Circuit Integrity Test: Clamp both the Air and O2 tubing.. Both Air and O2 lines are clamped or pinched. . | User does not clamp lines completely off.
Pre-Use Checks Gas Circuit Integrity Test: Wait 60 seconds for the Console to detect flow errors: - High Priority Alarm "Air Flow Out of Range" - High Priority Alarm "Oxygen Concentrator Flow Out of Range" If either flow error was NOT detected, this indicates there is a leak in the gas circuit.. Console detected flow errors: - High Priority Alarm "Air Flow Out of Range" - High Priority Alarm "Oxygen Concentrator Flow Out of Range". | UI Alarm "Air Flow Out of Range" or UI Alarm "Oxygen Concentrator Flow Out of Range" is not observed
Pre-Use Checks Gas Circuit Integrity Test: Verify both gas tubing lines are firmly attached to the Console and repeat the test. If either flow error was still NOT detected, the OXY-1 System should be removed from use.. gas tubing lines are firmly attached to the Console Console detected flow errors: - High Priority Alarm "Air Flow Out of Range" - High Priority Alarm "Oxygen Concentrator Flow Out of Range". | User does not replace the console
Pre-Use Checks Gas Circuit Integrity Test: If either flow error was detected, release clamps on Air and O2 tubing. Wait for the Console to clear the flow errors.. Clamp or pinch on both Air and O2 lines are released. . | User does not release Air and O2 lines
Pre-Use Checks Gas Circuit Integrity Test: Set Internal Gas, Air Flow to 0 LPM. Set Internal Gas, O2 Flow to 0 LPM. Reconnect to external gas source.. Internal Gas, Air Flow is 0 LPM. Internal Gas, O2 Flow is 0 LPM. External gas source is reconnected.. | User does not set Internal Gas, Air Flow to 0 LPM or Internal Gas, O2 Flow to 0 LPM or reconnect External Gas source.
Alarms: Powering Down –Press Power Button to Cancel. User waits 8 seconds for the console to power down if the console should be turned off.. | User presses the power button (standby switch) again stop the power down and the console remains on.
Alarms: Powering Down –Press Power Button to Cancel. User presses the power button (standby switch) again if the console should remain on. . | damage to PLU Disposable
Alarms: Powering Down –Press Power Button to Cancel. User presses the power button (standby switch) again if the console should remain on. . | User presses the touch screen to attempt to keep the console on.
Alarms: Unexpected System Restart. The Console is replaced at next opportunity.. | Console is not replaced at next opportunity
Alarms: Unexpected System Restart -Unable to Recover Settings. The Console is replaced at next opportunity.. | Console is not replaced at next opportunity
Alarms: Unexpected System Restart -Unable to Recover Settings. User returns console to intended settings. | User does not change settings to intended settings.
Alarms: System Failure and Restart. The user evaluates the patient condition and switches to the Emergency Drive if therapy is inadequate or replaces the Console.. | User switches to the emergency drive instead of directly replacing the console.
Alarms: System Failure and Restart. The user evaluates the patient condition and switches to the Emergency Drive if therapy is inadequate or replaces the Console.. | User directly replaces the console instead of switching to the emergency drive.
Alarms: No Internal Oxygen Set. The user evaluates the patient condition and determines whether air without oxygen should be supplied to the patient.. | Air without oxygen is used to supply sweep gas
Alarms: No Internal Oxygen Set. The user evaluates the patient condition and determines whether air without oxygen should be supplied to the patient.. | Air with oxygen is used to supply sweep gas
Alarms: Lost External Gas–Confirm Internal Settings. The user reconnects the system to external gas.. | Air with oxygen is used to supply sweep gas
Alarms: Lost External Gas–Confirm Internal Settings. The user evaluates the patient condition and determines whether air with oxygen should be supplied to the patient.. | Air without oxygen is used to supply sweep gas
Alarms: Lost External Gas–Confirm Internal Settings. The user evaluates the patient condition and determines whether air with oxygen should be supplied to the patient.. | Air with oxygen is used to supply sweep gas
Alarms: Lost External Gas –No Internal Oxygen Set. The user reconnects the system to external gas.. | Air without oxygen is used to supply sweep gas
Alarms: Lost External Gas –No Internal Oxygen Set. The user evaluates the patient condition and determines whether air without oxygen should be supplied to the patient.. | Air without oxygen is used to supply sweep gas
Alarms: Lost External Gas –No Internal Oxygen Set. The user evaluates the patient condition and determines whether air without oxygen should be supplied to the patient.. | Air with oxygen is used to supply sweep gas
Emergency: High Pressure Side Port (Pigtail stopcock) crossthreaded or disconnected/broken from being caught on an external object. Replace pigtail stopcock. | User does not replace pigtail stopcock
Emergency: Pigtail Removal from PLU Housing Tab . Pigtail is removed from PLU Housing connection tab. | User breaks PLU housing tab

## pFMEA_Pool
Input_statement | Process fm | Failure Mode | Hazard
Assembling of luer connector. Winded PUR tube. Purge tube. | Cannula is not cleaned with benzine
Assembling of luer connector. Winded PUR tube. Purge tube. | Exposure to UV light is too long
Assembling of luer connector. Winded PUR tube. Purge tube. | Exposure to UV light is too short
Assembling of luer connector. Winded PUR tube. Purge tube. | Kink protection purge lumen positioned in the wrong direction on the purge tube - conus shows in direction of the windings
Assembling of luer connector. Winded PUR tube. Purge tube. | Shelf life of Loctite AA3321 has expired
Assembling of luer connector. Winded PUR tube. Purge tube. | Shrink tube inner is not assembled
Assembling of luer connector. Winded PUR tube. Purge tube. | Temperature of heat gun is too high (> 150°C)
Assembling of luer connector. Winded PUR tube. Purge tube. | Temperature of heat gun is too low (< 150°C)
Assembling of luer connector. Winded PUR tube. Purge tube. | Too little Loctite AA3321 is used
Assembling of luer connector. Winded PUR tube. Purge tube. | Too much Loctite AA3321 is used
Assembling of luer connector. Winded PUR tube. Purge tube. | Tube is shortened distally too little (> 90mm)
Assembling of luer connector. Winded PUR tube. Purge tube. | Tube is shortened distally too much (< 80mm)
Assembling of luer connector. Winded PUR tube. Purge tube. | Wrong or missing labeling of the part
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Cannula of the purge tube shows in the wrong direction - not the same side than the ring of the pigtail
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Coil support is not positioned in the appropriate groove of the handle
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Coil support not fixed to the handle with Loctite 4062
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Compression of pump head by the pneumatic valve of the pressure tester during the flow test
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Distal end of the purge tube is cut too long
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Distal end of the purge tube is cut too short
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Exposure to UV light is too long
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Exposure to UV light is too short
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Purge tube is not positioned under the connector
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Sealing washer not assembled
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Shelf life of Loctite AA3321 has expired
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Shrink tube catheter is not correctly positioned over the purge tube
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to connect the purge tube to the motor is too high (> 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to connect the purge tube to the motor is too low (< 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to shrink the shrink tube catheter is too high (> 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to shrink the shrink tube catheter is too low (< 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to shrink the shrink tube motor is too high (> 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Temperature of the heat gun to shrink the shrink tube motor is too low (< 150°C)
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Too little Loctite AA3321 is used
Assembling of purge tube. Pigtail, Step 2, Drive, Sealing Washer and Purge Tube. Purge tube is connected to the drive and catheter. | Too much Loctite AA3321 is used
Cleaning process. Drive shaft. Drive shaft ready for molding. | Concentration of acidic water (hydrochloric acid) is too high
Cleaning process. Drive shaft. Drive shaft ready for molding. | Drive Shaft kinked during cleaning process
Cleaning process. Drive shaft. Drive shaft ready for molding. | Drive Shaft not dryed after cleaning process
Cleaning process. Drive shaft. Drive shaft ready for molding. | Drive Shaft not or insufficient cleaned with water after the acidic water process
Cleaning process. Drive shaft. Drive shaft ready for molding. | Drive shafts not cleaned
Cleaning process. Drive shaft. Drive shaft ready for molding. | Temperature during drying process too high
Cleaning process. Drive shaft. Drive shaft ready for molding. | Temperature during drying process too low
Cleaning. Pigtail. Cleaned pigtail. | Pigtail is not cleaned/wiped with benzine
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Bottom and top part of the handle are badly connected to each other
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Bottom and top part of the handle are not connected to each other
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Cable ties are removed too early - glue not fully cured
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Exposure to UV light is too long
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Exposure to UV light is too short
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Kink protection catheter is not glued to the catheter
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Kink protection catheter is not glued to the handle
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Kink protection purge tube is not glued to the handle
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Kink protection purge tube is not glued to the purge tube
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | PCA is not glued to the handle
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Shelf life of Loctite 4062 and/or AA3321 has expired
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Shelf life of Loctite 4062 has expired
. Closing the handle. Purge tube is connected to the drive and catheter. Handle is closed. | Wrong or missing labeling of the part
. Crimping. Outflow tube is connected to the catheter. Square bar is attached to drive shaft. | Crimp diameter is too small
. Crimping. Outflow tube is connected to the catheter. Square bar is attached to drive shaft. | Reinforcement not assembled
. Crimping. Outflow tube is connected to the catheter. Square bar is attached to drive shaft. | Square bar is positioned too close to the catheter - not flush to the proximal end of the drive shaft; drive shaft protrudes
. Crimping. Outflow tube is connected to the catheter. Square bar is attached to drive shaft. | Square bar is positioned too far away from the catheter - not flush to the proximal end of the drive shaft; square bar protrudes
. Cutting of outflow tube. Outflow tube. Outflow Tube, slotted. | Cut edges are frayed
. Cutting of outflow tube. Outflow tube. Outflow Tube, slotted. | Small diameter is not slotted
. Cutting of outflow tube. Outflow tube. Outflow Tube, slotted. | Struts are damaged
. Cutting of outflow tube. Outflow tube. Outflow Tube, slotted. | V-cut is too big
. Cutting of outflow tube. Outflow tube. Outflow Tube, slotted. | V-cut is too small
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Cut not deep enough into the Elastosil form
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Cut too deep into the Elastosil form
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Initial curing temperature in oven is too high (> 70°C)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Initial curing temperature in oven is too low (< 70°C)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Initial curing time in oven is too long (> 1h)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Initial curing time in oven is too short (< 1h)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Retention pins not removed
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Silicone band not removed
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Tempering temperature in oven is too high (> 90°C)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Tempering temperature in oven is too low (< 90°C)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Tempering time in oven is too long (> 35h)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Tempering time in oven is too short (< 35h)
. Demolding. Rotor is molded. Rotor is demolded, cured and tempered. | Wrong or missing labeling of the rotor during all process steps
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | OD of inserted Pin is too small
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Pin is not inserted into the arrester and spiral bearing during the welding process
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much proximally
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation line "Horizontal 2" is too high or position of "Horizontal 1" is too low on the monitor - Arrester and catheter are too close together
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much distally
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation line "Horizontal 2" is too low or position of "Horizontal 1" is too high on the monitor - Distance between areester and catheter is too big
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation lines "Horizontal 2" and "Horizontal 1" are too high on the monitor - welding area is too much distally
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Position of orientation lines "Horizontal 2" and "Horizontal 1" are too low on the monitor - welding area is too much proximally
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Shrink tube is located in the wrong direction on the catheter - slits show in distal direction (welding zone)
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Shrink tube not removed after welding process
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Spiral bearing becomes kinked during insertion
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Welding temperature is too high
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Welding temperature is too high or time is too long
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Welding temperature is too low
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Welding temperature is too low or time is too short
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong cavity is chosen - 1 instead of 3 (ID of cavity is too small)
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong cavity is chosen - 2 instead of 3 (cavity is too short)
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong heating unit installed
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong position of Arrester - is not deep enough in the catheter
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong position of Arrester - is too deep in the catheter
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong position of spiral bearing - distance between arrester and spiral bearing is too big
. Distal Welding step. Positioning of arrester and spiral bearing type L in the catheter. Arrester and spiral bearing are welded into the distal catheter end. | Wrong shrink tube is used - white instead of transparent
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Deairing ring not removed
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Molding tool falls down
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Placeholder for Biresin not pushed deeper into the negative form to puncture the Elastosil skin between injection port and rotor cavity
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Retention pins removed
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Safety screw not removed
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Safty screw not reassembled
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Scalpel brokes during the cuting process
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Scalpel cut in negative form not deep enough
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Scalpel cut in negative form too deep
. Finishing of negative form. Negative form is molded and cured. Master rotor is removed and negative form is ready for the Biresin molding. | Wrong roof part of molding tool assembled afterwards
. Finishing. Tempered rotor. Rotor is finished. | Cut into the drive shaft during removing the deairing reservoir with a pincer
. Finishing. Tempered rotor. Rotor is finished. | Cut into the drive shaft during removing the deairing reservoir with a scalpel
. Finishing. Tempered rotor. Rotor is finished. | cut too deep into the injection point on the drive shaft
. Finishing. Tempered rotor. Rotor is finished. | Missing labeling of the rotor
. Finishing. Tempered rotor. Rotor is finished. | Proximal distortion not fully removed - frayed edges
. Finishing. Tempered rotor. Rotor is finished. | Proximally too much material removed
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Delo cured not long enough (< 16h)
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Delo is not properly mixed
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Delo is too old - shelf life has expired
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Excess Delo not removed from bearing sleeve
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Excess Delo not removed from drive shaft
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Missing labeling of the part
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Too little Delo applied
. Fixation of Bearing Sleeve. Drive shaft sealed. Bearing sleeve attached. | Too much Delo applied
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Catheter is pushed not deep enough into the housing
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Catheter is pushed too deep into the housing
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Delo cured not long enough (< 16h)
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Delo is not properly mixed
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Delo is too old - shelf life has expired
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | The pump head is positioned in the wrong direction on the rack after gluing (pump head down)
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Too little Delo applied
. Gluing step arrester-housing. Circumferential gap between catheter and arrester is filled with Delo AD895. Arrester is glued into the housing. | Too much Delo applied
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Delo is not properly mixed
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Delo is too old - shelf life has expired
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Parts are not cleaned with benzine
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Protection tube positioned too far distally
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Protection tube positioned too far proximally
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Too little Delo applied
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Too much Delo applied
. Gluing. Spiral bearing and protection tube. Spiral bearing is glued into the protection tube . | Wrong or missing labeling of the part
. Inspection. Motor, Step 2. Pump inspected. | Pump kinked due to bad handling
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is cured and master rotor removed. | Wrong mixing ratio of Elastosil - too much Elastosil A
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is cured and master rotor removed. | Wrong mixing ratio of Elastosil - too much Elastosil B
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Argon gas is not connected
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Bottom of mixing cup punctured in the middle
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Curing temperature in oven under pressure is higher than 70°C
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Curing temperature in oven under pressure is lower than 70°C
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Curing time in oven under pressure is longer than 10 min
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Curing time in oven under pressure is shorter than 10 min
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Elastosil Tube is not connected to the injection pot and/or molding tool
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Injection port of Elastosil not not closed with blind plug after injection
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Lid of "Injection Pot, Elastosil" not closed
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Lid of mixing cup not punctured
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Lid of speedmixer pot not arrested
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Missed to turn on the vacuum pump
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Negative form not further cured without pressure at 70°C
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Pressure in pressure pot too low (< 1,8 bar)
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Pressure of dispenser is too high (> 1,7 bar)
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Pressure of dispenser is too low (< 1,3 bar)
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Pressure pot not closed
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Shelf life of Elastosil A and B has expired
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Wrong mixing ratio of Elastosil - too much Elastosil B
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Wrong speedmixer program - time too long and rotation speed too high
. Molding of negative form. Assembled tool/negative form is ready for the Elastosil molding. Negative form is molded and cured. | Wrong speedmixer program - time too short and rotation speed too low
. Packaging. Programmed pump. Packaged pump. | Catheter and handle not cleaned with isopropanol
. Packaging. Programmed pump. Packaged pump. | Ionization blower not used
. Packaging. Programmed pump. Packaged pump. | No ESD label
. Packaging. Programmed pump. Packaged pump. | Parts of the packaging not assembled
. Packaging. Programmed pump. Packaged pump. | Pump not arrested in the packaging
. Packaging. Programmed pump. Packaged pump. | Pump not correctly positioned in the tyvek pouch
. Packaging. Programmed pump. Packaged pump. | Pumps not dryed after cleaning
. Packaging. Programmed pump. Packaged pump. | Sealing machine not correctly used
. Packaging. Programmed pump. Packaged pump. | Tyvek pouch is labeled in the wrong position
. Packaging. Programmed pump. Packaged pump. | Tyvek pouch not labeled
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Ceramic bearing is positioned too far distally in the housing
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Delo cured not long enough (< 16h)
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Delo is not properly mixed
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Delo is too old - shelf life has expired
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | The pump head is positioned in the wrong direction on the rack after gluing (pump head up)
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Too little Delo applied
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Too much Delo applied
. Positioning and Gluing. Parts cleaned and prepared. Ceramic bearing is glued into the housing. | Wrong or missing labeling of the part
. Preparation for next working process. After the tempering process the Welding Sleeve must be removed. The housing must be packaged and the packaging must be labeled correctly. The housing has to relax for minimum 3 days. Housing can be inspected . | Packaging is not correctly labeled - no traceability
. Preparation for next working process. After the tempering process the Welding Sleeve must be removed. The housing must be packaged and the packaging must be labeled correctly. The housing has to relax for minimum 3 days. Housing can be inspected . | Welding sleeve is not removed - housing cannot relax
. Preparation of housing for thermoforming process. Alignment of internal and external coating to each other and the housing and positioning in the Thermoforming, Forming Tool. Housing is prepared for the thermoforming process in the oven. | Components are misaligned
. Preparation of housing for thermoforming process. Alignment of internal and external coating to each other and the housing and positioning in the Thermoforming, Forming Tool. Housing is prepared for the thermoforming process in the oven. | Housing and/or coating is damaged
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Circumferential collar is not removed/cut off
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Overlapping of outflow tube - housing is too long
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Overlapping of outflow tube - housing is too short
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Silicone tube is too long
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Silicone tube is too short
. Preparation of mesh-housing connection. Mesh and housing with attached outflow tube. Mesh and housing are positioned/aligned to each other. | Welding Sleeve wrongly positioned
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Deairing ring not or not correctly positioned
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Master rotor not cleaned with benzine
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Master rotor not correctly positioned in negative form - distal end shows in wrong direction
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Molding tool not cleaned
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Molding tool not secured against dislocation
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Passability of the holes in the molding tools not checked with the intended drive shaft
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Placeholder for Biresin injection not assembled
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Pressure pot is not preheated
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Retention pins not in negative form assembled
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Roof part of the molding tool reversed assembled
. Preparation of negative form. Alignment of all components (e.g. molding tool, master rotor, negative form). Assembled tool/negative form is ready for the Elastosil molding. | Wrong orientation between placeholder for Biresin injection and master rotor
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | No welding sleeve inserted
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | Overlapping of outflow tube - housing is too long
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | Overlapping of outflow tube - housing is too short
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | Silicone tube is too long
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | Silicone tube is too short
. Preparation of outflow tube-housing connection. Outflow tube and housing. Outflow tube and housing are positioned/aligned to each other. | Welding Sleeve wrongly positioned
. Preparation. Catheter, Step 1. Part prepared. | Compression of pump head by the pneumatic valve of the pressure tester during the flow test
. Preparation. Catheter, Step 1. Part prepared. | Silicone tube not removed
. Preparation. Catheter, Step 1. Part prepared. | Silicone tube not slitted
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Bearing sleeve in wrong direction positioned on drive shaft - smaller OD shows in proximal direction
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Ceramic bearing in wrong direction positioned on drive shaft - flushing grooves shows in distal direction
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Ceramic bearing not fixed on drive shaft with silicone tube
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Drive shaft not cleaned with benzine
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Position of sleeve bearing marked too far distally on drive shaft (< 9,8mm)
. Preparation. Drive Shaft, bearing sleeve and Ceramic bearing. Parts prepared. | Position of sleeve bearing marked too far proximally on drive shaft (> 10mm)
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Ceramic bearing and/ or inside of housing sleeve not cleaned with benzine
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Ceramic bearing not fixed on bearing sleeve with silicone tube
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Drive Shaft kinked during rotor insertion in housing
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Mesh ripped by removing the welding sleeve
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Outflow tube not turned on left
. Preparation. Drive Shaft, Step 3 and Outflow Tube Housing Connection. Parts cleaned and prepared. | Welding sleeve already removed
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Catheter and connector not cleaned with benzine
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Connector is not positioned in the correct radial orientation
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Delo cured not long enough (< 16h)
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Delo is not properly mixed
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Delo is too old - shelf life has expired
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Kink protection catheter not positioned or positioned in the wrong direction on the catheter (cone shows in proximal direction)
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Peel away not positioned or positioned in the wrong direction on the catheter (wings shows in distal direction)
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Shrink tube catheter not positioned on the catheter
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Too little Delo applied
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Too much Delo applied
. Prepositioning and connector. Square bar is attached to drive shaft. Parts prepositioned and connector glued to the catheter. | Wrong or missing labeling of the part
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Both shrink tubes not assembled
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Catheter is pushed not deep enough into the housing
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Catheter is pushed too deep into the housing
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Delo is not properly mixed
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Delo is too old - shelf life has expired
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Shrink tube inner is located in the wrong direction on the catheter - slits show in distal direction (welding zone)
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Shrink tube outer is located in the wrong direction on the catheter - slits show in distal direction (welding zone)
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Too little Delo applied
. Prepositioning of Catheter and gluing step arrester-catheter. PU Tube, Outer is prepostioned on the housing. Circumferential gap between catheter and arrester is filled with Delo AD895. | Too much Delo applied
. Prepositioning of PU Tube, Outer. PU Tube, Outer and housing. PU Tube, Outer is prepostioned on the housing. | PU Tube, Outer is positioned too much distally
. Programming. Inspected pump. Programmed pump. | No write protection programmed
. Programming. Inspected pump. Programmed pump. | Wrong program file is used
. Programming. Inspected pump. Programmed pump. | Wrong pump number is programmed
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Catheter is shortened too little
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Catheter is shortened too much
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | OD of inserted Pin is too small
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Pin is not inserted into the arrester and spiral bearing during the welding process
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Shrink tube is located in the wrong direction on the catheter - slits show in distal direction (welding zone)
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Shrink tube not removed after welding process
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Welding temperature is too high
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Welding temperature is too high or time is too long
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Welding temperature is too low
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Welding temperature is too low or time is too short
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Wrong cavity is chosen - 1 instead of 3 (ID of cavity is too small)
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Wrong cavity is chosen - 2 instead of 3 (cavity is too short)
. Proximal Welding step. Arrester and spiral bearing are welded into the distal catheter end. Catheter is shortened to the correct length and spiral bearing is welded into the proximal catheter end. | Wrong shrink tube is used - white instead of transparent
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Position of orientation line "Horizontal 2" is too high or position of "Horizontal 1" is too low on the monitor - PU tube is positioned too close to the arrester
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Position of orientation line "Horizontal 2" is too low or position of "Horizontal 1" is too high on the monitor - Distance between arrester and catheter is too big
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Position of orientation lines "Horizontal 2" and "Horizontal 1" are too high on the monitor - welding area is too much distally
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Position of orientation lines "Horizontal 2" and "Horizontal 1" are too low on the monitor - welding area is too much proximally
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Shrink tube is located in the wrong direction on the catheter - slits show in distal direction (welding zone)
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Welding temperature is too high
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Welding temperature is too high or time is too long
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Welding temperature is too low
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Welding temperature is too low or time is too short
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Wrong cavity is chosen - 1 instead of 3 (ID of cavity is too small)
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Wrong cavity is chosen - 2 instead of 3 (cavity is too short)
. PU Tube, Inner. Arrester and spiral bearing are welded into the catheter. PU Tube, Inner is welded onto the distal catheter end. | Wrong shrink tube is used - white instead of transparent
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | 3-way-valve of the storage-filling tool (component A) not removed at the end of the day and reused
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | 3-way-valve of the storage-filling tool (component B) not removed at the end of the day and reused
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Argon gas is not connected
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Argon gas is not connected for degasing
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Biresin Tube is not connected to the injection pot
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Biresin Tube is not connected to the molding tool
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Bottom of mixing cup punctured in the middle
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Degasing step of Biresin A not performed
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | First ejected material not rejected
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Lid of "Injection Pot, Elastosil" not closed
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Lid of mixing cup not punctured
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Lid of speedmixer pot not arrested
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Missed to turn on the vacuum pump for degasing
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Mixing cup not closed with lid
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Pressure of dispenser is too high (> 1,2 bar)
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Pressure of dispenser is too low (< 0,8 bar)
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Pressure on storage-filling tool of the component A and/or B is too high (> 1,2 bar)
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Pressure on storage-filling tool of the component A and/or B is too low (< 0,8 bar)
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Shelf life of Biresin A and B has expired
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Storage-Filling Tool of the component A and/or B is defect
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Tube clamp not closed
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong end of drive shaft inserted in the molding tool
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong mixing ratio of Biresin - too much Biresin A
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong mixing ratio of Biresin - too much Biresin B
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong speedmixer program - time too long and rotation speed too high
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong speedmixer program - time too short and rotation speed too low
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong speedmixer program for degasing - time too long and rotation speed too high
. Rotor molding. Drive shaft ready for molding. Rotor is molded. | Wrong speedmixer program for degasing - time too short and rotation speed too low
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | EPO-TEK 301 in a wrong position on drive shaft applied
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | EPO-TEK 301 not or not sufficient mixed
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | Mixing ratio of EPO-TEK 301 is wrong - too much component A
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | Mixing ratio of EPO-TEK 301 is wrong - too much component B
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | Too little EPO-TEK 301 on drive shaft applied (< 1µl)
. Sealing of Drive Shaft. Parts prepared. Drive shaft sealed. | Too much EPO-TEK 301 on drive shaft applied (> 1µl)
. Shortening (if required/ordered). Spiral bearing type U. Spiral bearing is shortened. | Spiral bearing is bended or shows burr
. Shortening (if required/ordered). Spiral bearing type U. Spiral bearing is shortened. | Spiral Bearing is shortened too much
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Cables connected to the wrong PCA port
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Cables not twisted
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Exposure to UV light is too long
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Exposure to UV light is too short
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Inductivities positioned in the wrong direction - solder contacts rotated by 90 degrees
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Inductivities positioned in the wrong direction - solder contacts show to the wall of the coil support
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | No Loctite AA3321 applied at the bottom of the inductivities
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Not all solder contacts and joints are cover with Loctite AA3321
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Shelf life of Loctite 4062 has expired
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Some cables badly connected (cold solder joints)
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Temperatur of soldering station is too high (>> 370°C)
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Temperatur of soldering station is too low (<< 370°C)
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Too little solder applied on the contact surfaces of the inductivities
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Too much solder applied on the contact surfaces of the inductivities
. Soldering and Covering. Motor with cables, inductivities, PCA and coil support. Parts soldered and all contacts covered. | Wrong or missing labeling of the part
. Tempering process in the oven. The housings are placed in the oven without the Thermoforming, Forming Tools for 2h at 95°C and a 10h cooling phase. Coated and tempered housing. | Housings are taken out of the oven too early - not fully tempered
. Tempering process in the oven. The housings are placed in the oven without the Thermoforming, Forming Tools for 2h at 95°C and a 10h cooling phase. Coated and tempered housing. | Housings are taken out of the oven too late
. Tempering process in the oven. The housings are placed in the oven without the Thermoforming, Forming Tools for 2h at 95°C and a 10h cooling phase. Coated and tempered housing. | Temperature is too high
. Tempering process in the oven. The housings are placed in the oven without the Thermoforming, Forming Tools for 2h at 95°C and a 10h cooling phase. Coated and tempered housing. | Temperature is too low
. Thermoforming process in the oven. All 9 Tools are placed in the oven at 205°C until the "dummy" tool with the temperature sensor and data logger shows 175°C (+ 1K) . Coated housing. | Housings are too long in the oven - temperature > 176°C
. Thermoforming process in the oven. All 9 Tools are placed in the oven at 205°C until the "dummy" tool with the temperature sensor and data logger shows 175°C (+ 1K) . Coated housing. | Housings are too short in the oven - temperature reached not 175°C
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much distally
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Position of orientation lines "Horizontal 2" is too low on the monitor - welding area is too much proximally
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | PU tube outer is positioned too far distally
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | PU tube outer is positioned too far proximally
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Second welding step not performed
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Second welding step too far distally performed
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Second welding step too far proximally performed
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Shrink tube not peeled off
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Welding temperatur is too high
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Welding temperatur is too high or time is too long
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Welding temperatur is too low
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Welding temperatur is too low or time is too short
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Wrong cavity is chosen - 1 instead of 3 (ID of cavity is too small)
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Wrong cavity is chosen - 2 instead of 3 (cavity is too short)
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Wrong heating unit installed
. Welding of catheter-housing connection. Part prepared. Catheter is connected to the housing. | Wrong or missing labeling of the part
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Drive shaft cover isn't pushed deep enough into the Pigtail
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Pebax tube inner is missing/not positioned
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much distally
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much proximally
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Shrink tube inner is located in the wrong direction on the pigtail - slits show in proximal direction (welding zone)
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Welding temperatur is too high
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Welding temperatur is too high or time is too long
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Welding temperatur is too low
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Welding temperatur is too low or time is too short
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Wrong cavity is chosen - 1 instead of 1 (ID of cavitiy is too small)
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Wrong cavity is chosen - 3 instead of 2 (cavitiy is longer)
. Welding of drive shaft cover-pigtail connection (step 1). Clean pigtail and drive shaft cover. Drive shaft cover is initally connected to the pigtail. | Wrong heating unit installed
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much distally
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much proximally
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Second welding step not performed
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Shrink tube outer is located in the wrong direction on the pigtail - slits show in proximal direction (welding zone)
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Welding temperatur is too high
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Welding temperatur is too high or time is too long
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Welding temperatur is too low
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Welding temperatur is too low or time is too short
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Wrong cavity is chosen - 1 instead of 1 (ID of cavitiy is too small)
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Wrong cavity is chosen - 3 instead of 2 (cavitiy is longer)
. Welding of drive shaft cover-pigtail connection (step 2). Drive shaft cover is initally connected to the pigtail. Drive shaft cover is finally connected to the pigtail. | Wrong or missing labeling of the part
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Drive shaft not positioned in drive shaft cover
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Pebax tube outer is missing/not positioned
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Pebax tube outer is positioned under the mesh
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Pigtail is not positioned in the correct radial orientation
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Pigtail is positioned too close to the rotor (< 3 windings of drive shaft)
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Pigtail is positioned too far away from the rotor (> 5 windings of drive shaft)
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much distally
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much proximally
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Shrink tube outer is located in the wrong direction on the pigtail - slits show in proximal direction (welding zone)
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Shrink tube outer not removed after welding step
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Welding temperatur is too high
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Welding temperatur is too high or time is too long
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Welding temperatur is too low
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Welding temperatur is too low or time is too short
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Wrong cavity is chosen - 1 instead of 1 (ID of cavitiy is too small)
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Wrong cavity is chosen - 3 instead of 2 (cavitiy is longer)
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Wrong heating unit installed
. Welding of housing-pigtail connection. Catheter, Step 3 and Pigtail, Step 1. Pigtail is connected to the housing. | Wrong or missing labeling of the part
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | 180° rotation between 1. and 2. welding step is missing
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | 2. welding step not performed
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Benzine is not completely vaporized
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Missing adjustment of the orientation line "Horizontal 2" in distal direction between the 1. and 2. welding step
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much proximally
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much distally
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Welding temperature is too high
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Welding temperature is too high or time is too long
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Welding temperature is too low
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Welding temperature is too low or time is too short
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Wrong cavity is chosen - 1 instead of 2 (cavity is not existing)
. Welding of mesh-housing connection. Mesh and housing are positioned/aligned to each other. Mesh and outflow tube are connected to the housing. | Wrong cavity is chosen - 3 instead of 2 (cavity is not existing)
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much distally
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much proximally
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Shrink tube inner is positioned too far distally on the catheter/outflow tube
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Shrink tube inner is positioned too far proximally on the catheter/outflow tube
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Welding temperatur is too high
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Welding temperatur is too high or time is too long
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Welding temperatur is too low
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Welding temperatur is too low or time is too short
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Wrong cavity is chosen - 1 instead of 1 (ID of cavity is too small)
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Wrong cavity is chosen - 3 instead of 2 (cavity is longer)
. Welding of outflow tube-catheter connection. Catheter, Step 2. Outflow tube is connected to the catheter. | Wrong heating unit installed
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | 180° rotation between 1. and 2. welding step is missing
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | 2. welding step not performed
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Benzine is not completely vaporized
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Position of orientation line "Horizontal 2" is too high on the monitor - welding area is too much proximally
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Position of orientation line "Horizontal 2" is too low on the monitor - welding area is too much distally
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Welding temperature is too high
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Welding temperature is too high or time is too long
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Welding temperature is too low
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Welding temperature is too low or time is too short
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Wrong cavity is chosen - 1 instead of 2 (cavity is not existing)
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Wrong cavity is chosen - 3 instead of 2 (cavity is not existing)
. Welding of outflow tube-housing connection. Outflow tube and housing are positioned/aligned to each other. Outflow tube is connected to the housing. | Wrong heating unit installed
. Winding. PUR tube. Winded PUR tube. | Exposure time of heat gun is distally too long (< 55s)
. Winding. PUR tube. Winded PUR tube. | Exposure time of heat gun is too short
. Winding. PUR tube. Winded PUR tube. | PEEK sleeve is not removed from the shrink tube motor
. Winding. PUR tube. Winded PUR tube. | Removing of tube before tool is cooled down
. Winding. PUR tube. Winded PUR tube. | Shrink tube motor not applied
. Winding. PUR tube. Winded PUR tube. | Temperature of heat gun is too high (> 160°C)
. Winding. PUR tube. Winded PUR tube. | Temperature of heat gun is too low (< 160°C)
. Winding. PUR tube. Winded PUR tube. | Tube is not pushed to the end of the axial groove of the winding tool
. Winding. PUR tube. Winded PUR tube. | Wrong winding direction
. Xantopren imprint. After the sub-process step the housing has to relax for minimum 3 days. Housing is ready for the Xantopren imprint. | Housing has not relaxed for minimum 3 days
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Housing and Xantopren imprint not correctly marked - missing traceability
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Housing not correctly cleaned from Xantopren residues
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Housing packed in wrong container - wrong traceability
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Improper or wrong name allocation of Keyence data file - missing traceability
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Mixed mass not correctly applied in the housing
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Too early performance of Keyence measurement
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Wrong mixing ratio for the Xantopren is used - too much Xantopren L
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Wrong mixing ratio for the Xantopren is used - too little Activator
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Wrong mixing ratio for the Xantopren is used - too little Xantopren L
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Wrong mixing ratio for the Xantopren is used - too much Activator
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Xantopren imprint from Thermoforming Mandrel removed
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Xantopren imprint too early removed from the housing
. Xantopren imprint. The mixed Xantopren (Xantopren L + Activator) must be injected into the housing, cured and removed. Measurement of the minimal ID of the housing by measuring the minimal OD of the imprint. | Xantopren L and Activator are not correctly/homogenously mixed
5.5 Template Packaging and Inspection (0550-7044). Place template in pouch. pouch and template. pouched template. | template graphic not visible in pouch
5.5 Template Packaging and Inspection (0550-7044). Seal Pouch . pouch sealer and pouched template. sealed pouched template. | pouch not properly sealed
AIC Assembly. Front Housing Assembly. Connection of both optical sockets to both optical benches. both optical socket connected. | mixed up connections
AIC Assembly. Front Housing Assembly. ferrite and fuse on PBM. Ferrite and fuse soldered on PBM. | No connection leads to no power supply for the double optical carrier board
AIC Assembly. Front Housing Assembly. Installation of Double Optical Carrier Board, Kapton tape. installed Double Optical carrier Board and protected PBM. | No kapton tape on PBM leads to damage to the PBM
AIC Assembly. Front Housing Assembly. Second optical socket hole drill, front housing. Second optical socket hole, front housing. | wrong position, pump connector cover cap can not be fully opened
AIC Assembly. Front Housing Assembly. Second optical socket, Dust cap. Mounted second optical socket in front housing. | dust in second optical socket
AIC Assembly. Front Housing Assembly. Second optical socket, glue. Mounted second optical socket in front housing. | not enough glue leeds to insufficient bond strength
AIC Assembly. Front Housing Assembly. Second optical socket. Mounted second optical socket in front housing. | incorrct alignment of socket and nut leeds to impropper connection of socket and connector
AU Catheter. Anneal catheters. Oven (time and temperatures), catheters. Annealed catheters. | variation in bond strength of catheters leading to failure
AU Catheter. Apply adhesion promoter HPA 41 on bonding surface of outlet cage and dry it in oven. HPA 41, outlet cage assembly, oven (time and temperature), holding fixture, swabs. outlet cage with promotor. | Weak bond
AU Catheter. Apply adhesion promoter HPA 41 on bonding surface of outlet cage and dry it in oven. HPA 41, outlet cage assembly, oven (time and temperature), holding fixture, swabs. outlet cage with promotor. | Weak bond due to leftover promotor
AU Catheter. Apply adhesion promoter HPA 41 on bonding surfaces on the cannula and inflow cage, then dry it in oven. Componets with adhesion promotor. Cool components. | damage inner surface of outflow cage
AU Catheter. Apply adhesion promoter HPA 41 on bonding surfaces on the cannula and inflow cage, then dry it in oven. Promotor (HPA 41), Cannula, Inflow cage, swab (woven), oven (time and temperature), cannula holding fixture, . Componets with adhesion promotor. | Weak bond
AU Catheter. Apply Delo Duopox to bonding surface of inflow cage and slide cannula onto bonding surface . Delo, inflow cage, cannula, dosing needle, cup. Inflow cage and cannula bonded together with delo. | Delo protrudes bond site/inflow cage to cannula transition not smooth leading to insertion issues
AU Catheter. Apply Delo Duopox to bonding surface of inflow cage and slide cannula onto bonding surface . Delo, inflow cage, cannula, dosing needle, cup. Inflow cage and cannula bonded together with delo. | Gap between inflow cage and cannula leading to thrombus formation
AU Catheter. Apply Delo Duopox to bonding surface of inflow cage and slide cannula onto bonding surface . Delo, inflow cage, cannula, dosing needle, cup. Inflow cage and cannula bonded together with delo. | Glue migrates into inflow cage/cannula
AU Catheter. Apply Delo Duopox to bonding surface of inflow cage and slide cannula onto bonding surface . Delo, inflow cage, cannula, dosing needle, cup. Inflow cage and cannula bonded together with delo. | Misaligned pigtail, leading to positioning issues
AU Catheter. Apply Delo Duopox to bonding surface of inflow cage and slide cannula onto bonding surface . Delo, inflow cage, cannula, dosing needle, cup. Inflow cage and cannula bonded together with delo. | Weak bond
AU Catheter. Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | incorrect orientation of cannula to pigtail (positioning issues)
AU Catheter. Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | OD too large, delivery issues
AU Catheter. Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | Weak bond
AU Catheter. Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | delo migrates to inside of cannula
AU Catheter. Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | tilted cage leading to gap between cage and cannula
AU Catheter. Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | Weak bond
AU Catheter. Clean components. End-Cap Catheter Jacket Petroleum ether Wipe/Swab. Cleaned components. | insufficient bond
AU Catheter. Clean inflow cage and cannula at the bonding area with petroleum ether. inflow cage, petroluem ether, swab, cannula. clean components. | Weak bond due surface contamination
AU Catheter. Clean outflow cage and cannula at the bonding areas with acetone. Cannula, outflow cage assembly, acetone, swabs, beaker, holding fixture, time. cleaned components and assemblies. | Weak bond
AU Catheter. Clean outflow cage and cannula at the bonding areas with acetone. Cannula, outflow cage assembly, acetone, swabs, beaker, holding fixture, time. cleaned components and assemblies. | Weak bond due to contamination
AU Catheter. Clean transition ring, pump-sided inner catheter end, and support ring. petroleum ether, catheter, tranisition ring, swabs, tweezers, clean room wipes. Cleaned components. | Weak bond
AU Catheter. Cool down parts to room temperature. Cannula, cannula holding fixture, inflow cage. Cool components. | Speed up delo curing process leading to weak bond
AU Catheter. Cool down parts to room temperature. outlet cage with promotor, holding fixture. outlet cage cooled down to room temp. | Weak bond
AU Catheter. Crimp catheter onto transition ring and then remove excessive. crimping pliers, swaps, glued together catheter and tranistion ring, support ring. crimped together assembly. | Damage catheter
AU Catheter. Crimp catheter onto transition ring and then remove excessive. crimping pliers, swaps, glued together catheter and tranistion ring, support ring. crimped together assembly. | unable assemble support ring (scrap/rework)
AU Catheter. Crimp catheter onto transition ring and then remove excessive. crimping pliers, swaps, glued together catheter and tranistion ring, support ring. crimped together assembly. | Weak bond
AU Catheter. Cure at room temperature and then place in oven to cure, then clean components. Cannula tray, bonded together cannula and inflow cage, oven (time and temperature), timer. AU cannula. | Gap between inflow cage and cannula leading to thrombus formation
AU Catheter. Cure at room temperature and then place in oven to cure, then clean components. Cannula tray, bonded together cannula and inflow cage, oven (time and temperature), timer. AU cannula. | Misaligned pigtail, leading to positioning issues
AU Catheter. Cure at room temperature and then place in oven to cure, then clean components. Cannula tray, bonded together cannula and inflow cage, oven (time and temperature), timer. AU cannula. | Weak bond
AU Catheter. Cut catheter with cutting fixture to specific length. Catheter, steel ruler, scapel. Catheter cut to length. | Unable to package device (scrap)
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | exposed metal
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | incomplete tipping
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | insufficient bond
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | insufficient bond, air leaking at catheter junction
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | pits/voids/gaps
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | pits/voids/gaps leading to scrap
AU Catheter. Initiate Tipping process. properly seated endcap, properly aligned Catheter, Catheter Tipping Machine, (Temp, Force, Time) Guide sleeve Catheter clamp air pressure. Tipped catheter assembly. | smeared ink markings on catheter tip
AU Catheter. Load catheter jacket into machine, push onto first bard of end cap. Catheter, Catheter Tipping Machine. Properly aligned and loaded catheter jacket. | incomplete tipping
AU Catheter. Load catheter jacket into machine, push onto first bard of end cap. Catheter, Catheter Tipping Machine. Properly aligned and loaded catheter jacket. | incorrect alignment of catheter jacket
AU Catheter. Load endcap onto catheter tipping mandrel. Endcap, Notch Mandrel (30 degree), Catheter Tipping Machine. Properly seated endcap. | incomplete tipping
AU Catheter. Load endcap onto catheter tipping mandrel. Endcap, Notch Mandrel (30 degree), Catheter Tipping Machine. Properly seated endcap. | incorrect alignment of catheter jacket
AU Catheter. Load Tipping Recipe. Access Level Password. correct recipe loaded. | insufficient bond
AU Catheter. Load Tipping Recipe. Access Level Password. correct recipe loaded. | insufficient bond, air leaking at catheter junction
AU Catheter. Package and ship assembly . Tipped Catheter subassembly in heat sealed bag, bubble wrap, cardboard box. Packaged units. | Damaged catheter assembly (unusable/scrap)
AU Catheter. Package and ship assembly . Tipped Catheter subassembly in heat sealed bag, bubble wrap, cardboard box. Packaged units. | Reduced joint strength
AU Catheter. Remove flash ring, clean flash. tipped catheter flash removal tool. tipped catheter with no flash/flaking material. | damaged to distal end of catheter jacket
AU Catheter. Removed tipped catheter form machine. tipped catheter bushing gripping fixture. tipped catheter with acceptable print markings. | print markings damaged on catheter
AU Catheter. Removed tipped catheter form machine. tipped catheter bushing gripping fixture. tipped catheter with no damage/skive marks. | clamp marks on catheter jacket
AU Catheter. Removed tipped catheter form machine. tipped catheter bushing gripping fixture. tipped catheter with no damage/skive marks. | damage/skive marks on catheter
AU Catheter. Removed tipped catheter form machine. tipped catheter bushing gripping fixture. tipped catheter with no damage/skive marks. | Discolored end cap
AU Catheter. Removed tipped catheter form machine. tipped catheter bushing gripping fixture. tipped catheter without contamination/residue. | contamination/residue on catheter jacket
AU Catheter. subassembly/storage. tipped catheter subassembly, heat sealed bag. tipped catheter subassembly ready for next level assembly. | insufficient bond strength
AU Catheter. subassembly/storage. tipped catheter subassembly, heat sealed bag. tipped catheter subassembly ready for next level assembly. | kinked catheter jacket
AU Catheter. subassembly/storage. tipped catheter subassembly, heat sealed bag. tipped catheter subassembly ready for next level assembly. | scratched end caps
AU Catheter. verify machine alignment. inductance coil, mandrel, guide sleeve, carriage. machine components properly aligned. | insufficient bond
AU Catheter. verify machine alignment. inductance coil, mandrel, guide sleeve, carriage. machine components properly aligned. | insufficient bond, air leaking at catheter junction
AU Catheter-Lumen Bonding. after curing, cool down motor housing in a magazine; and preform patency and tightness check in pyrogene free water. beaker, pyrogen free water, cured, twin pin to motor housing assembly, leakage/patency test system . Tested twin pin to motor housing assembly. | leaks/occlusions are allowed to pass
AU Catheter-Lumen Bonding. after curing, cool down motor housing in a magazine; patency and tightness check in qualified water. beaker, qualified water, cured, twin pin to motor housing assembly, leakage test system, motor housing sealing plugs . Tested twin pin to motor housing assembly. | Damage shrink tube
AU Catheter-Lumen Bonding. after curing, cool down motor housing in a magazine; patency and tightness check in qualified water. beaker, qualified water, cured, twin pin to motor housing assembly, leakage test system, motor housing sealing plugs . Tested twin pin to motor housing assembly. | Insufficient cooling time
AU Catheter-Lumen Bonding. Apply catheter transition with Delo Duopox 1895 and slide catheter above transition. delo-doupox 1985, twin pin and motor housing and catheter. Twin pin, catheter with end cap, and motor housing with glued transition. | Bond Failure
AU Catheter-Lumen Bonding. Apply catheter transition with Delo Duopox 1895 and slide catheter above transition. delo-doupox 1985, twin pin and motor housing and catheter. Twin pin, catheter with end cap, and motor housing with glued transition. | Tranistion with weak seal leading to fluid leak path
AU Catheter-Lumen Bonding. Catheter leakage test. Cured assembly, leakage test system. tested cured assembly. | Kink catheter
AU Catheter-Lumen Bonding. Catheter leakage test. Cured assembly, leakage test system. tested cured assembly. | unit with leak path in transition passes
AU Catheter-Lumen Bonding. clean catheter metal end cap with petroleum ether and slide on end cap onto catheter with o'ring. catheter, end cap, o'ring, petroleum ether, swab, gloves, time. cleaned catheter and endcap with o'ring. | Poor bond strength
AU Catheter-Lumen Bonding. clean catheter metal end cap with petroleum ether and slide on end cap onto catheter with o'ring. catheter, end cap, o'ring, petroleum ether, time. cleaned catheter and endcap with o'ring. | Poor bond strength
AU Catheter-Lumen Bonding. Clean catheter transition area, apply catheter transition with Adhesive Epoxy Duopox 1895 and slide catheter above transition. Adhesive Epoxy Duopox 1895, twin pin and motor housing and catheter, petroleum ether. Twin pin, catheter with end cap, and motor housing with glued transition. | bond failure at catheter joint
AU Catheter-Lumen Bonding. Clean catheter transition area, apply catheter transition with Adhesive Epoxy Duopox 1895 and slide catheter above transition. Adhesive Epoxy Duopox 1895, twin pin and motor housing and catheter, petroleum ether. Twin pin, catheter with end cap, and motor housing with glued transition. | Missaligned catheter to pressure port leading to insertion issues
AU Catheter-Lumen Bonding. Clean catheter transition area, apply catheter transition with Adhesive Epoxy Duopox 1895 and slide catheter above transition. Adhesive Epoxy Duopox 1895, twin pin and motor housing and catheter, petroleum ether. Twin pin, catheter with end cap, and motor housing with glued transition. | too little evaporation time leading to bond failure
AU Catheter-Lumen Bonding. clean end of motor housing with brush, clean assembly twin pin with swab using petroleum ether. brush, motor housing, petroleum ether, twin pin assembly. cleaned motor housing, and twin pin assembly. | damage motor cables
AU Catheter-Lumen Bonding. clean end of motor housing with brush, clean assembly twin pin with swab using petroleum ether. brush, motor housing, petroleum ether, twin pin assembly. cleaned motor housing, and twin pin assembly. | Leak due to poor bond adhesion at bond site
AU Catheter-Lumen Bonding. clean end of motor housing with brush, clean assembly twin pin with swab using petroleum ether. brush, motor housing, petroleum ether, twin pin assembly. Cleaned motor housing, and twin pin assembly. | Motor wires (red, white, green) kink/damaged
AU Catheter-Lumen Bonding. clean end of motor housing with brush, cleanassembly twin pin, the end of the purge and pressure lumen and nitinol wire with petroleum ether. brush, motor housing, petroleum ether, twin pin assembly, swab, time, gloves. cleaned motor housing, and twin pin assembly. | Poor bond strength
AU Catheter-Lumen Bonding. Clean tranistion area of catheter. Catheter, petroleum ether, swabs. Cleaned catheter transition area. | bond failure due to contimination
AU Catheter-Lumen Bonding. Cure in oven and cool down. Oven (time and temp), racks, bend protector, Twin pin, catheter with end cap, and motor housing with glued transition. Cured assembly. | bond failure at catheter joint
AU Catheter-Lumen Bonding. Cure in oven. Oven (time and temp), ranks, bend protector, Twin pin, catheter with end cap, and motor housing with glued transition. Cured assembly. | Bond Failure
AU Catheter-Lumen Bonding. Cure in oven. Oven (time and temp), ranks, bend protector, Twin pin, catheter with end cap, and motor housing with glued transition. Cured assembly. | Tranistion with weak seal leading to fluid leak path
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | catheter too long (unable to assemble)
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | catheter too short (positioning issue)
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | Creased/kink end of catheter
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | cut incorrect end of catheter leading to incorrect position identified
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | Debry accumulation, leading to weak bond
AU Catheter-Lumen Bonding. Cut catheter to length. catheter, cutting machine. Catheter at correct length. | misangled cut, gap between catheter and motor housing
AU Catheter-Lumen Bonding. cut PTFE mandrel, clean with petroleum ether and slide into the twin pin. PTFE mandrel, petroleum ether, twin pin. cleaned mandrel inside twin pin . | mandrel not cleaned leading to contimination issues
AU Catheter-Lumen Bonding. cut PTFE mandrel, clean with petroleum ether and slide into the twin pin. PTFE mandrel, petroleum ether, twin pin. cleaned mandrel inside twin pin . | mandrel not full seated lead to purge line blockage
AU Catheter-Lumen Bonding. dry motor housing in furnace. tested twin pint to motor housing assembly, oven (time and temp). Dried twin pin to motor housing assembly. | pump not fully dried leading to contimination issues
AU Catheter-Lumen Bonding. fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, Vacuum pump. Twin pin assembly in support fixture with catheter with endcap on. | Damage/kink nitinol wire
AU Catheter-Lumen Bonding. fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, Vacuum pump. Twin pin assembly in support fixture with catheter with endcap on. | Damaged motor cables leading to pump stop
AU Catheter-Lumen Bonding. fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, Vacuum pump. Twin pin assembly in support fixture with catheter with endcap on. | damaged/kinked pressure lumen
AU Catheter-Lumen Bonding. fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, Vacuum pump. Twin pin assembly in support fixture with catheter with endcap on. | damaged/kinked purge lumen
AU Catheter-Lumen Bonding. fix pressure lumen, purge lumen, nitinol wire and cable with adhesive tape and apply Adhesive Epoxy Duopox 1895 to the connection point between twin pin assembly and motor housing . Adhesive Epoxy Duopox 1895, adhesive tape, positioned twin pin assembly inside motor housing. twin pin assembly inside motor housing with glued transition. | Inability to assemble catheter
AU Catheter-Lumen Bonding. fix pressure lumen, purge lumen, nitinol wire and cable with adhesive tape and apply Adhesive Epoxy Duopox 1895 to the connection point between twin pin assembly and motor housing . Adhesive Epoxy Duopox 1895, adhesive tape, positioned twin pin assembly inside motor housing. twin pin assembly inside motor housing with glued transition. | Leakage due to poor bond strength
AU Catheter-Lumen Bonding. fix pressure lumen, purge lumen, nitinol wire and cable with adhesive tape and apply delo to the connection point between twin pin assembly and motor housing . Delo-Duopox, adhesive tape, positioned twin pin assembly inside motor housing. Twin pin assembly inside motor housing with glued transition. | Delo residue on motor cable
AU Catheter-Lumen Bonding. fix pressure lumen, purge lumen, nitinol wire and cable with adhesive tape and apply delo to the connection point between twin pin assembly and motor housing. Then apply additional delo on the notch of the twin pin. . Delo-Duopox 1895, adhesive tape, positioned twin pin assembly inside motor housing. twin pin assembly inside motor housing with additional glue on transition. | Leakage at bond site due to microchannels
AU Catheter-Lumen Bonding. fix pressure lumen, purge lumen, nitinol wire and cable with adhesive tape and apply delo to the connection point between twin pin assembly and motor housing. Then apply additional delo on the notch of the twin pin. . Delo-Duopox 1895, adhesive tape, positioned twin pin assembly inside motor housing. twin pin assembly inside motor housing with additional glue on transition. | Poor bond strength
AU Catheter-Lumen Bonding. position motor housing on a drying board and cure. drying board, oven (time and temp), twin pin assembly inside motor housing with additional glued transition. cured twin pin to motor housing assembly. | Poor bond strength
AU Catheter-Lumen Bonding. position motor housing on a drying board and cure. drying board, oven (time and temp), twin pin assembly inside motor housing with glued transition. cured twin pin to motor housing assembly. | Leakage due poor bond strength
AU Catheter-Lumen Bonding. position stainless steel mandrel in motor housing, apply adhesive Epoxy Duopox 1895 to twin pin and then push mandrel into twin pin and slide assembly into the catheter junction. Twin pin assembly, Adhesive Epoxy Duopox 1895, motor housing, stainless steel mandrel, dosing needle. Twin pin assembly with Adhesive Epoxy Duopox 1895 inserted into catheter junction. | Leak due to poor bond
AU Catheter-Lumen Bonding. position stainless steel mandrel in motor housing, apply Adhesive Epoxy Duopox 1895 to twin pin and then push mandrel into twin pin and slide assembly into the catheter junction. Twin pin assembly, Adhesive Epoxy Duopox 1895, motor housing, stainless steel mandrel, dosing needle. Twin pin assembly with Adhesive Epoxy Duopox 1895 inserted into catheter junction. | Damage shrink tube
AU Catheter-Lumen Bonding. position stainless steel mandrel in motor housing, apply Delo Duopox on purge pin and slide assembly into the catheter junction. Twin pin assembly, delo-duopox 1895, motor housing, auxillary tool, PTFE mandrel, Stainless Steel Mandrel. assembly with delo covered inside motor housing. | Air gap between motor housing and lumen leading to purge leak through catheter
AU Catheter-Lumen Bonding. position stainless steel mandrel in motor housing, apply Delo Duopox on purge pin and slide assembly into the catheter junction. Twin pin assembly, delo-duopox 1895, motor housing, auxillary tool, PTFE mandrel, Stainless Steel Mandrel. assembly with delo covered inside motor housing. | Poor bond strength
AU Catheter-Lumen Bonding. position stainless steel mandrel in motor housing, apply Delo Duopox to twin pin and then push mandrel into twin pin and slide assembly into the catheter junction. Twin pin assembly, delo-duopox, motor housing, stainless steel mandrel, dosing needle. Twin pin assembly with delo inserted into catheter junction. | Delo residue on motor cable
AU Catheter-Lumen Bonding. Prepare PTFE mandrel cutting it length, then clean with petroleum ether and slide into the connection pin. PTFE mandrel, petroleum ether, connection pin, oven (time and temperature), calipers, pliers. cleaned mandrel inside connection pin . | Incorrect diameter of mandrel (unable to assemble)
AU Catheter-Lumen Bonding. remove mandrel out of the motor housing, slide assembly onto auxillary tool until dead stop then removed. Twin pin assembly with Adhesive Epoxy Duopox 1895 inserted into catheter junction, auxillary tool. positioned twin pin assembly inside motor housing. | Damage shrink tube
AU Catheter-Lumen Bonding. remove mandrel out of the motor housing, slide assembly onto auxillary tool until dead stop then removed. Twin pin assembly with Adhesive Epoxy Duopox 1895 inserted into catheter junction, auxillary tool. positioned twin pin assembly inside motor housing. | incorrect positioning leading to leakage purge
AU Catheter-Lumen Bonding. remove mandrel out of the motor housing, slide assembly onto auxillary tool until dead stop then removed. Twin pin assembly with Adhesive Epoxy Duopox 1895 inserted into catheter junction, auxillary tool. positioned twin pin assembly inside motor housing. | incorrect positioning of twin pin in the motor housing leading to interference with ball bearing
AU Catheter-Lumen Bonding. remove mandrel out of the motor housing, slide onto auxillary tool, then remove from tool. assembly with delo covered inside motor housing, auxillary tool, PTFE mandrel. positioned twin pin assembly inside motor housing. | leak due to unit not fully seated
AU Catheter-Lumen Bonding. remove mandrel out of the motor housing, slide onto auxillary tool, then remove from tool. assembly with delo covered inside motor housing, auxillary tool, PTFE mandrel. positioned twin pin assembly inside motor housing. | Poor bond strength
AU Catheter-Lumen Bonding. Remove tape on lumens, fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring. Twin pin assembly in support fixture with catheter with endcap on. | Damaged motor cables leading to pump stop
AU Catheter-Lumen Bonding. Remove tape on lumens, fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, scapel. Twin pin assembly in support fixture with catheter with endcap on. | damaged/kinked placement lumen
AU Catheter-Lumen Bonding. Remove tape on lumens, fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, scapel. Twin pin assembly in support fixture with catheter with endcap on. | damaged/kinked purge lumen
AU Catheter-Lumen Bonding. Remove tape on lumens, fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, scalpel. Twin pin assembly in support fixture with catheter with endcap on. | Exposed motor wire conductor
AU Catheter-Lumen Bonding. Remove tape on lumens, fix catheter in twin pin support fixture, insert both lumen, motor cable and nitinol wire with vacuum into the catheter. twin pin support fixture, twin pin to motor housing assembly, catheter with endcap and o'ring, scapel. Twin pin assembly in support fixture with catheter with endcap on. | Motor wire shielding exposed
AU Catheter-Lumen Bonding. rinse motor housing with Isoproponal ; dry with compressed air, then place in oven. Oven (time and temp), isoproponal, tested cured assembly, . Cured catheter with end cap, twin pin, and motor housing assembly. | unit not cleaned leading to contimination issues
AU Catheter-Lumen Bonding. rinse motor housing with Isoproponal ; dry with compressed air. Oven (time and temp), isoproponal, tested cured assembly, . Dried assembly. | purge path not cleaned leading to contimination issues
AU Catheter-Lumen Bonding. Shorten catheter to correct length. Ruler, flexible tubing cutter, catheter. Catheter at correct length. | Too short catheter (scrap)
AU Inflow Cage and Pigtail. Evaluate pigtail for kinkability, clean inner and outer surface with petroleum ether, clean inflow cage with petroleum ether. Pigtail, inflow cage, petroleum ether, HFE wipe. Cleaned pigtail and inflow cage. | bond failure due to contiminated surface
AU Inflow Cage and Pigtail. Flare end of pigtail. Hakko controller (teemp and time), pigtail flaring fixture, thread forming mandrel . Flared pigtail. | unable to assemble
AU Inflow Cage and Pigtail. Flare end of pigtail. Hakko controller (temp and time), pigtail flaring fixture, thread forming mandrel . Flared pigtail. | Bond failure to end having poor contact with inflow cage
AU Inflow Cage and Pigtail. Mix epotek 301, adhere pigtail to inflow cage top. Epo-Tek 301, sleeve tube, inflow cage, flared pigtail, inflow cage assembly fixture, dosing needle. Flared pigtail adhered to top of inflow cage. | incorrectly seated pigtail leading to bond failure
AU Inflow Cage and Pigtail. Mix Epotek 302 together, apply to struts of Inflow cage, and cure in oven. cured pigtail to inflow cage, Epo-tek 302, holding fixture, oven (time and temperature), dosing needle, syringe, speed mixer. cured inflow cage to pigtail assembly. | rough struts and stagnat locations on inflow cage
AU Inflow Cage and Pigtail. Place Epo-Tek 301 at transition area and place in oven for curing. Flared pigtail adhered to top of inflow cage, Epo-tek 301, dosing needle, Oven (time and temperature). cured pigtail to inflow cage with transition. | bond failure
AU Inflow Cage and Pigtail. Place Epo-Tek 301 at transition area and place in oven for curing. Flared pigtail adhered to top of inflow cage, Epo-tek 301, dosing needle, Oven (time and temperature). cured pigtail to inflow cage with transition. | Gap between pigtail and inflow cage leading thrombus formation
AU Inflow Cage Assy to CP Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | Bonding failure at inflow cage to cannula bond
AU Inflow Cage Assy to CP Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | Increase scrap
AU Inflow Cage Assy to CP Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | protrusion of Delo into the window leading to material flake off during insertion/operation
AU Inflow Cage Assy to CP Cannula. Clean outer surface of inflow cage and inner diameter of cannula with petroleum. Then apply HPA 41 to those same areas. Then dry.. Petroleum ether, Cannula, pigtail-inflow cage assembly, swab, HPA 41, Oven (time and temp). Cleaned and prepped cannula and pigtail-inflow cage assembly. | Bonding failure at inflow cage to cannula bond
AU Inflow Cage Assy to CP Cannula. Clean outer surface of inflow cage and inner diameter of cannula with petroleum. Then apply HPA 41 to those same areas.. Petroleum ether, Cannula, pigtail-inflow cage assembly, swab, HPA 41, Oven (time and temp). Cleaned and prepped cannula and pigtail-inflow cage assembly. | Bonding failure at inflow cage to cannula bond
AU Inflow Cage Assy to CP Cannula. Place assembly into fixture and then into oven for curing. Oven (time and temperature), glued together cannula and inflow cage-pigtail, CP Inflow Cage Holding fixture. Assy, Cannula. Long Pump 350. | Bonding failure at inflow cage to cannula bond
AU Inflow Cage. Clean end of pigtail with petroleum ethe . pigtail, petroleum ether, compressed air, swabs, wipes. Cleaned pigtail. | contamination due to particles left on pigtail leading to weak bond
AU Inflow Cage. Clean thread of the ball drop with petroleum ether . inflow cage, petroluem ether, wipe. cleaned inflow cage. | contamination due to particles leading to bond failure of pigtail
AU Inflow Cage. Clean welding rod. Welding rod, petroleum ether, swabs, wipe, pressurized air. Cleaned fixture. | Scratches on inside of inflow cage (hemolysis/thrombus formation)
AU Inflow Cage. Cleaning of the polished inflow cage. polished inflow cage, cleaning fixture, beaker, isopropyl, ultra sonic bath, storage box. cleaned inflow cage. | Damage inflow cage (scratches)
AU Inflow Cage. Cleaning of the polished inflow cage. polished inflow cage, cleaning fixture, beaker, isopropyl, ultra sonic bath, storage box. cleaned inflow cage. | insufficient cleaning (biocompatibility)
AU Inflow Cage. Cure adhesive . pigtail and inflow cage assembly, oven (time and temperature), holding fixture. cured assembly. | Bond failure of pigtail to inflow cage
AU Inflow Cage. Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | Damage to pigtail leading thrombos formation
AU Inflow Cage. Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | Damage to pigtail leading to prolapsed pigtail
AU Inflow Cage. Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | pigtail incorrect length leading to position issues
AU Inflow Cage. Enlarge pigtail end with heat gun, threaded mandrel, and fan. Cool with compressed air . pigtail, threaded mandrel, heat gun, compressed air. Threaded pigtail. | Failure to fully engage pigtail to cage (thrombus formation)
AU Inflow Cage. Enlarge pigtail end with heat gun, threaded mandrel, and fan. Cool with compressed air . pigtail, threaded mandrel, heat gun, compressed air. Threaded pigtail. | Failure to fully engage pigtail to cage weak bond
AU Inflow Cage. Enlarge pigtail end with heat gun, threaded mandrel, and fan. Cool with compressed air . pigtail, threaded mandrel, heat gun, compressed air. Threaded pigtail. | Unable to assemble pigtail to cage due to incomplete threads
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Bond failure of pigtail to inflow cage
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Damage inflow cage (deformation/bending)
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Damage pigtail (prolapse)
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Failure to fully engage pigtail to cage (thrombus formation)
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Failure to fully engage pigtail to cage weak bond
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Occluded pigtail (unable to assemble red lumen)
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | Over stress of the pigtail material leading to weak bond
AU Inflow Cage. Mix Epotek. Apply Epo-tek 301 to thread and screw on pigtail. . Epo-tek 301, pigrail, inflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1). Pigtail bonded with Epoxy to inflow cage. | too much engagement leading to pigtail goemetry change (Positioning issue)
AU Inflow Cage. Perform Final visual inspection for cured unit and store assembly. finished subassembly (inflow cage + pigtail), blister + cover. Inspected and package subassembly. | Contamination of unit (biocompatibility)
AU Inflow Cage. Perform Final visual inspection for cured unit and store assembly. finished subassembly (inflow cage + pigtail), blister + cover. Inspected and package subassembly. | Damage inflow cage/pigtail
AU Inflow Cage. Polish inflow cage with the burnishers. Welded inflow cage, burnisher (RPM), Brush made out of goat hair, transportation box, interdental brush, isopropyl, cup, scapel, pressurized air. polished inflow cage. | Cage damaged leading to sharp edges
AU Inflow Cage. Polish inflow cage with the burnishers. Welded inflow cage, burnisher (RPM), Brush made out of goat hair, transportation box, interdental brush, isopropyl, cup, scapel, pressurized air. polished inflow cage. | Scratches on inside of inflow cage (hemolysis/thrombus formation)
AU Inflow Cage. Separate bars of cage by cutting ring. inflow cage, side cutter. inflow cage with bars separated. | Cage damaged leading to sharp edges
AU Inflow Cage. Separate bars of cage by cutting ring. inflow cage, side cutter. inflow cage with bars separated. | Weld failure
AU Inflow Cage. Set welding parameters. Laser welder, Laser welder program (adjusting frequency). Laser welder setup with correct parameters. | Weak weld strength
AU Inflow Cage. Setup laser welder. laser welder, argon, laser table, for manual laser welding, positioning nosol for gas. Laser welder setup. | corrosion at weld site leading to failure
AU Inflow Cage. Setup laser welder. laser welder, argon, laser table, for manual laser welding, positioning nosol for gas. Laser welder setup. | Equipment failure leading to poor weld
AU Inflow Cage. Slide inflow cage and ball on welding rod. Weld bars to ball. welding rod, ball, laser, inflow cage, laser table, positioning nosol. Welded together inflow cage and ball. | Weak weld strength
AU Inflow Cage. Slide inflow cage and ball on welding rod. Weld bars to ball. welding rod, ball, laser, inflow cage, laser table, positioning nozzle. Welded together inflow cage and ball. | Misalignment of cage and ball (insertion/positioning issues)
AU Inflow Cage. Slide inflow cage and ball on welding rod. Weld bars to ball. welding rod, ball, laser, inflow cage, laser table, positioning nozzle. Welded together inflow cage and ball. | rough transition of ball to cage leading to insertion issues
AU Inflow Cage. Slide inflow cage and ball on welding rod. Weld bars to ball. welding rod, ball, laser, inflow cage, laser table, positioning nozzle. Welded together inflow cage and ball. | Weak weld strength
AU Inflow-Cage to Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | Bonding failure at inflow cage to cannula bond
AU Inflow-Cage to Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | delo migrates into ID of cannula and inflow cage
AU Inflow-Cage to Cannula. Apply delo to the inflow cage and place cannula over designated area. Delo, cleaned and prepped cannula and pigtail-inflow cage assembly, timer. Glued together cannula and inflow cage-pigtail. | Increase scrap
AU Inflow-Cage to Cannula. Clean outer surface of inflow cage and inner diameter of cannula with petroleum. Then apply HPA 41 to those same areas. Then dry.. Petroleum ether, Cannula, pigtail-inflow cage assembly, swab, HPA 41, Oven (time and temp). Cleaned and prepped cannula and pigtail-inflow cage assembly. | Bonding failure at inflow cage to cannula bond
AU Inflow-Cage to Cannula. Place assembly into fixture and then into oven for curing. Oven (time and temperature), glued together cannula and inflow cage-pigtail, CP Inflow Cage Holding fixture. Assy, Cannula. Long Pump 350. | Bonding failure at inflow cage to cannula bond
AU MABS Check Valve to Smiths Rsvr. Assemble Check Valve to Rsvr, apply Loctite 3321 and cure rsvr to check valve. Rsvr Assy w/ no valve, check valve assy, Loctite 3321, dispensing needle, UV curing equipment. Reservoir with loctite. | Bond failure
AU MABS Check Valve to Smiths Rsvr. Clean and prep all bonding surfaces. pertoleum ether, swab, rsvr assy w/ no valve, Check Valve Assy. clean and prepped components. | Contamination Leading to Scrap
AU Motor Housing. Add UV adhesive to yoke slot. Yoke with shrink tube, UV glue, UV light source (time and distance. Yoke with shrink tube and UV adhesive. | Unable to assemble optical fiber
AU Motor Housing. Apply Frekote to the tip of the insertion pin. Insertion pin assembly, Frekote, swabs, timer. Insertion pin assembly with Frekote. | Frekote migrates to wires leading to improper adhesion of epotek which causes leakage current
AU Motor Housing. Apply Frekote to the tip of the insertion pin. Insertion pin assembly, Frekote, swabs, timer. Insertion pin assembly with Frekote. | Molding issue leading to scrap
AU Motor Housing. Apply heat shrink to yoke and trim remaining material after cycle. Oven (time and temperature), Assy magnetic yoke stand, scapel. Yoke with shrink tube. | Damage yoke during trimming leading to leakage path
AU Motor Housing. Apply release agent to insertion pin, let dry. Slide on mold insert.. Release agent, insertion pin, magnetic yoke, mold insert, hand press. Stator prepared for molding. | parts do not separate from mold leading to scrap
AU Motor Housing. Apply release agent to insertion pin, let dry. Slide on mold insert.. Release agent, insertion pin, magnetic yoke, mold insert, hand press. Stator prepared for molding. | Radial offest of insert leading to unable to route sensor cable (scrap)
AU Motor Housing. Apply release agent to insertion pin, let dry. Slide on mold insert.. Release agent, insertion pin, magnetic yoke, mold insert, hand press. Stator prepared for molding. | Release agent on winding/cable leading to corrosion due to moisture ingress
AU Motor Housing. Apply release agent to insertion pin, let dry.. Release agent, AU Stator. Stator prepared for molding. | parts do not separate from mold leading to scrap
AU Motor Housing. Apply release agent to insertion pin, let dry.. Release agent, AU Stator. Stator prepared for molding. | Release agent on winding/cable leading to corrosion due to moisture ingress
AU Motor Housing. Apply release agent to mold consistently over mold surface and dry with air pressure, then preheat mold. Mold, Release agent, oven (time and temperature), Air pressure, swab. Preheated mold release agent. | cosmetic defects
AU Motor Housing. Apply release agent to mold consistently over mold surface and dry with air pressure, then preheat mold. Mold, Release agent, oven (time and temperature), Air pressure, swab. Preheated mold release agent. | Inconsistent geometry of epotek leading to voids causing leakage current issues
AU Motor Housing. Apply release agent to mold consistently over mold surface and dry with air pressure, then preheat mold. Mold, Release agent, oven (time and temperature), Air pressure, swab. Preheated mold release agent. | parts do not separate from mold leading to scrap
AU Motor Housing. Assemble yoke and peek form inserter onto stator . Peek form inserter, stator, yoke, placement fixture. Assembled stator with yoke and peek form inserter. | Damaged motor cables leading to pump stop
AU Motor Housing. Assemble yoke and peek insert onto stator . Peek insert, stator, yoke, placement fixture. Assembled stator with yoke and peek insert. | Damaged motor cables leading to pump stop
AU Motor Housing. AU Motor Housing (WI 0048-7000). Screw the cast mold; assemble the riser in the appropriate hole. Riser, Assembled stator with yoke and peek form inserter in molded halves. | Cast mold with assembled riser in appropriate hole
AU Motor Housing. AU Motor Housing (WI 0048-7000). Screw the cast mold; assemble the riser tube in the appropriate hole. Riser tube, Assembled stator with yoke and peek form inserter in molded halves. | Cast mold with assembled riser in appropriate hole
AU Motor Housing. Clamp in syringe, feed casting tube through chamber valve and connect to mold and then lock chamber. vacuum chamber, clamps, syringe with casting tube, mold halves. Molding chamber prepared for process. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Clean mold using sand blaster, scapel, and clean room swab dipped in isopropanol.. Mold, sand blaster, scapel, swab, Isopropanol. clean mold. | Short shots/incorrect geometry
AU Motor Housing. Clean mold using sand blaster, scapel, and clean room swab dipped in isopropanol.. Mold, sand blaster, scapel, swab, Isopropanol. clean mold. | Surface defects on finshed motor housing
AU Motor Housing. Clean yoke with petroleum ether. Swab, AU Stator, petroleum ether. Clean assembly. | Uncleaned yoke leading to corrosion due to moisture ingress
AU Motor Housing. Clean yoke with petroleum ether. Swab, stator, petroleum ether. Clean assembly. | Uncleand yoke leading to corrosion due to moisture ingress
AU Motor Housing. Connect casting tube (syringe) with adaptor tube and mold. Cast mold with assembled riser in appropriate hole, with seal off outlet with Xantropen, disposable syringe, adaptor tube. Cast mold with assembled riser in appropriate hole, with seal off outlet with Xantropen with connected syringe with adaptor tube. | syringe/tubes not connected leading to increase in scrap
AU Motor Housing. Connect disposable syringe to the casting tube. syringe, casting tube. Connected syringe to casting tube. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Connect motor cable to resistance checker and test resistance. molded motor housing, measurement system setup. tested molded motor housing. | components with unacceptable leakage current pass (false positive)
AU Motor Housing. Connect motor cable to resistance checker and test resistance. molded motor housing, measurement system setup. tested molded motor housing. | Good components fail (False negatives)
AU Motor Housing. connect motor cable with checking device and check resistance . Molded motor housing, resistance test device. Molded motor housing. | Rejects are accepted
AU Motor Housing. evacuate chamber with the vacuum pump. Epotek 301, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. Epotek 301, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. | Chamber failed to reach vacuum leading to scrap
AU Motor Housing. evacuate chamber with the vacuum pump. Epotek 301, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. Epotek 301, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. | Chamber not evacuated properly leading to scrap
AU Motor Housing. evacuate chamber with the vacuum pump. Epotek 301, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. Epotek 301, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. | Internal voids in molding area around PCB leading to corrosion from heat/leakage current
AU Motor Housing. evacuate chamber with the vacuum pump. Epotek 301, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. Epotek 301, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. | Internal voids leading to weak structure for purge connection leading to leak
AU Motor Housing. evacuate chamber with the vacuum pump. Grouting, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. Grouting, Cast mold with connected syringe with coupling in Vacuum (vacuum) chamber. | Chamber not evacuated properly leading to scrap
AU Motor Housing. Evacuate chamber with vacuum pump. Epo-tek (301), molding chamber prepared for process. Epo-tek mixture in syringe and casting tube extended beyond valve with vacuum present. | Part has internal bubbles or cracks
AU Motor Housing. Evacuate chamber with vacuum pump. Epo-tek (301), molding chamber prepared for process. Epo-tek mixture in syringe and casting tube extended beyond valve with vacuum present. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Fill sensor notch on insert with UV glue, apply UV light. AU Stator, UV glue, UV light. Filled sensor notch on insert. | Excess UV glue on stator
AU Motor Housing. Fill sensor notch on insert with UV glue, apply UV light. AU Stator, UV glue, UV light. Filled sensor notch on insert. | Glue not hardened sufficiently
AU Motor Housing. Fill sensor notch on insert with UV glue, apply UV light. AU Stator, UV glue, UV light. Filled sensor notch on insert. | Incompletely filled notch
AU Motor Housing. fill syringe with Epotek, pinch off syringe adaptor with the pneumatic clamp when the Epotek hits the clamp. Cast mold with connected syringe with adaptor tube in Vacuum chamber, Epotek 301. Epotek, Cast mold with connected syringe with adaptor tube in Vacuum chamber. | clamp not closed leading to scrap
AU Motor Housing. fill syringe with Epotek, pinch off syringe adaptor with the pneumatic clamp when the Epotek hits the clamp. Cast mold with connected syringe with adaptor tube in Vacuum chamber, Epotek 301. Epotek, Cast mold with connected syringe with adaptor tube in Vacuum chamber. | not enough Epotek in syringe leading to scrap
AU Motor Housing. Fill syringe with Epo-Tek, pinch off syringe coupling with pneumatic valve. Epo-tek (301), molding chamber prepared for process. Epo-tek mixture in syringe and casting tube extended beyond valve. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Insert alignment fixture, then screw the cast mold; assemble the riser in the appropriate hole. Riser, Assembled stator with yoke and peek form inserter in molded halves, Alignment fixture. Cast mold with assembled riser in appropriate hole. | mold not tight leading to scrapped unit
AU Motor Housing. Insert alignment fixture, then screw the cast mold; assemble the riser in the appropriate hole. Riser, Assembled stator with yoke and peek form inserter in molded halves, Alignment fixture. Cast mold with assembled riser in appropriate hole. | riser not mounted correctly leading to scrap
AU Motor Housing. Insert alignment fixture, then screw the cast mold; assemble the riser in the appropriate hole. Riser, Assembled stator with yoke and peek form inserter in molded halves, Alignment fixture. Cast mold with assembled riser in appropriate hole. | unable to assemble
AU Motor Housing. insert cast tube (syringe) into vacuum chamber, clamp syringe in bracket and tighten nob to squeeze O'ring. Cast mold with assembled riser in appropriate hole, with seal off outlet with Xantropen with connected syringe with adaptor tube, vacuum chamber. Cast mold in Vacuum chamber. | syringe/tubes not sealed properly leading to unable reach vacuum
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | Damage motor cables/cable insulation
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | Eccentricity of stator in mold
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | incorrect geometry of molded components leading to scrap
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | Incorrect mold depth leading to poor rotor engagement to ball bearing
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | Incorrect mold depth leading to too much engagement with ball bearing (purge flow issues or mechanical engagement with bearing and magnet)
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | PTFE wire incorrect position causing incorrect geometry of nitinol wire notche leading to unable to assemble (scrap)
AU Motor Housing. Insert stator into lower mold half with PTFE wire, adjust its end position, press on upper mold half. PTFE wire, scapel, stator, upper and lower mold halves. Stator within upper and lower mold halves. | Sensor notch incorrect position causing incorrect geometry of sensor wire notch leading to unable to assemble (scrap)
AU Motor Housing. Insert stator into the lower cast mold half, position end stop up to insertion pin, and assemble top half of mold. Assembled stator with yoke and peek form inserter, mold halves. Assembled stator with yoke and peek form inserter in molded halves. | damage to insulation of motor cable leading to leakage current
AU Motor Housing. Insert stator into the lower cast mold half, position end stop up to insertion pin, and assemble top half of mold. Assembled stator with yoke and peek form inserter, mold halves. Assembled stator with yoke and peek form inserter in molded halves. | Damage to yoke due incorrect assembly of mold halves leading to leakage current
AU Motor Housing. Insert stator into the lower cast mold half, position end stop up to insertion pin, and assemble top half of mold. Assembled stator with yoke and peek form inserter, mold halves. Assembled stator with yoke and peek form inserter in molded halves. | end position not fixed leading to motor housing depth incorrect
AU Motor Housing. Insert stator into the lower cast mold half, position end stop up to insertion pin, and assemble top half of mold. Assembled stator with yoke and peek form inserter, mold halves. Assembled stator with yoke and peek form inserter in molded halves. | Mold not fixed properly leading to scrap
AU Motor Housing. Mix Epotek Resin and aluminum oxide, then stir by hand and heat in oven, add hardener and stir in vacuum mixer. Resin, hardener, scale, aluminum oxide, vacuum mixer, oven (time and temperature), Mixing Ratio spreadsheet. Epotek mixture. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Mix Epotek Resin and aluminum oxide, then stir by hand and heat in oven, add hardener and stir in vacuum mixer. Resin, hardener, scale, aluminum oxide, vacuum mixer, oven (time and temperature), Mixing Ratio spreadsheet. Epotek mixture. | Uncured material reduced mechanical properties of epotek molding
AU Motor Housing. open gripper with master switch, when Epotek reaches the correct level in riser tube aerate chamber. Epotek 301, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. Grouting, Cast mold with connected syringe with adaptor tube in Vacuum (vacuum) chamber. | gripper opened too early leading to scrap (short shot)
AU Motor Housing. Open mold, remove mold insert from insertion pin and remove insertion pin completely from motor housing. Cured motor housing within mold, insertion pin extraction fixture, pliers. molded motor housing out of mold. | Cracks in motor housing
AU Motor Housing. Open mold, remove mold insert from insertion pin and remove insertion pin completely from motor housing. Cured motor housing within mold, insertion pin extraction fixture, pliers. molded motor housing out of mold. | Damage motor cables
AU Motor Housing. Open mold, remove mold insert from insertion pin and remove insertion pin completely from motor housing. Cured motor housing within mold, insertion pin extraction fixture, pliers. molded motor housing out of mold. | Damage to internal shrink tubing
AU Motor Housing. Perform leakage current test. molded motor housing, measurement system setup. tested molded motor housing. | components with unacceptable leakage current pass (false positive)
AU Motor Housing. Perform leakage current test. molded motor housing, measurement system setup. tested molded motor housing. | Good components fail (False negatives)
AU Motor Housing. Perform leakage current testing. Molded motor housing, leakage current tester. Molded motor housing. | falls positives leading to scrap
AU Motor Housing. Perform leakage current testing. Molded motor housing, leakage current tester. Molded motor housing. | Rejects are accepted
AU Motor Housing. Perform leakage current testing. Molded motor housing, leakage current tester. Molded motor housing. | wrong leakage current displayed leading to scrap
AU Motor Housing. Perform molding process by openign clamps allowing Epo-tek to flow into mold. Epo-tek (301), molding chamber prepared for process. Mold filled with Epotek. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. place cast mold into oven. Cast mold with Epotek 301, heating cabinet (temperature). Completed heating cycle with mold with grout. | Under curred assembly leading to structural integriy issues resulting pump stop
AU Motor Housing. Place filled mold on hot plate. aluminum mold plate, mold filled with Epotek, hot plate (temperature), timer. Cured motor housing within mold. | Part has internal bubbles or cracks
AU Motor Housing. Place filled mold on hot plate. aluminum mold plate, mold filled with Epotek, hot plate (temperature), timer. Cured motor housing within mold. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Place filled mold on hot plate. aluminum mold plate, mold filled with Epotek, hot plate (temperature), timer. Cured motor housing withing mold. | Part has internal bubbles or cracks
AU Motor Housing. Place filled mold on hot plate. aluminum mold plate, mold filled with Epotek, hot plate (temperature), timer. Cured motor housing withing mold. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. place insertion pin on fixture 0046-65178. Stator with assembly with no outer shrink tube. Stator placed in fixture and labled. | Mix up of lots or AU types
AU Motor Housing. Place plexi-glass door on chamber. Cast mold with connected syringe with coupling in Vacuum chamber. Cast mold with connected syringe with coupling in Vacuum chamber. | Door not sealed leading to unable to achieve vacuum
AU Motor Housing. Plasma activate stator with motor cable. Holding board, plasma chamber, stator with no shrink tube. Plasma activated stator placed on fixture and labled. | Shrink delamination/protrusion leading to pump stop
AU Motor Housing. Plasma activate stator with motor cable. Holding board, plasma chamber, stator with no shrink tube. Plasma activated stator. | Shrink delamination/protrusion leading to pump stop
AU Motor Housing. Plasma activate stator with motor cable. Plasma activated stator placed on fixture and labled. Plasma activated stator. | Shrink delamination/protrusion leading to pump stop
AU Motor Housing. Preheat mold. Assembled mold halves with riser and feeder and Xantopren, oven (time and temperature). Preheated mold with stator assembly. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Prepare syringe assembly. Syringe, adaptor tube, casting tube. Syringe assembly for molding. | syringe/tubes not assembled correctly, leaking epotek out of syringe leading to molding shorts (scrap)
AU Motor Housing. Prepare the Epotek mixture. Gloves, Resin, hardener, and filler, speedmixer, stirrer. Prepared Epotek. | incorrect mixture of Epo-Tek leading to pump stop (structural integrity loss/soft epotek)
AU Motor Housing. Prepare the Epotek mixture. Gloves, Resin, hardener, and filler, speedmixer, stirrer. Prepared Epotek. | non-homogenious mixture of Epo-Tek leading to pump stop (structural integrity loss/soft epotek)
AU Motor Housing. Remove casting tubes, loosen screws on mold and then back out insertion pin. Then pull motor housing out of mold.. Casting mold, tools for opening mold, Pin extraction fixture. Molded motor housing with peek mold inlay. | insertion pin not backed out leading scrap
AU Motor Housing. Remove heat shrink from yoke. Molded motor housing with heat shrink on yoke, scapel, tweezers, microscope. Molded motor housing without heat shrink. | Damage yoke during removal leading to leakage path
AU Motor Housing. remove mold inlay from insertion pin from housing and remove all casting residual. molded motor housing with peek mold inlay, scapel. Molded motor housing. | damage motor cables
AU Motor Housing. remove mold inlay from insertion pin from housing and remove all casting residual. molded motor housing with peek mold inlay, scapel. Molded motor housing. | Damage yoke during removal leading to leakage path
AU Motor Housing. remove mold inlay from insertion pin from housing and remove all casting residual. molded motor housing with peek mold inlay, scapel. Molded motor housing. | engagement pin not removed leading to unable to assemble
AU Motor Housing. remove mold inlay from insertion pin from housing and remove all casting residual. molded motor housing with peek mold inlay, scapel. Molded motor housing. | motor housing damaged during pulling pin leading to pump stop
AU Motor Housing. Remove outer shrink tube from stator . Stator, scapel, tweezers. Stator with assembly with no outer shrink tube. | Damage to winding during removal of shrink tube leading to pump stop
AU Motor Housing. Remove outer shrink tube from stator . Stator, scapel, tweezers. Stator with assembly with no outer shrink tube. | shrink tube not removed leading to unable to mold (scrap)
AU Motor Housing. Rinse motor housing with isopropanol, dry with compressed air and in oven. Isopropanol, molded motor housing, compressed air, oven (time and temperature), glass beaker. rinsed and dried motor housing. | contamination of fluid path
AU Motor Housing. Rinse motor housing with pyrogen-free water and dry with compressed air. mold motor housing, pyrogen-free water, compressed air, syringe. rinsed and dried motor housing. | corrosion/crystal formation inside motor housing
AU Motor Housing. Screw the cast mold; assemble the riser in the appropriate hole. Riser, Assembled stator with yoke and peek form inserter in molded halves. Cast mold with assembled riser in appropriate hole. | riser not mounted correctly leading to scrap
AU Motor Housing. Screw the cast mold; assemble the riser tube in the appropriate hole. Riser tube, Assembled stator with yoke and peek form inserter in molded halves. Cast mold with assembled riser in appropriate hole. | mold not tight leading to scrapped unit
AU Motor Housing. Seal off motor cable outlet with Xantopren. Xantopren, assembled molded halves with riser and feeder. Assembled mold halves with riser and feeder and Xantopren. | Epotek on motor cables leading to scrap
AU Motor Housing. Seal off motor cable outlet with Xantopren. Xantopren, assembled molded halves with riser and feeder. Assembled mold halves with riser and feeder and Xantopren. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Seal off motor cable outlet with Xantopren. Xantopren, assembled molded halves with riser and feeder. Assembled mold halves with riser and feeder and Xantopren. | PTFE wire incorrect position causing incorrect geometry of nitinol wire notche leading to unable to assemble (scrap)
AU Motor Housing. Seal off the outlet of the motor cable with Xantropen. Xantropen, Cast mold with assembled riser in appropriate hole. Cast mold with assembled riser in appropriate hole, with seal off outlet with Xantropen. | outlet not sealed and/or not properly sealed leading to scrap
AU Motor Housing. Secondary cure of the motor housing . Oven (time and temperature), Molded motor housing with peek mold inlay. Cure cycle completed for assembly. | Under curred assembly leading to structural integriy issues resulting pump stop
AU Motor Housing. Take sample of Epo-tek from syringe and conduct hardness test. Epotek sample, hardness measurement system. Tested sample. | false negative
AU Motor Housing. Take sample of Epo-tek from syringe and conduct hardness test. Epotek sample, hardness measurement system. Tested sample. | false positive
AU Motor Housing. Tighten mold halves, assemble riser and feeder in appropriate hole. Stator within upper and lower mold halves, torque wrench, riser, feeder. Assembled mold halves with riser and feeder. | Partially molded/incorrect geometry leading to scrap
AU Motor Housing. Transport of stator utilizing fixture 0046-65178. Stator placed in fixture and labled. stator on fixture and labled placed in pasma chamber. | Damage of stator or cable due to catching motor cable
AU Motor Housing. Transport of stator utilizing fixture 0046-65178. Stator placed in fixture and labled. stator on fixture and labled placed in pasma chamber. | Damage of stator/ cable due to touching nearby stator
AU Motor Housing. Wet the cast mold with a release agent. Mold and release agent. Mold with release agent. | no release agent utilized leading to cosmetic defect
AU Motor Housing. Wet the cast mold with a release agent. Mold and release agent. Mold with release agent. | too little agent utilized leading to increased scrap
AU Motor Housing. Wet the cast mold with a release agent. Mold and release agent. Mold with release agent. | too much agent utilized leading to cosmetic defect
AU Motor Pump. Clean and prep inside of motor of pump catheter assembly (non-welded) and the ball bearing. Swab, petroleum ether, pump catheter assembly (non-welded), and ball bearing. Cleaned and prepped ball bearing and pump catheter assembly (non-welded). | bond failure leading to bearing wear or failure at bond site
AU Motor Pump. Cure Assembly with glued in ball bearing. Oven (time, temperature), drying rack. Cured assembly with pump catheter assembly (non-welded) with ball bearing. | bond failure leading to bearing wear or failure at bond site
AU Motor Pump. Cure assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device, oven (time and temperature). AU Pump Motor (non-welded). | bond failure leading to bearing wear or failure at bond site
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | bearing wear leading to pump stop
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | damage coil during insertion leading to scrap
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | damage coil during insertion leading to short
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to bearing wear
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to unable to assemble rest of assembly
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to bond failure or bearing wear
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to unable to install rotor
AU Motor Pump. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab. Pump Catheter assembly (non-welded) with glued in ball bearing . | unable to seat rotor leading to scrap
AU Motor Pump. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | bond failure leading to bearing wear or failure at bond site (too shallow insertion)
AU Motor Pump. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | Damage Impeller blades/tip
AU Motor Pump. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to scrap
AU Motor Pump. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to short
AU Motor Pump. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | purge blockage due to glue migration into purge area
AU Motor Pump. Prep and clean rotor assembly. Rotor assembly, petroleum assembly, swab. Cleaned and prepped rotor assembly. | bond failure leading to bearing wear or failure at bond site
AU Motor. Adjust parameters and start activation, remove retainer plate after process. Assembly, Activation Chamber (Gas, Vacuum, Time, Mass flow of gas). Plasma activated assembly. | Reduced mechanical properties of adhesive bond for purge lumen
AU Motor. Adjust parameters and start activation, remove retainer plate after process. Assembly, Activation Chamber (Gas, Vacuum, Time, Mass flow of gas). Plasma activated assembly. | Shrink tube delamination from coil
AU Motor. Apply Delo-Duopox to end of impeller shaft and hub, slide rotor assembly into motor housing and position with shim in the fixture. Delo-Duopox, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Bond failure leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and hub, slide rotor assembly into motor housing and position with shim in the fixture. Delo-Duopox, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Damage coil windings during insertion
AU Motor. Apply Delo-Duopox to end of impeller shaft and hub, slide rotor assembly into motor housing and position with shim in the fixture. Delo-Duopox, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Delo applied to incorrect location leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and hub, slide rotor assembly into motor housing and position with shim in the fixture. Delo-Duopox, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Shim not utlized leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and sleeve bearing, slide rotor assembly into motor housing and position with shim in the fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Bond failure leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Bond failure leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Bond failure of rotor to bearing leading to increased engagement which results in pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Damage coil windings
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Delo applied to incorrect location leading to pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Shim not utlized leading increased engagement of rotor to ball bearing which results in pump stop
AU Motor. Apply Delo-Duopox to end of impeller shaft and slide bearing, slide rotor assembly into motor housing and position with shim fixture. Delo, Impeller, sleeve bearing, motor housing, shim, Motor Assembly fixture . Assembly in fixture with glued in impeller. | Shim not utlized leading to pump stop
AU Motor. Clean and prep inside of motor of pump catheter assembly (non-welded) and the ball bearing. Swab, petroleum ether, pump catheter assembly (non-welded), and ball bearing. Cleaned and prepped ball bearing and pump catheter assembly (non-welded). | bond failure leading to bearing wear or failure at bond site
AU Motor. Clean impeller shaft end and hub with petroleum ether and dry with compressed air. petroleum ether, compressed air, rotor assembly, motor housing assembly. Cleaned assemblies. | Contiminated bonding surface leading to bond failure leading to pump stop
AU Motor. Clean impeller shaft end and sleeve bearing with petroleum ether and dry with compressed air. Impeller shaft, sleeve bearing, petroleum ether, compressed air, time. Cleaned assemblies. | Contiminated bonding surface leading to bond failure leading to pump stop
AU Motor. Clean impeller shaft end and sleeve bearing with petroleum ether and dry with compressed air. Impeller shaft, sleeve bearing, petroleum ether, compressed air. Cleaned assemblies. | Contiminated bonding surface leading to bond failure leading to pump stop
AU Motor. Clean motor housing interior, stiffening pin, and purge lumen with petroleum ether and dry with compressed air. Motor housing, petroleum ether, AU purge lumen, Air supply, swabs, wipes. Cleaned assemblies. | Contiminated bonding surface leading to leak
AU Motor. Cure adhesive at room temperature. Then cure adhesive in oven.. Oven (time and temperature), Glued assembly in fixture, timer. Cured assembly. | Bond failure leading to pump stop
AU Motor. Cure adhesive. Bonded Motor housing to AU purge lumen, Oven (time and temperature), Magazine Rack, Tape. Cured adhesive on assembly. | bond failure leading to leak
AU Motor. Cure adhesive. Oven (time and temperature), Glued assembly in fixture. Cured assembly. | Bond failure leading to pump stop
AU Motor. Cure adhesive. Oven (time and temperature), Glued assembly in fixture. Cured assembly. | Bond failure of rotor to bearing leading to increased engagement which results in pump stop
AU Motor. Cure Assembly with glued in ball bearing. Oven (time, temperature), drying rack, pump catheter assembly. Cured assembly with pump catheter assembly (non-welded) with ball bearing. | bond failure leading to bearing wear or failure at bond site
AU Motor. Cure assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device, oven (time and temperature). AU Pump Motor (non-welded). | bond failure leading to bearing wear or failure at bond site
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Bearing failure leading to pump stop
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | damage coil during insertion
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | glue on coil leading to short
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to bearing wear
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to unable to assemble rest of assembly
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to bond failure or bearing wear
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to unable to install rotor
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | unable rotate bearing leading to unable to start pump
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Adhesive Epoxy Duopox 1895, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | unable to seat rotor leading to scrap
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Bearing failure leading to pump stop
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | damage coil during insertion
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | glue on coil leading to short
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to bearing wear
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | Insertion tilted leading to unable to assemble rest of assembly
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to bond failure or bearing wear
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | poor insertion leading to unable to install rotor
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | unable rotate bearing leading to unable to start pump
AU Motor. Glue in ball bearing into pump catheter assembly (non-welded) and clean excess glue. Delo, ball bearing holder tool, swab, dosing needle. Pump Catheter assembly (non-welded) with glued in ball bearing . | unable to seat rotor leading to scrap
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | failure at bond site (too shallow / Rotor assembly detaches
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | Damage Impeller blades/tip
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to scrap
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to short
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | eccentricity issues/tilted rotor
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | purge blockage due to glue migration into purge area
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Adhesive Epoxy Duopox 1895, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | Too much engagement leading to purge issues
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | failure at bond site (too shallow / Rotor assembly detaches
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | Damage Impeller blades/tip
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to scrap
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | damage to coil leading to short
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | eccentricity issues/tilted rotor
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | purge blockage due to glue migration into purge area
AU Motor. Glue in rotor assembly into the pump catheter assembly (non-welded) with ball bearing. Delo, rotor assembly, pump catheter assembly (non-welded) with ball bearing, shim, motor clamping device. Glued assembly with rotor, pump catheter assembly (non-welded) with ball bearing and shim in place resting in clamping device. | Too much engagement leading to purge issues
AU Motor. Match the rotational position to the motor housing notch. AU Rotor, AU Motor housing. Matched unit. | Rotor stiffeness issues leading to high motor current
AU Motor. Match the sensor notch position to the motor housing notch. AU Rotor, AU Motor housing. Matched unit. | Sensor cannot be mounted in later step
AU Motor. Perform dry fit test of rotor assembly into motor housing. rotor assembly, motor housing with glued purge lumen, . Tested assembly. | damage coating of magnet
AU Motor. Perform dry fit test of rotor assembly into motor housing. rotor assembly, motor housing with glued purge lumen, . Tested assembly. | stiff impeller (scrap)
AU Motor. Perform dry mounting of the rotor assembly . AU Rotor, Motor housing with glued purge lumen,. Paired assembly. | Stiff rotor issue
AU Motor. Perform dry mounting of the rotor assembly . AU Rotor, Motor housing with glued purge lumen,. Paired assembly. | Unable to assemble device (scrap)
AU Motor. Perform Tightness at hub. Cured assembly, Tightness tet setup. Measured tightness of assembly. | High rate of units to MRB
AU Motor. Perform Tightness test at hub. AU Motor, Tightness Test Setup, glass beaker, pyrogen free water. Measured assembly (Tightness Level). | Fail acceptable units for untighness (excess purge fluid)
AU Motor. Prep and clean rotor assembly. Rotor assembly, petroleum assembly, swab, rotor assembly tray fixture. Cleaned and prepped rotor assembly. | failure at bond site due to contimination
AU Motor. Prepare Epo-Tek 301 adhesive, stir in speed mixer, apply to stiffening pin and bond to motor housing with mandrel, then tape motor cable to purge lumen for correct positioning. Speed mixer, AU purge assembly, motor housing, Epotek 301, mandrel, dosing needle, tweezer. Bonded Motor housing to AU purge lumen. | bond failure leading to leak
AU Motor. Prepare Epo-Tek 301 adhesive, stir in speed mixer, apply to stiffening pin and bond to motor housing with mandrel, then tape motor cable to purge lumen for correct positioning. Speed mixer, AU purge assembly, motor housing, Epotek 301, mandrel, dosing needle, tweezer. Bonded Motor housing to AU purge lumen. | Contamination of ball bearing
AU Motor. Prepare Epo-Tek 301 adhesive, stir in speed mixer, apply to stiffening pin and bond to motor housing with mandrel, then tape motor cable to purge lumen for correct positioning. Speed mixer, AU purge assembly, motor housing, Epotek 301, mandrel, dosing needle, tweezer. Bonded Motor housing to AU purge lumen. | incorrect position of assembly (scrap)
AU Motor. Purge flow rate test. Cured assembly, Purge rate test setup (monometer, pressure supply, tubings). Measured assembly (Airflow). | Pass failed units for airflow (high purge pressure)
AU Motor. Remove shim, check for easy movement of rotor, and cool down.. Final cured assembly with shim. AU Motor. | Damage impeller during removal of shim
AU Motor. Remove shim, check for easy movement of rotor, and cool down.. Final cured assembly with shim. AU Motor. | Shim not utlized leading to pump stop
AU Motor. Remove unit from test setup, clean with Isoproponal, and dry unit. Oven (time and temperature), Isoproponal, timer, AU Motor Assembly. Clean and Dried Assembly. | Contamination of components leading unsterile product
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | components not dry leading to sterility issue
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Damage causing mechanical interference with magnet
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Damage to inner cavity leading to moisture ingress into winding
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Fail acceptable components for patency
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Fail acceptable components for tightness
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Passed failed units for patency
AU Motor. Run tightness Test and patency test, then dry in oven. Tightness test setup, patency test setup, plug, cured assembly, pyrogen free water, compressed air. Tested assembly. | Passed failed units for tightness
AU Motor. Switch on plasma device, position motor housing in retainer plate and move retainer plate into plasma device. Plasma device, motor housing, retainer plate. assembly in plasma device. | Reduced mechanical properties of adhesive bond for purge lumen
AU Motor. Switch on plasma device, position motor housing in retainer plate and move retainer plate into plasma device. Plasma device, motor housing, retainer plate. assembly in plasma device. | Shrink tube delamination from coil
AU Motor-Pump Housing. Apply silicone around circumference of attached sensor. silicone, pump housing with sensor cable asssembly, dosing needle. pump housing with sensor cable subassembly with silicone on circumference. | moisture ingress due to silicone bond failure leading to sensor leakage current
AU Motor-Pump Housing. Apply thin layer of primer HPA41 into the sensor slot, cure primer. Primer HPA41, pump housing asembly, swabs, timer. prepped assembly. | moisture ingress due to silicone bond failure leading to sensor corrosion
AU Motor-Pump Housing. Attach klickadapter to repositioning unit . klickadapter, repositioning unit, repositioning unit modification fixture . Repo with adapter . | klickadapter damages repo
AU Motor-Pump Housing. Attach klickadapter to repositioning unit . klickadapter, repositioning unit, repositioning unit modification fixture . Repo with adapter . | klickadapter not fully attached
AU Motor-Pump Housing. Bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU Motor Pump Housing with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU with sensor and nitinol positioned and tacked. | adhesive not added - fiber breaks - failed SNR
AU Motor-Pump Housing. Bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU Motor Pump Housing with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU with sensor and nitinol positioned and tacked. | Fiber breaks - failed SNR
AU Motor-Pump Housing. Bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU Motor Pump Housing with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU with sensor and nitinol positioned and tacked. | Nitinol too far toward proximal end - unable to assemble
AU Motor-Pump Housing. Check assembly for Eccentricity. assembly with eptoek on weld seams, feeler gauge, prepping needle. acceptable impeller pump housing gap. | gap size too small is accepted
AU Motor-Pump Housing. Check assembly for Eccentricity. assembly with eptoek on weld seams, feeler gauge, prepping needle. acceptable impeller pump housing gap. | reject good components
AU Motor-Pump Housing. Check SNR of sensor. Sensor, SNR test box, 1-click cleaner. Checked Sensor. | Broken Sensor - fails SNR
AU Motor-Pump Housing. Check SNR of sensor. Sensor, SNR test box, 1-click cleaner. Checked Sensor. | Contamination of ferrule or SNR test box - false failure
AU Motor-Pump Housing. Clean flow straightener and slot of pump housing with Isopropyl. Flow straigtener, motor and pump housing assembly, isopropyl, swabs. Cleaned assemblies. | Damage flow straigthner/Pump housing internal surface
AU Motor-Pump Housing. Clean interior of pump housing dummy, hub, peek sleeve and impeller with Isopropanol. Motor subassembly, Isopropanol, swabs, pump housing dummy, timer, peek sleeve. cleaned components. | particulate contamination leading to scratches on impeller
AU Motor-Pump Housing. Clean interior of pump housing, collar transition and impeller with benzine. pump housing, Impeller, Isopropyl, swabs, centering sleeve. Cleaned assemblies. | Damage Impeller/Pump housing internal surface
AU Motor-Pump Housing. Clean interior of pump housing, hub, peek sleeve and impeller with Isoproponal. Motor subassembly, Isoproponal, swabs, pump housing, timer, peek sleeve. cleaned components. | damage/deform struts of pump housing
AU Motor-Pump Housing. Clean interior of pump housing, hub, peek sleeve and impeller with Isoproponal. Motor subassembly, Isoproponal, swabs, pump housing, timer, peek sleeve. cleaned components. | particulate contamination leading to scratches on pump housing surface
AU Motor-Pump Housing. Clean peek sleeve and impeller with Isopropanol. Motor subassembly, Isopropanol, swabs, timer, peek sleeve. cleaned components. | damage/deform struts of pump housing
AU Motor-Pump Housing. Clean peek sleeve and impeller with Isopropanol. Motor subassembly, Isopropanol, swabs, timer, peek sleeve. cleaned components. | particulate contamination leading to scratches on pump housing surface
AU Motor-Pump Housing. Clean pump housing in the area of the sensor slot with petroleum ether. pump housing assembly, petroleum ether, swabs. cleaned assembly. | moisture ingress due to silicone bond failure leading to sensor corrosion
AU Motor-Pump Housing. Clean weld seams with Isproponal and apply HPA41. Then cure in oven. completed welded assembly, HPA41, Ispoproponal, Oven (time and temperature), timer, swabs. cleaned and prepped assembly. | epotek delamination on weld seams leading to skiving of guidewire
AU Motor-Pump Housing. Cure adhesive. Oven (time and temperature), Assembly with Epotek, sensor cable hold down fixture. Cured assembly. | Epotek breaks off leading to exposed cables which damages/kinks during use/ Transition not smooth leading to kink/damage of cable during use
AU Motor-Pump Housing. Cure adhesive. Oven (time and temperature), Assembly with Epotek, sensor cable hold down fixture. Cured assembly. | Weak adhesive leading to sensor cable exposure leading to damage (i.e. sensor failure)
AU Motor-Pump Housing. Cure adhesive. Oven (time and temperature), Assembly with Epotek. Cured assembly. | Bond failure leading thrombus formation
AU Motor-Pump Housing. Feed the carrier lumen up the catheter toward the pump. AU Motor Pump Housing loaded on sensor loading fixture, sensor. AU loaded on sensor loading fixture with sensor loaded. | Pull lumen disengages from strain relief - damages fiber
AU Motor-Pump Housing. Fill void between pump housing and horshoe with 1:1 silicone. protector horshoe on pump assembly, silicone and cyclohexane (1:1), dispenser unit, silicone fill area. protector horseshoe on pump assembly with silicone filled locations. | asymmetrical fill of sensor cavity with silicone leading to sensor issues (i.e. dampening of signal)
AU Motor-Pump Housing. Fill void between pump housing and horshoe with 1:1 silicone. protector horshoe on pump assembly, silicone and cyclohexane (1:1), dispenser unit, silicone fill area. protector horseshoe on pump assembly with silicone filled locations. | fill sensor cavity with silicone leading to sensor issues (i.e. dampening of signal)
AU Motor-Pump Housing. Fill void between pump housing and horshoe with 1:1 silicone. protector horshoe on pump assembly, silicone and cyclohexane (1:1), dispenser unit, silicone fill area. protector horseshoe on pump assembly with silicone filled locations. | silicone contamination in the cannula bonding area
AU Motor-Pump Housing. Flow straightener into the slot of the pump housing. Flow straigtener, motor and pump housing assembly. Assembled units. | Damage Impeller/Pump housing internal surface
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing dummy, peek sleeve. welded hub to insert. | damage purge lumen
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing dummy, peek sleeve. welded hub to insert. | damage to motor cables
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing dummy, peek sleeve. welded hub to insert. | Discoloration/visual defects on surface (scrap)
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing dummy, peek sleeve. welded hub to insert. | Pump housing dummy welded to insert/hub
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing dummy, peek sleeve. welded hub to insert. | weak weld
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing, peek sleeve. welded hub and pump housing to insert. | damage purge lumen
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing, peek sleeve. welded hub and pump housing to insert. | damage to motor cables
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing, peek sleeve. welded hub and pump housing to insert. | Discoloration/visual defects on surface (scrap)
AU Motor-Pump Housing. Load subassembly into combi fixture and perform welding. combi fixture with RP nasol, laser welder with program, pre-assembled motor with pump housing, peek sleeve. welded hub and pump housing to insert. | weak weld
AU Motor-Pump Housing. Load the catheter with sensor cable. Motor pump housing assembly with sensor cable, pull lumen. pump catheter with sensor cable. | Sensor Cable damage
AU Motor-Pump Housing. Load the catheter with support ring, motor cable and purge lumen. Catheter, support ring, motor pump housing assembly, order of insertion of lumen and cable. pump catheter with motor cable and lumen. | Motor Cable damage
AU Motor-Pump Housing. Load the catheter with support ring, motor cable and purge lumen. Catheter, support ring, motor pump housing assembly, order of insertion of lumen and cable. pump catheter with motor cable and lumen. | Purge lumen damage/kink
AU Motor-Pump Housing. Manually weld the flanks of each strut to hub. hand rest, laster welder, pre-welded motor with pump housing assembly. completed welded assembly. | weak weld
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | Damage catheter
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | Damage purge lumen
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | Incorrect position of assemblies to be welded in relation to each other
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | Motor cable twisted/ Motor cable damage
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | weak weld
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly, Laser welding fixture. completed welded assembly. | Weld point over sensor - sensor damage - failed SNR
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly. completed welded assembly. | weak weld
AU Motor-Pump Housing. Manually weld the struts to hub. hand rest, laster welder, pre-welded motor with pump housing assembly. completed welded assembly. | Weld point over sensor - sensor damage - failed SNR
AU Motor-Pump Housing. Mix Epo-Tek 302 and cover pump housing weld seams. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | Damage guidewire (Skiving)
AU Motor-Pump Housing. Mix Epo-Tek 302 and cover pump housing weld seams. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | epotek delamination on weld seams leading to skiving of guidewire
AU Motor-Pump Housing. Mix Epo-Tek 302 and cover pump housing weld seams. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | Visually unacceptable component (scrap)
AU Motor-Pump Housing. Mix silicone and fill into syringe. Silicone and cyclohexane (1:1 ratio), beaker, de-air cartridge. mixed silicone. | sensor protrudes leading to scrap
AU Motor-Pump Housing. Mix silicone and fill into syringe. Silicone and cyclohexane (1:1 ratio), beaker, de-air cartridge. mixed silicone. | Silicone migrates into other areas leading to visual defects (scrap)
AU Motor-Pump Housing. Perform functional test of sensor resistances, offset, ESD criteria. Test setup, pump assembly. Tested pump assembly. | Failure acceptable units
AU Motor-Pump Housing. Perform functional test of sensor resistances, offset, ESD criteria. Test setup, pump assembly. Tested pump assembly. | Unacceptable sensor is accepted as good
AU Motor-Pump Housing. Place fiber in bag. AU Motor Pump Housing loaded on sensor loading fixture with sensor loaded. AU loaded on sensor loading fixture with sensor loaded and carrier lumen pulled - fiber in bag. | Damage to ferrule - failed SNR
AU Motor-Pump Housing. Place pump in Sensor Loading fixture. AU Motor Pump Housing, sensor loading fixture. AU loaded on sensor loading fixture. | AU positioned incorrectly - unable to assemble
AU Motor-Pump Housing. Place pump in Sensor Loading fixture. AU Motor Pump Housing, sensor loading fixture. AU loaded on sensor loading fixture. | Scratch assembly (scrapped)
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | contamination of cannula bonding area (bond failure)
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | Damage to sensor
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | moisture ingress due to silicone bond failure leading to sensor corrosion
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | position of the sensor incorrect leading to sensor failure
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | pressing down on sensor during curing leading to pre-stress failure (scrap)
AU Motor-Pump Housing. Place silicone on pump housing, apply loctite 4061 and then place sensor. Cure at room temperature.. Silicone, pump housing assembly, cyclohexane, sensor cable subassembly, dosing needle, fixation block, tweezers, loctite 4061. pump housing with sensor cable subassembly. | voids/bubbles underneath sensor membrane leading to incorrect sensor performance (dampening)
AU Motor-Pump Housing. Place unit into laser welder. Then manually laser weld horshoe to pump housing.. laser welder, pump assembly with horshoe, inert gas, hand rest fixture. pump assembly with horeshoe welded. | bond failure of horeshoe
AU Motor-Pump Housing. Place unit into laser welder. Then manually laser weld horshoe to pump housing.. laser welder, pump assembly with horshoe, inert gas, hand rest fixture. pump assembly with horeshoe welded. | Damage internal surface of pump housing
AU Motor-Pump Housing. Place unit into laser welder. Then manually laser weld horshoe to pump housing.. laser welder, pump assembly with horshoe, inert gas, hand rest fixture. pump assembly with horeshoe welded. | Damage sensor during welding process
AU Motor-Pump Housing. Position sensor protector (horseshoe) onto silicone border . protector horshoe, pump housing with sensor cable subassembly with silicone on circumference, position fixture. protector horseshoe on pump assembly. | Axial position on pump housing incorrect impact bonding area of cannula
AU Motor-Pump Housing. Position sensor protector (horseshoe) onto silicone border . protector horshoe, pump housing with sensor cable subassembly with silicone on circumference, position fixture. protector horseshoe on pump assembly. | Axial position on pump housing incorrect unprotected sensor
AU Motor-Pump Housing. Position sensor protector (horseshoe) onto silicone border . protector horshoe, pump housing with sensor cable subassembly with silicone on circumference, position fixture. protector horseshoe on pump assembly. | sensor collet/horshoe protruding to high
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Apply to pump housing ring transition and behind pump housing struts. Epo-Tek, pump housing with motor housing assembly, dosing needle, . Assembly with Epotek applied to ring and struts. | Epotek bond failure leading thrombus formation
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Form ramp for sensor cable transition. Fill and seal strut from ramp up to horseshoe. Epo-Tek, pump housing with motor housing assembly, dosing needle. Assebly with Epo-tek. | Epotek breaks off leading to exposed cables which damages/kinks during use
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Form ramp for sensor cable transition. Fill and seal strut from ramp up to horseshoe. Epo-Tek, pump housing with motor housing assembly, dosing needle. Assebly with Epo-tek. | Epotek breaks off leading to exposed cables which damages/kinks during use/ Transition not smooth leading to kink/damage of cable during use
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Form ramp for sensor cable transition. Fill and seal strut from ramp up to horseshoe. Epo-Tek, pump housing with motor housing assembly, dosing needle. Assebly with Epo-tek. | Transition not smooth leading to kink/damage of cable during use
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Form ramp for sensor cable transition. Fill and seal strut from ramp up to horseshoe. Epo-Tek, pump housing with motor housing assembly, dosing needle. Assemebly with Epo-tek. | Incorrect material on surface
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Lay sensor cable in epoxy bed on strut and encapsulate. Epo-Tek, pump housing with motor housing assembly, dosing needle, sensor cable hold down fixture. Assembly with sensor cable encapuslated with Epotek on strut. | inadequate amount of adhesive leading to sensor cable damage (i.e. sensor failure)
AU Motor-Pump Housing. Prepare Epo-Tek adhesive. Lay sensor cable in epoxy bed on strut and encapsulate. Epo-Tek, pump housing with motor housing assembly, dosing needle, sensor cable hold down fixture. Assembly with sensor cable encapuslated with Epotek on strut. | Weak adhesive leading to sensor cable exposure leading to damage (i.e. sensor failure)
AU Motor-Pump Housing. Prepare laser welder for manual welding. Laser welder, inert gas, hand rest. prepped laser welder. | unable to perform process as intended (fixturing incorrect)
AU Motor-Pump Housing. Pull carrier lumen until removed. AU Motor Pump Housing loaded on sensor loading fixture with sensor loaded. AU loaded on sensor loading fixture with sensor loaded and carrier lumen pulled. | Damage to ferrule - failed SNR
AU Motor-Pump Housing. Seal sensor cable in area under sensor protector with silicone mixture and cure at room temperature. pump assembly with horeshoe, silicone and cyclohexane (1:1), dosing needle and syringe. pump assembly with silicone on sensor cable. | compromised adhesion of delo leading to moisture ingress causing corrision of sensor (i.e. sensor failure)
AU Motor-Pump Housing. Seal sensor cable in area under sensor protector with silicone mixture and cure at room temperature. pump assembly with horeshoe, silicone and cyclohexane (1:1), dosing needle and syringe. pump assembly with silicone on sensor cable. | moisture ingress leading to corrosion causing sensor failure
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding . Welding machine, auxiliary table, pump housing assembly, argon gas, motor assembly, Welding Program. Equipment and assembly setup. | incorrect position leading to weak weld
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding . Welding machine, pump housing assembly, argon gas, motor assembly, Welding Program. Equipment and assembly setup. | corrosion on weld site
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding . Welding machine, pump housing assembly, argon gas, motor assembly, Welding Program. Equipment and assembly setup. | corrosion on weld site leading weak bond
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding, then weld units together . Welding machine, auxiliary table, pump housing assembly, argon gas, motor assembly, Welding Program. Welded assembly. | corrosion on weld site leading weak bond
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding, then weld units together . Welding machine, auxiliary table, pump housing assembly, argon gas, motor assembly, Welding Program. Welded assembly. | Weak weld strength
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding, then weld units together . Welding machine, pump assembly with flow straigtener, argon gas, Welding Program. Welded assembly. | increased friction leading to displacement and reduction of structural integrity
AU Motor-Pump Housing. Setup machine establishing parameters and position pump housing for manual welding, then weld units together . Welding machine, pump assembly with flow straigtener, argon gas, Welding Program. Welded assembly. | Misalignment leading to displacement and reduction of structural integrity
AU Motor-Pump Housing. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Assemble unit incorrect order (scrap)
AU Motor-Pump Housing. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage protection tube
AU Motor-Pump Housing. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage tip of repositioning unit
AU Motor-Pump Housing. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage/kink catheter
AU Motor-Pump Housing. Slide peek sleeve into motor housing and then slide assembly onto hub.. Motor subassembly, pump housing, peek sleeve. pre-assembled motor with pump housing. | damage pump housing struts
AU Motor-Pump Housing. Slide peek sleeve into pump housing and then slide assembly onto hub.. Motor subassembly, pump housing, peek sleeve. pre-assembled motor with pump housing. | damage pump housing struts
AU Motor-Pump Housing. Slide pump housing ring over hub transition. Clamp assembly into welding fixture.. pump housing, Impeller, centering sleeve, welding fixture. Assebled units in welding fixture. | Damage sensor cable during placement
AU Motor-Pump Housing. Slide pump housing ring over hub transition. Clamp assembly into welding fixture.. pump housing, Impeller, centering sleeve, welding fixture. Assebled units in welding fixture. | Damage sensor leads during placement
AU Motor-Pump Housing. Slide pump housing ring over hub transition. Clamp assembly into welding fixture.. pump housing, Impeller, centering sleeve, welding fixture. Assebled units in welding fixture. | pump housing ring damaged during placement
AU Motor-Pump Housing. Slide pump housing ring over hub transition. Clamp assembly into welding fixture.. pump housing, Impeller, centering sleeve, welding fixture. Assebled units in welding fixture. | pump housing ring damaged during placement leading to scrap
AU Motor-Pump Housing. Start laser welder. Open argon supply. Measure laser energy. Load laser program. Laser program is executed automatically.. Laser welder, laser program, measurement device (mulitmeter + recepticle), inert gas, excel sheet to document energy measurements, pulse energy laser program. prepped laser welder. | weak weld
AU Motor-Pump Housing. Start laser welder. Open inert gas supply. Measure laser energy. Load laser program. Laser program is executed automatically.. Laser welder, laser program, measurement device (mulitmeter + recepticle), inert gas, excel sheet to document energy measurements, pulse energy laser program. prepped laser welder. | weak weld
AU Motor-Pump Housing. Tack down sensor cable with Loctite 4061. pump assembly with silicone on sensor cable, Loctite 4061. pump assembly with tack down sensor cable. | Cable tacked down in catheter transition leading to cable damage by transition ring
AU Motor-Pump Housing. Tack down sensor cable with Loctite 4061. pump assembly with silicone on sensor cable, Loctite 4061. pump assembly with tack down sensor cable. | Unable to assemble device (scrap)
AU Motor-Pump Housing. Tack down sensor cable with Loctite 4061. pump assembly, Loctite 4061. pump assembly with tack down sensor cable. | Cable tacked down in catheter transition leading to cable damage by transition ring
AU Motor-Pump Housing. Tack down sensor cable with Loctite 4061. pump assembly, Loctite 4061. pump assembly with tack down sensor cable. | Unable to assemble device (scrap)
AU Motor-Pump Housing. Tack sensor and fiber using UV to lay flat into slot, cure. AU Motor Pump Housing with sensor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor positioned and tacked. | not fully seated in channel - unable to assemble pump housing
AU Motor-Pump Housing. Tack sensor and fiber using UV to lay flat into slot, cure. AU Motor Pump Housing with sensor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor positioned and tacked. | Sensor positioned incorrectly in axial direction - inaccurate sensor signals
AU Motor-Pump Housing. Tack sensor and fiber using UV to lay flat into slot, cure. AU Motor Pump Housing with sensor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor positioned and tacked. | under-cured UV - sensor moves - unable to assemble
AU Motor-Pump Housing. Unpackage sensor. Packaged sensor. Unpackaged Sensor. | Broken Sensor - fails SNR
AU Motor-Pump Housing. Unwind and prepare the sensor cable. Motor pump housing assembly with sensor cable, scapel. Motor pump housing assembly with sensor cable unwound. | Sensor cable damage
AU Motor-Pump Housing. use carrier lumen to locate sensor location. carrier lumen, pump assembly, optical sensor, optical sensor fixture. correctly positioned optical sensor. | incorrectly positioned optical sensor
AU Motor-Pump Housing. Visually and functionally check motor subassembly. Motor subassembly. acceptable motor subassembly. | Misalignment of hub to motor housing notch leading to unable to assemble sensor cable (scrap)
AU Motor-Pump Housing. Visually and functionally check motor subassembly. Motor subassembly. acceptable motor subassembly. | Stiff Impeller
AU Outlet Cage (10002172). Apply adhesion promoter HPA 41 on bonding surface of outlet cage and dry it in oven. HPA 41, outlet cage assembly, oven (time and temperature), holding fixture, swabs. outlet cage with promotor. | Weak bond
AU Outlet Cage (10002172). Apply adhesion promoter HPA 41 on bonding surface of outlet cage and dry it in oven. HPA 41, outlet cage assembly, oven (time and temperature), holding fixture, swabs. outlet cage with promotor. | Weak bond due to leftover promotor
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | incorrect orientation of cannula to pigtail (positioning issues)
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | OD too large, delivery issues
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outflow cage and slide cannula onto bonding surface . Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | Weak bond
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | delo migrates to inside of cannula
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | tilted cage leading to gap between cage and cannula
AU Outlet Cage (10002172). Apply Delo Duopox to bonding surface of outlet cage and slide cannula onto bonding surface. Delo Duopox, outlet cage assembly, dosing needle, cannula, curing fixture. Cannula and outlet cage bonded together. | Weak bond
AU Outlet Cage (10002172). Clean end of pigtail with petroleum ether . pigtail, petroleum ether, compressed air, swabs, wipes. Cleaned pigtail. | contamination due to particles left on pigtail leading to weak bond
AU Outlet Cage (10002172). Clean outflow cage and cannula at the bonding areas with acetone. Cannula, outflow cage assembly, acetone, swabs, beaker, holding fixture, time. cleaned components and assemblies. | Weak bond
AU Outlet Cage (10002172). Clean outflow cage and cannula at the bonding areas with acetone. Cannula, outflow cage assembly, acetone, swabs, beaker, holding fixture, time. cleaned components and assemblies. | Weak bond due to contamination
AU Outlet Cage (10002172). Cool down parts to room temperature. outlet cage with promotor, holding fixture. outlet cage cooled down to room temp. | Weak bond
AU Outlet Cage (10002172). Cure first at room temperature. Then cure adhesive in oven.. Oven (time and temperature), bonded together assembly, holding fixture, Room Temperature. AU Cannula assembly. | Weak bond
AU Outlet Cage (10002172). Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | Damage to pigtail leading thrombos formation
AU Outlet Cage (10002172). Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | Damage to pigtail leading to prolapsed pigtail
AU Outlet Cage (10002172). Cut pigtail to length with catheter trimmer . pigtail, catheter trimmer fixture. pigtail to length. | pigtail incorrect length leading to position issues
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Damage outflow cage (deformation/bending)
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Damage pigtail (prolapse)
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Failure to fully engage pigtail to cage (thrombus formation)
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Failure to fully engage pigtail to cage weak bond
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Over stress of the pigtail material leading to weak bond
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | too much epotek leading to occluded pigtail
AU Outlet Cage (10002172). Mix Epotek. Apply Epo-tek 301 to to pigtail and slide into outflow cage ball. Then cure in oven . Epo-tek 301, pigtail,outflow cage, dosing needle, beaker, Ratio of Epoxy to Hardener (4:1), speed mixer, primer, Oven (time and temperature). Pigtail bonded with Epoxy to outflow cage. | Weak bond
AU Outlet Cage (10002172). Perform Final visual inspection for cured unit and store assembly. finished subassembly (outflow cage + pigtail), blister + cover. Inspected and package subassembly. | Contamination of unit (biocompatibility)
AU Outlet Cage (10002172). Perform Final visual inspection for cured unit and store assembly. finished subassembly (outflow cage + pigtail), blister + cover. Inspected and package subassembly. | Damage outlet cage/pigtail
AU Outlet Cage (10002172). Plasma Activation of pigtail. pigtail, plasma activation chamber, holding fixture, oxygen and helium. Plasma activated pigtail. | weak bond
AU PH - Sensor Cable. Apply silicone around circumference of attached sensor. silicone, pump housing with mounted sensor, dosing needle, dispenser. pump housing with sensor mounted with silicone around. | moisture ingress leading to sensor failure
AU PH - Sensor Cable. Apply silicone around circumference of attached sensor. silicone, pump housing with mounted sensor, dosing needle, dispenser. pump housing with sensor mounted with silicone around. | Sensor leakage current
AU PH - Sensor Cable. Apply thin layer of primer MED160 into the sensor slot and around the inside surface of the hole, cure primer. pump housing, room temperature, MED160, swab. prepped assembly. | Poor adhesion of silicone leading to moisture ingress into sensor
AU PH - Sensor Cable. Clean pump housing in the area of the sensor slot with petroleum ether. pump housing, swab, petroleum ether. Cleaned assembly. | Poor adhesion of silicone leading to moisture ingress into sensor
AU PH - Sensor Cable. Mix silicone glue Nusil MED 1511 with cyclohexane. Fill in border of sensor with silicone mixture 2x. Seal sensor cable in area under sensor protector with silicone mixture and cure at room temperature. Apply ring of silicone around cavity and cure at room temperature. Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet, pump housing with horseshoe welded on. pump housing with sensor covered with silicone. | moisture ingress leading to sensor failure
AU PH - Sensor Cable. Mix silicone glue Nusil MED 1511 with cyclohexane. Fill in border of sensor with silicone mixture 2x. Seal sensor cable in area under sensor protector with silicone mixture and cure at room temperature. Apply ring of silicone around cavity and cure at room temperature. Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet, pump housing with horseshoe welded on. pump housing with sensor covered with silicone. | Protruding silicone over horseshoe leading to damage sensor during insertion of pump
AU PH - Sensor Cable. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture, then fix sensor cable with the superglue. Cure at room temperature. silicone, pump housing, sensor mounting fixture, Loctite 4061, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Bubbles trapped in sensor area leading to incorrect sensor single
AU PH - Sensor Cable. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture, then fix sensor cable with the superglue. Cure at room temperature. silicone, pump housing, sensor mounting fixture, Loctite 4061, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Damage sensor
AU PH - Sensor Cable. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture, then fix sensor cable with the superglue. Cure at room temperature. silicone, pump housing, sensor mounting fixture, Loctite 4061, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Incorrect position of sensor on pump housing leading to incorrect signal/sensor failure
AU PH - Sensor Cable. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture, then fix sensor cable with the superglue. Cure at room temperature. silicone, pump housing, sensor mounting fixture, Loctite 4061, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Poor adhesion of silicone leading to moisture ingress into sensor
AU PH - Sensor Cable. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | incorrect position of horshoe leading to cannula bonding issue
AU PH - Sensor Cable. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | incorrect position of horshoe leading to sensor damage
AU PH - Sensor Cable. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | Sensor protruding from device leading to insertion issues
AU PH - Sensor Cable. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | Sensor protruding from device/exposed leading to sensor damage
AU PH - Sensor Cable. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Damage inner surface of pump housing
AU PH - Sensor Cable. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Damage sensor
AU PH - Sensor Cable. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Unacceptable sensor accepted
AU PH - Sensor Cable. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isproponal.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isproponal. welded on horseshoe onto pump housing. | Damage to pump housing leading to thrombus foramtiono
AU PH - Sensor Cable. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isproponal.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isproponal. welded on horseshoe onto pump housing. | horseshoe bond failure
AU PH - Sensor Cable. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isproponal.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isproponal. welded on horseshoe onto pump housing. | horseshoe bond failure leading to sensor damage
AU PH - Sensor Cable. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isproponal.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isproponal. welded on horseshoe onto pump housing. | Sensor protector bond failure
AU Plug-Sensor. apply bottom half of plug in the rear cylindric area with Loctite 4062 and glue in the interface plug; then store in ESD bag or . Bottom handle half, cured asssembly, Loctite 4062, dosing unit, ESD bag or storage device for plug and sensor. Completed AU plug-sensor subassembly. | Board exposed to ESD
AU Plug-Sensor. apply bottom half of plug in the rear cylindric area with Loctite 4062 and glue in the interface plug; then store in ESD bag or . Bottom handle half, cured asssembly, Loctite 4062, dosing unit, ESD bag or storage device for plug and sensor. Completed AU plug-sensor subassembly. | poor bond leading to damaged PCB board
AU Plug-Sensor. bend the pins of the pressure sensor . pressure sensor, bend fixture. pressure sensor with pins bent. | damage sensor
AU Plug-Sensor. bend the pins of the pressure sensor . pressure sensor, bend fixture. pressure sensor with pins bent. | damage sensor leading to scrap
AU Plug-Sensor. Measure sensor resistance and all phases and then place Loctite 3321 on solder joints and holes then cure. Mulitmeter, Soldered together sensor and PCB board, measureing box, Bracket BG pump plug, Loctite 3321, UV lamp (time and distance). Tested and cured assembly. | Unit with poor resistance or phases passes
AU Plug-Sensor. place PCB and plug into production aid and insert legs from sensor into holes in PCB board and solder connection. Production aid, PCB board, pressure sensor with pins bent, solder material, soldering iron. Soldered together sensor and PCB board. | Damage of board during insertion due to mechanical stress (unable to program)
AU Plug-Sensor. place PCB and plug into production aid and insert legs from sensor into holes in PCB board and solder connection. Production aid, PCB board, pressure sensor with pins bent, solder material, soldering iron. Soldered together sensor and PCB board. | poor soldering quality
AU Pressure Reservoir Set for relocated sidearm. Anneal Filter. Oven (time and temperature), Filter. annealed filter. | reduced strength of filter neck causing crack/leak
AU Pressure Reservoir Set for relocated sidearm. Apply dichloromethane to filter, rsvr assy, assemble components. Filter, Rsvr Assy, Dicholoromethane. Rsvr assy connected to filter. | Bond failure
AU Pressure Reservoir Set for relocated sidearm. Apply dichloromethane to filter, rsvr assy, assemble components. Filter, Rsvr Assy, Dicholoromethane. Rsvr assy connected to filter. | damage filter due to contact with dichloromethane
AU Pressure Reservoir Set for relocated sidearm. Apply dichloromethane to filter, rsvr assy, assemble components. Filter, Rsvr Assy, Dicholoromethane. Rsvr assy connected to filter. | Filter facing the wrong way
AU Pressure Reservoir Set for relocated sidearm. Apply dichloromethane to filter, rsvr assy, assemble components. Filter, Rsvr Assy, Dicholoromethane. Rsvr assy connected to filter. | Leaking at Locking Nut
AU Pressure Reservoir Set for relocated sidearm. Apply Dymax 108-MSK to locking nut and cure. PSS, Dymax 108-MSK, UV Curing Equipment, Dispensing needle. PSS with glue at locking nut. | Bond failure
AU Pressure Reservoir Set for relocated sidearm. Apply Dymax 108-MSK to locking nut and cure. PSS, Dymax 108-MSK, UV Curing Equipment, Dispensing needle. PSS with glue at locking nut. | Leaking at Locking Nut
AU Pressure Reservoir Set for relocated sidearm. Assemble Protective Cap, Apply Loctite 3321 and cure adhesive. Rsvr Assy, Loctite 3321, Protective Cap (Optical, Non-Optical), UV surface lamp. Reservoir with bonded protective cap. | Misaligned PSS
AU Pressure Reservoir Set for relocated sidearm. Assemble Protective Cap, Apply Loctite 3321 and cure adhesive. Rsvr Assy, Loctite 3321, Protective Cap (Optical, Non-Optical), UV surface lamp. Reservoir with bonded protective cap. | Reduced Bond Strength
AU Pressure Reservoir Set for relocated sidearm. Assemble Protective Cap, Apply Loctite 3321 and cure adhesive. Rsvr Assy, Loctite 3321, Protective Cap (Optical, Non-Optical), UV surface lamp. Reservoir with bonded protective cap. | Reduced elasticity of Membrane, unable to deliver purge fluid
AU Pressure Reservoir Set for relocated sidearm. Check PSS alignment in fuctional gage. PSS, functional gage. aligned PSS. | Misfit in PSS retainer
AU Pressure Reservoir Set for relocated sidearm. Clean and prep all surfaces. pertoleum ether, swab, Rsvr Assy w/ no valve, protective cap (optical/non-optical), filter. clean and prepped components. | Contamination Leading to Scrap
AU Pressure Reservoir Set for relocated sidearm. Clean and prep locking nut bonding surface. Filter, Rsvr Assy, Petroleum Ether. clean and prepped components. | Contamination Leading to Scrap
AU Pressure Reservoir Set for relocated sidearm. Clean and prep protective cap bonding surface. Rsvr Assy, Petroleum Ether. clean and prepped components. | Contamination Leading to Scrap
AU Pressure Reservoir Set for relocated sidearm. Reinforce inflow and outflow neck with UV adhesive and cure. Dymax 108-MSK UV adeshive, 20-302 GA syringe tip, UV light source (time and distance), filter. Filter with UV enforced filter inflow and outflow neck. | Bond failure
AU Pressure Reservoir Set for relocated sidearm. Reinforce inflow and outflow neck with UV adhesive and cure. Dymax 108-MSK UV adeshive, 20-302 GA syringe tip, UV light source (time and distance), filter. Filter with UV enforced filter inflow and outflow neck. | Component damage causing crack/leak
AU Pressure Reservoir Set for relocated sidearm. Rest minimum of 12 hours to allow for evaporation. PSS. PSS with evaporated dicholormethane. | Leaking at Locking Nut
AU Pressure Reservoir Set for relocated sidearm. Seal seams of filter with UV glue. Dymax 108-MSK UV adeshive, 20-302 GA syringe tip, UV light source (time and distance), filter. Filter with seams sealed. | Filter Leaking
AU Pressure Reservoir Set for relocated sidearm. Seal seams of filter with UV glue. Dymax 108-MSK UV adeshive, 20-302 GA syringe tip, UV light source (time and distance), filter. Filter with seams sealed. | Visual Defect
AU Pressure Reservoir Set for relocated sidearm. USON Leak Test. Rsvr Assy, USON, Leak Test Fixture. Reservoir with no leaks. | Leaking at check valve
AU Protection Sleeve. Apply Miglyol 812 oil on silicone disk and then store assembly unit. assembled contamination protection unit, Miglyol 812. AU Protection Unit. | leakage of system leading bleeding issues
AU Protection Sleeve. Apply Miglyol 812 oil on silicone disk and then store assembly unit. assembled contamination protection unit, Miglyol 813. AU Protection Unit. | leakage of system leading bleeding issues
AU Protection Sleeve. Apply Miglyol 812 oil on silicone disk and then store assembly unit. assembled contamination protection unit, Miglyol 814. AU Protection Unit. | Damage contamination sleeve during storage
AU Protection Sleeve. Clean components with Isoproponal. Then assemble modified caps, new silicone ring and POM ring. Modified components, Isoproponal, silicone ring, POM ring, non-modified components of initial assembly. re-assembled cap. | incorrect order of assembly resulting in failure of compression ring leading contamination issues
AU Protection Sleeve. Clean components with Isoproponal. Then assemble modified caps, new silicone ring and POM ring. Modified components, Isoproponal, silicone ring, POM ring, non-modified components of initial assembly. re-assembled cap. | incorrect order of assembly resulting in failure of compression ring leading positioning issues
AU Protection Sleeve. Cut catheter protection sheath to length and assemble on spikes . catheter protection sheath, re-assembled cap, clamping ring, scapel, contamination sleeve cutting board. assembled contamination protection unit. | sheath length incorrect/positioned incorrectly resulting in failure of compression ring leading contamination issues
AU Protection Sleeve. Cut catheter protection sheath to length and assemble on spikes . catheter protection sheath, re-assembled cap, clamping ring, scapel, contamination sleeve cutting board. assembled contamination protection unit. | sheath length incorrect/positioned incorrectly resulting in failure of compression ring leading positioning issues
AU Protection Sleeve. Disassemble repositioning unit, scrap contamination protection sleeve and inner silicone ring, then modify caps in machine shop . contamination protection sleeve . dissassembled and modified unit. | Bubbles inside cap (damage to catheter/Unable to tighten tuohy)
AU Protection Sleeve. Disassemble repositioning unit, scrap contamination protection sleeve and inner silicone ring, then modify caps in machine shop . contamination protection sleeve . dissassembled and modified unit. | damage components during dissassembly/modified leading to contamination issues
AU Protection Sleeve. Disassemble repositioning unit, scrap contamination protection sleeve and inner silicone ring, then modify caps in machine shop . contamination protection sleeve . dissassembled and modified unit. | Damage components leading to unable to assembly units downstream
AU Protection Sleeve. Disassemble repositioning unit, scrap contamination protection sleeve and inner silicone ring, then modify caps in machine shop . contamination protection sleeve . dissassembled and modified unit. | Scrap incorrect components/modified incorrect component (unable to utilize)
AU Pump Cannula. Bond cannula assembly to pump sensor assembly using delo. Pump sensor assembly with adhesion promotor, delo, cannula assembly. Pump cannula assembly. | Cannula misaligned on assembly causing issues with placement
AU Pump Cannula. Bond cannula assembly to pump sensor assembly using delo. Pump sensor assembly with adhesion promotor, delo, cannula assembly. Pump cannula assembly. | Weak bond
AU Pump Cannula. Clean cannula bonding surfaces and prime with adhesion promoter. Pump sensor assembly with cured Epo-Tek on welds, petroleum ether, adhesion promotor HPA 41, oven (time and temperature). Pump sensor assembly with adhesion promotor. | Weak bond
AU Pump Cannula. Clean weld surface and prime with adhesion promoter. Pump sensor assembly, petroleum ether, Adhesion promotor HPA 41, Oven (Time and Temperature). Pump sensor assembly with adhesion promotor. | Epo-Tek delamination on weld transitions leading to corrosion at site
AU Pump Cannula. Cure assembly in oven and store on rack. Pump cannula assembly, oven (time and temperature). Pump cannula assembly. | OD of bond joint large leading to insertion issues
AU Pump Cannula. Cure assembly in oven and store on rack. Pump cannula assembly, oven (time and temperature). Pump cannula assembly. | Weak bond
AU Pump Cannula. Cure unit with Epo-Tek on welds. Oven (Time and Temperature), Pump sensor assembly with Epo-Tek on welds, rack. Pump sensor assembly with cured Epo-Tek on welds. | Epo-Tek delamination on weld transitions leading to corrosion at site
AU Pump Cannula. Prep Epo-Tek mixture and apply to weld seams. Epo-Tek 302-3M mixture, speed mixer, scale, beaker, hot plate, syringe, pump sensor assembly. Pump sensor assembly with Epo-Tek on welds. | Epo-Tek delamination on weld transitions leading to corrosion at site
AU Pump Cannula. Prep Epo-Tek mixture and apply to weld seams. Epo-Tek 302-3M mixture, speed mixer, scale, beaker, hot plate, syringe, pump sensor assembly. Pump sensor assembly with Epo-Tek on welds. | Voids in Epo-Tek on weld leading thrombus formation
AU Pump Catheter. Clean and prep single pin bonding site on motor housing. brush, molded motor housing assembly, microscope, scalpel, petroleum ether. Cleaned and prepped molded motor housing assembly. | bond failure (purge fluid leakage)
AU Pump Catheter. Clean and prep single pin bonding site on motor housing. brush, molded motor housing assembly, microscope, scalpel, petroleum ether. Cleaned and prepped molded motor housing assembly. | damage Epo-Tek transition area leading to unable to assemble
AU Pump Catheter. Clean and prep single pin bonding site on motor housing. brush, molded motor housing assembly, microscope, scalpel, petroleum ether. Cleaned and prepped molded motor housing assembly. | Particulates causing blockage of purge fluid
AU Pump Catheter. Cure Assembly and straighten lumens. Oven (time, temperature), drying board. Cured assembly with motor housing and single pin and straighten lumens. | bond failure (purge fluid leakage)
AU Pump Catheter. Glue single pin to motor housing bonding site. Single pin assembly, PTFE mandrel, petroleum ether, swab, delo,. molded motor housing with single pin with delo. | bond failure (purge fluid leakage)
AU Pump Catheter. Glue single pin to motor housing bonding site. Single pin assembly, PTFE mandrel, petroleum ether, swab, delo,. molded motor housing with single pin with delo. | single pin in incorrect position leading to purge issues
AU Pump Catheter. Place tipped catheter on end of assembly and align to yoke slot. Catheter assembly device, Cured assembly with motor housing and single pin and straighten lumens, prepped and cleaned tipped-catheter assembly, vacuum, pull lumen, . Pump Catheter assembly (non-welded). | Kink catheter
AU Pump Catheter. Place tipped catheter on end of assembly and align to yoke slot. Catheter assembly device, Cured assembly with motor housing and single pin and straighten lumens, prepped and cleaned tipped-catheter assembly, vacuum, pull lumen, . Pump Catheter assembly (non-welded). | kink of pull lumen causing unable to assemble sensor
AU Pump Catheter. Place tipped catheter on end of assembly and align to yoke slot. Catheter assembly device, Cured assembly with motor housing and single pin and straighten lumens, prepped and cleaned tipped-catheter assembly, vacuum, pull lumen, . Pump Catheter assembly (non-welded). | kink of purge lumen
AU Pump Catheter. Place tipped catheter on end of assembly and align to yoke slot. Catheter assembly device, Cured assembly with motor housing and single pin and straighten lumens, prepped and cleaned tipped-catheter assembly, vacuum, pull lumen, . Pump Catheter assembly (non-welded). | Motor wire shielding exposed
AU Pump Catheter. Place tipped catheter on end of assembly and align to yoke slot. Catheter assembly device, Cured assembly with motor housing and single pin and straighten lumens, prepped and cleaned tipped-catheter assembly, vacuum, pull lumen, . Pump Catheter assembly (non-welded). | slot not aligned leading to unable assemble unit
AU Pump Catheter. Prep and clean tipped catheter assembly. Petroleum ether, swab, tipped-catheter assembly. Prepped and cleaned tipped-catheter assembly. | component contamination leading to corrosion
AU Pump Housing. Apply silicone around circumference of attached sensor. silicone, pump housing with mounted sensor, dosing needle, dispenser. pump housing with sensor mounted with silicone around. | moisture ingress leading to sensor failure
AU Pump Housing. Apply silicone around circumference of attached sensor. silicone, pump housing with mounted sensor, dosing needle, dispenser. pump housing with sensor mounted with silicone around. | Sensor leakage current
AU Pump Housing. Apply thin layer of primer MED160 into the sensor slot and around the inside surface of the hole, cure primer. pump housing, room temperature, MED160, swab. prepped assembly. | Poor adhesion of silicone leading to moisture ingress into sensor
AU Pump Housing. Apply thin layer of primer Nusil MED160 into the sensor slot and around the inside surface of the hole, cure primer. Primer Nusil MED160, pump housing, oven control sheet, swabs, timer. pump housing with primer. | moisture ingress due to silicone bond failure leading to sensor corrosion
AU Pump Housing. Apply thin layer of primer Nusil MED160 into the sensor slot and around the inside surface of the hole, cure primer. Primer Nusil MED160, pump housing, oven control sheet, swabs, timer. pump housing with primer. | primer migrates into pump housing leading to biocompatibility issue
AU Pump Housing. Apply thin layer of primer Nusil MED160 into the sensor slot and around the inside surface of the hole, cure primer. Primer Nusil MED160, pump housing, oven control sheet, swabs, timer. pump housing with primer. | primer migrates to bonding area leading to weak bond
AU Pump Housing. Clean components. Petroleum ether, swabs, motor housing, endcap/tipped catheter. All components prepped for welding. | Low weld strength
AU Pump Housing. Clean pump housing in the area of the sensor slot with petroleum ether. pump housing, petroleum ether, swabs, timer. cleaned component. | moisture ingress due to silicone bond failure leading to sensor corrosion
AU Pump Housing. Clean pump housing in the area of the sensor slot with petroleum ether. pump housing, swab, petroleum ether. Cleaned assembly. | Poor adhesion of silicone leading to moisture ingress into sensor
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1511. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | hole not filled leading to scrap
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1512. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | too much silicone leading to pre-tension on sensor membrane leading to failure (scrap)
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1513. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | protruding silicone into pump housing
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1514. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | protruding silicone into pump housing
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1515. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | Bond failure leading to mositure ingress into sensor
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1515. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | Rough surface/voids/bubbles leading to thrombus formation
AU Pump Housing. Fill hole of pump housing with silicone Nusil MED1515. pump housing, silicone Nusil MED1511, dosing needle, cyclohexane, holding fixture. pump housing with silicone applied. | too much silicone leading to pre-tension on sensor membrane leading to failure (scrap)
AU Pump Housing. Key alignment between motor housing and pump housing fiber slots (manually). Combi fixture, pump housing, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder. Aligned pump housing to motor housing. | Damage components/unusable
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder, end cap. Components aligned for welding. | Low weld strength
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Damage to pump housing inner surface
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Damage to pump housing inner/outer surfaces
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Deformation of pump housing causing blockage
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Low weld strength
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Pump housing out of concentricity
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Components aligned for welding. | Tip/Blades of impeller are damaged
AU Pump Housing. Load components into machine (manual positioning). Combi fixture, pump housing, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder. Components aligned for welding. | Damage components/unusable
AU Pump Housing. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture. Cure at room temperature. silicone, pump housing, sensor mounting fixture, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Bubbles trapped in sensor area leading to incorrect sensor single
AU Pump Housing. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture. Cure at room temperature. silicone, pump housing, sensor mounting fixture, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Damage sensor
AU Pump Housing. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture. Cure at room temperature. silicone, pump housing, sensor mounting fixture, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Incorrect position of sensor on pump housing leading to incorrect signal/sensor failure
AU Pump Housing. Mount pump housing into sensor mounting fixture and seal hole from inside with clamp. Fill hole and bed of pump housing with silicone. Apply one drop of super-glue behind the silicone bed. Remove sensor / cable subassembly from storage and position into the silicone bed and fix with wire of sensor mounting fixture. Cure at room temperature. silicone, pump housing, sensor mounting fixture, AU sensor cable, dosing needle, dispenser, cartridge. pump housing with sensor mounted. | Poor adhesion of silicone leading to moisture ingress into sensor
AU Pump Housing. Nitinol Wire Shortening and Welding Rework. Combi fixture, Laser Welder pump housing, endcap/tipped catheter, nitinol wire, purge lumen, fiber lumen, motor cable, protective sleeve. . shortened nitinol wire with smooth welded end. | Damage lumens/motor cable
AU Pump Housing. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | incorrect position of horshoe leading to cannula bonding issue
AU Pump Housing. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | incorrect position of horshoe leading to sensor damage
AU Pump Housing. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | Sensor protruding from device leading to insertion issues
AU Pump Housing. Position sensor protector (horseshoe) onto silicone border . sensor protector (horseshoe), pump housing with sensor mounted, tweezer. positioned horseshoe on pump housing. | Sensor protruding from device/exposed leading to sensor damage
AU Pump Housing. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Damage inner surface of pump housing
AU Pump Housing. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Damage sensor
AU Pump Housing. Remove assembly from fixture. Then perform visual and electrical test. pump housing within fixture, electrical test setup, excel data collection sheet, storage fixture. Tested and accepted AU PH- Sensor cable assembly. | Unacceptable sensor accepted
AU Pump Housing. Run machine and weld endcap to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Welded endcap to motor housing. | Incomplete weld (ex. sink holes/distortions)
AU Pump Housing. Run machine and weld endcap to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Welded endcap to motor housing. | Incorrect alignment of endcap
AU Pump Housing. Run machine and weld endcap to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Welded endcap to motor housing. | Low weld strength
AU Pump Housing. Run machine and weld endcap to motor housing. Combi fixture, pump housing, motor housing, endcap/tipped catheter, peek alignment sleeve , Laser welder. Welded endcap to motor housing. | Low weld strength
AU Pump Housing. Run machine and weld endcap to motor housing. Combi fixture, pump housing, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder. Welded endcap to motor housing. | Damage components/unusable
AU Pump Housing. Run machine and weld outflow cage to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder, outflow cage. Welded outflow cage to motor housing as well as motor housing to endcap. | hemolysis
AU Pump Housing. Run machine and weld outflow cage to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder, outflow cage. Welded outflow cage to motor housing as well as motor housing to endcap. | Incomplete weld (ex. sink holes/distortions)
AU Pump Housing. Run machine and weld outflow cage to motor housing. Combi fixture, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder, outflow cage. Welded outflow cage to motor housing as well as motor housing to endcap. | Low weld strength
AU Pump Housing. Run machine and weld pump housing. Combi fixture, pump housing, motor housing, endcap/tipped catheter, peek alignment sleeve, Laser welder. Welded pump housing to motor housing as well as pump housing to endcap. | Damage components/unusable
AU Pump Housing. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isopropanol.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isopropanol. welded on horseshoe onto pump housing. | Damage to pump housing leading to thrombus foramtiono
AU Pump Housing. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isopropanol.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isopropanol. welded on horseshoe onto pump housing. | Sensor protector bond failure
AU Pump Housing. Set up laser welder. Turn on inert gas. Position pump housing and protector sub assembly in the laser welder, set 3 welding points per side. Then clean with Isopropanol.. Laser welder, inert gas, pump housing with positioned horseshoe, hand welding table, laser program (time, spot diameter, power), swab, Isopropanol. welded on horseshoe onto pump housing. | Sensor protector bond failure leading to sensor damage
AU Pump Housing. Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program. Laser Welder operational and tested for power. | Damage components/unusable
AU Pump Housing. Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program. Laser Welder operational and tested for power. | Low weld strength
AU Pump Housing. Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program. Laser Welder operational and tested for power. | Machine faults out
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Deformation of pump housing causing blockage
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | hemolysis
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Impeller unable to rotate/rotation impeded
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Particulate generation/delamination of Epo-Tek in body
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Sharp edge exposed leading to red lumen breakage
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Tip of impeller is damaged
AU Pump Housing. Unload pump from machine and put Epo-Tek around struts. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap/catheter-tip, Epo-Tek , petroleum ether, swabs, syringe tip, automated dispensers. Welded outflow cage to motor housing as well as motor housing to endcap with Epo-Tek on struts. | Tip of impeller is damaged/pump housing damaged
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | Insufficient adhesive at catheter attachment - fails tightness test
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | Insufficient adhesive at catheter attachment - tensile strength comprimised
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | Insufficient adhesive at catheter attachment - tensile strength compromised
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | Insufficient adhesive not encapsulating fiber - fiber breaks in field
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | Insufficient adhesive not encapsulating UV - visor may detach
AU Pump Sensor. Apply adhesive to sensor visor, channel, and catheter transition - encapsulating UV glue. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked, adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. | too much adhesive - sensor damaged - failed SNR
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | Leak path into catheter jacket
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | pressure port not backfilled with enough delo
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | Rough delo surface/sharp edges
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | Sensor not functioning
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | too little delo applied (bond failure)
AU Pump Sensor. Apply final delo coat. Delo Catheter assembly Microscope Loctite dispenser Pressure Suck-back setting. Final delo coat applied. | too much delo applied, assembly exceeds OD spec
AU Pump Sensor. Assemble Repo unit onto assembly. Repo Unit components, Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths . Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit. | During assembly Repo. Unit Damaged
AU Pump Sensor. Assemble Repo unit unto assembly. Repo Unit components, Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths . Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit. | Unable to reposition pump or difficulty repositioning unit
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispenser, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | adhesive not added - fiber breaks - failed SNR
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispenser, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | Fiber breaks - failed SNR
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispenser, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | Nitinol too far toward proximal end - unable to assemble
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | adhesive not added - fiber breaks - failed SNR
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | Fiber breaks - failed SNR
AU Pump Sensor. bond nitinol tube into yoke at catheter transition and tack and cure with UV. AU pump catheter motor with sensor, Adhesive, Adhesive dispensor, adhesive, dosing needle, mixing tip, swabs, UV point source. AU pump catheter motor with sensor and nitinol positioned and tacked. | Nitinol too far toward proximal end - unable to assemble
AU Pump Sensor. Check SNR of sensor. Sensor, SNR test box, 1-click cleaner. Checked Sensor. | Broken Sensor - fails SNR
AU Pump Sensor. Check SNR of sensor. Sensor, SNR test box, 1-click cleaner. Checked Sensor. | Contamination of ferrule or SNR test box - false failure
AU Pump Sensor. Clean bonding areas. AU pump catheter motor, petroleum ether, swabs. AU pump catheter motor with cleaned bonding areas. | Insufficient bond strength
AU Pump Sensor. Clean bonding region. AU Pump Sensor with glued catheter and trimmed and tinned motor cable, wipes, petroleum ether. AU Pump Sensor with glued catheter and trimmed and tinned motor cable - cleaned bonding region. | Insufficient bond strength between kink protector and catheter
AU Pump Sensor. Clean PSS plug outlet, then bond on PSS plug outlet onto assembly. Petroleum ether, 23GA dosing needle, PSS plug outlet assembly, THF, IR unit, mandrel, Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, kink protector, and bonded PSS plug outlet assembly . | Unable to assemble leading to scrap
AU Pump Sensor. Clean PSS plug outlet, then bond on PSS plug outlet onto assembly. Petroleum ether, 23GA dosing needle, PSS plug outlet assembly, THF, IR unit, mandrel, Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, kink protector, and bonded PSS plug outlet assembly . | Weak bond to PSS plug outlet/occluded PSS plug outlet
AU Pump Sensor. Cure adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured with adhesive bead, oven (time and temp). AU Pump Sensor. | insufficient curing - lower bond strength
AU Pump Sensor. Cure final delo application. Oven Oven Racks Temperature Time Pump Sensor Assembly. Pump sensor assembly with, cured final delo coat. | delo not cured (Partial bond failure during use)
AU Pump Sensor. Cure final delo application. Oven Oven Racks Temperature Time Pump Sensor Assembly. Pump sensor assembly with, cured final delo coat. | OD of sensor out of specification
AU Pump Sensor. Feed sensor into carrier lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. Sensor feeds into catheter, stop feeding at correct distance. | fiber kinks
AU Pump Sensor. Feed sensor into carrier lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. Sensor feeds into catheter, stop feeding at correct distance. | Mixup Sensor Serial number
AU Pump Sensor. Feed sensor into carrier lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. Sensor feeds into catheter, stop feeding at correct distance. | sensor damaged during handling at work station
AU Pump Sensor. Feed sensor into carrier lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. sensor protruding from pressure port, pull lumen removed . | Damage sensor head during assembly
AU Pump Sensor. Feed sensor into pull lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. Sensor feeds into catheter, stop feeding at correct distance. | fiber breaks during removal from Opsens package
AU Pump Sensor. Feed sensor into pull lumen to specified distance. Sensor and Optical Fiber sub assembly Pump housing subassembly, Alignment/ Positioning Fixture. Sensor feeds into catheter, stop feeding at correct distance. | sensor stuck in pull lumen
AU Pump Sensor. Feed the carrier lumen up the catheter toward the pump. AU pump catheter motor loaded on sensor loading fixture, sensor. AU pump catheter motor loaded on sensor loading fixture with sensor loaded. | Pull lumen disengages from strain relief - damages fiber
AU Pump Sensor. Fill voids in adhesive on channel, cover welds with adhesive, and apply a smooth transition of adhesive around the catheter transition. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured, adhesive, dispensing tips. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured with adhesive bead. | rough transitions
AU Pump Sensor. Glue kink protector to catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable, Adhesive 4062. AU Pump Sensor with glued catheter and glued kink protector. | Glued in incorrect location - pump orientation incorrect
AU Pump Sensor. Glue kink protector to catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable, Adhesive 4062. AU Pump Sensor with glued catheter and glued kink protector. | Glued in incorrect location - unable to assemble
AU Pump Sensor. Glue kink protector to catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable, Adhesive 4062. AU Pump Sensor with glued catheter and glued kink protector. | Insufficient glue - low bond strength
AU Pump Sensor. Laser weld tacking of coil. AU Pump Catheter Motor with Sensor UV tacked, laser, manual weld supporting fixture. AU Pump Catheter Motor with Sensor UV tacked with Coil laser welded on yoke. | Damage purge lumen - purge leak
AU Pump Sensor. Laser weld tacking of coil. AU Pump Catheter Motor with Sensor UV tacked, laser, manual weld supporting fixture. AU Pump Catheter Motor with Sensor UV tacked with Coil laser welded on yoke. | incomplete laser weld - unable to assemble catheter
AU Pump Sensor. Measure sensor SNR. Sensor properly positioned on pump housing, secured by fixture, SNR Fixture. Measured Sensor improperly positioned on pump housing, secured by fixture. | Scrap product due to low SNR failure
AU Pump Sensor. Measure sensor SNR. Sensor properly positioned on pump housing, secured by fixture, SNR Fixture. Measured Sensor improperly positioned on pump housing, secured by fixture. | Sensor failure/low sensor signal
AU Pump Sensor. Mount coil on to yoke area. AU Pump Catheter Motor with Sensor UV tacked. AU Pump Catheter Motor with Sensor UV tacked with Coil mounted on yoke. | Coil deformed - unable to assemble
AU Pump Sensor. Mount coil on to yoke area. AU Pump Catheter Motor with Sensor UV tacked. AU Pump Catheter Motor with Sensor UV tacked with Coil mounted on yoke. | Damage to fiber - failed SNR
AU Pump Sensor. place fiber in bag with purge sidearm. AU pump catheter motor loaded on sensor loading fixture with sensor loaded and carrier lumen pulled, plastic bag. AU pump catheter motor loaded on sensor loading fixture with sensor loaded and carrier lumen pulled - fiber in bag. | Damage to ferrule - failed SNR
AU Pump Sensor. place fiber in bag. AU pump catheter motor loaded on sensor loading fixture with sensor loaded and carrier lumen pulled, plastic bag. AU pump catheter motor loaded on sensor loading fixture with sensor loaded and carrier lumen pulled - fiber in bag. | Damage to ferrule - failed SNR
AU Pump Sensor. Place pump in Sensor Loading fixture and allow catheter to sit flush with the edge of the window. AU pump catheter motor, sensor loading fixture. AU pump catheter motor loaded on sensor loading fixture. | Catheter slid too far - unable to assemble
AU Pump Sensor. Place pump in Sensor Loading fixture and allow catheter to sit flush with the edge of the window. AU pump catheter motor, sensor loading fixture. AU pump catheter motor loaded on sensor loading fixture. | PH not pushed to end stop - unable to assemble
AU Pump Sensor. Place pump in Sensor Loading fixture and allow catheter to sit flush with the edge of the window. AU pump catheter motor, sensor loading fixture. AU pump catheter motor loaded on sensor loading fixture. | Scratch assembly (scrapped)
AU Pump Sensor. Place Ziploc around sensor connector, place pump on rack. Pump Housing Assembly with sensor, Ziploc, Rack. Ziploc around connector, pump on rack. | Damaged connector/fiber
AU Pump Sensor. Place Ziploc or Blister around sensor connector, place pump on rack. Pump Housing Assembly with sensor, Ziploc or Blister, Rack. Ziploc around connector or connector in Blister, pump on rack. | Damaged connector/fiber
AU Pump Sensor. Position sensor onto pump housing, clamp in place with holding fixture. Pump housing, Sensor, microscope. Sensor properly positioned on pump housing, secured by fixture. | Sensor damaged during positioning on pump housing
AU Pump Sensor. Prep Epo-Tek mixture and apply to weld seams. Epo-Tek 302-3M mixture, speed mixer, scale, beaker, hot plate, syringe, pump sensor assembly. Pump sensor assembly with Epo-Tek on welds. | Voids in Epo-Tek on weld leading scrap/unacceptable pump
AU Pump Sensor. Prepare (cut) purge lumen and motor cables. AU Pump Sensor with glued catheter, trimming fixture, side cutters, scalpel. AU Pump Sensor with glued catheter and trimmed. | Unable to assemble
AU Pump Sensor. Pull carrier lumen until removed. AU pump catheter motor loaded on sensor loading fixture with sensor loaded. AU pump catheter motor loaded on sensor loading fixture with sensor loaded and carrier lumen pulled. | Damage to ferrule - failed SNR
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Bending of sensor visor
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Deep Scratch on Sensor Visor, where blood is contacting
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Indentation on sensor visor with no protrusion & sharp edges
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Indentation, Protrusion & sharp edge on sensor visor where blood is contacting
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Indentation, Protrusion & sharp edge on sensor visor, where blood is contacting
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Sensor face protruding
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . Sensor fiber loading. | Unable to seat visor on sensor bump
AU Pump Sensor. Pull sensor through pressure port, remove carrier lumen. Sensor and Optical Fiber sub assembly , Pump Housing subassembly, Sensor Loading Fixture . sensor protruding from pressure port, pull lumen removed . | Damage sensor head during assembly
AU Pump Sensor. remove excess adhesive, and cure. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place, swab, oven (time and temp), pump magazine. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured. | cosmetic defects
AU Pump Sensor. remove excess adhesive, and cure. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place, swab, oven (time and temp), pump magazine. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured. | insufficient cure - bond strength comprimised
AU Pump Sensor. remove excess adhesive, and cure. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place, swab, oven (time and temp), pump magazine. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place - cured. | insufficient cure - bond strength compromised
AU Pump Sensor. Remove pump from rack, load into positioning fixture . Pump assembly , Sensor Loading Fixture . Catheter loaded into positioning fixture. | Over constraining catheter during assembly
AU Pump Sensor. Remove pump from rack, load into positioning fixture . Pump assembly , Sensor Loading Fixture . Catheter loaded into positioning fixture. | Sensor damage during handling
AU Pump Sensor. Slide catheter over catheter transition and align marking with sensor. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place. | Catheter not aligned - difficulty placing pump
AU Pump Sensor. Slide catheter over catheter transition and align marking with sensor. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place. | Catheter not fully pushed on - insufficient bonding area - tensile strength comprimised
AU Pump Sensor. Slide catheter over catheter transition and align marking with sensor. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked and covered with adhesive, with catheter in place. | Catheter not fully pushed on - insufficient bonding area - tensile strength compromised
AU Pump Sensor. Strip and tin motor cables. AU Pump Sensor with glued catheter and trimmed, hot tweezers, soldering iron, solder, protective catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable. | Damage purge lumen - purge leak
AU Pump Sensor. Strip and tin motor cables. AU Pump Sensor with glued catheter and trimmed, hot tweezers, soldering iron, solder, protective catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable. | Damage to purge lumen - unable to assemble side arm kink protector
AU Pump Sensor. Strip and tin motor cables. AU Pump Sensor with glued catheter and trimmed, hot tweezers, soldering iron, solder, protective catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable. | Damaged motor cables
AU Pump Sensor. Strip and tin motor cables. AU Pump Sensor with glued catheter and trimmed, hot tweezers, soldering iron, solder, protective catheter. AU Pump Sensor with glued catheter and trimmed and tinned motor cable. | Unable to assemble
AU Pump Sensor. Strip wires and glue on grey kink protector onto catheter end. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit, grey kink protector, Loctite, wire stripping tool. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. | Damage of wire leads
AU Pump Sensor. Strip wires and glue on grey kink protector onto catheter end. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit, grey kink protector, Loctite, wire stripping tool. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. | Damage purge lumen causing leak
AU Pump Sensor. Strip wires and glue on grey kink protector onto catheter end. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit, grey kink protector, Loctite, wire stripping tool. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. | Unable to assemble unit leading to scrap
AU Pump Sensor. Strip wires and glue on grey kink protector onto catheter end. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths and repo unit, grey kink protector, Loctite, wire stripping tool. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths, repo unit, and glued on kink protector. | Weak bond of kink protector to catheter
AU Pump Sensor. tack nitinol to fiber transition into catheter transition area, cure UV. AU pump catheter motor with sensor and nitinol and sensor visor positioned and tacked, Adhesive, UV point source. AU pump catheter motor with sensor and nitinol to fiber transition and sensor visor positioned and tacked. | damage to fiber - failed SNR
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump catheter motor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor and nitinol and sensor visor positioned and tacked. | not fully seated in channel - unable to cover with adhesive
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump catheter motor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor and nitinol and sensor visor positioned and tacked. | Positioning incorrect- sensor rolled
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump catheter motor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser. AU pump catheter motor with sensor and nitinol and sensor visor positioned and tacked. | under-cured UV - sensor moves - unable to assemble
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump sensor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser, UV LED. AU pump sensor with sensor and nitinol and sensor visor positioned and tacked. | not fully seated in channel - unable to cover with adhesive
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump sensor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser, UV LED. AU pump sensor with sensor and nitinol and sensor visor positioned and tacked. | Positioning incorrect- sensor rolled
AU Pump Sensor. tack visor and fiber using UV to lay flat into slot, cure. AU pump sensor with sensor and nitinol and sensor visor positioned, Adhesive, Adhesive dispenser, UV LED. AU pump sensor with sensor and nitinol and sensor visor positioned and tacked. | under-cured UV - sensor moves - unable to assemble
AU Pump Sensor. Trim Motor Cable, Purge Lumen, Optical Lumen. Trim Fixture, Cutting tools, Pump Housing Assembly. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths. | Damage lumens
AU Pump Sensor. Trim Motor Cable, Purge Lumen, Optical Lumen. Trim Fixture, Cutting tools, Pump Housing Assembly. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths. | Damage Wires
AU Pump Sensor. Trim Motor Cable, Purge Lumen, Optical Lumen. Trim Fixture, Cutting tools, Pump Housing Assembly. Pump Housing Assembly with cables, purge lumen, optical lumen at correct lengths. | Lumens cut to incorrect lengths leading to unable to assemble
AU Pump Sensor. Unpackage sensor. Packaged sensor. Unpackaged Sensor. | Broken Sensor - fails SNR
AU Pump/Plug-Sensor. apply bottom half of plug in the rear cylindric area with Loctite 4062 and glue in the interface plug; then store in ESD bag or . Bottom handle half, cured asssembly, Loctite 4062, dosing unit, ESD bag or storage device for plug and sensor. Completed AU plug-sensor subassembly. | Board exposed to ESD
AU Pump/Plug-Sensor. apply bottom half of plug in the rear cylindric area with Loctite 4062 and glue in the interface plug; then store in ESD bag or . Bottom handle half, cured asssembly, Loctite 4062, dosing unit, ESD bag or storage device for plug and sensor. Completed AU plug-sensor subassembly. | poor bond leading to damaged PCB board
AU Pump/Plug-Sensor. bend the pins of the pressure sensor . pressure sensor, bend fixture. pressure sensor with pins bent. | damage sensor
AU Pump/Plug-Sensor. bend the pins of the pressure sensor . pressure sensor, bend fixture. pressure sensor with pins bent. | damage sensor leading to scrap
AU Pump/Plug-Sensor. Measure sensor resistance and all phases and then place Loctite 3321 on solder joints and holes then cure. Mulitmeter, Soldered together sensor and PCB board, measureing box, Bracket BG pump plug, Loctite 3321, UV lamp (time and distance). Tested and cured assembly. | Unit with poor resistance or phases passes
AU Pump/Plug-Sensor. place PCB and plug into production aid and insert legs from sensor into holes in PCB board and solder connection. Production aid, PCB board, pressure sensor with pins bent, solder material, soldering iron. Soldered together sensor and PCB board. | Damage of board during insertion due to mechanical stress (unable to program)
AU Pump/Plug-Sensor. place PCB and plug into production aid and insert legs from sensor into holes in PCB board and solder connection. Production aid, PCB board, pressure sensor with pins bent, solder material, soldering iron. Soldered together sensor and PCB board. | poor soldering quality
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven, then cool down parts. HPA 41, pump plug assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven, then cool down parts. HPA 41, pump sensor assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump plug assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump plug assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump plug assembly, Oven (time and Temperature). assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump sensor assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump sensor assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump sensor assembly, Oven (time and Temperature). assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump sesnor assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. Apply adhesion promoter HPA 41 on bonding surface and dry it in oven. HPA 41, pump sesnor assembly, Oven (time and Temperature), cannula. assembly with promoter. | Weak bond
AU Pump-Cannula. apply Adhesive Epoxy Duopox 1895 on rough area of pump housing and slide the cannula on it. Adhesive Epoxy Duopox 1895, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | Bonding failure at cannula to outflow cage bond
AU Pump-Cannula. apply Adhesive Epoxy Duopox 1895 on rough area of pump housing and slide the cannula on it. Adhesive Epoxy Duopox 1895, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | gap between cannula and pump housing
AU Pump-Cannula. apply Adhesive Epoxy Duopox 1895 on rough area of pump housing and slide the cannula on it. Adhesive Epoxy Duopox 1895, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | incorrect orientation of cannula to pump assembly leading to insertion issues
AU Pump-Cannula. apply Adhesive Epoxy Duopox 1895 on rough area of pump housing and slide the cannula on it. Adhesive Epoxy Duopox 1895, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | incorrect orientation of cannula to pump assembly leading to unable to package
AU Pump-Cannula. Apply adhesive on bonding region of the pump housing and slide on the cannula. Pump sensor assembly, cannula, adhesive. glued on assembly. | Cannula not aligned
AU Pump-Cannula. Apply adhesive on bonding region of the pump housing and slide on the cannula. Pump sensor assembly, cannula, adhesive. glued on assembly. | Step at transition
AU Pump-Cannula. Apply adhesive on bonding region of the pump housing and slide on the cannula. Pump sensor assembly, cannula, adhesive. glued on assembly. | Weak bond
AU Pump-Cannula. Apply adhesive ring to pump housing and cannula transition. Then cure in oven.. Oven (time and temperature), rack, and glued assembly. assembly with adhesive ring. | Unsmooth transition/steps in transition (thrombus formation)
AU Pump-Cannula. Apply Delo Duopox 1895 on bonding region of the pump housing and slide on the cannula. Pump plug assembly, cannula, delo. glued on assembly. | Weak bond
AU Pump-Cannula. apply Delo Duopox 1895 on rough area of pump housing and slide the cannula on it. delo, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | Bonding failure at cannula to outflow cage bond
AU Pump-Cannula. apply Delo Duopox 1895 on rough area of pump housing and slide the cannula on it. delo, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | gap between cannula and pump housing
AU Pump-Cannula. apply Delo Duopox 1895 on rough area of pump housing and slide the cannula on it. delo, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | incorrect orientation of cannula to pump assembly leading to insertion issues
AU Pump-Cannula. apply Delo Duopox 1895 on rough area of pump housing and slide the cannula on it. delo, cleaned and prepped pump and cannula assemblies. glued together cannula-pigtail-inflow cage and pump assemblies. | incorrect orientation of cannula to pump assembly leading to unable to package
AU Pump-Cannula. Apply Delo Duopox on bonding region of the pump housing and slide on the cannula. Pump plug assembly, cannula, delo. glued on assembly. | Weak bond
AU Pump-Cannula. Apply delo ring to pump housing and cannula transition. Then cure in oven.. Oven (time and temperature), rack, and glued assembly. assembly with delo ring. | OD too large, delivery issues
AU Pump-Cannula. Apply delo ring to pump housing and cannula transition. Then cure in oven.. Oven (time and temperature), rack, and glued assembly. assembly with delo ring. | Unsmooth transition/steps in transition (thrombus formation)
AU Pump-Cannula. clean bonding surface cannula and housing with petroluem ether; then apply adhesion promotor HPA 41 on bonding surface and dry it. Petroleum ether, Cannula-pigtail-inflow cage assembly, swab, HPA 41, petroleum ether, Oven (time and temperature). cleaned and prepped assemblies. | Bonding failure at cannula to outflow cage bond
AU Pump-Cannula. clean bonding surface cannula and housing with petroluem ether; then apply adhesion promotor HPA 41 on bonding surface and dry it. Petroleum ether, Cannula-pigtail-inflow cage assembly, swab, HPA 41, petroleum ether, Oven (time and temperature). cleaned and prepped assemblies. | promoter on the impeller leading to rough surface on component
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with aceton and dry with compressed air. Pump plug assembly, cannula, aceton, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with aceton and dry with compressed air. Pump plug assembly, cannula, aceton, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with aceton and let dry. Pump sensor assembly, cannula, aceton, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with aceton and let dry. Pump sensor assembly, cannula, aceton, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with aceton and let dry. Pump sesnor assembly, cannula, aceton, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with acetone and let dry. Pump sensor assembly, cannula, acetone, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with acetone and let dry. Pump sensor assembly, cannula, acetone, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with petroleum ether and dry with compressed air. Pump plug assembly, cannula, petroleum ether, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean bonding surface of cannula and pump housing with petroleum ether and dry with compressed air. Pump plug assembly, cannula, petroleum ether, swabs. Cleaned assembly. | Weak bond
AU Pump-Cannula. Clean end of endcap with petroleum ether, fill endcap with Epo-Tek and apply to endcap-catheter transition then cure. Repeat 1x.. Petroleum ether, swabs, Epotek 302-3M, dispenser, racks, oven (time and temperature), pump assembly. pump assembly with partially filled endcap and catheter tranisition with epotek. | epo-tek falling off at transitions (thrombus formation)
AU Pump-Cannula. Clean end of endcap with petroleum ether, fill endcap with Epo-Tek and apply to endcap-catheter transition then cure. Repeat 1x.. Petroleum ether, swabs, Epotek 302-3M, dispenser, racks, oven (time and temperature), pump assembly. pump assembly with partially filled endcap and catheter tranisition with epotek. | bubbles on the surface of the epo-tek
AU Pump-Cannula. Clean end of endcap with petroleum ether, fill endcap with Epo-Tek and apply to endcap-catheter transition then cure. Repeat 1x.. Petroleum ether, swabs, Epotek 302-3M, dispenser, racks, oven (time and temperature), pump assembly. pump assembly with partially filled endcap and catheter tranisition with epotek. | uncured epo-tek leading to it falling off at transitions (thrombus formation)
AU Pump-Cannula. Clean end of endcap with petroleum ether, fill endcap with Epo-Tek and apply to endcap-catheter transition then cure. Repeat 1x.. Petroleum ether, swabs, Epotek 302-3M, dispenser, racks, oven (time and temperature), pump assembly. pump assembly with partially filled endcap and catheter tranisition with epotek. | white spots in epotek (component separation) leading to rough surfaces
AU Pump-Cannula. Clean with petroleum ether and apply epotek to sleeve bearing. petroleum ether, epotek 302-3M, dispenser, swabs, oven (time and temperature). pump assembly with epo-tek 302-3M on sleeve bearing. | delamination of epot-tek a transition
AU Pump-Cannula. Clean with petroleum ether and apply epotek to sleeve bearing. petroleum ether, epotek 302-3M, dispenser, swabs, oven (time and temperature). pump assembly with epo-tek 302-3M on sleeve bearing. | epo-tek blocks purge gap
AU Pump-Cannula. Clean with petroleum ether and apply epotek to sleeve bearing. Then cure in oven.. petroleum ether, epotek 302-3M, dispenser, swabs, oven (time and temperature). pump assembly with epo-tek 302-3M on sleeve bearing. | delamination of epot-tek a transition
AU Pump-Cannula. Clean with petroleum ether, apply Adhesive Epoxy Duopox 1895 bead and cure.. pump cannula assembly, Adhesive Epoxy Duopox 1895 , dosing needle, petroleum ether, swabs. Finished pump cannula assmebly. | OD too large
AU Pump-Cannula. Clean with petroleum ether, apply Adhesive Epoxy Duopox 1895 bead and cure.. pump cannula assembly, Adhesive Epoxy Duopox 1895 , dosing needle, petroleum ether, swabs. Finished pump cannula assmebly. | Too little Adhesive Epoxy Duopox 1895 leading to step formation
AU Pump-Cannula. Clean with petroleum ether, apply delo bead and cure.. pump cannula assembly, delo, dosing needle, petroleum ether, swabs. Finished pump cannula assmebly. | OD too large
AU Pump-Cannula. Clean with petroleum ether, apply delo bead and cure.. pump cannula assembly, delo, dosing needle, petroleum ether, swabs. Finished pump cannula assmebly. | Too little delo leading to step formation
AU Pump-Cannula. Conduct tightness and passability test. Tightness test fixture, compressed air, pressure chamber, pressure storage clamping fixture, pressure manometer. Tested assembly. | Accept failed assembly for tightness and passability
AU Pump-Cannula. Cure adhesive. Oven (time and temperature), holding fixture, and glued assembly. cured assembly. | Weak bond
AU Pump-Cannula. Cure assembly in oven. glued together cannula-pigtail-inflow cage and pump assemblies, oven (time and temp), hanging rack. Cured Pump cannula assembly. | Bonding failure at cannula to outflow cage bond
AU Pump-Cannula. Cure assembly in oven. glued together cannula-pigtail-inflow cage and pump assemblies, oven (time and temp), hanging rack. Cured Pump cannula assembly. | incorrect orientation of cannula to pump assembly leading to insertion issues
AU Pump-Cannula. Cure assembly in oven. glued together cannula-pigtail-inflow cage and pump assemblies, oven (time and temp), hanging rack. Cured Pump cannula assembly. | incorrect orientation of cannula to pump assembly leading to unable to package
AU Pump-Cannula. Mix Epotek 302-3M. Clean weld seams with petroleum ether and apply HPA41 to seams. Place in oven. Then apply epo-tek to seams at transition of end-cap to pressure port. Then cure a second time. Oven (time, temperature), HPA41, Epo-tek 302-3M (no aluminum), petroleum ether, dispenser, swabs. pump assembly with epo-tek on weld seams and at transition of end cap to pressure port. | blocked pressure port (too much epotek or promotor)
AU Pump-Cannula. Mix Epotek 302-3M. Clean weld seams with petroleum ether and apply HPA41 to seams. Place in oven. Then apply epo-tek to seams at transition of end-cap to pressure port. Then cure a second time. Oven (time, temperature), HPA41, Epo-tek 302-3M (no aluminum), petroleum ether, dispenser, swabs. pump assembly with epo-tek on weld seams and at transition of end cap to pressure port. | epo-tek delamination/falling promotor
AU Pump-Cannula. Mix Epotek 302-3M. Clean weld seams with petroleum ether and apply HPA41 to seams. Place in oven. Then apply epo-tek to seams at transition of end-cap to pressure port. Then cure a second time. Oven (time, temperature), HPA41, Epo-tek 302-3M (no aluminum), petroleum ether, dispenser, swabs. pump assembly with epo-tek on weld seams and at transition of end cap to pressure port. | lack of promotor leads to epo-tek delamination/falling off
AU Pump-Cannula. Mix epo-tek 302-3M. Resin, hardener, Aluminum Oxide, Speed mixer, scale, hot plate. Mixed epo-tek. | uncured epo-tek leading to it falling off at transitions (hemolysis)
AU Pump-Cannula. Mix epo-tek 302-3M. Resin, hardener, Aluminum Oxide, Speed mixer, scale, hot plate. Mixed epo-tek. | uncured epo-tek leading to it falling off at transitions (thrombus formation)
AU Pump-Cannula. Mix epo-tek 302-3M. Resin, hardener, Aluminum Oxide, Speed mixer, scale, hot plate. Mixed epo-tek. | white spots in epotek (component separation) leading to rough surfaces
AU Pump-Cannula. Mix silicone glue Nusil MED 1511 with cyclohexane, fill gap between sensor and cannula as well as cover sensor with silicone mix, cure at room temperature. Nus MED 1511, Cyclohexane, pump-cannula assembly. Pump assembly with full embedded sensor. | Compromise encapsulation of sensor
AU Pump-Cannula. Mix silicone glue Nusil MED 1511 with cyclohexane, fill gap between sensor and cannula as well as cover sensor with silicone mix, cure at room temperature. Nus MED 1511, Cyclohexane, pump-cannula assembly. Pump assembly with full embedded sensor. | Cross contaimination with uncure silicone gets into other bonding work stations
AU Pump-Cannula. Mix silicone glue Nusil MED 1511 with cyclohexane, fill gap between sensor and cannula as well as cover sensor with silicone mix, cure at room temperature. Nus MED 1511, Cyclohexane, pump-cannula assembly. Pump assembly with full embedded sensor. | Cross contaimination with uncure silicone gets into work other bonding work stations
AU Pump-Cannula. Mix silicone glue Nusil MED 1511 with cyclohexane, fill gap between sensor and cannula as well as cover sensor with silicone mix, cure at room temperature. Nus MED 1511, Cyclohexane, pump-cannula assembly. Pump assembly with full embedded sensor. | Unable to apply properly (scrap)
AU Pump-Cannula. Pull out PTFE mandrel and remove excess epotek with scapel. pump assembly, scapel, mandrel. Pump assembly with fully formed pressure port. | blocked pressure port
AU Pump-Cannula. Pull out PTFE mandrel and remove excess epotek with scapel. pump assembly, scapel, mandrel. Pump assembly with fully formed pressure port. | Scratch endcap surface
AU Pump-Cannula. Room temperature cure. Then cure assembly in oven.. Oven (time and temperature), holding fixture, and glued assembly. cured assembly. | Weak bond
AU Pump-Catheter. Align catheter marking to the sensor. pump assembly with glued catheter and Nitinol wire. Aligned glued pump assembly. | Misalignment of cannula to handle (unable to fit in package)
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, position pump in fixture. Apply Delo onto pump housing ring to fixate the sensor cable. pump catheter with loaded cables, lumen, and nitinol, Delo, fixation block, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | Bond failure of nitinol wire, leading to kink of catheter
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, position pump in fixture. Apply Delo onto pump housing ring to fixate the sensor cable. pump catheter with loaded cables, lumen, and nitinol, Delo, fixation block, dosing needle, wipes, catheter support ring, swabs. pump assembly with glued catheter and Nitinol wire. | Too much delo on pump housing (scrap)
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, position pump in fixture. Apply Delo onto pump housing ring to fixate the sensor cable. pump catheter with loaded cables, lumen, and nitinol, Delo, fixation block, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | Weak bond
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, then slide support ring into position and then position pump in fixture. Apply Delo onto pump housing strut end. pump catheter with loaded cables, lumen, and nitinol, Delo,curing fixture, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | Bond failure of nitinol wire, leading to kink of catheter
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, then slide support ring into position and then position pump in fixture. Apply Delo onto pump housing strut end. pump catheter with loaded cables, lumen, and nitinol, Delo,curing fixture, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | Too much delo on pump housing (scrap)
AU Pump-Catheter. Apply delo to notch and slide in nitinol wire. Apply Delo on the catheter transition and slide catheter onto motorhousing, then slide support ring into position and then position pump in fixture. Apply Delo onto pump housing strut end. pump catheter with loaded cables, lumen, and nitinol, Delo,curing fixture, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | Weak bond
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover strut welds of sensor strut with delo. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered sensor strut welds. | Delo contaminated optical sensor membrane
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover the sensor cable up to the pump delo dot. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered horshoe transition and sensor cable on pump housing. | Exposed sensor cable leading to damage
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover the sensor cable up to the pump housing ring. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered horshoe transition and sensor cable on pump housing. | Damage sensor cable during application
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover the sensor cable up to the pump housing ring. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered horshoe transition and sensor cable on pump housing. | delo migrates to pump housing rough surface leading to a fitting issue with the cannula (Scrap)
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover the sensor cable up to the pump housing ring. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered horshoe transition and sensor cable on pump housing. | Exposed sensor cable leading to damage
AU Pump-Catheter. Apply delo to the pump housing horeshoe transition and cover the sensor cable up to the pump housing ring. cleaned pump catheter assembly, delo, dosing needle, swabs. pump catheter assembly with delo covered horshoe transition and sensor cable on pump housing. | Exposed sharp edges/rough surfaces leading to insertion issues
AU Pump-Catheter. Attach klickadapter to repositioning unit . klickadapter, repositioning unit, repositioning unit modification fixture . Repo with adapter . | klickadapter damages repo
AU Pump-Catheter. Attach klickadapter to repositioning unit . klickadapter, repositioning unit, repositioning unit modification fixture . Repo with adapter . | klickadapter not fully attached
AU Pump-Catheter. Check assembly for Eccentricity. assembly with eptoek on weld seams, feeler gauge, prepping needle. acceptable impeller pump housing gap. | gap size too small is accepted
AU Pump-Catheter. Check assembly for Eccentricity. assembly with eptoek on weld seams, feeler gauge, prepping needle. acceptable impeller pump housing gap. | reject good components
AU Pump-Catheter. Clean catheter end and catheter support ring with petrol ether. Shortened catheter, support ring, petrolether, swabs, time, glove. cleaned components. | Weak bond
AU Pump-Catheter. Clean catheter end and catheter support ring with petrol ether. Shortened catheter, support ring, petrolether, swabs, time. cleaned components. | Weak bond
AU Pump-Catheter. Clean motor housing transition with swabs and petroleum ether. petroleum ether, pump catheter with loaded cables, lumen, and nitinol wire, swabs. cleaned components. | Weak bond
AU Pump-Catheter. Clean pump catheter horseshoe and sensor cable notch with petrol ether. cured pump catheter, petrol ether, swabs, time. cleaned pump catheter. | Reduced adhesion of delo exposing sensor cable to damage
AU Pump-Catheter. Clean pump catheter horseshoe and sensor cable notch with petrol ether. cured pump catheter, petrol ether, swabs, time. cleaned pump catheter. | Reduced adhesion of delo leading to particulates in patient
AU Pump-Catheter. Clean weld seams with Isopropanol and apply HPA41. Then cure in oven. completed welded assembly, HPA41, Ispoproponal, Oven (time and temperature), timer, swabs. cleaned and prepped assembly. | epotek delamination on weld seams leading to skiving of guidewire
AU Pump-Catheter. Cure adhesive. pump with glued catheter and Nitinol wire, oven (time and temperature), curing fixture, rack. cured pump catheter . | Weak bond
AU Pump-Catheter. Cure adhesive. pump with glued catheter and Nitinol wire, oven (time and temperature), fixation block, rack. cured pump catheter . | Weak bond
AU Pump-Catheter. Load the catheter with support ring, sensor cable, motor cable and purge lumen and nitinol wire. Catheter, support ring, motor pump housing assembly with sensor cable, vacuum pull bench, order of insertion of lumen and cable. pump catheter with loaded cables, lumen, and nitinol wire. | Motor Cable damage
AU Pump-Catheter. Load the catheter with support ring, sensor cable, motor cable and purge lumen and nitinol wire. Catheter, support ring, motor pump housing assembly with sensor cable, vacuum pull bench, order of insertion of lumen and cable. pump catheter with loaded cables, lumen, and nitinol wire. | Purge lumen damage/kink
AU Pump-Catheter. Load the catheter with support ring, sensor cable, motor cable and purge lumen and nitinol wire. Catheter, support ring, motor pump housing assembly with sensor cable, vacuum pull bench, order of insertion of lumen and cable. pump catheter with loaded cables, lumen, and nitinol wire. | Sensor cable damage
AU Pump-Catheter. Mix Epo-Tek 302 and cover pump housing weld seams and backfll 2 of 3 struts. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | Damage guidewire (Skiving)
AU Pump-Catheter. Mix Epo-Tek 302 and cover pump housing weld seams and backfll 2 of 3 struts. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | epotek delamination on weld seams leading to skiving of guidewire
AU Pump-Catheter. Mix Epo-Tek 302 and cover pump housing weld seams and backfll 2 of 3 struts. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | sensor covered with epotek - unable to calibrate at final qual
AU Pump-Catheter. Mix Epo-Tek 302 and cover pump housing weld seams and backfll 2 of 3 struts. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. assembly with epotek on weld seams. | Visually unacceptable component (scrap)
AU Pump-Catheter. Mix Epo-Tek 302 and cover pump housing weld seams and backfll 2 of 3 struts. Then cure in oven.. Epotek 302, completed welded assembly, dosing needle, speed mixer. No epotek on sensor strut welds. | epotek contamination of optical sensor
AU Pump-Catheter. Mix Epotek 302. Then cover sensor cable notch with epotek and as well as apply on both ends of catheter support ring. Then cure at room temperature and then in oven. . Pump catheter assembly, epotek 302 (50% aluminum oxide), oven (time and temperature), dosing needle, syringe, curing board. pump catheter with epotek covered sensor cable. | bear sensor cable leading to damage
AU Pump-Catheter. Mix Epotek 302. Then cover sensor cable notch with epotek and as well as apply on both ends of catheter support ring. Then cure at room temperature and then in oven. . Pump catheter assembly, epotek 302 (50% aluminum oxide), oven (time and temperature), dosing needle, syringe, curing board. pump catheter with epotek covered sensor cable. | missing adhesive leading to rough surface (thrombus formation)
AU Pump-Catheter. Mix Epotek 302. Then cover sensor cable notch with epotek and as well as apply on both ends of catheter support ring. Then cure at room temperature and then in oven. . Pump catheter assembly, epotek 302 (50% aluminum oxide), oven (time and temperature), dosing needle, syringe, curing board. pump catheter with epotek covered sensor cable. | Rough surface leading to thrombus formation
AU Pump-Catheter. Mix Epotek 302. Then cover sensor cable notch with epotek and as well as apply on both ends of catheter support ring. Then cure at room temperature and then in oven. . Pump catheter assembly, epotek 302 (50% aluminum oxide), oven (time and temperature), dosing needle, syringe, curing board. pump catheter with epotek covered sensor cable. | uncured epoxy leading to rough surface (thrombus formation)
AU Pump-Catheter. Mix Epotek 302. Then cover sensor cable notch with epotek and as well as apply on both ends of catheter support ring. Then cure at room temperature and then in oven. . Pump catheter assembly, epotek 302 (50% aluminum oxide), oven (time and temperature), dosing needle, syringe, curing board. pump catheter with epotek covered sensor cable. | uncured epoxy leading/incorrect hardness/viscosity leading to rough surface (thrombus formation)
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | catheter too long (unable to assemble)
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | catheter too short (positioning issue)
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | Creased/kink end of catheter
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | cut incorrect end of catheter leading to incorrect position identified
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | Debry accumulation, leading to weak bond
AU Pump-Catheter. Shorten catheter. Catheter, catheter cutting device, defined length. Shortened catheter. | misangled cut, gap between catheter and motor housing
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Assemble unit incorrect order (scrap)
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage motor cables
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage protection tube
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage sensor cable
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage tip of repositioning unit
AU Pump-Catheter. Slide on repositioning unit and protection tube. Then the catheter kink protector. Pump catheter assembly, repositioning unit, auxillary catheter, catheter kink protector, protection tube. AU pump catheter. | Damage/kink catheter
AU Pump-Catheter. Tack fiber using UV to lay flat into slot, cure. pump catheter with loaded cables, lumen, and nitinol, Delo,curing fixture, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | under-cured UV - sensor moves - unable to assemble
AU Pump-Catheter. tack visor and fiber using UV to lay flat into slot, cure. pump catheter with loaded cables, lumen, and nitinol, Delo,curing fixture, dosing needle, wipes, catheter support ring. pump assembly with glued catheter and Nitinol wire. | not fully seated in channel - unable to cover with adhesive
AU Pump-Catheter. Unwind and prepare the sensor cable. Motor pump housing assembly with sensor cable, scapel. Motor pump housing assembly with sensor cable unwound. | Sensor cable damage
AU Pump-Plug. apply Loctite 3321 on sensor and glue on pressure cube. The cure with UV. . Assembly with catheter kink protector bonded into handle, loctite 3321, dosing needle, UV point source. Assembly with pressure cube glued into handle. | Glue migrates onto dome of pressure cube leading to inaccurate pressure readings
AU Pump-Plug. apply Loctite 3321 on sensor and glue on pressure cube. The cure with UV. . Assembly with catheter kink protector bonded into handle, loctite 3321, dosing needle, UV point source. Assembly with pressure cube glued into handle. | Kink pressure line
AU Pump-Plug. apply Loctite 3321 on sensor and glue on pressure cube. The cure with UV. . Assembly with catheter kink protector bonded into handle, loctite 3321, dosing needle, UV point source. Assembly with pressure cube glued into handle. | Leak of pressure line
AU Pump-Plug. Apply Loctite 4062 into the rectangle opening and slide kink protector near the end of the catheter . Kink protector, Pump assembly with catheter, Loctite 4062, rotational position of assembly, axial position of assembly, time. glued on kink protector onto pump assembly. | Additional stress on pigtail
AU Pump-Plug. Apply Loctite 4062 into the rectangle opening and slide kink protector near the end of the catheter . Kink protector, Pump assembly with catheter, Loctite 4062, rotational position of assembly, axial position of assembly, time. glued on kink protector onto pump assembly. | Bond Failure at kink protector to catheter
AU Pump-Plug. Apply Loctite 4062 into the rectangle opening and slide kink protector near the end of the catheter . Kink protector, Pump assembly with catheter, Loctite 4062, rotational position of assembly, axial position of assembly, time. glued on kink protector onto pump assembly. | Unable to package unit
AU Pump-Plug. Apply Loctite 4062 into the rectangle opening and slide kink protector near the end of the catheter align window of kink protector with markings on catheters . Kink protector, Pump assembly with catheter, Loctite 4062, rotational position of assembly, axial position of assembly, time. glued on kink protector onto pump assembly. | Bond Failure at kink protector to catheter
AU Pump-Plug. apply Loctite 4062 on interface plug, kink protections, fixing pins and –bushes; then press top plug half on and fix it with device; then cure. Upper handle half, assembly with microtube and nitinol glued down, loctite 4062, dosing needle, stop watch, clamping fixture. Finished AU pump plug. | bond failure of the upper half
AU Pump-Plug. apply Loctite 4062 on interface plug, kink protections, fixing pins and –bushes; then press top plug half on and fix it with device; then cure. Upper handle half, assembly with microtube and nitinol glued down, loctite 4062, dosing needle, stop watch, clamping fixture. Finished AU pump plug. | Damage outside of pump halves leading to scrap
AU Pump-Plug. Apply Loctite 4062 on interface pump plug half, kink protectors, and fixing pins, then press top plug half on and fix it with clamp; harden at room temperature and remove clamp . Loctite 4062, top pump plug half, press fixture. glued together assembly. | bond failure leading to exposed wiring/board leading to electrical failure of system
AU Pump-Plug. Apply Loctite 4062 on interface pump plug half, kink protectors, and fixing pins, then press top plug half on and fix it with clamp; harden at room temperature and remove clamp . Loctite 4062, top pump plug half, press fixture. glued together assembly. | Damage/Crack pump halfs leading to damage internal electronics
AU Pump-Plug. Apply Loctite 4062 on the kink protector side arm, then rest it in pump plug half and press the kink protection to the pump plug half. Loctite 4062, UV cured assembly. Assembly with glued in side arm. | Bond failure leading to strain on purge line leading to leak
AU Pump-Plug. Apply Loctite 4062 on the plug half bond plug Interface and catheter kink protection to the plug half. Loctite 4062, Grey connector with PCB, pump plug shell. Gluled together assemblies. | Bond failure stress on lumen/wires
AU Pump-Plug. Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection blocked
AU Pump-Plug. Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection bond failure
AU Pump-Plug. Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
AU Pump-Plug. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm damaged
AU Pump-Plug. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge system occluded or break not detected
AU Pump-Plug. Clean Catheter and Kink protector with Petroleum ether. Petroleum ether, kink protector, pump assembly with catheter. Cleaned assemblies. | Bond Failure at kink protector to catheter
AU Pump-Plug. Clean mandrel, inert into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | bond failure at lumen joint leading to leak
AU Pump-Plug. Clean mandrel, inrt into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | bond failure at lumen joint leading to leak
AU Pump-Plug. Clean mandrel, insert into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | bond failure at lumen joint leading to leak
AU Pump-Plug. Clean mandrel, insert into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | insufficient engagement leading to bond leakage
AU Pump-Plug. Clean pump plug shell petroleum ether. Pump plug shell, petroleum ether, swab. cleaned component. | Bond failure
AU Pump-Plug. Clean strain relief, cable and blue handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3325. Handle half with serial number and glued in patient cable assembly. | Damage sensor wires
AU Pump-Plug. Clean strain relief, cable and blue handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3324. Handle half with serial number and glued in patient cable assembly. | Damage to patient cable wires
AU Pump-Plug. Clean strain relief, cable and blue handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3326. Handle half with serial number and glued in patient cable assembly. | Damage unit/leading scrap
AU Pump-Plug. Clean strain relief, cable and blue handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3322. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
AU Pump-Plug. Clean strain relief, cable and blue handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3323. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
AU Pump-Plug. Coat boring and soldering joints with Loctite 3321 and harden under UV- light. Loctite 3321, Grey connector with PCB assembly, UV light (distance and intensity), Holding Fixture. Coated grey connected with PCB. | uncured loctite leading short circuit leading to pump stop
AU Pump-Plug. Coat boring and soldering joints with Loctite 3321 and harden under UV- light. Loctite 3321, Grey connector with PCB assembly, UV light (distance and intensity), Holding Fixture. Coated grey connected with PCB. | uncured loctite leading short circuit leading to pump stop
AU Pump-Plug. Coat soldering joints and contacts with Loctite 3321 and harden under UV- light. Loctite 3321, Grey connector with PCB assembly, UV light (distance and intensity), Holding Fixture. Coated grey connected with PCB. | uncured loctite leading short circuit leading to pump stop
AU Pump-Plug. Conduct Resistance measurement on assmebly. Soldered Assembly, Resistance fixture, Ohm meter equipment with software (DAQ) or switch box w/ multimeter. Tested assembly. | Accept failed assembly for motor cables/sensor connection
AU Pump-Plug. Conduct Resistance measurement on assmebly. Soldered Assembly, Resistance fixture, Ohm meter equipment with software (DAQ). Tested assembly. | Accept failed assembly for motor cables/sensor connection
AU Pump-Plug. Cut microtube to length, slide into the catheter, slide protection catheter above Nitinol wire and then glue opening with Loctite 3321 and harden . Pump assembly with side arm, Loctit 3321, protection catheter, UV light source. UV cured assembly. | Microtube occlude leading insufficient sterilization of internal catheter
AU Pump-Plug. cut to length microtube, slide into the catheter, place nitinol protective tubing over catheter then glue nitinol and microtube at juntion with Loctite 3321 and cure with UV equipment. Cutting board with lengths, assembly with UV glue in kink protector windows, scapel, UV equipment, loctite 3321, microtube, protective tubing. assembly with microtube and nitinol glued down. | difficulty with assembly, potential scrap
AU Pump-Plug. cut to length microtube, slide into the catheter, place nitinol protective tubing over catheter then glue nitinol and microtube at juntion with Loctite 3321 and cure with UV equipment. Cutting board with lengths, assembly with UV glue in kink protector windows, scapel, UV equipment, loctite 3321, microtube, protective tubing. assembly with microtube and nitinol glued down. | Occluded/kinked microtube
AU Pump-Plug. cut to length motor cable. Assembly with glued in kink pressure and purge lumen kink protectors, wire cutters. Assembly with cut motor cable. | unable to assesmble
AU Pump-Plug. Glue kink protector which is connect to catheter onto the pump handle assembly. Plug sensor assembly, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, DSS assembly, and PL plug outlet assembly, loctite 4062. Assembly with catheter kink protector bonded into handle. | Bond failure at handle to kink protector joint
AU Pump-Plug. gluing rectangle opening of kink protection pressure and purge lumen with Loctite 3321 and cure UV equipment. assembly with motor cables soldered to eeprom, loctite 3321, UV equipment, dosing needle. assembly with UV glue in kink protector windows. | bond failure at pressure kink protector leading to internal joint stress (leak)
AU Pump-Plug. gluing rectangle opening of kink protection pressure and purge lumen with Loctite 3321 and cure UV equipment. assembly with motor cables soldered to eeprom, loctite 3321, UV equipment, dosing needle. assembly with UV glue in kink protector windows. | bond failure at purge kink protector leading to internal joint stress (leak)
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4062. Assembly with solder joints with loctite, loctite 4062. Assembly with glued in kink pressure and purge lumen kink protectors. | Kink pressure line
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4063. Assembly with solder joints with loctite, loctite 4062. Assembly with glued in kink pressure and purge lumen kink protectors. | Kink pressure line
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4064. Assembly with solder joints with loctite, loctite 4062. Assembly with glued in kink pressure and purge lumen kink protectors. | Kink purge line
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4065. Assembly with solder joints with loctite, loctite 4063. Assembly with glued in kink pressure and purge lumen kink protectors. | stretching of purge/pressure lumens leading to leak
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4066. Assembly with solder joints with loctite, loctite 4064. Assembly with glued in kink pressure and purge lumen kink protectors. | stretching of purge/pressure lumens leading to leak
AU Pump-Plug. insert kink protecters of pressure and purge lumen into handle slots, fill the gap with Loctite 4067. Assembly with solder joints with loctite, loctite 4065. Assembly with glued in kink pressure and purge lumen kink protectors. | stretching of purge/pressure lumens leading to leak
AU Pump-Plug. Insert PTFE mandrel through pressure cube, apply Loctite 3321 at the end of the pressure lumen; join parts and fill the gap with Loctite 3321,cure parts using UV light, remove mandrel. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly, pressure cube assembly, PTFE mandrel, loctite 3321, UV light. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, DSS assembly, and cured on PL plug outlet assembly. | bond failure at lumen joint leading to leak
AU Pump-Plug. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Damage component/unusable
AU Pump-Plug. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Logo and/or serial number printed on incorrect handle half
AU Pump-Plug. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Smugged/illegible printing
AU Pump-Plug. Measure SNR. Blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Scrap product due to low SNR failure
AU Pump-Plug. Measure SNR. Completed AU Pump-Plug, SNR Tester. Tested Assembly Unit. | Scrap product due to low SNR failure
AU Pump-Plug. Place delrin tip on the auxillary end of catheter, then mount repositioning unit, catheter sleeve, slotted silicon stopper onto the catheter.. Pump cannula assembly, repositioning unit, catheter sleeve, silicon stopper, delrin tip. pump cannula assembly with repositioning unit and catheter protection sleeve. | unable to utilize repostion unit appropriately (scrap)
AU Pump-Plug. Place kink protector on end of assembly, clean area beneath, align so that window is 90 degrees from markings, and glue on with Loctite 4062. pump cannula assembly with repositioning unit and catheter protection sleeve, petroleum ether, loctite 4062, kink protector. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector. | bond failure at kink protector
AU Pump-Plug. Place kink protector on end of assembly, clean area beneath, align so that window is 90 degrees from markings, and glue on with Loctite 4062. pump cannula assembly with repositioning unit and catheter protection sleeve, petroleum ether, loctite 4062, kink protector. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector. | incorrect alignment leading difficulty placing pump in optimal position
AU Pump-Plug. Place kink protector on end of assembly, clean area beneath, align so that window is 90 degrees from markings, and glue on with Loctite 4063. pump cannula assembly with repositioning unit and catheter protection sleeve, petroleum ether, loctite 4062, kink protector. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector. | bond failure at kink protector
AU Pump-Plug. Place kink protector on end of assembly, clean area beneath, align so that window is 90 degrees from markings, and glue on with Loctite 4064. pump cannula assembly with repositioning unit and catheter protection sleeve, petroleum ether, loctite 4062, kink protector. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector. | too much catheter beyond kink protector leading to assembly issues
AU Pump-Plug. Place POM tip on the auxillary end of catheter, then mount repositioning unit, catheter sleeve, slotted silicon stopper onto the catheter.. Pump cannula assembly, repositioning unit, catheter sleeve, silicon stopper, POM tip. pump cannula assembly with repositioning unit and catheter protection sleeve. | unable to utilize repostion unit appropriately/damage reposition unit
AU Pump-Plug. Place POM tip on the auxillary end of catheter, then mount repositioning unit, catheter sleeve, slotted silicon stopper onto the catheter.. Pump cannula assembly, repositioning unit, catheter sleeve, silicon stopper. pump cannula assembly with repositioning unit and catheter protection sleeve. | unable to lock protective sleeve in place during use/damage sleeve
AU Pump-Plug. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Loctite 4062, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | kink protector window orientation to catheter incorrect
AU Pump-Plug. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Loctite 4062, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | Low bond strength leading to component separation
AU Pump-Plug. Position sensor wires above the soldering pads of the PCB, add UV glue to restrain units, tin solder pads and solder them to pad. UV glue (Loctite (3321), Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly, Wire Stripper/Cutter, UV light source (time and distance). Soldered sensor cable. | Damage solder leads/incomplete solder leading to pressure sensor failure
AU Pump-Plug. remove jacket of motor cable with hot weezer; turn and skin the separate strands. hot weezer, Assembly with cut motor cable, metal ruler. assembly with motor cable jacket removed. | damage motor cables
AU Pump-Plug. Seal all soldering joints with Loctite 3321 and harden with UV. UV glue (Loctite (3321), tested assembly, UV cure equipment (Time and distance). UV cured assembly. | open bubbles or voids in UV adhesive
AU Pump-Plug. Seal all soldering joints with Loctite 3321 and harden with UV. UV glue (Loctite (3321), tested assembly, UV cure equipment (Time and distance). UV cured assembly. | uncured loctite leading short circuit leading to pump stop
AU Pump-Plug. seal all soldering joints with Loctite 3321 and harden. Assembly with pressure cube glued into handle, loctite 3321, UV flood lab (distance and intensity). Assembly with solder joints with loctite. | Unacceptable sealant leading to scrap
AU Pump-Plug. Shorten end of motor cable and Sensor cable, strip off insulation off each cable wires and position them above the soldering pads of the PCB, add UV glue to restrain units, tin solder pads and solder them to pad. UV glue (Loctite (3321), Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly, Wire Stripper/Cutter, UV light source (time and distance). Soldered assembly. | Damage cables/incomplete solder leading to pump stop
AU Pump-Plug. Shorten end of motor cable and Sensor cable, strip off insulation off each cable wires and position them above the soldering pads of the PCB, add UV glue to restrain units, tin solder pads and solder them to pad. UV glue (Loctite (3321), Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly, Wire Stripper/Cutter, UV light source (time and distance). Soldered assembly. | Damage solder leads/incomplete solder leading to pressure sensor failure
AU Pump-Plug. Shorten end of motor cable and Sensor cable, strip off insulation off each motor cable wires and position them above the soldering pads of the PCB, add UV glue to restrain units, tin solder pads and solder them to pad. UV glue (Loctite (3321), Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly, Wire Stripper/Cutter, UV light source (time and distance). Soldered assembly. | Damage cables/incomplete solder leading to pump stop
AU Pump-Plug. Shorten end of motor cable and Sensor cable, strip off insulation off each motor cable wires and position them above the soldering pads of the PCB, add UV glue to restrain units, tin solder pads and solder them to pad. UV glue (Loctite (3321), Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly, Wire Stripper/Cutter, UV light source (time and distance). Soldered assembly. | Damage solder leads/incomplete solder leading to pressure sensor failure
AU Pump-Plug. shorten pressure and purge lumen, Clean mandrel, insert into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | bond failure at lumen joint leading to leak
AU Pump-Plug. shorten pressure and purge lumen, Clean mandrel, insert into DSS plug outlet, gluing between purge lumen and kink protection purge lumen by use of THF, curein infrared oven. Mandrel, THF, pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector with lumens cut to length, dosing needle, petroleum ether, DSS plug assembly, cutting fixture. pump cannula assembly with repositioning unit, catheter protection sleeve, glued on kink protector, and cured on DSS assembly. | unable to assesmble to lengths cut incorrectly
AU Pump-Plug. Shorten purge lumen, solvent bond kink protection and purge lumen with THF, process assembly in IR-oven. IR Oven (time and temperature), pump assembly with kink protector, scapel, cutting board, mandrel, THF. Solvent bonded assembly. | Bond failure leading to leak
AU Pump-Plug. Shorten purge lumen, solvent bond kink protection and purge lumen with THF, process assembly in IR-oven. IR Oven (time and temperature), pump assembly with kink protector, scapel, cutting board, mandrel, THF. Solvent bonded assembly. | purge lumen too short unable to assemble leading to scrap
AU Pump-Plug. Solvent bond kink protection and purge lumen with THF, process assembly in IR-oven. IR Oven (time and temperature), pump assembly with kink protector, mandrel, THF. Solvent bonded assembly. | Bond failure leading to leak
AU Pump-Plug. Solvent bond kink protection and purge lumen with THF, process assembly in IR-oven. IR Oven (time and temperature), pump assembly with kink protector, mandrel, THF. Solvent bonded assembly. | purge lumen too short unable to assemble leading to scrap
AU Pump-Plug. Strip insulation off each motor cable wires and tin wire ends. Solder station, Solder material, Hot weezer, Pump plug fixture, pump plug assembly. Tinned wires. | Damage cables/incomplete solder leading to pump stop
AU Pump-Plug. Tin solder pads and solder motor cables to pad. Solder station, Solder material, Pump plug fixture, pump plug assembly,. Soldered motor cable. | Damage cables/incomplete solder leading to pump stop
AU Pump-Plug. tin strands, shield and soldering pads; than soldering wires, measure resistance, then seal all soldering joints with Loctite 3321 and cure with UV. soldering iron (temp) , solder material, assembly with motor cable jacket removed, mulitmeter, UV light source (time and surface). assembly with motor cables soldered to eeprom. | damage eeprom/wires
AU Pump-Plug. tin strands, shield and soldering pads; than soldering wires, measure resistance, then seal all soldering joints with Loctite 3321 and cure with UV. soldering iron (temp) , solder material, assembly with motor cable jacket removed, mulitmeter, UV light source (time and surface). assembly with motor cables soldered to eeprom. | poor soldering quality
AU Pump-Plug. tin strands, shield and soldering pads; than soldering wires, measure resistance, then seal all soldering joints with Loctite 3321 and cure with UV. soldering iron (temp) , solder material, assembly with motor cable jacket removed, mulitmeter, UV light source (time and surface). assembly with motor cables soldered to eeprom. | poor soldering quality/incorrect connections
AU Pump-Plug. Trim Motor Cable and Purge Lumen. Trim Fixture, Cutting tools, Pump Plug Assembly. Pump Assembly with cables and purge lumen at correct lengths. | Lumens cut to incorrect lengths leading to unable to assemble
AU Pump-Plug. Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | parts printed with incorrect logo
AU Pump-Plug. Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | Unable to print components with logo
AU Pump-Plug. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Low SNR leading to inaccurate placement signal information
AU Pump-Plug. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Placement signal failure
AU Pump-Plug. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Scrap product due to low SNR failure
AU Pump-Plug. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested blue handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Sensor failure/low sensor signal
AU Purge Lumen. Apply Adhesive Epoxy Duopox 1895 on area over lumen, then pull into inside twin pin. Adhesive Epoxy Duopox 1895, purge lumen, pressure lumen, twin pin unit, mandrel. glued in pressure and purge lumen with mandrel. | Kink purge lumen
AU Purge Lumen. Apply Adhesive Epoxy Duopox 1895 on area over lumen, then pull into inside twin pin. Adhesive Epoxy Duopox 1895, purge lumen, pressure lumen, twin pin unit, mandrel. glued in pressure and purge lumen with mandrel. | Poor bond strength
AU Purge Lumen. Apply Adhesive Epoxy Duopox 1895 on area over lumen, then pull into inside twin pin. Adhesive Epoxy Duopox 1895, purge lumen, pressure lumen, twin pin unit, mandrel. glued in pressure and purge lumen with mandrel. | Poor bond strength/Glue in lumen
AU Purge Lumen. Apply Delo Duopox adhesive to overhanging purge lumen. Delo, mandrel stiffening pin, and purge lumen assembly, microscope. Delo applied to assembly. | occluded purge line
AU Purge Lumen. Apply Delo Duopox adhesive to overhanging purge lumen. Delo, mandrel stiffening pin, and purge lumen assembly, microscope. Delo applied to assembly. | Weak bond leading to leak
AU Purge Lumen. Apply Delo-Duopox on area over lumen, then pull into inside twin pin. Delo-Duopox 1895, purge lumen, pressure lumen, twin pin unit, mandrel. glued in pressure and purge lumen with mandrel. | Poor bond strength
AU Purge Lumen. Apply Delo-Duopox on area over lumen, then pull into inside twin pin. Delo-Duopox 1895, purge lumen, pressure lumen, twin pin unit, mandrel. glued in pressure and purge lumen with mandrel. | Poor bond strength/Glue in lumen
AU Purge Lumen. bond nitinol to single pin. single pin with bonded purge lumen, nitinol wire, delo, dispensing machine, oven (Time, Temperature). nitinol bonded to single pin. | insufficient bond
AU Purge Lumen. Bond nitinol with Adhesive Epoxy Duopox 1895 into assembly and cure. Nitinol wire, Adhesive Epoxy Duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to pressure lumen
AU Purge Lumen. Bond nitinol with Adhesive Epoxy Duopox 1895 into assembly and cure. Nitinol wire, Adhesive Epoxy Duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to purge lumen
AU Purge Lumen. Bond nitinol with Adhesive Epoxy Duopox 1895 into assembly and cure. Nitinol wire, Adhesive Epoxy Duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to pressure lumen
AU Purge Lumen. Bond nitinol with Adhesive Epoxy Duopox 1895 into assembly and cure. Nitinol wire, Adhesive Epoxy Duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to purge lumen
AU Purge Lumen. Bond nitinol with delo into assembly and cure. Nitinol wire, delo-duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to pressure lumen
AU Purge Lumen. Bond nitinol with delo into assembly and cure. Nitinol wire, delo-duopox 1895, oven (time and temperature), pressure and purge lumen in twin pin unit. Bonded twin pin with nitinol wire, purge lumen, and pressure lumen assembly (Twin pin assembly). | bond failure of nitinol leading to damage to purge lumen
AU Purge Lumen. Bond optical lumen to single pin assembly. Single pin with bonded purge lumen and nitinol, delo, oven (Temp and Time), PTFE Mandrel, Optical lumen, dispensing machine (pressure, suck-back setting). Single pin assembly with bonded optical lumen. | delo migrates to single pin surface causing it to be unable to be seated properly
AU Purge Lumen. Bond optical lumen to single pin assembly. Single pin with bonded purge lumen and nitinol, delo, oven (Temp and Time), PTFE Mandrel, Optical lumen, dispensing machine (pressure, suck-back setting). Single pin assembly with bonded optical lumen. | Unable to assemble optical fiber
AU Purge Lumen. Bond optical lumen to single pin assembly. Single pin with bonded purge lumen and nitinol, delo, oven (Temp and Time), PTFE Mandrel, Optical lumen, dispensing machine (pressure, suck-back setting). Single pin assembly with bonded optical lumen. | unable to assembly catheter jacket
AU Purge Lumen. clean bonding surfaces and mandrel with petroleum ether. purge lumen, mandrel, petroleum ether, microscope, tape, stiffening pin. Clean Purge lumen, mandrel, and stiffening pin. | Particles within the purge lumen
AU Purge Lumen. clean bonding surfaces and mandrel with petroleum ether. purge lumen, mandrel, petroleum ether, microscope, tape, stiffening pin. Clean Purge lumen, mandrel, and stiffening pin. | Weak bond due to contaminated surfaces
AU Purge Lumen. clean mandrel and twin pin with petroleum ether. mandrel, twin pin unit, petroleum ether. cleaned twin pin unit and mandrel. | Poor bond strength
AU Purge Lumen. Clean nitinol wire with petroleum ether. cured pressure and purge lumen in twin pin unit with mandrel, scapel, petroleum ether. clean pressure and purge lumen in twin pin unit. | contaminated surface leading to bond failure leading to kink catheter
AU Purge Lumen. Cure adhesive. Oven (time and temperature), Assembly with purge lumen with correct pin positioning. . | Weak bond leading to leak
AU Purge Lumen. Cure adhesive. Oven (time and temperature), Assembly with purge lumen with correct pin positioning. Cured assembly. | Weak bond leading to leak
AU Purge Lumen. Cure assembly. glued in pressure and purge lumen with mandrel, oven (time and temperature). cured pressure and purge lumen in twin pin unit with mandrel. | Poor bond strength
AU Purge Lumen. Cut excess purge tubing from reinforcing pin end with scalpel. Cured assembly, scapel. Trimmed assembly. | Too much engagement with ball bearing leading to increase motor current
AU Purge Lumen. Cut optical lumen to length. Length fixture, optical lumen, blade, operator. Coiled optical lumen. | kinked optical lumen, unable to pass through
AU Purge Lumen. Cut optical lumen to length. Length fixture, optical lumen, blade, operator. Coiled optical lumen. | Optical Lumen Cut too Long leading to rework
AU Purge Lumen. Cut optical lumen to length. Length fixture, optical lumen, blade, operator. Coiled optical lumen. | Optical Lumen Cut too Short leading to unable to assemble
AU Purge Lumen. Cut optical lumen to length. Length fixture, optical lumen, blade, operator. Coiled optical lumen. | Optical Lumen damaged leading purge lumen leak
AU Purge Lumen. Cut purge lumen to length, . purge lumen, mandrel, microscope, tape, catheter cutting bench. cut to legnth purge lumen. | Damage purge lumen leading to leak
AU Purge Lumen. Cut purge lumen to length, . purge lumen, mandrel, microscope, tape, catheter cutting bench. cut to legnth purge lumen. | Lumen cut too short leading to scrap
AU Purge Lumen. Cut purge lumen to length. Length fixture, purge lumen, blade, operator. Coiled purge lumen. | Purge Lumen Cut too Long leading to rework
AU Purge Lumen. Cut purge lumen to length. Length fixture, purge lumen, blade, operator. Coiled purge lumen. | Purge Lumen Cut too Short leading to unable to assemble
AU Purge Lumen. Cut purge lumen to length. Length fixture, purge lumen, blade, operator. Coiled purge lumen. | Purge Lumen damaged leading purge lumen leak
AU Purge Lumen. cut to length purge and pressure lumen; clean bonding surface . Scale, pressure lumen, purge lumen, petroleum ether, blade. Cut and cleaned pressure lumen. | Damage pressure lumen
AU Purge Lumen. cut to length purge and pressure lumen; clean bonding surface . Scale, pressure lumen, purge lumen, petroleum ether, blade. Cut and cleaned pressure lumen. | incorrect length of components (scrap/rework)
AU Purge Lumen. cut to length purge and pressure lumen; clean bonding surface . Scale, pressure lumen, purge lumen, petroleum ether, blade. Cut and cleaned pressure lumen. | Poor bond strength
AU Purge Lumen. Insert mandrel into purge lumen and then add stiffening pin onto purge lumen . mandrell, purge lumen, stiffen pin. mandrel, stiffening, pin and purge lumen assembly. | occluded purge line
AU Purge Lumen. insert mandrell in connection pin; apply seam of UV glue Loctite 3321 at the end of the connection pin and insert pressure lumen. pressure lumen, UV glue, mandrel, twin pin unit, curing system (intensity and distance and time). glued in pressure lumen with mandrel. | Poor bond strength/Glue in lumen
AU Purge Lumen. Package units and label. Package, trimmed assembly, label. Packaged and labeled assembly. | Mix-up with different product types
AU Purge Lumen. Place single pin on mandrel and bond purge lumen, cure in oven. Nitinol wire, Single pin, purge lumen, delo, Oven (Temp and Time). Single pin with bonded purge lumen. | delo blocks lumen
AU Purge Lumen. Place single pin on mandrel and bond purge lumen, cure in oven. Nitinol wire, Single pin, purge lumen, delo, Oven (Temp and Time). Single pin with bonded purge lumen. | insufficient bond
AU Purge Lumen. Place single pin on mandrel and bond purge lumen, cure in oven. Nitinol wire, Single pin, purge lumen, delo, Oven (Temp and Time). Single pin with bonded purge lumen. | insufficient bond
AU Purge Lumen. Place single pin on mandrel and bond purge lumen, cure in oven. Nitinol wire, Single pin, purge lumen, delo, Oven (Temp and Time). Single pin with bonded purge lumen. | mandrel puncture purge lumen
AU Purge Lumen. Place single pin on mandrel and bond purge lumen, cure in oven. Nitinol wire, Mandrel, Single pin, purge lumen, delo, Oven (Temp and Time). Single pin with bonded purge lumen. | purge lumen protrudes distally out of single pin
AU Purge Lumen. Pull purge lumen back until almost flush with end of stiffening pin. Mandrel, stiffening pin, and purge lumen assembly with delo. Assembly with purge lumen with correct pin positioning. | Purge lumen incorrect position leading to leak
AU Purge Lumen. subassembly shipping. finished single pin subassembly, heat sealed bag, shipping container, shipping conditions. single pin subassembly ready for use at next assembly level. | insufficient bond strength
AU Purge Lumen. subassembly shipping. finished single pin subassembly, heat sealed bag, shipping container, shipping conditions. single pin subassembly ready for use at next assembly level. | lumens punctured by nitinol wire
AU Purge Lumen. subassembly storage. finished single pin subassembly, heat sealed bag, storage conditions. single pin subassembly ready for use at next assembly level. | insufficient bond strength
AU Purge Lumen. subassembly storage. finished single pin subassembly, heat sealed bag, storage conditions. single pin subassembly ready for use at next assembly level. | lumens punctured by nitinol wire
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 10 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Butterfuly in incorrect position/not positioned leading to leak (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 11 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Damage repositioning unit assembly leading to leak (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 12 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Unable to fix butterfly to skin (stabilize pump position)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 3 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | bond failure leading to leakage of the system (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 4 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | bond failure leading to leakage of the system (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 5 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Positioning issues (unable to reposition in the field)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 6 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Positioning issues (unable to reposition in the field)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 7 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | bond failure leading to leakage of the system (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 8 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | bond failure leading to leakage of the system (hematoma)
AU Repositioning. Replace blue butterfly then clean and thenslide on 3 rib butterfly above the lock on the haemostatic valve and snap it, fixed with one drop of Loctite 4061. . 9 rib butterfly, repositioning unit, petroleum ether, wipes, loctite 4061, dosing needle. Repositioning unit assembly. | Butterfuly in incorrect position/not positioned leading to leak (hematoma)
AU Rotor. Clean shaft, magnet and sleeve bearing with petroleum ether, slide bearing on shaft. Impeller shafts, sleeve bearings, petroleum ether, wipes, swabs, magnet brush, compressed air, time. Cleaned components. | Bond failure leading to pump stop
AU Rotor. Clean shaft, magnet and sleeve bearing with petroleum ether, slide bearing on shaft. Impeller shafts, sleeve bearings, petroleum ether, wipes, swabs, magnet brush, compressed air, time. Cleaned components. | Damage magnet surface
AU Rotor. Clean shaft, magnet and sleeve bearing with petroleum ether, slide bearing on shaft. Impeller shafts, sleeve bearings, petroleum ether, wipes, swabs, magnet brush, compressed air. Cleaned components. | Contimination of shaft leading to bond failure
AU Rotor. Clean shaft, magnet and sleeve bearing with petroleum ether, slide bearing on shaft. Impeller shafts, sleeve bearings, petroleum ether, wipes, swabs, magnet brush, compressed air. Cleaned components. | Contimination of shaft/magnet leading to bond failure
AU Rotor. Apply Delo Duopox adhesive to impeller shaft, slide magnet onto shaft and bond in place. Delo, Impeller, sleeve bearing, magnet, dosing needle. Bonded on Magnet to shaft. | Bond failure leading to pump stop
AU Rotor. Apply delo to hub and mount sleeve bearing. Then cure adhesive.. hub, sleeve bearing, alignment fixture, delo, dosing needle, Oven (time and temperature). sleeve bearing glued into hub. | Bond failure leading to loose sleeve bearing interfering with rotating shaft/Impeller
AU Rotor. clean assembly with Isopropanol and dry with compressed air. rotor assembly, Isopropanol . cleaned rotor assembly. | contamination leading to bioburden issue
AU Rotor. clean shaft and bearing with isopropanol, sleeve bearing on shaft. paired sleeve bearing and shaft, Isopropanol. cleaned paired sleeve bearing and shaft. | Debry on rotor leading to shaft wear
AU Rotor. clean shaft and bearing with isopropanol, sleeve bearing on shaft. paired sleeve bearing and shaft, Isopropanol. cleaned paired sleeve bearing and shaft. | Unit unable to be utilized leading to scrap
AU Rotor. Clean shaft and sleeve bearing with petroleum ether, slide on mounting fixture. shaft, sleeve bearing, mounting fixture, petroleum ether, brush, swabs, compressed air. Cleaned components mounted on fixture. | Bond failure leading to pump stop
AU Rotor. Clean shaft and sleeve bearing with petroleum ether, slide on mounting fixture. shaft, sleeve bearing, mounting fixture, petroleum ether, brush, swabs, compressed air. Cleaned components mounted on fixture. | Contimination of shaft leading to bond failure
AU Rotor. Clean shaft and sleeve bearing with petroleum ether, slide on mounting fixture. shaft, sleeve bearing, mounting fixture, petroleum ether, brush, swabs, compressed air. Cleaned components mounted on fixture. | Contimination of shaft/magnet leading to bond failure
AU Rotor. Clean shaft and sleeve bearing with petroleum ether, slide on mounting fixture. shaft, sleeve bearing, mounting fixture, petroleum ether, brush, swabs, compressed air. Cleaned components mounted on fixture. | Thrombus formation
AU Rotor. Cure adhesive. . Uncured Rotor assembly with shim in fixture, Oven (time and temperature). Cured Rotor. | Bond failure leading to pump stop
AU Rotor. Cure adhesive. Then clean assembly with petroleum ether and dry with compressed air. Then package in box.. Uncured Rotor assembly with shim in fixture, Oven (time and temperature), petroleum ether, compressed air, swab, box. Cured Rotor assembly in box. | Bond failure leading to pump stop
AU Rotor. Cure adhesive. Then lean assembly with isopropanol and dry with compressed air. Uncured Rotor assembly with shim in fixture, Oven (time and temperature), petroleum ether, compressed air, swab, box. Cured Rotor assembly in box. | Bond failure leading to pump stop
AU Rotor. Cure adhesive.. Uncured Rotor assembly with shim in fixture, Oven (time and temperature),. Cured Rotor. | Bond failure leading to pump stop
AU Rotor. Cure adhesive.. Uncured Rotor assembly with shim in fixture, Oven (time and temperature). Cured Rotor. | Bond failure leading to pump stop
AU Rotor. curing of assembly and functional test. rotor assembly with shim in fixture, oven (time and temperature). cured rotor assembly. | bond failure of magnet to shaft
AU Rotor. dose ring of glue with Adhesive Epoxy Duopox 1895 on the shaft and slide the shaft rotary in the magnet. paired sleeve bearing and shaft, Adhesive Epoxy Duopox 1895, magnet. glued on magnet onto shaft and sleeve bearing. | bond failure to pump stop
AU Rotor. dose ring of glue with Delo Duopox 1895 on the shaft and slide the shaft rotary in the magnet. paired sleeve bearing and shaft, delo, magnet. glued on magnet onto shaft and sleeve bearing. | bond failure to pump stop
AU Rotor. dose ring of glue with Delo Duopox on the shaft and slide the shaft rotary in the magnet. paired sleeve bearing and shaft, delo, magnet. glued on magnet onto shaft and sleeve bearing. | bond failure to pump stop
AU Rotor. Fill chamfer of the magnet with adhesive. Cure adhesive. Clean Assembly with petroleum Ether. Dry with comepressesd air. Pack in box. Rotor assembly in back filling fixture, adhesive, Oven (time and temperature), petroleum ether, compressed air, swab, box. Cured assembly in box. | expired adhesive used
AU Rotor. Fill chamfer of the magnet with adhesive. Cure adhesive. Clean Assembly with petroleum Ether. Dry with comepressesd air. Pack in box. Rotor assembly in back filling fixture, adhesive, Oven (time and temperature), petroleum ether, compressed air, swab, box. Cured assembly in box. | insufficient adhesive - corrosion of magnet - pump stop
AU Rotor. Fill chamfer of the magnet with adhesive. Cure adhesive. Clean Assembly with petroleum Ether. Dry with comepressesd air. Pack in box. Rotor assembly in back filling fixture, adhesive, Oven (time and temperature), petroleum ether, compressed air, swab, box. Cured assembly in box. | too much adhesive - pump stop
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match. . Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft and Match results data list. | Low purge pressure
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match.. Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft. | High purge pressure
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match.. Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft and Match results data list. | High purge pressure
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match.. Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft and Match results data list. | Low purge pressure
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match.. Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft and Match results data list. | Match failes incorrectly fails requirement
AU Rotor. Inspect for damage. Match impeller shaft diameter to sleeve bearing by test fitting pieces together. Conduct purge rate test to ensure match.. Impeller shaft, sleeve bearing, Bubble Test fixture,Smart Manometer, Pyrogen Free Water, Compressed Air, Clean glass beakers, Timer, Magic Tape (Scotch), Bubble collector fixture for bubble test, Isopropyl. Matched pair sleeve bearing and Impeller shaft and Match results data list. | match fails incorrectly fails requirement
AU Rotor. Package units in marked boxes and attach work order. Plastic box with segements, matched pair sleeve bearings and shaft, work order. Units in box. | Damage units
AU Rotor. Package units in marked boxes and attach work order. Plastic box with segements, matched pair sleeve bearings and shaft, work order. Units in box. | High purge pressure
AU Rotor. Package units in marked boxes and attach work order. Plastic box with segements, matched pair sleeve bearings and shaft, work order. Units in box. | Low purge pressure
AU Rotor. packing combination bearing – shaft together. bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, qualified water, stop watch. paired slide bearing and shaft. | wrong combination assembled leading to pump stop
AU Rotor. packing combination bearing – shaft together. bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, water, stop watch. paired slide bearing and shaft. | wrong combination assembled leading to pump stop
AU Rotor. Perform concentricity check of sleeve bearing. sleeve bearing glued into hub, concentricity gauge. co-axial sleeve bearing glued into hub assembly. | stiff impeller (unable rotate)
AU Rotor. Perform concentricity check of sleeve bearing. sleeve bearing glued into hub, concentricity gauge. co-axial sleeve bearing glued into hub assembly. | stiff impeller (unable rotate)/high motor current
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, water, stop watch. Setup bubble test. | bearing not fixed
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, qualified water, stop watch. Setup bubble test. | bearing not fixed
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, water, stop watch. Setup bubble test. | Damaged sealing of sleeve bearing
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, qualified water, stop watch. Setup bubble test. | Damaged sealing of sleeve bearing
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, water, stop watch. Setup bubble test. | Incorrect position of sleeve bearing
AU Rotor. position and fix slide bearing into collet, put shaft into bearing, fix it, connect pressure reducer and gauge and identify purge flow . bubble test fixture, bubble device for pump 350, three way stopcock, Impeller shaft, slide bearing, purge flow beaker, qualified water, stop watch. Setup bubble test. | Incorrect position of sleeve bearing
AU Rotor. Position impeller shaft and magnet with bushing, then place in fixture. Bushing, bondded magnet to shaft assembly, Rotor assembly fixture. Uncured Rotor assembly with bushing in fixture. | Damage magnet surface
AU Rotor. Position impeller shaft and magnet with bushing, then place in fixture. Bushing, bondded magnet to shaft assembly, Rotor assembly fixture. Uncured Rotor assembly with bushing in fixture. | Misalignment of magnet to shaft leading to bearing issues
AU Rotor. Position impeller shaft and magnet with shim, then place in fixture. Shim, bondded magnet to shaft assembly, Rotor assembly fixture. Uncured Rotor assembly with shim in fixture. | Damage magnet surface
AU Rotor. Position impeller shaft and magnet with shim, then place in fixture. Shim, bondded magnet to shaft assembly, Rotor assembly fixture. Uncured Rotor assembly with shim in fixture. | Misalignment of magnet to shaft leading to bearing issues
AU Rotor. position rotor in the fixture with the shim. glued on magnet onto shaft and sleeve bearing, shim, rotor assembly fixture. rotor assembly with shim in fixture. | damaged Impeller tip
AU Rotor. position rotor in the fixture with the shim. glued on magnet onto shaft and sleeve bearing, shim, rotor assembly fixture. rotor assembly with shim in fixture. | damaged sleeve bearing
AU Rotor. position rotor in the fixture with the shim. glued on magnet onto shaft and sleeve bearing, shim, rotor assembly fixture. rotor assembly with shim in fixture. | incorrect rotor engagement to bearing leading to pump stop
AU Rotor. Slide on sleeve bearing-hub assembly on shaft. Then apply Delo Duopox adhesive to impeller shaft, slide magnet onto shaft and bond in place. shaft, Delo, sleeve bearing-hub assembly, curing fixture. glued on magnet sleeve bearing-hub, and shaft assembly. | Bond failure leading to pump stop
AU Sensor-Sensor Cable. Apply the Glob Top epoxy (Epotek H70E-2) with a small needle. Epotek H70E-2, dosing needle, cable/sensor assembly, dipping/glob topping fixture, scale, cup, spatula. Sensor/cable assembly with epoxy . | Inadequate amount of epoxy leading to sensor corrosion due to moisture
AU Sensor-Sensor Cable. Apply the Glob Top epoxy (Epotek H70E-2) with a small needle. Epotek H70E-2, dosing needle, cable/sensor assembly, dipping/glob topping fixture, scale, cup, spatula. Sensor/cable assembly with epoxy . | Incorrect ratio of epoxy leading to sensor corrosion due to moisture
AU Sensor-Sensor Cable. Apply thin layer of silicone (Ratio 1:1) onto the sensor's connection manually. Cure and dip sensor with second silicone (Ratio 1:3), then cure. Then dip the sensor with third layer of silicone (Ratio 1:4) and cure.. Mixed silicone glue ratios, beaker, dipping glob-topping fixture, dosing needle, dispensor, sensor assembly, syringe, tweezers, timer. Sensor with three layers of cured silicone. | Damage sensor during application
AU Sensor-Sensor Cable. Apply thin layer of silicone (Ratio 1:1) onto the sensor's connection manually. Cure and dip sensor with second silicone (Ratio 1:3), then cure. Then dip the sensor with third layer of silicone (Ratio 1:4) and cure.. Mixed silicone glue ratios, beaker, dipping glob-topping fixture, dosing needle, dispensor, sensor assembly, syringe, tweezers, timer. Sensor with three layers of cured silicone. | Inadequate amount of silicone leading to moisture ingress leading to sensor failure
AU Sensor-Sensor Cable. Apply thin layer of silicone (Ratio 1:1) onto the sensor's connection manually. Cure and dip sensor with second silicone (Ratio 1:3), then cure. Then dip the sensor with third layer of silicone (Ratio 1:4) and cure.. Mixed silicone glue ratios, beaker, dipping glob-topping fixture, dosing needle, dispensor, sensor assembly, syringe, tweezers, timer. Sensor with three layers of cured silicone. | moisture ingress leading to sensor failure
AU Sensor-Sensor Cable. Apply thin layer of silicone (Ratio 1:1) onto the sensor's connection manually. Cure and dip sensor with second silicone (Ratio 1:3), then cure. Then dip the sensor with third layer of silicone (Ratio 1:4) and cure.. Mixed silicone glue ratios, beaker, dipping glob-topping fixture, dosing needle, dispensor, sensor assembly, syringe, tweezers, timer. Sensor with three layers of cured silicone. | Sensor cavity is filled with silicone leading to sensor failure
AU Sensor-Sensor Cable. Clean sensor pocket with petroleum ether. Petroleum ether, bonder, bond-fixture, swabs. Cleaned equipment. | contamination leading to sensor failure
AU Sensor-Sensor Cable. Clean sensor pocket with petroleum ether. Petroleum ether, bonder, bond-fixture, swabs. Cleaned equipment. | Damage sensor
AU Sensor-Sensor Cable. Clean sensor pocket with petroleum ether. Petroleum ether, bonder, bond-fixture, swabs. Cleaned equipment. | Sensor failure due to incorrect bond
AU Sensor-Sensor Cable. Clean sensor with Aceton, Isopropyl, and pyrogen free water and then place in oven.. Cleaning fixture, Aceton, Isopropyl, pyrogen free water, beaker, dispensor, sensor/cable assembly, oven (time and temperature). Cleaned sensor/cable assembly. | Clean not clean Glob Top epoxy fails leading to moisture damage
AU Sensor-Sensor Cable. Clean sensor with Aceton, Isopropyl, and pyrogen free water and then place in oven.. Cleaning fixture, Aceton, Isopropyl, pyrogen free water, beaker, dispensor, sensor/cable assembly, oven (time and temperature). Cleaned sensor/cable assembly. | Clean not/discoloration clean Glob Top epoxy fails leading to moisture damage
AU Sensor-Sensor Cable. Clean sensor with Aceton, Isopropyl, and pyrogen free water and then place in oven.. Cleaning fixture, Aceton, Isopropyl, pyrogen free water, beaker, dispensor, sensor/cable assembly, oven (time and temperature). Cleaned sensor/cable assembly. | Damage sensor/cable assembly
AU Sensor-Sensor Cable. Clean sensor with Aceton, Isopropyl, and pyrogen free water and then place in oven.. Cleaning fixture, Aceton, Isopropyl, pyrogen free water, beaker, dispensor, sensor/cable assembly, oven (time and temperature). Cleaned sensor/cable assembly. | Glob top epoxy fails leading to moisture damage
AU Sensor-Sensor Cable. Cure epoxy in oven. Remove from oven and then apply second layer of epoxy. Then cure in oven again.. Oven (time and temperature), sensor cable with epoxy, Epotek H70E-2, dipping/glob topping fixture, dosing needle, cup. Sensor/cable assembly with double layered cured epoxy. | Corrosion of sensor leading to sensor failure
AU Sensor-Sensor Cable. Cure epoxy in oven. Remove from oven and then apply second layer of epoxy. Then cure in oven again.. Oven (time and temperature), sensor cable with epoxy, Epotek H70E-2, dipping/glob topping fixture, dosing needle, cup. Sensor/cable assembly with double layered cured epoxy. | Damage sensor during application
AU Sensor-Sensor Cable. Cure epoxy in oven. Remove from oven and then apply second layer of epoxy. Then cure in oven again.. Oven (time and temperature), sensor cable with epoxy, Epotek H70E-2, dipping/glob topping fixture, dosing needle, cup. Sensor/cable assembly with double layered cured epoxy. | Inadequate amount of epoxy leading to sensor corrosion due to moisture
AU Sensor-Sensor Cable. Cure epoxy in oven. Remove from oven and then apply second layer of epoxy. Then cure in oven again.. Oven (time and temperature), sensor cable with epoxy, Epotek H70E-2, dipping/glob topping fixture, dosing needle, cup. Sensor/cable assembly with double layered cured epoxy. | Only one layer of epoxy leading to sensor corrosion
AU Sensor-Sensor Cable. Cure epoxy in oven. Remove from oven and then apply second layer of epoxy. Then cure in oven again.. Oven (time and temperature), sensor cable with epoxy, Epotek H70E-2, dipping/glob topping fixture, dosing needle, cup. Sensor/cable assembly with double layered cured epoxy. | Sensor failure
AU Sensor-Sensor Cable. Dismantle sensor out of the fixture and insert sensor/cable assembly into dipping/glob-topping fixture.. Scissors (tongs), sensor/cable assembly, dipping/glob-topping fixture. sensor/cable assembly in dipping/glob-topping fixture. | Damage bonding between sensor/cable or damage sensor or damage sensor cable
AU Sensor-Sensor Cable. Dismantle sensor out of the fixture, cut upper part of PCB, and then insert sensor/cable assembly into dipping/glob-topping fixture.. Scissors (tongs), sensor/cable assembly, dipping/glob-topping fixture. sensor/cable assembly in dipping/glob-topping fixture. | Damage bonding between sensor/cable or damage sensor or damage sensor cable
AU Sensor-Sensor Cable. Mix epoxy (Epotek H70E-2). Apply the Glob Top epoxy with a small needle. Epotek H70E-2, dosing needle, cable/sensor assembly, dipping/glob topping fixture, scale, cup, spatula. Sensor/cable assembly with epoxy . | Inadequate amount of epoxy leading to sensor corrosion due to moisture
AU Sensor-Sensor Cable. Mix primer solution. Then apply to assembly.. Sensor cable assembly, N-hexane, primer, beaker, dipping glop-topping fixture, syringe. Primer applied to assembly. | Bond failure of silicone leading to moisture damaging sensor
AU Sensor-Sensor Cable. Mix solution of silicone glue Nusil MED 1511 with cyclohexane. (3 different ratios). Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet. Mixed silicone glue ratios (3x). | Bond failure of silicone leading to moisture damaging sensor
AU Sensor-Sensor Cable. Mix solution of silicone glue Nusil MED 1511 with cyclohexane. (3 different ratios). Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet. Mixed silicone glue ratios (3x). | Rough/uneven surface
AU Sensor-Sensor Cable. Mix solution of silicone glue Nusil MED 1511 with cyclohexane. (3 different ratios). Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet. Mixed silicone glue ratios (3x). | Silicone cracks leading to sensor failure
AU Sensor-Sensor Cable. Mix solution of silicone glue Nusil MED 1511 with cyclohexane. (3 different ratios). Cyclohexane, Silicone glue (Nusil MED 1511), beaker, scale, speed mixer, centrifuge, dosing needle, syringe, pasteur pipet. Mixed silicone glue ratios (3x). | Too viscous leading to cavity of sensor being filled leading to sensor failure
AU Sensor-Sensor Cable. Perform bonding for all four strands. Sensor cable and sensor in bond fixture, bonder (temperature, Ultra sonic power, Utra sonic time, Utra sonic weight, Ultra sonic frequency, Ultra sonic generator), humidity meter, Bonding wedge. Cable and sensor bonded together. | Bond failure leading to pump stop
AU Sensor-Sensor Cable. Perform functional test of sensor resistances, offset, ESD criteria. Multimeter, power supply, adaptor cable, sensor/cable assembly, Acceptance Testing excel sheet, Black box. Tested assembly. | Damage sensor
AU Sensor-Sensor Cable. Perform functional test of sensor resistances, offset, ESD criteria. Multimeter, power supply, adaptor cable, sensor/cable assembly, Acceptance Testing excel sheet, Black box. Tested assembly. | Unacceptable sensor accepted
AU Sensor-Sensor Cable. Remove sensor cable from packaging. Carefully fix sensor cable to the sensor bond fixture. Bond fixture, sensor cable, sensor, bonder, tweezer, packaging of the sensor cable. sensor cable and sensor in bond fixture. | Damage sensor cable/sensor cable
AU Sensor-Sensor Cable. Remove sensor with static safe tweezers from wafer or waffle-pack and insert sensor in the opened sensor pocket of the bond-fixture.. Static safe tweezers, bond-fixture, sensor, wafer or waffle-pack, bonder. Sensor in bond-fixture. | Damage sensor
AU Sensor-Sensor Cable. Setup bonder for sensor to the sensor cable. Bonder (temperature, Ultra sonic power, Utra sonic time, Utra sonic weight, Ultra sonic frequency, Ultra sonic generator), humidity meter, thermocouple, Vacuum, Height . Bonder properly setup. | ESD failure during bonding process
AU Sensor-Sensor Cable. Setup bonder for sensor to the sensor cable. Bonder (temperature, Ultra sonic power, Utra sonic time, Utra sonic weight, Ultra sonic frequency, Ultra sonic generator), humidity meter, thermocouple, Vacuum, Height . Bonder properly setup. | Sensor failure due to incorrect bond
AU Sensor-Sensor Cable. Storage of assembly unit in an ESD bag. Sensor assembly with silicone coating, ESD bag. Sensor assembly in ESD bag. | ESD leading to sensor failure
AU Sensor-Sensor Cable. Trim end of sensor cable behind the bond position at the pull linkage with a tweezers and remove excess. cable and sensor bonded together, tweezers, fixture needle. removed cable and sensor bonded together with excess trimmed. | Damage bonding between sensor/cable or damage sensor
AU Sensor-Sensor Cable. Trim end of sensor cable behind the bond position at the pull linkage with a tweezers and remove excess. cable and sensor bonded together, tweezers, fixture needle. removed cable and sensor bonded together with excess trimmed. | Damage bonding between sensor/cable or damage sesnor
AU Sensor-Sensor Cable. Trim end of sensor cable behind the bond position at the pull linkage with a tweezers and remove excess. cable and sensor bonded together, tweezers, fixture needle. removed cable and sensor bonded together with excess trimmed. | Sensor cables trimmed incorrect length leading to scrap
AU Sensor-Sensor Cable. Verify positioning of sensor and cable. Stripped end of cable should be above the pads.. Sensor cable and sensor in bond fixture, Fixture needle, color coding of cables. Cable positioned over sensor. | Improper sensor position leading to sensor failure
AU Stator. Align notch on insert with correct flattened surface on yoke. Yoke, insert. Insert correctly aligned with yoke. | Incorrectly aligned insert
AU Stator. Align the center litz to a solder pad with the pin in the holder of solder fixturing, then utilizing a tweezer to position the lids to the pads following the shape of the peak cap, then use the soldering iron solder all three lids to the pads. prepared insertion pin with bearing unit and stator in correct orientation, solder, solder iron (temperature), holding fixture, microscope, tweezers. Stator soldered/bonded to the bearing unit on the insertion pin with protuding litz. | Damage bearing unit/melt peak leading to pump failure (too high temp)
AU Stator. Align the center litz to a solder pad with the pin in the holder of solder fixturing, then utilizing a tweezer to position the lids to the pads following the shape of the peak cap, then use the soldering iron solder all three lids to the pads. prepared insertion pin with bearing unit and stator in correct orientation, solder, solder iron (temperature), holding fixture, microscope, tweezers. Stator soldered/bonded to the bearing unit on the insertion pin with protuding litz. | Damage soldering area (too high temp)
AU Stator. Align the center litz to a solder pad with the pin in the holder of solder fixturing, then utilizing a tweezer to position the lids to the pads following the shape of the peak cap, then use the soldering iron solder all three lids to the pads. prepared insertion pin with bearing unit and stator in correct orientation, solder, solder iron (temperature), holding fixture, microscope, tweezers. Stator soldered/bonded to the bearing unit on the insertion pin with protuding litz. | incomplete soldering (too low temperature)
AU Stator. Clean insertion pins with Isopropanol (IPA 99%) and air blow to dry. lint free wipe, IPA 99%, compressed air. cleaned insertion pin with no damage. | inconsistent wiping residual on pin leading damage of coil
AU Stator. Clean insertion pins with petroleum ether and air blow to dry. lint free wipe, petroleum ether, compressed air. cleaned insertion pin with no damage. | inconsistent wiping residual on pin leading damage of shrink tube
AU Stator. Clean insertion pins with scapel and visual inspect for damage. insertion pin, scapel, microscope. pre-cleaned and inspected insertion pin with no damage. | bent pin leading to rotor misalignment
AU Stator. Clean insertion pins with scapel and visual inspect for damage. insertion pin, scapel, microscope. pre-cleaned and inspected insertion pin with no damage. | Pin damage leading to damage of bearing cap
AU Stator. Clean insertion pins with scapel and visual inspect for damage. insertion pin, scapel, microscope. pre-cleaned and inspected insertion pin with no damage. | Pin damage leading to damage of shrink tube
AU Stator. Clean insertion pins with scapel and visual inspect for damage. insertion pin, scapel, microscope. pre-cleaned and inspected insertion pin with no damage. | Pin not cleaned leading to damage of the coil
AU Stator. Clean UV seals with swab with petroleum ether. Then let dry in fixture.. board fixture, swab, petroleum ether, insertion pin with bearing unit and shrink tube with cured loctite 3321. Cleaned UV seal on insertion pin with bearing unit and shrink tube. | epotek not adhering to UV leading to leak path for purge fluid due to delamination
AU Stator. Clean UV seals with swab with petroleum ether. Then let dry in fixture.. board fixture, swab, petroleum ether, insertion pin with bearing unit and shrink tube with cured loctite 3321. Cleaned UV seal on insertion pin with bearing unit and shrink tube. | Inconsistent seal leading to coil short from glucose during pump operation
AU Stator. Coil motor cable and fixate with tape. Assembly Unit Stator with with cable aligned with O'ring with trimmed shrink tube, tape. Assembly Unit Stator with with cable aligned with O'ring with trimmed shrink tube and taped motor cable. | Overhandling leading to damage soldering points
AU Stator. Cut shrink tube to length. razor blade, shrink tube, scale, cutting fixture. prepared insertion pin with cap and shrink tube. | Cut incorrect length (unable to utilize)
AU Stator. Cut shrink tube to length. Scissors, shrink tube, scale. shrink tube to correct length. | Cut incorrect length (unable to utilize)
AU Stator. Cut the protuding shrink tube on the upper end of the pin shaft using a scapel, visual inspect that release agent was applied properly. insertion pin with shrink tube applied, scapel, microscope. insertion pin with shrink tube trimmed at upper end. | Damage shrink tube insulation leading to leakage current
AU Stator. Cut the protuding shrink tube on the upper end of the pin shaft using a scapel, visual inspect that release agent was applied properly. insertion pin with shrink tube applied, scapel, microscope. insertion pin with shrink tube trimmed at upper end. | Damage shrink tube material unable to utilize
AU Stator. Cut the shrink tube down as close to the coil as possible. Assembly Unit Stator with with cable aligned with O'ring, scapel. Assembly Unit Stator with with cable aligned with O'ring with trimmed shrink tube. | Cut incorrect length (long) collide with mold inlay leading wrinkles in shrink tube
AU Stator. Cut the shrink tube down as close to the coil as possible. Assembly Unit Stator with with cable aligned with O'ring, scapel. Assembly Unit Stator with with cable aligned with O'ring with trimmed shrink tube. | Cut incorrect length (short) damage coil
AU Stator. Perform coil resistance check. Coil resistance checker, Assembly Unit Stator, resistance specification. Acceptable Assembly Unit Stator. | Reject good pumps
AU Stator. Place in board fixture in oven and dry. board fixture, insertion pin, oven (time and temperature). dry insertion pin with frekote. | Not dry leading to demoulding issues
AU Stator. Protruding litz on the unit, remove with scapel. Scapel, Stator soldered/bonded to the bearing unit on the insertion pin with protuding litz. Stator soldered/bonded to the bearing unit on the insertion pin with no protuding litz. | exposed litz leading to leakage current
AU Stator. Pull unit off the pin and visually inspect for burrs and remove any found with tweezers. assembled bearing unit on the insertion pin, tweezers, microscope. insertion pin and bearing unit separate with no burrs. | burrs not removed, leading to bearing blockage
AU Stator. Seal bearing unit with UV loctite 3321 between front cap and pin shaft, and then move shrink tube over bearing unit ensuring bearing is not shifted. block fixture, press tool, insertion pin with shrink applied and correct length, front cap, bearing, cap with four soldering pads. assembled bearing unit on the insertion pin with no burrs. | UV migrates into bearing leading to failure
AU Stator. Seal bearing unit with UV loctite 3321 between front cap and pin shaft, and then move shrink tube over bearing unit ensuring bearing is not shifted. Insertion pin with bearing unit and shrink tube, Loctite 3321, gloves. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied. | Inadequate seal leading to coil short from glucose during pump operation
AU Stator. Seal bearing unit with UV loctite 3321 between front cap and pin shaft, and then move shrink tube over bearing unit ensuring bearing is not shifted. Insertion pin with bearing unit and shrink tube, Loctite 3321, gloves. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied. | Inconsistent seal leading to coil short from glucose during pump operation
AU Stator. Seal bearing unit with UV loctite 3321 between front cap and pin shaft, and then move shrink tube over bearing unit ensuring bearing is not shifted. Insertion pin with bearing unit and shrink tube, Loctite 3321, gloves. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied. | UV migrates into bearing leading to failure
AU Stator. Seal needle side of pin and the bearing ring with Loctite 3321, then place in board fixture cure under UV light. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied, UV light, gloves, board fixture. insertion pin with bearing unit and shrink tube with cured loctite 3322. | Inadequate seal leading to coil short from glucose during pump operation
AU Stator. Seal needle side of pin and the bearing ring with Loctite 3321, then place in board fixture cure under UV light. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied, UV light, gloves, board fixture. insertion pin with bearing unit and shrink tube with cured loctite 3323. | Inadequate seal leading to coil short from glucose during pump operation
AU Stator. Seal needle side of pin and the bearing ring with Loctite 3321, then place in board fixture cure under UV light. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied, UV light, gloves, board fixture. insertion pin with bearing unit and shrink tube with cured loctite 3323. | Inconsistent seal leading to coil short from glucose during pump operation
AU Stator. Seal needle side of pin and the bearing ring with Loctite 3321, then place in board fixture cure under UV light. insertion pin with bearing unit and shrink tube with loctite 3321 seal applied, UV light, gloves, board fixture. insertion pin with bearing unit and shrink tube with cured loctite 3321. | UV migrates into bearing leading to failure
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke and litz. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Molding defects leading to scrap
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke and litz. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Unable to solder litz to pads
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke, litz, and shrink tub cutting edge. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Molding defects leading to scrap
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke, litz, and shrink tub cutting edge. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Molding housing depth incorrect leading to engagement issues between rotor shaft and bearing (purge line blockage)
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke, litz, and shrink tub cutting edge. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Molding housing depth incorrect leading to engagement issues between rotor shaft and bearing (reduced mechanical integrity of shaft and bearing connection)
AU Stator. Shift stator from the backside over the pin, with light turning movement, then align with yoke, litz, and shrink tub cutting edge. insetion pin with bearing unit and shrink tube, stator, alignment of yoke, litz, and stator. prepared insertion pin with bearing unit and stator in correct orientation. | Unable to solder litz to pads
AU Stator. Shrink the shrink tube to the pin utilizing a heat gun to form a smooth consistent layer, then visually inspect. heat gun, pin and needle holder, insertion pin with Frekote, shrink tube. insertion pin with shrink tube applied. | Damage shrink tube material/degredation unable to utilize
AU Stator. Shrink the shrink tube to the pin utilizing a heat gun to form a smooth consistent layer, then visually inspect. heat gun, pin and needle holder, insertion pin with Frekote, shrink tube. insertion pin with shrink tube applied. | Insufficient shrinkage, wrinkles leading unable assemble stator
AU Stator. Shrink the shrink tube to the pin utilizing a heat gun to form a smooth consistent layer, then visually inspect. heat gun, pin and needle holder, insertion pin with Frekote, shrink tube. insertion pin with shrink tube applied. | No shrinkage, wrinkles leading unable assemble stator
AU Stator. Start laser welder. Open inert gas supply. Measure laser energy. Load laser program. Laser program is executed automatically.. Laser welder, laser program, measurement device (mulitmeter + recepticle), inert gas, excel sheet to document energy measurements, pulse energy laser program. Welded insert to yoke. | Corrosion of weld seam leading to scrap
AU Stator. Start laser welder. Open inert gas supply. Measure laser energy. Load laser program. Laser program is executed automatically.. Laser welder, laser program, measurement device (mulitmeter + recepticle), inert gas, excel sheet to document energy measurements, pulse energy laser program. Welded insert to yoke. | Damage winding
AU Stator. Start laser welder. Open inert gas supply. Measure laser energy. Load laser program. Laser program is executed automatically.. Laser welder, laser program, measurement device (mulitmeter + recepticle), inert gas, excel sheet to document energy measurements, pulse energy laser program. Welded insert to yoke. | unable to weld components (scrap)
AU Stator. Unshorten litz from the cable and re-tin, then align cables to the right position and solder to pads. Scapel, cable, stator holder fixture, Stator soldered/bonded to the bearing unit on the insertion pin with no protuding litz, solder, solder iron (temperature). Assembly Unit Stator. | Damage soldering area (too high temp)
AU Stator. Unshorten litz from the cable and re-tin, then align cables to the right position and solder to pads. Scapel, cable, stator holder fixture, Stator soldered/bonded to the bearing unit on the insertion pin with no protuding litz, solder, solder iron (temperature). Assembly Unit Stator. | incomplete soldering (too low temperature)
AU Stator. Unshorten litz from the cable and re-tin, then align cables to the right position and solder to pads. Scapel, cable, stator holder fixture, Stator soldered/bonded to the bearing unit on the insertion pin with no protuding litz, solder, solder iron (temperature). Assembly Unit Stator. | reverse flow of pump
AU Stator. Using O-ring to fixate the litz of the cable to pin in the order that they were soldered. O'ring, Assembly Unit Stator. Assembly Unit Stator with with cable aligned with O'ring. | difficulty assembling units in molding cap
AU Stator. Utilizing a block fixture, place pin within fixture and then place ball bearing cap, ball bearing, then second cap, lastly press together to form bearing unit. . block fixture, press tool, insertion pin with shrink applied and correct length, front cap, bearing, cap with four soldering pads. assembled bearing unit on the insertion pin with no burrs. | Incorrect comprehsion force leading unable to assemble
AU Stator. Utilizing a block fixture, place pin within fixture and then place ball bearing cap, ball bearing, then second cap, lastly press together to form bearing unit. . block fixture, press tool, insertion pin with shrink applied and correct length, front cap, bearing, cap with four soldering pads. assembled bearing unit on the insertion pin with no burrs. | UV migrates into bearing leading to failure
AU Stator. Utilizing a block fixture, place pin within fixture and then place ball bearing cap, ball bearing, then second cap, lastly press together to form bearing unit. block fixture, press tool, insertion pin with shrink applied and correct length, front cap, bearing, cap with four soldering pads. assembled bearing unit on the insertion pin. | Incorrect comprehsion force leading unable to assemble
AU Stator. With gloves apply release agent (frekote 770NC) on shaft of pin execept last 5mm and then place in holding fixture. holding fixture, insertion pin, release agent (frekote 770NC), gloves, microscope, swab. insertion pin with frekote 770NC. | inconsitent application of release agent leading to pin separation after molding
AU Stator. With gloves apply release agent (frekote 770NC) on shaft of pin execept last 5mm and then place in holding fixture. holding fixture, insertion pin, release agent (frekote 770NC), gloves, microscope, swab. insertion pin with frekote 770NC. | Pin separation after molding
AU Stator. With gloves apply release agent (frekote 770NC) on shaft of pin execept last 5mm and then place in holding fixture. holding fixture, insertion pin, release agent (frekote 770NC), gloves, microscope, swab. insertion pin with frekote 770NC. | too little unable to remove insertion pin from motor housing (scrap)
AU Stator. With gloves apply release agent (frekote 770NC) on shaft of pin execept last 5mm and then place in holding fixture. holding fixture, insertion pin, release agent (frekote 770NC), gloves, microscope, swab. insertion pin with frekote 770NC. | too much leading to migration of release agent (scrap)
Back Loading Lumen. Cut guide lumen to length and attach label. Lumen label application fixture, backloading lumen cutter/flexible tubing cutter, tweexers, rounded end spatula, bar sealer . Backloading lumen with label in bag. | Damage lumen leading to break upon removal leading to pump stop
Back Loading Lumen. Cut guide lumen to length and attach label. Lumen label application fixture, backloading lumen cutter/flexible tubing cutter, tweexers, rounded end spatula, bar sealer . Backloading lumen with label in bag. | Damage lumen leading to scrap
Balancing Manifold Mfg Process (MFG-3034). Assemble Balancing Manifold. . ASSY, Balancing Manifold. | Balancing manifold not properly assembled
Balancing Solenoid Mfg Process (MFG-3032). Assemble Balancing Solenoid. . ASSY, Balancing Solenoid. | Balancing solenoid not properly assembled
Battery Bracket Assembly Mfg Process (MFG-3009). Assemble Battery Bracket. . Battery Holder Bracket Assembly. | Batteries are not assembled into bracket
Battery Bracket Assembly Mfg Process (MFG-3009). Select Batteries. . Battery Holder Bracket Assembly. | Batteries have mismatched voltages
Bearing to Sleeve (0550-7026). Apply adhesive into notch of proximal bearing. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted, adhesive, dispensing tip. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive. | Insufficient adhesive - fluid path out of motor
Bearing to Sleeve (0550-7026). Apply adhesive into notch of proximal bearing. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted, adhesive, dispensing tip. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive. | Insufficient adhesive - weak bond
Bearing to Sleeve (0550-7026). Apply adhesive into notch of proximal bearing. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted, adhesive, dispensing tip. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive. | too much adhesive - ingress into sleeve
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | Alignment pin damaged
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | ceramics damaged
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | Coil Damaged
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | Parts not aligned correctly
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | Parts not cleaned - insufficient bond strength
Bearing to Sleeve (0550-7026). Clean proximal bearing, ceramic sleeve, and bearing alignment pin. Proximal bearing, ceramic sleeve with coil, bearing alignment pin, petroleum ether, lint free wipes, swab. cleaned parts and fixture. | Parts not dry - insufficient bond strength
Bearing to Sleeve (0550-7026). Partially Mount proximal bearing onto bearing alignment pin with tweezers. Bearing alignment pin with ceramic sleeve coil mounted, proximal bearing, tweezers. Bearing alignment pin with ceramic sleeve and proximal bearing mounted. | ceramics damaged
Bearing to Sleeve (0550-7026). Partially Mount proximal bearing onto bearing alignment pin with tweezers. Bearing alignment pin with ceramic sleeve coil mounted, proximal bearing, tweezers. Bearing alignment pin with ceramic sleeve and proximal bearing mounted. | Proximal bearing too far onto pin - unable to apply adhesive into notch
Bearing to Sleeve (0550-7026). Partially Mount proximal bearing onto bearing alignment pin with tweezers. Bearing alignment pin with ceramic sleeve coil mounted, proximal bearing, tweezers. Bearing alignment pin with ceramic sleeve and proximal bearing mounted. | proximal bearing upside down - unable to glue
Bearing to Sleeve (0550-7026). Place in Sleeve Proximal bearing assembly in transport tray. Cured assembly, transport tray. Bearing to sleeve coil assembly. | Ceramic Sleeve Damaged
Bearing to Sleeve (0550-7026). Place in Sleeve Proximal bearing assembly in transport tray. Cured assembly, transport tray. Bearing to sleeve coil assembly. | Coil winding damaged
Bearing to Sleeve (0550-7026). Place in Sleeve Proximal bearing assembyl in transport tray. Cured assembly, transport tray. Bearing to sleeve coil assembly. | Ceramic Sleeve Damaged
Bearing to Sleeve (0550-7026). Place in Sleeve Proximal bearing assembyl in transport tray. Cured assembly, transport tray. Bearing to sleeve coil assembly. | Coil winding damaged
Bearing to Sleeve (0550-7026). Place in storage tray and cure. Cleaned uncured assembly on pin, storage tray, oven (time and temp). Bearing to sleeve coil assembly. | Bearing moves out of position - unable to assemble
Bearing to Sleeve (0550-7026). Place in storage tray and cure. Cleaned uncured assembly on pin, storage tray, oven (time and temp). Bearing to sleeve coil assembly. | Coil winding glued to Adhesive
Bearing to Sleeve (0550-7026). Place in storage tray and cure. Cleaned uncured assembly on pin, storage tray, oven (time and temp). Bearing to sleeve coil assembly. | Unable to assemble
Bearing to Sleeve (0550-7026). Remove alignement pin. Bearing to sleeve coil assembly. Bearing to sleeve coil assembly without alignment pin. | Ceramic Sleeve Damaged
Bearing to Sleeve (0550-7026). Remove alignement pin. Bearing to sleeve coil assembly. Bearing to sleeve coil assembly without alignment pin. | Coil winding damaged
Bearing to Sleeve (0550-7026). Remove alignment pin. Bearing to sleeve coil assembly. Bearing to sleeve coil assembly without alignment pin. | Ceramic Sleeve Damaged
Bearing to Sleeve (0550-7026). Remove alignment pin. Bearing to sleeve coil assembly. Bearing to sleeve coil assembly without alignment pin. | Coil winding damaged
Bearing to Sleeve (0550-7026). remove excess adhesive with swab. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted with adhesive - proximal bearing in place, swab. Cleaned uncured assembly on pin. | Coil Damaged
Bearing to Sleeve (0550-7026). Slide ceramic sleeve onto bearing alignment pin. Ceramic sleeve with coil, bearing alignment pin. Bearing alignment pin with ceramic sleeve mounted. | ceramics damaged
Bearing to Sleeve (0550-7026). Slide proximal bearing all the way to ceramic sleeve. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted with adhesive. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive - proximal bearing in place. | Alignment pin out of specification
Bearing to Sleeve (0550-7026). Slide proximal bearing all the way to ceramic sleeve. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted with adhesive. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive - proximal bearing in place. | not fully seated - assembly out of aligment - eccentricity
Bearing to Sleeve (0550-7026). Slide proximal bearing all the way to ceramic sleeve. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted with adhesive. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive - proximal bearing in place. | not fully seated - assembly out of alignment - eccentricity
Bearing to Sleeve (0550-7026). Slide proximal bearing all the way to ceramic sleeve. Bearing alignment pin with ceramic sleeve coil and proximal bearing mounted with adhesive. Bearing alignment pin with ceramic sleeve and proximal bearing mounted with adhesive - proximal bearing in place. | not fully seated - unable to assemble
Blower Branch Mfg Process (MFG-3030). Assemble Blower Branch. . ASSY, Blower Branch. | Blower branch not properly assembled
Blower Check Valve Mfg Process (MFG-3029). Assemble Blower Check Valve. . ASSY, Blower Check Valve. | Blower Check valve not properly assembled
Bottom Shell Mfg Process (MFG-3004). Assemble Bottom Shell. . ASSY, Bottom Shell. | Improper routing of IEC cable
Bottom Shell Mfg Process (MFG-3004). Assemble Bottom Shell. . ASSY, Bottom Shell. | Power Supply and Filter Insulating material improperly installed
Bottom Shell Mfg Process (MFG-3004). Attach Caster Grounding Plate. . ASSY, Bottom Shell. | Casters are not grounded
Bottom Shell Mfg Process (MFG-3004). Place Bottom bracket into bottom shell. . ASSY, Bottom Shell. | Casters are not grounded
Cannula Assembly (0550-7039). Anneal cannula for 4 hours min at 70C. cured cannula, oven. annealed cannula. | Shape out of specification
Cannula Assembly (0550-7039). Anneal raw material tubing. raw tubing material. annealed tubing (removed moisture and stress release). | not annealed tubing
Cannula Assembly (0550-7039). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Biocompatibility comprimised
Cannula Assembly (0550-7039). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Biocompatibility compromised
Cannula Assembly (0550-7039). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Bond strength comprimised
Cannula Assembly (0550-7039). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Bond strength compromised
Cannula Assembly (0550-7039). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. mandrel with cleaned ends. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly (0550-7039). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. more than just ends cleaned. | too little acmos unable to remove from mandrel
Cannula Assembly (0550-7039). Cool cannula. Thermoformed First Layer with heat shrink, time at room temp. Cooled first layer with heat shrink. | cosmetic defects
Cannula Assembly (0550-7039). let cool. Thermoformed cannula with heat shrink tubing, time at room temp. cooled Thermoformed cannula with heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (0550-7039). let cool. Thermoformed first and reinforcement layers with heath shrink tubing, time at room temp. cooled Thermoformed first and reinforcement layers with heath shrink tubing. | cosmetic defects
Cannula Assembly (0550-7039). let dry in oven. first layer with reinforcement tubing, time temp. dried first layer with reinforcement tubing. | Material deterioration leads to delamination
Cannula Assembly (0550-7039). let dry in oven. first layer with reinforcement tubing, time temp. dried first layer with reinforcement tubing. | Material does not re-flow
Cannula Assembly (0550-7039). Load cannula onto coil machine. thermoformed first and reinforcement layer, coil machine. thermoformed first and reinforcement layer loaded on machine. | Loaded incorrectly
Cannula Assembly (0550-7039). load coil onto machine. thermoformed first and reinforcement layer, coil machine. coil loaded on machine. | Coil loaded incorrectly
Cannula Assembly (0550-7039). lubricate inner layer with HFE. reinforcement tubing, HFE. lubricated reinforcement layer. | Unable to continue with build
Cannula Assembly (0550-7039). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos
Cannula Assembly (0550-7039). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly (0550-7039). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos
Cannula Assembly (0550-7039). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | too much acmos dried off unable to remove from mandrel
Cannula Assembly (0550-7039). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly (0550-7039). Place bushings on ends of cannula. trimmed cannula, bushings. cannula with bushings. | deformed transition area
Cannula Assembly (0550-7039). place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (0550-7039). place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layer with outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | bubbles in cannula
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layer with outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layer with outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | bubbles in cannula
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (0550-7039). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly (0550-7039). place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Material deterioration leads to delamination
Cannula Assembly (0550-7039). place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Material does not re-flow
Cannula Assembly (0550-7039). place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Unable to continue with assembly
Cannula Assembly (0550-7039). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, time, temp. Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Cannula Assembly (0550-7039). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, time, temp. Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Cannula Assembly (0550-7039). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too cold
Cannula Assembly (0550-7039). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too hot - yields degraded cannula
Cannula Assembly (0550-7039). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too long of time - yields degraded cannula
Cannula Assembly (0550-7039). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too short or time
Cannula Assembly (0550-7039). Remove cannula from mandrel. Cured cannula on bend mandrel. cured cannula. | unable to assemble
Cannula Assembly (0550-7039). Remove cannula from mandrel. trimmed cannula on mandrel. trimmed cannula. | ID roughness
Cannula Assembly (0550-7039). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | cut cannula
Cannula Assembly (0550-7039). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | Unable to continue with assembly
Cannula Assembly (0550-7039). remove heat shrink. cooled Thermoformed cannula with heath shrink tubing, razor. Thermoformed cannula. | razor cuts cannula
Cannula Assembly (0550-7039). remove heat shrink. cooled Thermoformed first and reinforcement layers with heath shrink tubing, razor. thermoformed first and reinforcement layer. | cut cannula
Cannula Assembly (0550-7039). remove one bushing. cannula with bushings warmed. Cannula with one bushing. | unable to assemble
Cannula Assembly (0550-7039). run program. thermoformed first and reinforcement layer loaded on machine, coil loaded on machine, coil program. thermoformed first and reinforcement layer with coil. | Program corrupted or incorrect
Cannula Assembly (0550-7039). select program. thermoformed first and reinforcement layer, coil machine. coil machine with program selected. | wrong program selected
Cannula Assembly (0550-7039). slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Cannula Assembly (0550-7039). slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . cannula, bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Cannula Assembly (0550-7039). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (0550-7039). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (0550-7039). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | slid on to incorrect location
Cannula Assembly (0550-7039). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | Unable to thermoform
Cannula Assembly (0550-7039). slide outer layer. thermoformed first and reinforcement with coil, outer layer. thermoformed first and reinforcement with coil with outer layer. | cannula does not recover from kinking
Cannula Assembly (0550-7039). slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | cannula does not recover from kinking
Cannula Assembly (0550-7039). slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | OD too large
Cannula Assembly (0550-7039). soften cannula in oven. cannula with bushings, oven (time, temp). cannula with bushings warmed. | delamination of cannula
Cannula Assembly (0550-7039). Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material deterioration leads to delamination
Cannula Assembly (0550-7039). Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material does not re-flow
Cannula Assembly (0550-7039). Trim cannula. Cannula, Cannula Trimming Fixture (SPEC-2474). trimmed cannula. | Too long - step at transition to PH or IC
Cannula Assembly (0550-7039). Trim cannula. Cannula, Cannula Trimming Fixture (SPEC-2474). trimmed cannula. | Too short - weak bond
Cannula Assembly (0550-7039). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | coil trimmed too short
Cannula Assembly (0550-7039). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | Improper step formed at PH end
Cannula Assembly (0550-7039). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | reduced bond strength
Cannula Assembly (0550-7039). use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | Adhesive not applied
Cannula Assembly (0550-7039). use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | too much Adhesive applied leads to delamination
Cannula Assembly (10002389). Anneal cannula for 8 hours min at 70C. cured cannula, oven. annealed cannula. | Shape out of specification
Cannula Assembly (10002389). Anneal raw material tubing. raw tubing material. annealed tubing (removed moisture and stress release). | not annealed tubing
Cannula Assembly (10002389). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Biocompatibility comprimised
Cannula Assembly (10002389). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Bond strength comprimised
Cannula Assembly (10002389). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. mandrel with cleaned ends. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly (10002389). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. more than just ends cleaned. | too little acmos unable to remove from mandrel
Cannula Assembly (10002389). Cool cannula. Thermoformed assembly with heat shrink, time at room temp. Cooled assembly with heat shrink. | cosmetic defects
Cannula Assembly (10002389). Cool cannula. Thermoformed First Layer with heat shrink, time at room temp. Cooled first layer with heat shrink. | cosmetic defects
Cannula Assembly (10002389). Cool cannula. Thermoformed First Layers with heat shrink, time at room temp. Cooled first layers with heat shrink. | cosmetic defects
Cannula Assembly (10002389). End Forming. trimmed cannula off mandrel. end formed cannula. | ID too small
Cannula Assembly (10002389). let cool. Thermoformed cannula with heat shrink tubing, time at room temp. cooled Thermoformed cannula with heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (10002389). load coil onto machine. thermoformed first layers, coil machine. coil loaded on machine. | Coil loaded incorrectly
Cannula Assembly (10002389). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos
Cannula Assembly (10002389). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos
Cannula Assembly (10002389). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | too much acmos dried off unable to remove from mandrel
Cannula Assembly (10002389). Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly (10002389). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | bubbles in cannula
Cannula Assembly (10002389). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (10002389). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly (10002389). place in oven. dried first/reinforcement layer with outer tubing and double heat shrink. Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Cannula Assembly (10002389). place in oven. dried first/reinforcement layer with outer tubing and double heat shrink. Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Cannula Assembly (10002389). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, time, temp. Thermoformed cannula with heat shrink tubing. | incorrect pitch
Cannula Assembly (10002389). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, time, temp. Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Cannula Assembly (10002389). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, time, temp. Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Cannula Assembly (10002389). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too cold
Cannula Assembly (10002389). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too hot - yields degraded cannula
Cannula Assembly (10002389). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too long of time - yields degraded cannula
Cannula Assembly (10002389). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, time, temp. thermoformed first and reinforcement coil. | too short or time
Cannula Assembly (10002389). place second heat shrink. dried first/reinforcement layer with outer tubing and heat shrink, second heat shrink. dried first/reinforcement layer with outer tubing and double heat shrink. | bubbles in cannula
Cannula Assembly (10002389). place second heat shrink. dried first/reinforcement layer with outer tubing and heat shrink, second heat shrink. dried first/reinforcement layer with outer tubing and double heat shrink. | Improper step formed at PH end
Cannula Assembly (10002389). place second heat shrink. dried first/reinforcement layer with outer tubing and heat shrink, second heat shrink. dried first/reinforcement layer with outer tubing and double heat shrink. | Unable to continue with assembly
Cannula Assembly (10002389). pre-shrink heat shrink. mandrel with first layer and heat shrink tubing, time, temperature. mandrel with first layer and heat shrink tubing. | surface imperfections (bubbles)
Cannula Assembly (10002389). Remove cannula from mandrel. trimmed cannula on mandrel. trimmed cannula. | ID roughness
Cannula Assembly (10002389). remove heat shrink. Cooled assembly with heat shrink, razor. cooled assembly. | cut cannula
Cannula Assembly (10002389). remove heat shrink. Cooled assembly with heat shrink, razor. cooled assembly. | Unable to continue with assembly
Cannula Assembly (10002389). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | cut cannula
Cannula Assembly (10002389). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | Unable to continue with assembly
Cannula Assembly (10002389). remove heat shrink. Cooled first layers with heat shrink, razor. cooled first layers. | cut cannula
Cannula Assembly (10002389). remove heat shrink. Cooled first layers with heat shrink, razor. cooled first layers. | Unable to continue with assembly
Cannula Assembly (10002389). remove heat shrink. cooled Thermoformed cannula with heath shrink tubing, razor. Thermoformed cannula. | razor cuts cannula
Cannula Assembly (10002389). run program. thermoformed first and reinforcement layer loaded on machine, coil loaded on machine, coil program. thermoformed first and reinforcement layer with coil. | Program corrupted or incorrect
Cannula Assembly (10002389). select program. thermoformed first and reinforcement layer, coil machine. coil machine with program selected. | wrong program selected
Cannula Assembly (10002389). slide filler layers over inner layer. cooled first layer, filler layers. mandrel with first layers and filler layers. | slid on to incorrect location
Cannula Assembly (10002389). slide filler layers over inner layer. cooled first layer, filler layers. mandrel with first layers and filler layers. | Unable to thermoform
Cannula Assembly (10002389). Slide Heat Shrink tubing. cooled first layer, filler layers, heat shrink. mandrel with first layers and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (10002389). Slide Heat Shrink tubing. cooled first layer, filler layers, heat shrink. mandrel with first layers and heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (10002389). Slide Heat Shrink tubing. mandrel with assembly, heat shrink tubing. mandrel assembly and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (10002389). Slide Heat Shrink tubing. mandrel with assembly, heat shrink tubing. mandrel assembly and heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (10002389). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly (10002389). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Unable to continue with assembly
Cannula Assembly (10002389). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | slid on to incorrect location
Cannula Assembly (10002389). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | Unable to thermoform
Cannula Assembly (10002389). slide outer layer. thermoformed first and reinforcement with coil, outer layer. thermoformed first and reinforcement with coil with outer layer. | cannula does not recover from kinking
Cannula Assembly (10002389). Slide reinforcement layers over coil. coil on cannula, annealed tubing. coil on cannula with with reinforcement layers. | slid on to incorrect location
Cannula Assembly (10002389). Slide reinforcement layers over coil. coil on cannula, annealed tubing. coil on cannula with with reinforcement layers. | Unable to thermoform
Cannula Assembly (10002389). Thermoform cannula. mandrel with assembly and heat shrink tubing, time, temperature. Thermoformed assembly with heat shrink. | Material deterioration leads to delamination
Cannula Assembly (10002389). Thermoform cannula. mandrel with assembly and heat shrink tubing, time, temperature. Thermoformed assembly with heat shrink. | Material does not re-flow
Cannula Assembly (10002389). Thermoform cannula. mandrel with assembly and heat shrink tubing, time, temperature. Thermoformed assembly with heat shrink. | stiffness out of spec
Cannula Assembly (10002389). Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material deterioration leads to delamination
Cannula Assembly (10002389). Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material does not re-flow
Cannula Assembly (10002389). Thermoform cannula. mandrel with first layers and heat shrink tubing, time, temperature. Thermoformed First Layers with heat shrink. | Material deterioration leads to delamination
Cannula Assembly (10002389). Thermoform cannula. mandrel with first layers and heat shrink tubing, time, temperature. Thermoformed First Layers with heat shrink. | Material does not re-flow
Cannula Assembly (10002389). Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too long - step at transition to PH or IC
Cannula Assembly (10002389). Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too short - weak bond
Cannula Assembly (10002389). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | Improper step formed at PH end
Cannula Assembly (10002389). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | reduced tensile strength
Cannula Assembly (10002389). use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | Adhesive not applied
Cannula Assembly (10002389). use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | too much Adhesive applied leads to delamination
Cannula Assembly. Add polyeurethane 3mm layer to transitions areas at both ends and then place heat shrink over length. Cannula with 2 layers, 3mm long polyeurethane, heat shrink. Cannula with polyeurethane layer at both ends. | Weak bond at transition area/weak reinforcement area
Cannula Assembly. Anneal cannula for 4 hours min at 70C. cured cannula, oven. annealed cannula. | Shape out of specification
Cannula Assembly. Anneal raw material tubing. raw tubing material. annealed tubing (removed moisture and stress release). | not annealed tubing
Cannula Assembly. Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Biocompatibility comprimised
Cannula Assembly. Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Bond strength comprimised
Cannula Assembly. Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. mandrel with cleaned ends. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly. Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. more than just ends cleaned. | too little acmos unable to remove from mandrel
Cannula Assembly. Cool cannula. Thermoformed First Layer with heat shrink, time at room temp. Cooled first layer with heat shrink. | cosmetic defects
Cannula Assembly. let cool. block fixture, thermoformed cannula with heat shrink tubing. cooled assembly. | Unable to continue with assembly
Cannula Assembly. let cool. Thermoformed cannula with heat shrink tubing, time at room temp. cooled Thermoformed cannula with heat shrink tubing. | Unable to continue with assembly
Cannula Assembly. let cool. Thermoformed first and reinforcement layers with heath shrink tubing, time at room temp. cooled Thermoformed first and reinforcement layers with heath shrink tubing. | cosmetic defects
Cannula Assembly. let dry in oven. first layer with reinforcement tubing, time temp. dried first layer with reinforcement tubing. | Material deterioration leads to delamination
Cannula Assembly. let dry in oven. first layer with reinforcement tubing, time temp. dried first layer with reinforcement tubing. | Material does not re-flow
Cannula Assembly. Load cannula onto coil machine. thermoformed first and reinforcement layer, coil machine. thermoformed first and reinforcement layer loaded on machine. | Loaded incorrectly
Cannula Assembly. load coil onto machine. thermoformed first and reinforcement layer, coil machine. coil loaded on machine. | Coil loaded incorrectly
Cannula Assembly. lubricate inner layer with HFE. reinforcement tubing, HFE. lubricated reinforcement layer. | Unable to continue with build
Cannula Assembly. Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos
Cannula Assembly. Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly. Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos
Cannula Assembly. Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | too much acmos dried off unable to remove from mandrel
Cannula Assembly. Mandrel Preparation. lubricated mandrel, time, temperature. lubricated mandrel with carrier solvent removed. | Too much acmos on inner layer causes double step on ID of cannula
Cannula Assembly. Place bushings on ends of cannula. trimmed cannula, bushings. cannula with bushings. | deformed transition area
Cannula Assembly. place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly. place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly. place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | bubbles in cannula
Cannula Assembly. place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly. place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Unable to continue with assembly
Cannula Assembly. place in oven. Cannula with polyeurethane layer at both ends, block fixture, oven (time and temperature. Thermoformed cannula with heat shrink tubing. | Weak bond at transition area/weak reinforcement area
Cannula Assembly. place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Incompete forming of layers (bubbles/moving coil)
Cannula Assembly. place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage during insertion process
Cannula Assembly. place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage/kinking during insertion process
Cannula Assembly. place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Material does not re-flow
Cannula Assembly. place in oven. dried first layer with reinforcement tubing and heat shrink, time, temp. Thermoformed first and reinforcement layers with heath shrink tubing. | Unable to continue with assembly
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and double heat shrink, oven (time and temperature). Thermoformed cannula with heat shrink tubing. | Incompete forming of layers (bubbles/moving coil)
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and double heat shrink, oven (time and temperature). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and double heat shrink, oven (time and temperature). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage during insertion process
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and double heat shrink, oven (time and temperature). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage/kinking during insertion process
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and double heat shrink, oven (time and temperature). Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and heat shrink, Oven (time, temp). Thermoformed cannula with heat shrink tubing. | Incompete forming of layers (bubbles/moving coil)
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and heat shrink, Oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and heat shrink, Oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage during insertion process
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and heat shrink, Oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination/material too soft leading to damage/kinking during insertion process
Cannula Assembly. place in oven. dried first/reinforcement layer with outer tubing and heat shrink, Oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Cannula Assembly. place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too cold
Cannula Assembly. place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too hot - yields degraded cannula
Cannula Assembly. place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too long of time - yields degraded cannula
Cannula Assembly. place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too short or time
Cannula Assembly. place second heat shrink. dried first/reinforcement layer with outer tubing and heat shrink, second heat shrink. dried first/reinforcement layer with outer tubing and double heat shrink. | Improper step formed at PH end
Cannula Assembly. Remove cannula from mandrel. Cured cannula on bend mandrel. cured cannula. | unable to assemble
Cannula Assembly. Remove cannula from mandrel. trimmed cannula on mandrel. trimmed cannula. | ID roughness
Cannula Assembly. remove heat shrink and then measure. cooled Thermoformed cannula with heath shrink tubing, razor. Thermoformed cannula. | razor cuts cannula
Cannula Assembly. remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | cut cannula
Cannula Assembly. remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | Unable to continue with assembly
Cannula Assembly. remove heat shrink. cooled Thermoformed cannula with heath shrink tubing, razor. Thermoformed cannula. | razor cuts cannula
Cannula Assembly. remove heat shrink. cooled Thermoformed first and reinforcement layers with heath shrink tubing, razor. thermoformed first and reinforcement layer. | cut cannula
Cannula Assembly. remove one bushing. cannula with bushings warmed. Cannula with one bushing. | unable to assemble
Cannula Assembly. run program. thermoformed first and reinforcement layer loaded on machine, coil loaded on machine, coil program. thermoformed first and reinforcement layer with coil. | Program corrupted or incorrect
Cannula Assembly. select program. thermoformed first and reinforcement layer, coil machine. coil machine with program selected. | wrong program selected
Cannula Assembly. slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Cannula Assembly. slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . cannula, bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Cannula Assembly. Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Cannula Assembly. Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Unable to continue with assembly
Cannula Assembly. Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | slid on to incorrect location
Cannula Assembly. Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | Unable to thermoform
Cannula Assembly. slide outer layer. thermoformed first and reinforcement with coil, outer layer. thermoformed first and reinforcement with coil with outer layer. | cannula does not recover from kinking
Cannula Assembly. slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | cannula does not recover from kinking
Cannula Assembly. slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | OD too large
Cannula Assembly. soften cannula in oven. cannula with bushings, oven (time, temp). cannula with bushings warmed. | delamination of cannula
Cannula Assembly. Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material deterioration leads to delamination
Cannula Assembly. Thermoform cannula. mandrel with first layer and heat shrink tubing, time, temperature. Thermoformed First Layer with heat shrink. | Material does not re-flow
Cannula Assembly. Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too long - step at transition to PH or IC
Cannula Assembly. Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too short - weak bond
Cannula Assembly. trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | coil trimmed too short
Cannula Assembly. trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | Improper step formed at PH end
Cannula Assembly. trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | reduced bond strength
Cannula Assembly. use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | Adhesive not applied
Cannula Assembly. use Adhesive to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, Adhesive. thermoformed first and reinforcement layer with trimmed and secured coil. | too much Adhesive applied leads to delamination
Cannula Bent. Clean cannula in utltrasonic bath with Isporponal and water. Then dry with pressurized air.. Ultrasonic bath (time and temperature), Isproponal/Water mixture (1:1), bent cannula assembly, Air pressure. Cleaned cannula assembly. | debry still left on component
Cannula Bent. Then place into oven to remove remaining moisture. Then perform final inspection of unit and trimming.. cleaned cannula assembly, oven (time and temperature), THF solution, clean room swab. Final bent cannula assembly. | debry still left on component
Cannula Bent. Then place into oven to remove remaining moisture. Then perform final inspection of unit and trimming.. cleaned cannula assembly, oven (time and temperature), THF solution, clean room swab. Final bent cannula assembly. | Residual moisture still left in material leading to bond failure
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Bend angle incorrect
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Damage from debry on bending mandrel
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Deformation at bonding areas leading unable to assemble
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Deformation of surface (bubbles)
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Deformation of surface (impact structural)
Cannula Bent. Visually inspect cannula. Place cannula on bending mandrel. Then place in oven. Then place in water bath. Once cool perform visual inspection and dimensional check. (5.0 only). Straight cannula, bending mandrel, Oven (time and temperature), water bath, drawing template. Bent cannula with correct dimension and visually acceptable. | Stretching cause wall thickness to change (too thin) leading loss of structural integrity
Cannula Straight. Clean hydex rings and dry units. Isopropanol 99% and pyrogen free water (4:1), utlrasonic bath, Oven (time and temperature). Clean components. | reduced adhesion of taxen to hydex outer surface
Cannula Straight. Cut cannula to length. Remove cannula from mandrel and store. Then perform visual inspection.. Scapel, cannula on mandrel, oven (time and temperature). stored straight cannula. | damage to inner surface of cannula
Cannula Straight. Cut cannula to length. Remove cannula from mandrel and store. Then perform visual inspection.. Scapel, cannula on mandrel, oven (time and temperature). stored straight cannula. | Damage to outer surface of cannula
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with 2nd PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 3rd layer of PU. | incorrect thickness of cannula leading to integrity issues
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 2nd layer of PU. | Bubbles and voids on cannula surface
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 3rd layer of PU. | Bubbles and voids on cannula surface
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 2nd layer of PU. | Cannula does not flex as intended
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 3rd layer of PU. | Cannula does not flex as intended
Cannula Straight. Load mandrel into dispenser and dispense second layer of PU. Place into IR Oven. Then store in drying oven.. cannula mandrel with PU layer and nitinol wire, dispensing machine. IR oven (time and temperature), IR oven program, Oven (time and temperature). cannula mandrel with 2nd layer of PU. | incorrect thickness of cannula leading to integrity issues
Cannula Straight. Perform dimensional diameter inspection. cannula mandrel with 3rd layer of PU, ICL, ruler. Inspected cannula on mandrel. | OD not inspected correctly leading component being too flexible/flimsy
Cannula Straight. Perform dimensional diameter inspection. cannula mandrel with 3rd layer of PU, ICL, ruler. Inspected cannula on mandrel. | OD not inspected correctly leading not flexible/too stiff
Cannula Straight. Restart the machine and load with correct program. Add nitinol coil and mandrel with first PU layer. Run program to spool coil onto mandrel. Then cut nitinol wire. Then remove mandrel from machine store on a board.. 1st layer of PU dispensed on cannula mandrel with hydex, ninitol wire, dispensing machine, board fixture. cannula mandrel with PU layer and nitinol wire. | distance between turns too spaced (weak integrity of component)
Cannula Straight. Restart the machine and load with correct program. Add nitinol coil and mandrel with first PU layer. Run program to spool coil onto mandrel. Then cut nitinol wire. Then remove mandrel from machine store on a board.. 1st layer of PU dispensed on cannula mandrel with hydex, ninitol wire, dispensing machine, board fixture. cannula mandrel with PU layer and nitinol wire. | distance between turns too tight (too stiff component)
Cannula Straight. Store in oven until ready for use. Cannula mandrel with hydex rings, Oven (temperature). hydex rings on mandrel. | Damage hydex rings at high temperatures
Cannula Straight. Transport mandrel back to cleanroom, then disassemble mandrel, visually inspect hydex ring for burrs, orientate hydex rings unto mandrel so the small diameters each other on the ends, then re-assemble mandrel and secure with screw. Last ensure rings are shifted to the outside of the mandrels. Cannula mandrel, hydex rings. hydex rings on mandrel. | Unable to run cannula dispensing process
Cannula Straight. Turn on ventilation system, manually apply ACMOS 70-2406, then place it in oven. ACMOS 70-2406, Oven (time and temperature), ventiliation system, gloves, cannula mandrel. Cannula mandrel with release agent. | unable to remove cannula
Cannula Straight. Turn on/setup equipment, clean out PU from previous process and ensure fresh PU is available. Adjust expected pressure, then dispense first PU layer onto cannula mandrel. Then let dry for 2 hrs oven.. Cannual mandrel, dispensing machine, IR oven (time and temperature), IR oven program, computer, dispenser, Oven (time and Temperature). 1st layer of PU dispensed on cannula mandrel with hydex. | Cannula does not flex as intended
Cannula Straight. Turn on/setup equipment, clean out PU from previous process and ensure fresh PU is available. Adjust expected pressure, then dispense first PU layer onto cannula mandrel. Then let dry for 2 hrs.. Cannual mandrel, dispensing machine, IR oven (time and temperature), IR oven program, computer, dispenser, Oven (time and Temperature). 1st layer of PU dispensed on cannula mandrel with hydex. | Bubbles and voids on cannula surface
Cannula Straight. Turn on/setup equipment, clean out PU from previous process and ensure fresh PU is available. Adjust expected pressure, then dispense first PU layer onto cannula mandrel. Then let dry for 2 hrs.. Cannual mandrel, dispensing machine, IR oven (time and temperature), IR oven program, computer, dispenser, Oven (time and Temperature). 1st layer of PU dispensed on cannula mandrel with hydex. | Cannula does not flex as intended
Cannula Straight. Turn on/setup equipment, clean out PU from previous process and ensure fresh PU is available. Adjust expected pressure, then dispense first PU layer onto cannula mandrel. Then let dry for 2 hrs.. Cannual mandrel, dispensing machine, IR oven (time and temperature), IR oven program, computer, dispenser, Oven (time and Temperature). 1st layer of PU dispensed on cannula mandrel with hydex. | Contamination of remaining PU impact integrity of cannula
Cannula Straight. Turn on/setup equipment, clean out PU from previous process and ensure fresh PU is available. Adjust expected pressure, then dispense first PU layer onto cannula mandrel. Then let dry for 2 hrs.. Cannual mandrel, dispensing machine, IR oven (time and temperature), IR oven program, computer, dispenser, Oven (time and Temperature). 1st layer of PU dispensed on cannula mandrel with hydex. | incorrect thickness of cannula leading to integrity issues
Cannula Straight. Visual inspect cannula mandrel for damages, then clean mandrel with petroleum ether, then place on plate to transport to grey room. Cannula mandrel, petroleum ether, wipes, mandrel transport plate. Cleaned components. | Contamination of product
Cartridge Cap THF Bonding. Actuate fixture and bond assemblies together. Then remove the pump housing after pressing process and position it on a carrier to dry. assembly in THF press fixture, THF press fixture, retainer plate. Glass Cylinder & Pump Body Assembly. | damage glass cylinder leading to leak
Cartridge Cap THF Bonding. Actuate fixture and bond assemblies together. Then remove the pump housing after pressing process and position it on a carrier to dry. assembly in THF press fixture, THF press fixture, retainer plate. Glass Cylinder & Pump Body Assembly. | weak bond leading to leak
Cartridge Cap THF Bonding. Adjust pressure regulator to acceptable pressure on the THF press fixture. THF press fixture. setup press fixture. | cap bond inadequate leading purge leak
Cartridge Cap THF Bonding. Clean inner surface of glass cylinder. retainer plate, swab, petroleum ether, glass cylinder assembly. cleaned components. | blockage of piston due to particles
Cartridge Cap THF Bonding. Clean inner surface of glass cylinder. retainer plate, swab, petroleum ether, glass cylinder assembly. cleaned components. | contamination from particles
Cartridge Cap THF Bonding. Clean inner surface of glass cylinder. retainer plate, swab, petroleum ether, glass cylinder assembly. cleaned components. | particles/fibers on product
Cartridge Cap THF Bonding. Clean inner surface of glass cylinder. retainer plate, swab, petroleum ether, glass cylinder assembly. cleaned components. | residual particles damaging piston leading to leaks
Cartridge Cap THF Bonding. Compress both parts directly together in the correct position. glasss cylinder, cartridge cap, THF reservoirs, gloves, cartridge cap holder. pressed together assembly. | incorrect position of components leading to leak
Cartridge Cap THF Bonding. Compress both parts directly together in the correct position. glasss cylinder, cartridge cap, THF reservoirs, gloves, cartridge cap holder. pressed together assembly. | poor enagagement leading to leak
Cartridge Cap THF Bonding. Place glass cylinder assembly and cartridge cap into THF reservoirs, at the same time. glasss cylinder, cartridge cap, THF reservoirs, gloves, cartridge cap holder. glass cylinder and cartridge in glass holders. | poor engagement leading to leaks
Cartridge Cap THF Bonding. Place glass cylinder assembly and cartridge cap into THF reservoirs, at the same time. glasss cylinder, cartridge cap, THF reservoirs, gloves, cartridge cap holder. glass cylinder and cartridge in glass holders. | too much engagement leading to occlusion (unable to pass case start)
Cartridge Cap THF Bonding. Remove assembly from holder and place into THF press fixture. pressed together assembly, THF press fixture. assembly in THF press fixture. | damage bond leading to leak
Caster Installation Assy Mfg Process (MFG-3024). Secure casters to bottom of Handle Installation Assy. . ASSY, Caster Installation. | Casters are not secured
Caster Installation Assy Mfg Process (MFG-3024). Verify caster locks engage. . ASSY, Caster Installation. | casters can not lock
Caster Installation Assy Mfg Process (MFG-3024). Verify casters rotate freely. . ASSY, Caster Installation. | casters can not rotate freely
Catheter and Motor Installation (0550-7030). Bond motor housing to rotor. AU MH , AU Rotor PH - cleaned with PEEK sleeve, adhesive, dosing needles. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued. | Not fully seated - reduced bond strength
Catheter and Motor Installation (0550-7030). Bond motor housing to rotor. AU Pump Motor. AU Pump Motor - cooled (PEEK sleeve removed). | channel misaligned - unable to assemble
Catheter and Motor Installation (0550-7030). Check dry Rotor fit. AU MH , AU Rotor PH - cleaned . AU MH , AU Rotor PH - cleaned with. | Damage to magnet coating
Catheter and Motor Installation (0550-7030). Check dry Rotor fit. AU MH , AU Rotor PH - cleaned with. AU MH , AU Rotor PH - cleaned with. | Unable to assemble
Catheter and Motor Installation (0550-7030). Clean AU Motor housing and AU Rotor Pump Housing. AU MH , AU Rotor PH, swabs, petroleum ether. AU MH , AU Rotor PH - cleaned. | Insufficient bond strength between AU Rotor and Motor Housing
Catheter and Motor Installation (0550-7030). clean excess adhesive. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued, dry swabs. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued and cleaned. | unable to assemble
Catheter and Motor Installation (0550-7030). Cool . AU Pump Motor. AU Pump Motor - cooled. | concentricity comprimised
Catheter and Motor Installation (0550-7030). Cool. AU motor with straightened purge lumen and motor cable. AU motor with straightened purge lumen and motor cable cooled. | Purge lumen kinks - high purge pressure
Catheter and Motor Installation (0550-7030). encapsulation of shied with UV glue. AU Pump Motor, adhesive, dispenser, UV source. Motor cable shield encapsulated with UV glue. | in sufficient curing of glue resulting in contact between shield and coil causing leakage current
Catheter and Motor Installation (0550-7030). encapsulation of shied with UV glue. AU Pump Motor, adhesive, dispenser, UV source. Motor cable shield encapsulated with UV glue. | in sufficient encapsulation resulting in contact between shield and coil causing leakage current
Catheter and Motor Installation (0550-7030). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU , AU Rotor PH - cleaned with PEEK sleeve. | Concentricity not controlled
Catheter and Motor Installation (0550-7030). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU MH , AU Rotor PH - cleaned with PEEK sleeve. | Damage to impeller from PEEK sleeve
Catheter and Motor Installation (0550-7030). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU MH, AU Rotor PH - cleaned with PEEK sleeve. | Damage to impeller from PEEK sleeve
Catheter and Motor Installation (0550-7030). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Coil detached from catheter
Catheter and Motor Installation (0550-7030). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Components tangled in catheter - kinked purge lumen
Catheter and Motor Installation (0550-7030). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Kink purge lumen - high purge pressure
Catheter and Motor Installation (0550-7030). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU Pump Catheter Motor assembly. | Components tangled in catheter - unable to load sensor
Catheter and Motor Installation (0550-7030). Load Repo Unit and contamination sleeve. Repo unit, catheter, contamination sleeve. Catheter with repo and contamination sleeve loaded. | Repo unit loaded backwards
Catheter and Motor Installation (0550-7030). oven cure. AU Pump Motor - semi cured, oven (time and temp), pump magazine. AU Pump Motor. | concentricity comprimised
Catheter and Motor Installation (0550-7030). Place kink protector. Catheter with repo and contamination sleeve loaded, kink protector. Catheter with repo, contamination sleeve, and kink protector loaded. | unable to assemble pump handle
Catheter and Motor Installation (0550-7030). Place kink protector. Catheter with repo and contamination sleeve loaded, kink protector. Catheter with repo, contamination sleeve, and kink protector loaded. | unable to assembly pump handle
Catheter and Motor Installation (0550-7030). Remove PEEK sleeve and Check Concentricity. AU Pump Motor, swab, feeler gauge. AU Pump Motor - cooled (PEEK sleeve removed). | unable to continue
Catheter and Motor Installation (0550-7030). Remove PEEK sleeve and Check Concentricity. AU Pump Motor, swab, feeler gauge. AU Pump Motor - cooled (PEEK sleeve removed). | unacceptable gap
Catheter and Motor Installation (0550-7030). room temperature cure. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued and cleaned, magazine. AU Pump Motor - semi cured. | concentricity comprimised
Catheter and Motor Installation (0550-7030). Straighten purge lumen. AU motorwith coiled motor cable and purge lumen, clip, oven (time and temp), pump magazine. AU motor with straightened purge lumen amd motor cable. | Incorrect purge lumen length - unable to assemble to pump handle
Catheter and Motor Installation (0550-7030). Wipe Catheter. Catheter, petroleum ether, swabs, wipes. clean catheter outside. | debris/fibers on pump
Catheter and Motor Installation (10002254). Bond motor housing to rotor. AU MH , AU Rotor PH - cleaned with PEEK sleeve, adhesive, dosing needles. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued. | Not fully seated - reduced bond strength
Catheter and Motor Installation (10002254). Bond motor housing to rotor. AU Pump Motor. AU Pump Motor - cooled (PEEK sleeve removed). | channel misaligned - unable to assemble
Catheter and Motor Installation (10002254). Check dry Rotor fit. AU MH , AU Rotor PH - cleaned . AU MH , AU Rotor PH - cleaned with. | Damage to magnet coating
Catheter and Motor Installation (10002254). Check dry Rotor fit. AU MH , AU Rotor PH - cleaned with. AU MH , AU Rotor PH - cleaned with. | Unable to assemble
Catheter and Motor Installation (10002254). Clean AU Motor housing and AU Rotor Pump Housing. AU MH , AU Rotor PH, swabs, petroleum ether. AU MH , AU Rotor PH - cleaned. | Insufficient bond strength between AU Rotor and Motor Housing
Catheter and Motor Installation (10002254). clean excess adhesive. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued, dry swabs. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued and cleaned. | unable to assemble
Catheter and Motor Installation (10002254). Cool . AU Pump Motor. AU Pump Motor - cooled. | concentricity compromised
Catheter and Motor Installation (10002254). Cool. AU motor with straightened purge lumen and motor cable. AU motor with straightened purge lumen and motor cable cooled. | Purge lumen kinks - high purge pressure
Catheter and Motor Installation (10002254). encapsulation of shield with UV glue. AU Pump Motor, adhesive, dispenser, UV source. Motor cable shield encapsulated with UV glue. | insufficient curing of glue resulting in contact between shield and coil causing leakage current
Catheter and Motor Installation (10002254). encapsulation of shield with UV glue. AU Pump Motor, adhesive, dispenser, UV source. Motor cable shield encapsulated with UV glue. | insufficient encapsulation resulting in contact between shield and coil causing leakage current
Catheter and Motor Installation (10002254). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU , AU Rotor PH - cleaned with PEEK sleeve. | Concentricity not controlled
Catheter and Motor Installation (10002254). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU MH , AU Rotor PH - cleaned with PEEK sleeve. | Damage to impeller from PEEK sleeve
Catheter and Motor Installation (10002254). Insert PEEK sleeve. AU MH , AU Rotor PH - cleaned, PEEK sleeve. AU MH, AU Rotor PH - cleaned with PEEK sleeve. | Damage to impeller from PEEK sleeve
Catheter and Motor Installation (10002254). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Coil detached from catheter
Catheter and Motor Installation (10002254). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Components tangled in catheter - kinked purge lumen
Catheter and Motor Installation (10002254). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU motor assembled with catheter. | Kink purge lumen - high purge pressure
Catheter and Motor Installation (10002254). Load purge lumen and motor cables. Catheter with repo, contamination sleeve, and kink protector loaded, AU motor. AU Pump Catheter Motor assembly. | Components tangled in catheter - unable to load sensor
Catheter and Motor Installation (10002254). Load Repo Unit and contamination sleeve. Repo unit, catheter, contamination sleeve. Catheter with repo and contamination sleeve loaded. | Repo unit loaded backwards
Catheter and Motor Installation (10002254). Load Repo Unit and contamination sleeve. Repo unit, catheter, contamination sleeve. Catheter with repo and contamination sleeve loaded. | Repo unit loaded with misaligned butterfly
Catheter and Motor Installation (10002254). oven cure. AU Pump Motor - semi cured, oven (time and temp), pump magazine. AU Pump Motor. | concentricity compromised
Catheter and Motor Installation (10002254). Place kink protector. Catheter with repo and contamination sleeve loaded, kink protector. Catheter with repo, contamination sleeve, and kink protector loaded. | unable to assemble pump handle
Catheter and Motor Installation (10002254). Place kink protector. Catheter with repo and contamination sleeve loaded, kink protector. Catheter with repo, contamination sleeve, and kink protector loaded. | unable to assembly pump handle
Catheter and Motor Installation (10002254). Remove PEEK sleeve and Check Concentricity. AU Pump Motor, swab, feeler gauge. AU Pump Motor - cooled (PEEK sleeve removed). | unable to continue
Catheter and Motor Installation (10002254). Remove PEEK sleeve and Check Concentricity. AU Pump Motor, swab, feeler gauge. AU Pump Motor - cooled (PEEK sleeve removed). | unacceptable gap
Catheter and Motor Installation (10002254). room temperature cure. AU MH , AU Rotor PH - cleaned with PEEK sleeve - glued and cleaned, magazine. AU Pump Motor - semi cured. | concentricity compromised
Catheter and Motor Installation (10002254). Straighten purge lumen. AU motor with coiled motor cable and purge lumen, clip, oven (time and temp), pump magazine. AU motor with straightened purge lumen and motor cable. | Incorrect purge lumen length - unable to assemble to pump handle
Catheter and Motor Installation (10002254). Wipe Catheter. Catheter, petroleum ether, swabs, wipes. clean catheter outside. | debris/fibers on pump
Catridge Cap Tightness Test. Adjust pressure regulator to acceptable pressure. tightness test fixture (pressure). adjusted tightness test fixture regulator. | Unable to run test
Catridge Cap Tightness Test. Adjust pressure regulator to acceptable pressure. tightness test fixture. adjusted tightness test fixture regulator. | Damage components leading to leak in system and purge pressure loss (unable to run test)
Catridge Cap Tightness Test. Insert each production assembly and perform test in same manner as test part. pump body assembly, tightness test fixture. tested pump body assembly. | damage cylinder leading to leak in system and purge pressure loss leading to test failure (scrap)
Catridge Cap Tightness Test. Insert each production assembly and perform test in same manner as test part. pump body assembly, tightness test fixture. tested pump body assembly. | fail good units
Catridge Cap Tightness Test. Perform functional self test of Tightness Test Fixture and adjust pressure regulator until correct. pump body assembly, tightness test fixture. tested pump body assembly. | Unable to proceed with additional tests
Catridge Cap Tightness Test. Perform functional self test of Tightness Test Fixture and adjust pressure regulator until correct. tightness test fixture, reference part. setup up tightness fixture. | Unable to run test
Catridge Cap Tightness Test. Remove assembly from the device and store in carrier. tested pump body assembly, carrier. tested pump body assembly. | damage cylinder leading to leak in system and purge pressure loss
Cleaving Procedure. Cleave fibers. Oxygenator Asm Procedure. Oxygenator Asm. | Fibers not open
Cleaving Procedure. Cleave fibers. Oxygenator Asm Procedure. Oxygenator Asm. | Improper cut
Coil Forming (0550-7027). Clean forming tool. tool, 70% IPA, wipes. Cleaned tool. | Damage to ceramic sleeve
Coil Forming (0550-7027). insert tooling to machine and run machine. Proximal bearing/sleeve assembly with first formed winding loaded onto cleaned forming tool, forming machine, program 002, settings (temp and airflow). Proximal bearing/sleeve assembly with formed winding loaded onto forming tool. | improper forming
Coil Forming (0550-7027). insert tooling to machine and run machine. Proximal bearing/sleeve assembly with first formed winding loaded onto cleaned forming tool, forming machine, program 002, settings (temp and airflow). Proximal bearing/sleeve assembly with formed winding loaded onto forming tool. | machine will not cycle
Coil Forming (0550-7027). place first formed winding on sleeve/bearing assembly, pull shrink tubing over winding. Proximal bearing/sleeve assembly loaded onto cleaned forming tool, first formed winding. Proximal bearing/sleeve assembly with first formed winding loaded onto cleaned forming tool. | incorrect placement - winding not properly placed on sleeve
Coil Forming (0550-7027). place in storage tray. Formed winding on sleeve, storage tray. Formed winding on sleeve in storage tray. | leads damaged
Coil Forming (0550-7027). place proximal bearing and sleeve on forming tool, secure with screw cap assembly. Proximal bearing, sleeve, cleaned forming tool. Proximal bearing/sleeve assembly loaded onto cleaned forming tool. | improper forming
Coil Forming (0550-7027). Place shrink tube over Proximal bearing/sleeve assembly. Proximal bearing/sleeve assembly loaded onto cleaned forming tool, shrink tubing. Proximal bearing/sleeve assembly loaded onto cleaned forming tool with shrink tubing. | Coil forms incorrectly - coil swells over time - unable to assemble
Coil Forming (0550-7027). Place shrink tube over Proximal bearing/sleeve assembly. Proximal bearing/sleeve assembly loaded onto cleaned forming tool, shrink tubing. Proximal bearing/sleeve assembly loaded onto cleaned forming tool with shrink tubing. | Coil swells over time - unable to assemble
Coil Forming (0550-7027). Remove Proximal bearing. Formed winding on proximal bearing/sleeve assembly, scalpel. Formed winding on sleeve proximal bearin and shrink tubing removed. | Coil Damaged
Coil Forming (0550-7027). Remove Proximal bearing. Formed winding on proximal bearing/sleeve assembly, scalpel. Formed winding on sleeve proximal bearing and shrink tubing removed. | Coil Damaged
Coil Forming (0550-7027). Remove shrink tubing from coil. Formed winding on proximal bearing/sleeve assembly, scalpel. Formed winding on proximal bearing/sleeve assembly, shrink tubing removed. | Coil damaged
Coil Forming (0550-7027). remove tool from machine and remove screw cap from tool and remove assembly. Proximal bearing/sleeve assembly with formed winding loaded onto forming tool. Formed winding on proximal bearing/sleeve assembly. | ceramics damaged
Coil Forming (0550-7027). remove tool from machine and remove screw cap from tool and remove assembly. Proximal bearing/sleeve assembly with formed winding loaded onto forming tool. Formed winding on proximal bearing/sleeve assembly. | leads damaged
Coil Forming (0550-7027). remove tool from machine and remove screw cap from tool and remove assembly. Proximal bearing/sleeve assembly with formed winding loaded onto forming tool. Formed winding on proximal bearing/sleeve assembly. | unable to continue
Coil winding (0550-7035). Clean tooling. tool, 70% IPA, wipes. Cleaned tool. | Unable to wind
Coil winding (0550-7035). Identify phases and twist and solder. first formed coil on winding mandrel loaded with coil leads cut, soldering iron, solder. first formed coil on winding mandrel loaded with coil leads cut and twisted. | Motor would not run correctly
Coil winding (0550-7035). Insert tool with winding into forming machine. Winding on mandrel, forming machine. Forming machine with winding mandrel loaded. | unable to form
Coil winding (0550-7035). Load tooling into machine. Cleaned tool, winding machine. tool loaded into machine. | Unable to wind
Coil winding (0550-7035). Remove tool with first formed winding from forming machine. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. first formed coil on winding mandrel loaded with coil leads cut. | unable to continue
Coil winding (0550-7035). remove tool with winding from winding machine. Winding on mandrel in machine. Winding on mandrel. | coil damaged while removing - may affect resistance or other spec
Coil winding (0550-7035). remove winding from tooling. first formed coil on winding mandrel loaded with coil leads cut and twisted. first formed coil with coil leads cut and twisted and soldered. | insulation damaged - leakage current
Coil winding (0550-7035). run forming machine. Forming machine with winding mandrel loaded, settings (temp, airflow), Program 002. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. | Does not detach from mandrel
Coil winding (0550-7035). run forming machine. Forming machine with winding mandrel loaded, settings (temp, airflow), Program 002. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. | incorrect forming damage to insulation during second forming
Coil winding (0550-7035). run forming machine. Forming machine with winding mandrel loaded, settings (temp, airflow), Program 002. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. | incorrect forming damage to insullation during second forming
Coil winding (0550-7035). run forming machine. Forming machine with winding mandrel loaded, settings (temp, airflow), Program 002. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. | incorrect lead wire cut length
Coil winding (0550-7035). run forming machine. Forming machine with winding mandrel loaded, settings (temp, airflow), Program 002. Forming machine with first formed coil on winding mandrel loaded with coil leads cut. | may not be able to complete second forming
Coil winding (0550-7035). run machine. machine, program 002, settings (temp, airflow). Winding on mandrel in machine. | incorrect coil produced
Coil winding (0550-7035). run machine. machine, program 002, settings (temp, airflow). Winding on mandrel in machine. | produce coil that does not meet specification
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 10. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | coil damage during winding
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 3. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | Unfinished coil (scrap)
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 4. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | Crossovers leading to damaged coil (i.e. short)
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 5. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | Crossovers leading to damaged coil (i.e. short)
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 6. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | misaligned coil phase leading to resistance imbalance/Back EMF Imbalance
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 7. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | unacceptable coil resistance due to tension
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 8. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | incorrect coil length (unable to assemble)
Coil Winding and Compression. Clean tooling, setup equipment and begin winding Coil phases 1, 2, and 9. Coil winder, mandrel for coil winding fixture, allen wrench, Wipes, alcohol, wire brush, never-seez grease, swab, wire. Wound Coil. | coil damage during winding
Coil Winding and Compression. Form coil on insertion pin. Coil positioned on insertion pin within forming machine. formed coil on insertion pin. | Coil OD too large (unable to assemble)
Coil Winding and Compression. Form coil on insertion pin. Coil positioned on insertion pin within forming machine. formed coil on insertion pin. | Damaged coil
Coil Winding and Compression. Form coil on insertion pin. Coil positioned on insertion pin within forming machine. formed coil on insertion pin. | incorrect position leading to pump stop
Coil Winding and Compression. Place shrink tube onto winding. Then place back into machine and shrink the shrink tube onto winding. Shrink tube, formed coil on insertion pin, forming machine. Fully formed motor winding. | damage coil
Coil Winding and Compression. Place shrink tube onto winding. Then place back into machine and shrink the shrink tube onto winding. Shrink tube, formed coil on insertion pin, forming machine. Fully formed motor winding. | damage coil during insertion/handling
Coil Winding and Compression. Place shrink tube onto winding. Then place back into machine and shrink the shrink tube onto winding. Shrink tube, formed coil on insertion pin, forming machine. Fully formed motor winding. | damage the shrink tube leaving coil unprotected leading to damage to the coil
Coil Winding and Compression. Place shrink tube onto winding. Then place back into machine and shrink the shrink tube onto winding. Shrink tube, formed coil on insertion pin, forming machine. Fully formed motor winding. | Shrink tube not covering all areas coil leading to damage of coil
Coil Winding and Compression. Place wound coil into forming machine and execute forming process. Wound coil on mandrel, forming machine. Formed Coil on mandrel. | Coil OD too large (unable to assemble)
Coil Winding and Compression. Place wound coil into forming machine and execute forming process. Wound coil on mandrel, forming machine. Formed Coil on mandrel. | damage coil during forming process
Coil Winding and Compression. Place wound coil into forming machine and execute forming process. Wound coil on mandrel, forming machine. Formed Coil on mandrel. | Damaged coil
Coil Winding and Compression. Place wound coil into forming machine and execute forming process. Wound coil on mandrel, forming machine. Formed Coil on winding. | damage coil during forming process
Coil Winding and Compression. Place wound coil into forming machine and execute forming process. Wound coil on mandrel, forming machine. Formed Coil on winding. | Damaged coil
Coil Winding and Compression. Remove coil from winding mandrel and then load onto insertion pin in the forming machine. Formed Coil on mandrel, insertion pin, forming machine. Coil positioned on insertion pin within forming machine. | damage coil during insertion/removal handling
Coil Winding and Compression. Remove coil from winding mandrel and then load onto insertion pin in the forming machine. Formed Coil on mandrel, insertion pin, forming machine. Coil positioned on insertion pin within forming machine. | damage coil during insertion/removal handling (scrap)
Coil Winding and Compression. Remove coil from winding mandrel and then load onto insertion pin in the forming machine. Formed Coil on mandrel, insertion pin, forming machine. Coil positioned on insertion pin within forming machine. | damage inner shrink tube leading
Coil Winding and Compression. Remove coil from winding mandrel and then load onto insertion pin in the forming machine. Formed Coil on mandrel, insertion pin, forming machine. Coil positioned on insertion pin within forming machine. | incorrect position leading to pump stop
Coil Winding and Compression. Remove coil from winding mandrel and then load onto insertion pin in the forming machine. Formed Coil on mandrel, insertion pin, forming machine. Coil positioned on insertion pin within forming machine. | unable to assemble/difficulty assembling
Coil Winding and Compression. Sampling of coils for testing. WI, operator. Sampled Coils. | Incorrect sampling of coils for electrical testing
Coil Winding and Compression. Storage of Wound coils. Storage Fixture, Wounded Coil. Wound Coils. | damage coil during insertion/handling
Concentrator Exhaust Duct Mfg Process (MFG-3010). Assemble Concentrator Exhaust Duct. . ASSY, Concentrator Exhaust Duct. | Concentrator Exhaust Duct Assembly not properly fastened
Double Thermoform Cannula (0048-7064). Anneal cannula for 4 hours min at 70C. cured cannula, oven. annealed cannula. | Shape out of specification
Double Thermoform Cannula (0048-7064). Anneal raw material tubing. raw tubing material. annealed tubing (removed moisture and stress release). | not annealed tubing
Double Thermoform Cannula (0048-7064). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Biocompatibility comprimised
Double Thermoform Cannula (0048-7064). Clean cannula in 80% alcohol for 4 minutes. annealed cannula, alcohol, US bath. Double Thermoformed Cannula. | Bond strength comprimised
Double Thermoform Cannula (0048-7064). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. mandrel with cleaned ends. | Too much acmos on inner layer causes double step on ID of cannula
Double Thermoform Cannula (0048-7064). Clean Mandrel with alcohol wipe. lubricated mandrel with carrier solvent removed. more than just ends cleaned. | too little acmos
Double Thermoform Cannula (0048-7064). Cool cannula. Thermoformed First Layer with heat shrink, time at room temp. Cooled first layer with heat shrink. | cosmetic defects
Double Thermoform Cannula (0048-7064). inspect marker position. cannula with marker band. inspected cannula with marker band on mandrel. | marker band in incorrect position
Double Thermoform Cannula (0048-7064). let cool. cured cannula with marker band and heat shrink. cooled cannula with marker band and heat shrink. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). let cool. Thermoformed cannula with heat shrink tubing, time at room temp. cooled Thermoformed cannula with heat shrink tubing. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). let cool. Thermoformed first and reinforcement layers with heath shrink tubing, time at room temp. cooled Thermoformed first and reinforcement layers with heat shrink tubing. | cosmetic defects
Double Thermoform Cannula (0048-7064). let dry in oven. first layer with reinforcement tubing, oven (time, temp). dried first layer with reinforcement tubing. | Material deterioration leads to delamination
Double Thermoform Cannula (0048-7064). let dry in oven. first layer with reinforcement tubing, oven (time, temp). dried first layer with reinforcement tubing. | Material does not re-flow
Double Thermoform Cannula (0048-7064). Load cannula onto coil machine. thermoformed first and reinforcement layer, coil machine. thermoformed first and reinforcement layer loaded on machine. | Loaded incorrectly
Double Thermoform Cannula (0048-7064). load coil onto machine. thermoformed first and reinforcement layer, coil machine. coil loaded on machine. | Coil loaded incorrectly
Double Thermoform Cannula (0048-7064). load the marker band onto the cannula with the cannula marker inspection fixture. cannula on mandrel assembly with marker band, cannula marker inspection fixture. cannula with marker band. | difficulty placing pump in optimal position
Double Thermoform Cannula (0048-7064). lubricate inner layer with HFE. reinforcement tubing, HFE. lubricated reinforcement layer. | Unable to continue with build
Double Thermoform Cannula (0048-7064). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | too little acmos
Double Thermoform Cannula (0048-7064). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos
Double Thermoform Cannula (0048-7064). Mandrel Preparation. assembled Mandrel, acmos. lubricated mandrel. | Too much acmos on inner layer causes double step on ID of cannula
Double Thermoform Cannula (0048-7064). Mandrel Preparation. lubricated mandrel, oven (time, temperature). lubricated mandrel with carrier solvent removed. | Incorrect ID of cannula
Double Thermoform Cannula (0048-7064). Mandrel Preparation. lubricated mandrel, oven (time, temperature). lubricated mandrel with carrier solvent removed. | Too much acmos
Double Thermoform Cannula (0048-7064). Mandrel Preparation. lubricated mandrel, oven (time, temperature). lubricated mandrel with carrier solvent removed. | too much acmos dried off
Double Thermoform Cannula (0048-7064). Mandrel Preparation. lubricated mandrel, oven (time, temperature). lubricated mandrel with carrier solvent removed. | Too much acmos on inner layer causes double step on ID of cannula
Double Thermoform Cannula (0048-7064). Place heat shrink. cannula with marker band, heat shrink. cannula with marker band and heat shrink. | bubbles in cannula
Double Thermoform Cannula (0048-7064). Place heat shrink. cannula with marker band, heat shrink. cannula with marker band and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Double Thermoform Cannula (0048-7064). Place heat shrink. cannula with marker band, heat shrink. cannula with marker band and heat shrink. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Double Thermoform Cannula (0048-7064). place heat shrink. dried first layer with reinforcement tubing, heat shrink tubing. dried first layer with reinforcement tubing and heat shrink. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | bubbles in cannula
Double Thermoform Cannula (0048-7064). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Improper step formed at PH end causes thrombus or embolism
Double Thermoform Cannula (0048-7064). place heat shrink. dried first/reinforcement layerwith outer tubing, heat shrink tubing. dried first/reinforcement layer with outer tubing and heat shrink. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). place in oven. cannula with marker band and heat shrink, oven (time, temp). cured cannula with marker band and heat shrink. | Material deterioration leads to delamination
Double Thermoform Cannula (0048-7064). place in oven. cannula with marker band and heat shrink, oven (time, temp). cured cannula with marker band and heat shrink. | Material does not re-flow
Double Thermoform Cannula (0048-7064). place in oven. dried first layer with reinforcement tubing and heat shrink, oven (time, temp). Thermoformed first and reinforcement layers with heat shrink tubing. | Material deterioration leads to delamination
Double Thermoform Cannula (0048-7064). place in oven. dried first layer with reinforcement tubing and heat shrink, oven (time, temp). Thermoformed first and reinforcement layers with heat shrink tubing. | Material does not re-flow
Double Thermoform Cannula (0048-7064). place in oven. dried first layer with reinforcement tubing and heat shrink, oven (time, temp). Thermoformed first and reinforcement layers with heat shrink tubing. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material deterioration leads to delamination
Double Thermoform Cannula (0048-7064). place in oven. dried first/reinforcement layer with outer tubing and heat shrink, oven (time, temp). Thermoformed cannula with heat shrink tubing. | Material does not re-flow
Double Thermoform Cannula (0048-7064). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too cold
Double Thermoform Cannula (0048-7064). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too hot - yields degraded cannula
Double Thermoform Cannula (0048-7064). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too long of time - yields degraded cannula
Double Thermoform Cannula (0048-7064). place in oven. thermoformed first and reinforcement layer with trimmed and secured coil, oven (time, temp). thermoformed first and reinforcement coil. | too short or time
Double Thermoform Cannula (0048-7064). Remove cannula from mandrel. Cured cannula on bend mandrel. cured cannula. | unable to assemble
Double Thermoform Cannula (0048-7064). Remove cannula from mandrel. inspected cannula with marker band on mandrel. inspected cannula with marker band. | unable to assemble
Double Thermoform Cannula (0048-7064). remove heat shrink. cooled cannula with marker band and heat shrink. cannula with marker band. | razor cuts cannula
Double Thermoform Cannula (0048-7064). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | cut cannula
Double Thermoform Cannula (0048-7064). remove heat shrink. Cooled first layer with heat shrink, razor. cooled first layer. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). remove heat shrink. cooled Thermoformed cannula with heath shrink tubing, razor. Thermoformed cannula. | razor cuts cannula
Double Thermoform Cannula (0048-7064). remove heat shrink. cooled Thermoformed first and reinforcement layers with heath shrink tubing, razor. thermoformed first and reinforcement layer. | cut cannula
Double Thermoform Cannula (0048-7064). run program. thermoformed first and reinforcement layer loaded on machine, coil loaded on machine, coil program. thermoformed first and reinforcement layer with coil. | Program corrupted or incorrect
Double Thermoform Cannula (0048-7064). select program. thermoformed first and reinforcement layer, coil machine. coil machine with program selected. | wrong program selected
Double Thermoform Cannula (0048-7064). slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Double Thermoform Cannula (0048-7064). slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . cannula, bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Double Thermoform Cannula (0048-7064). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Improper step formed at PH end causes thrombus or embolism
Double Thermoform Cannula (0048-7064). Slide Heat Shrink tubing. mandrel with first layer, heat shrink tubing. mandrel with first layer and heat shrink tubing. | Unable to continue with assembly
Double Thermoform Cannula (0048-7064). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | slid on to incorrect location
Double Thermoform Cannula (0048-7064). Slide inner layer over mandrel. mandrel with cleaned ends, annealed tubing. mandrel with first layer. | Unable to thermoform
Double Thermoform Cannula (0048-7064). slide outer layer. thermoformed first and reinforcement with coil, outer layer. thermoformed first and reinforcement with coil with outer layer. | cannula does not recover from kinking
Double Thermoform Cannula (0048-7064). slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | cannula does not recover from kinking
Double Thermoform Cannula (0048-7064). slide reinforcement layer. lubricated reinforcement layer, cooled first layer. first layer with reinforcement tubing. | OD too large
Double Thermoform Cannula (0048-7064). Thermoform cannula. mandrel with first layer and heat shrink tubing, oven (time, temperature). Thermoformed First Layer with heat shrink. | Material deterioration leads to delamination
Double Thermoform Cannula (0048-7064). Thermoform cannula. mandrel with first layer and heat shrink tubing, oven (time, temperature). Thermoformed First Layer with heat shrink. | Material does not re-flow
Double Thermoform Cannula (0048-7064). Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too long - step at transition to PH or IC
Double Thermoform Cannula (0048-7064). Trim cannula. Cannula, bubble picking block, razer. trimmed cannula. | Too short - weak bond
Double Thermoform Cannula (0048-7064). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | coil trimmed too short
Double Thermoform Cannula (0048-7064). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | Improper step formed at PH end
Double Thermoform Cannula (0048-7064). trim coil. thermoformed first and reinforcement layer with coil, snips. thermoformed first and reinforcement layer with trimmed coil. | reduced bond strength
Double Thermoform Cannula (0048-7064). use loctite to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, loctite. thermoformed first and reinforcement layer with trimmed and secured coil. | loctite not applied
Double Thermoform Cannula (0048-7064). use loctite to secure ends of coil onto cannula. thermoformed first and reinforcement layer with trimmed coil, loctite. thermoformed first and reinforcement layer with trimmed and secured coil. | too much loctite applied leads to delamination
Double Thermoform Cannula. Submerge cannula marker bands and place in ultrasonic bath, let the piece dry. cannula on mandrel assembly, ultrasonic bath, marker band, lint free wipes. prepped marker band. | degredation of material leading to particles
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | Fly wheel installed at wrong height
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | Fly wheel shaft disconnection
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | Incorrect orientation of circuit to PCB
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | LED E-Drive Label not aligned
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | Pump cradle not installed too high
E-Drive Assembly (MFG-0029). Assemble E-Drive. . . | Pump cradle not installed too low
E-Drive Wheel Assembly (MFG-0028). Assemble E-drive wheel . . . | Incorrect magnetic orientation
E-Drive Wheel Assembly (MFG-0028). Assemble E-drive wheel . . . | Interference with e-drive housing
E-Drive Wheel Assembly (MFG-0028). Assemble E-drive wheel . . . | Magnets shifting
Electrical Safety Test. Test console and pump driver for electrical safety. . . | Incorrect test setup and test fails (incorrect jack and cable connections)
Electrical Safety Test. Test console and pump driver for electrical safety. . . | Incorrect test setup and test passes (incorrect jack and cable connections)
End Qualification Test. After test has concluded, note test result on screen. Once lot is complete, press "Close Lot" button to close the lot. Purge casseette, test station. Closed lot. | Component mix-up unacceptable units released
End Qualification Test. After test has concluded, note test result on screen. Once lot is complete, press "Close Lot" button to close the lot. Purge casseette, test station. Closed lot. | loss of component traceability
End Qualification Test. Choose EQT Recipe for Gen1 or Gen2. EQT, Purge Cassette, barcode scanner, test station. Tested purge cassette. | Unable to start test due to incorrect recipe
End Qualification Test. Prepare materials for packaging on Purge Cassette Tubing Assembly Fixture. Then Align Paper Tape and Pressure Line and wind tubing around fixture mandrels. Tested assembly, scotch, tube winding fixture, paper tape. Assembly ready for assembly. | tubing unravels leading to kink in system (unable to move pass case start)
End Qualification Test. Prepare materials for packaging on Purge Cassette Tubing Assembly Fixture. Then Align Paper Tape and Pressure Line and wind tubing around fixture mandrels. Tested assembly, scotch, tube winding fixture, paper tape. Assembly ready for assembly. | tubing unravels leading to kink in system (unable to move past case start)
End Qualification Test. Scan lot number from DHR with barcode scanner. Press the "Open Lot Code" button to open the lot.. Barcode Scanner, DHR. Scanned in lot number. | loss of component traceability
End Qualification Test. Scan purge cassette serial number with barcode scanner. Load purge cassette into test station and initiate test. Purge cassette, barcode scanner, test station. Tested purge cassette. | loss of component traceability
End Qualification Test. Scan purge cassette serial number with barcode scanner. Load purge cassette into test station and initiate test. Purge cassette, barcode scanner, test station. Tested purge cassette. | Unable to start test
End Qualification Test. Scan purge cassette serial number with barcode scanner. Load purge cassette into test station and initiate test. Purge cassette, barcode scanner, test station. Tested purge cassette. | Unacceptable units are passed leading purge failures in field
End Qualification Test. Secure Suction Line tubing with Paper Tape, apply Scotch Tape around Tape Overlap, and remove from mandrels. Tested assembly, scotch, tube winding fixture, paper tape. Assembly ready for assembly. | damage/kink of tubing (unable to move pass case start)
End Qualification Test. Secure Suction Line tubing with Paper Tape, apply Scotch Tape around Tape Overlap, and remove from mandrels. Tested assembly, scotch, tube winding fixture, paper tape. Assembly ready for assembly. | damage/kink of tubing (unable to move past case start)
End Qualification Test. Use IPA and swab to clean all EQT luer connections. IPA, swab, EQT luer connections. Cleaned equipment. | insufficient O'ring seal leading to false rejection
Final Assembly and Qualification. ​Clean the Leakage Current Test conductivity sensor and leakage electrodes. electrodes, alcohol, conductivity sensor. Cleaned electrodes and conductivity sensor. | Incorrect leakage reading
Final Assembly and Qualification. ​Clean the Leakage Current Test conductivity sensor and leakage electrodes. electrodes, alcohol, conductivity sensor. Cleaned electrodes and conductivity sensor. | Incorrectly cleaned electrodes and conductivity sensor
Final Assembly and Qualification. ​Fill the Rinse Test beaker with 8 liters of pyrogen-free water.. Beaker, Pyrogen free water. Fill beaker with pyrogen free water. | Pyrogen free water contaminated
Final Assembly and Qualification. Apply voltage to pump. Leakage current voltage source, Voltage selector, ABC software, ACS software. Correct voltage applied to pump. | Incorrect voltage applied to pump
Final Assembly and Qualification. Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
Final Assembly and Qualification. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm damaged
Final Assembly and Qualification. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge system occluded or break not detected
Final Assembly and Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection blocked
Final Assembly and Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection bond failure
Final Assembly and Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
Final Assembly and Qualification. Attach Pump cable to Cable holder. Pump cable, cable holder, Operator. Pump cables attached to holder. | Pump cables unable to be attached to holder
Final Assembly and Qualification. Attach Pump cable to Cable holder. Pump cable, cable holder, Operator. Pump cables attached to holder. | Pumps unable to be attached to holder
Final Assembly and Qualification. Attach Pump cable to Cable holder. Pump cables, holder, operator. Pump cables attached to holder. | Pump cables unable to be attached to holder
Final Assembly and Qualification. Attach Pump to Holder. Pump Holder, Staging area, Operator. Pumps attached to holder. | Pumps unable to be attached to holder
Final Assembly and Qualification. Attach the conductivity sensor to the beaker. Conductivity Sensor, Beaker, Operator. Conductivity sensor attached to the beaker. | Conductivity sensor not attached to the beaker
Final Assembly and Qualification. Attach the Rinse Test conductivity sensor to the beaker. Conductivity sensor, beaker. Conductivity sensor attached to the beaker. | Conductivity sensor not attached to the beaker
Final Assembly and Qualification. Attach thermometer to the Leakage current test beaker . Thermometer, Leakage test beaker. Thermometer attached to the leakage test beaker. | Thermometer not attached to the leakage test beaker
Final Assembly and Qualification. Automated Water heater . Water Heater, Controller, Central station. Heated water to correct temp.. | Water not heated to correct temp.
Final Assembly and Qualification. Automated Water heater. Water Heater, Controller, Central station. Heated water to correct temp.. | Water not heated to correct temp.
Final Assembly and Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Gage factors/G0 not loaded
Final Assembly and Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Low SNR leading to scrap
Final Assembly and Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Sensor failure/low sensor signal
Final Assembly and Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Wrong Gage factors/G0 offset loaded
Final Assembly and Qualification. Central station indication to remove pumps. ACS, Pump, Pump Holder, encoder, Operator. Central station indication to operator. | Central station does not indicate to operator
Final Assembly and Qualification. Central Station ready for next step. Central Station, ACS. Display message to operator. | "Ready" Message not displayed
Final Assembly and Qualification. Central Station Self test. Central station, bench controller (ABC), ACS software, Pump controller (APC), periferals. Self test successful. | Failed Self test
Final Assembly and Qualification. Check Pen-style cleaner. Pen-style cleaner, Operator. Pen style cleaner checked. | Unchecked Pen-style cleaner
Final Assembly and Qualification. Clean 2x with petroleum ether. Then attach Pressure Storage Set with filter to pump assembly. Then dry in Oven. Then clean with petroleum ether again and apply UV adhesive.. THF, Pressure storage set, pump assembly, dosing needle, petroleum ether, oven (time and temperature). Pump assembly with Pressure Storage Set attached. | Bond Failure
Final Assembly and Qualification. Clean 2x with petroleum ether. Then attach Pressure Storage Set with filter to pump assembly. Then dry in Oven. Then clean with petroleum ether again and apply UV adhesive.. THF, Pressure storage set, pump assembly, dosing needle, petroleum ether, oven (time and temperature). Pump assembly with Pressure Storage Set attached. | Kink/Bond failure
Final Assembly and Qualification. Clean Rinse Test conductivity sensor. Conductivity sensor, alcohol. Cleaned rinse conductivity sensor. | Incorrect rinse conductivity reading
Final Assembly and Qualification. Clean the Leakage Current Test conductivity sensor and leakage electrodes. electrodes, alcohol, conductivity sensor. Cleaned electrodes and conductivity sensor. | Incorrectly Conductivity reading
Final Assembly and Qualification. Cleaning of Blue connector. blue connector, Pen Cleaner. Cleaned blue connector. | Dirty blue connector
Final Assembly and Qualification. Cleaning of electodes- Rinse Beaker. electrodes, alcohol. Cleaned electrodes. | Incorrect Conductivity reading
Final Assembly and Qualification. Cleaning of electodes- Rinse Beaker. electrodes, alcohol. Cleaned electrodes. | Incorrectly cleaned electrodes
Final Assembly and Qualification. Connect Pump Purge catheter to Purger. Purger, Connector, Flow meters, Operator. Pumps connected to Purge line. | Incorrect purge flow reading for a pump
Final Assembly and Qualification. Connect Pump Purge catheter to Purger. Purger, Connector, Operator. Pumps connected to Purge line. | incorrect purge flow reading for a pump
Final Assembly and Qualification. Connect Pump Purge catheter to Purger. Purger, Connector, Operator. Pumps connected to Purge line. | Pumps unable to be attached to purger
Final Assembly and Qualification. Connect Pump Purge catheter to Purger. Purger, Connector, Operator. Pumps connected to Purge line. | Purger does not turn on
Final Assembly and Qualification. Connect Pump Purge catheter to Purger. Purger, Connector, Operator. Pumps connected to Purge line. | Purger system not pressurized properly
Final Assembly and Qualification. Connect pump to blue connector. Blue Connector, AFQ Controller. Pumps connected to AFQ Controller. | EEPROM communication failure
Final Assembly and Qualification. Connect pump to blue connector. Blue Connector, AFQ Controller. Pumps connected to AFQ Controller. | Leakage current test not properly performed
Final Assembly and Qualification. Connect pump to blue connector. Blue Connector, AFQ Controller. Pumps connected to AFQ Controller. | Pump Run test failure
Final Assembly and Qualification. Connect purge line to Purger (pinch valve line). Pinch valve line, Purger. Purge line connected to Purger. | Purge line not connected to Purger
Final Assembly and Qualification. Connect the purge line inside the canister to the bag. Purge line, canister, bag. Purge line connected inside the canister to the bag. | Purge line not connected inside the canister to the bag
Final Assembly and Qualification. Create work order. Print the pump serial number on handle. Finished AU pump plug, Domino Printer. Pump with serial number on handle. | incorrect serial/missing serial number
Final Assembly and Qualification. De-air all purge lines for all pumps. Purge line, pumps, AFQ system. All purge lines de-aired. | Loose line
Final Assembly and Qualification. DMM Setting. DMM, Controller, ABC software, APC Softwre, ACS Software. Calibrated DMM. | Incorrect Leakage Voltage Reading
Final Assembly and Qualification. DMM Setting. DMM, Controller, ABC software, APC Softwre, ACS Software. Correct Leakage Current Value. | Incorrect Leakage Voltage Reading
Final Assembly and Qualification. DMM Setting. DMM, Controller, ABC software, APC Softwre, ACS Software. Correct Leakage Current Value. | Unable to Read Leakage Current Values
Final Assembly and Qualification. Dry in Oven. Oven Oven Rack Temperature Time Air manifold Air pressure/flow rate Pump with serial number programmed. Dry Pump. | Pump not fully dry, leads to bacterial growth
Final Assembly and Qualification. Dry pumps in oven. Pump assembly, oven (time and temperature), hanging rank, air pressure. Dry pumps. | Pumps not dry leading to contamination issues
Final Assembly and Qualification. EEPROM UID Write. Operator, EEPROM, APC software, ACS software, pumps. EEPROM content verified and CRC checking performed. | Unable to write UID to Pump EEPROM
Final Assembly and Qualification. EEPROM Write protection. ACS software, Pump, EEPROM, APC software. write protected EEPROM. | Unable to turn on Write Protection
Final Assembly and Qualification. Ensure purger air-line is shut off. Purge line, Purger. Purge line is shut off. | Purge line is ON
Final Assembly and Qualification. Enter the lot number. AFQ System, Operator, Lot Number. Lot Number Entered . | Incorrect Lot Number entered
Final Assembly and Qualification. Fill Beaker with 100 mililiters of 0.025% Sodium Bicarbonate solution. 2% Sodium Bicarbonate, Pyrogen free water, beaker. filled beaker 0.025% Sodium Bicarbonate solution. | Pyrogen free water contaminated
Final Assembly and Qualification. Fill Beaker with 8 liters of 0.025% Sodium Bicarbonate solution. 2% Sodium Bicarbonate, Pyrogen free water, beaker. filled beaker 0.025% Sodium Bicarbonate solution. | Incorrect ratio of Sodium Bicarbonate solution
Final Assembly and Qualification. Fill Run Test beaker with pyrogen-free water up to the fill sensor. Beaker, Pyrogen free water. Fill beaker with pyrogen free water. | Incorrect water level in beaker
Final Assembly and Qualification. Fill Run Test beaker with pyrogen-free water up to the fill sensor. Beaker, Pyrogen free water. Filled beaker with pyrogen free water. | Pyrogen free water contaminated
Final Assembly and Qualification. Fill the Rinse Test beaker with 8 liters of pyrogen-free water.. Beaker, Pyrogen free water, level sensor. Filled beaker with pyrogen free water. | Incorrect water level in beaker
Final Assembly and Qualification. Final Programming. Data Structure, Pump, APC software, ACS software. Appropriate Bin selected on each pump. | incorrect Bin file selected on each pump
Final Assembly and Qualification. Final Programming. Data Structure, Pump, APC software, ACS software. Appropriate Bin selected on each pump. | Unable to write Bin file selected on each pump
Final Assembly and Qualification. Hang the pyrogen-free water bag in the purger canister. Pyrogen-free water bag, purger canister. Pyrogen-free water bag hanged correctly in the purger canister. | Incorrect placement of Pyrogen-free water bag in the purger canister
Final Assembly and Qualification. Holder position indicator. Holder, pumps, operator, sensor, ABC software. Holder at correct position. | Holder at incorrect position
Final Assembly and Qualification. Identify and isolate failed pumps. Operator, DMM, Controller, ABC software, ACS Software. Failed pumps identified and isolated. | Failed pumps not identified
Final Assembly and Qualification. If initial reading of Leakage Current Test fails, then identify failed pump(s). Failed pumps LEDs will turn RED​. DMM, Controller, ABC software, ACS Software. Correct initial readings and LED's turn red. | DMM not Zeroed prior to test
Final Assembly and Qualification. If initial reading of Leakage Current Test fails, then identify failed pump(s). Failed pumps LEDs will turn RED​. DMM, Controller, ABC software, ACS Software. Correct initial readings and LED's turn red. | DMM set incorrectly
Final Assembly and Qualification. If initial reading of Leakage Current Test fails, then identify failed pump(s). Failed pumps LEDs will turn RED​. DMM, Controller, ABC software, ACS Software. Correct initial readings and LED's turn red. | DMM reading error
Final Assembly and Qualification. If table is in Leakage Current test position, system activates locking magnet. AFQ system, Locking magnet, Table. Locking magnet is activated. | System fails to activate locking magnet
Final Assembly and Qualification. Leakage Current Water temp. Beaker, Temp. sensor, Solution, Heating plate, Stir bar,. Solution at correct temp.. | Solution at incorrect temp.
Final Assembly and Qualification. Leakage Current Water temp.. Beaker, Temp. sensor, Solution, Heating plate, Stir bar,. Solution at correct temp.. | Solution at incorrect temp.
Final Assembly and Qualification. Lifting the pumps from run beaker. Beaker, Pump, Holder, Operator. Pumps lifted from the beaker to the right height. | Pumps not lifted from the beaker to the right height
Final Assembly and Qualification. Location of pump on holder. ACS, Pump, Pump Holder, encoder. Pump location identified at Central station. | Pump unable to be located at central station
Final Assembly and Qualification. Lock the door with the hinges. Canister door, Operator, Hinges . Canister door locked. | Canister door not locked securely
Final Assembly and Qualification. Lower the bag into the canister and hook the bag on the canister door’s hook.. Purger Canister, Pyrogen-free water bag, hook. Bag lowered and hung on the canister door's hook. | Incorrect placement of bag
Final Assembly and Qualification. Lowering the pumps into run beaker. Beaker, Pump, Holder, Operator. Pumps lowered into the beaker to the right height. | Pumps interfere with beaker
Final Assembly and Qualification. Lowering the pumps into run beaker. Beaker, Pump, Holder, Operator. Pumps lowered into the beaker to the right height. | Pumps not lowered into the beaker
Final Assembly and Qualification. Lowering the pumps into run beaker. Beaker, Pump, Holder, Operator. Pumps lowered into the beaker to the right height. | Pumps not lowered into the beaker at the right height
Final Assembly and Qualification. Lowering the pumps into run beaker. Beaker, Pump, Holder, Operator. Pumps lowered into the beaker to the right height. | Pumps not lowered into the beaker to the right height
Final Assembly and Qualification. Measure conductivity of leakage current solution. Electrodes, Conductivity meter, operator, ABC software, ACS, solution. Conductivity reading > 0. | Incorrect reading
Final Assembly and Qualification. Measure conductivity of leakage current solution. Electrodes, Conductivity meter, operator, ABC software, ACS, solution. Conductivity reading > 0. | No conductivity reading
Final Assembly and Qualification. Measure motor leakage current of the pump. Pump with basic programming in it, high purity water, glass beaker, NAHCO3 solution, hot plate, Conductometer HANNA DiST3, purger, ten port pressure regulator, voltmeter, mounting device for lemo connector, leakage current measurement fixture, qualified water. Pump with measured leakage current. | Unable to detect pump with high leakage current
Final Assembly and Qualification. Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, mulitmeter, holding fixture. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Acceptable pumps fail
Final Assembly and Qualification. Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, multimeter, holding fixture. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Pumps with high leakage current pass
Final Assembly and Qualification. Measure purge flow for all pumps lines while running. Pumps . Purge flow measured . | Incorrect Purge flow reading
Final Assembly and Qualification. Measure purge flow for all pumps lines. Operator, Purge flow reading. Purge flow reading measured. | Incorrect measurement of Purge flow reading
Final Assembly and Qualification. Open Purger's air-line valve to the canister. Airline valve, AFQ system, Canister. Purger’s Airline valve is opened to the canister. | No purge
Final Assembly and Qualification. Operator lowers the fixture. Leakage current fixture, operator. Fixture lowered into the leakage current beaker. | Incorrect pump placement into the beaker
Final Assembly and Qualification. Perform functional testing of sensor. AIC, pump assembly, air pressure, gloves, run test plug holding fixture. Tested pump. | pump with malfunction sensor passes
Final Assembly and Qualification. Perform leakage test of the flush and pressure line. Pump with PSS attached, Tightness and pass ability test fixture, USON. Tested pump with PSS attached. | Damage purge/PSS due to fixture and/or handling
Final Assembly and Qualification. Perform leakage test of the flush and pressure line. Pump with PSS attached, Tightness and pass ability test fixture, USON. Tested pump with PSS attached. | Scrapped unit due to damage
Final Assembly and Qualification. Perform passability test. Pump assembly with Pressure Storage Set attached, ten port pressure regulator, USON Flow Tester, Tightness & Passability Test Fixture, Input and output pressure gauge. Tested Pump assembly with Pressure Storage Set attached. | pump with occlusions passes
Final Assembly and Qualification. Perform passability test. Pump assembly with Pressure Storage Set attached,Tightness & Passability Test Fixture, Input and output pressure gauge, smart monometer. Tested Pump assembly with Pressure Storage Set attached. | pump with occlusions passes
Final Assembly and Qualification. Perform pump run test. AIC console, pump assembly, AIC Pump Test Run DAQ Software, purge cassette, qualified water, beaker, run test fixture, runtest plug holding fixture. Tested pump. | Pump data becomes corrupted during test
Final Assembly and Qualification. Perform pump run test. AIC console, pump assembly, AIC Pump Test Run DAQ Software, purge cassette, qualified water, beaker, run test fixture, runtest plug holding fixture. Tested pump. | pump with errors passes (unable to run successfully)
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Inaccurate motor current measured
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Incorrect Average Max Motor Current
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Incorrect motor current value stored for pump
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Lost connectivity to database
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Low SNR leading to scrap
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Motor current is not properly written to database
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commercial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Sensor failure/low sensor signal
Final Assembly and Qualification. Perform pump run test. Tested pump with Serial Number Programmed and initial data structure, Commericial Optical AIC, Data Acquisition Software (DAQ), sterile water, purge cassette, one-click cleaner, beaker, high purity water. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Incorrect selection of bin
Final Assembly and Qualification. Perform purge, pressure, and PSS line pressure test. Pump assembly with Pressure Storage Set attached, USON Flow Tester, Tightness & Passability Test Fixture, Seal filter. Tested Pump assembly with Pressure Storage Set attached. | pump with weak leak passes
Final Assembly and Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump assembly with correct motor characterisitc curve. | File corrupted
Final Assembly and Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump assembly with correct motor characterisitc curve. | Incorrect curves uploaded leading to incorrect flow
Final Assembly and Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump with motor current curves programmed in it. | Corrupted data uploaded leading to device unable to run
Final Assembly and Qualification. Perform Tightness test of catheter junction. Pump with Serial Number Programmed, beaker, high purity water, pressure chamber, pressure supply regulator. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Acceptable pumps fail
Final Assembly and Qualification. Perform Tightness test of catheter junction. Pump with Serial Number Programmed, beaker, high purity water, pressure chamber, pressure supply regulator. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Damage pump (i.e. kink catheter)
Final Assembly and Qualification. Perform Tightness test of catheter junction. Pump with Serial Number Programmed, beaker, high purity water, pressure chamber, pressure supply regulator. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Damage pump/unusable
Final Assembly and Qualification. Perform Tightness test of catheter junction. Pump with Serial Number Programmed, beaker, high purity water, pressure chamber, pressure supply regulator. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Pumps with unacceptable leaks pass
Final Assembly and Qualification. Perform Tightness test of purge line. high purity water, beaker, pump with serial number, purge line pressure supply valve, timer. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Acceptable pumps fail
Final Assembly and Qualification. Perform Tightness test of purge line. high purity water, beaker, pump with serial number, purge line pressure supply valve, timer. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Damage pump/unusable
Final Assembly and Qualification. Perform Tightness test of purge line. high purity water, beaker, pump with serial number, purge line pressure supply valve, timer. Tested pump with Serial Number Programmed, G0-G3 loaded, and initial data structure. | Pumps with unacceptable low purge pressure pass
Final Assembly and Qualification. Place the electrode . Electrode, Operator, fixture. Electrode hanged. | Electrode placed incorrectly
Final Assembly and Qualification. Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Low SNR leading to scrap
Final Assembly and Qualification. Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Sensor failure/low sensor signal
Final Assembly and Qualification. Pre-stress PSS and pump joint.. Pump assembly with Pressure Storage Set attached, ten port regulator. pre-stressed pump assembly with PSS attached. | pump with leak
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, Computer, Operator. Fully programmed pump with write protection enabled. | Delayed process
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, Computer, Operator. Fully programmed pump with write protection enabled. | Programming step incomplete leading to incomplete data structure
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, Computer, Operator. Fully programmed pump with write protection enabled. | Pump fails to program (CRC error)
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, Computer, Operator. Fully programmed pump with write protection enabled. | Write protection fails to execute
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Write protection not enabled
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Pump with basic data structure. | 2nd Programming incomplete (CRC error)
Final Assembly and Qualification. Program EEPROM with final bin and enable write protection. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Pump with basic data structure. | Incorrect Bin number assigned to the pump (basic data structure)
Final Assembly and Qualification. Program pump serial number onto pump using the handle held programmer and labview software.. Handleheld programmer, pump with serial number on handle, PC, labview program . Pump with programmed serial number. | incorrect serial/corrupted data in handle
Final Assembly and Qualification. Program Serial Number and initial data structure. Pump Serial Number (from pump/WO) Jumper Box 2D Barcode Scanner Handheld 804. Pump with Serial Number Programmed and Initial Data Structure. | No Serial Number on pump
Final Assembly and Qualification. Program Serial Number and initial data structure. Pump Serial Number (from pump/WO) Jumper Box 2D Barcode Scanner Handheld 804. Pump with Serial Number Programmed and Initial Data Structure. | Unable to bin pump appropriately
Final Assembly and Qualification. Program Serial Number and initial data structure. Pump Serial Number (from pump/WO) Jumper Box 2D Barcode Scanner Handheld 804. Pump with Serial Number Programmed and Initial Data Structure. | Unable to perform run test
Final Assembly and Qualification. Program Serial Number and initial data structure. Pump Serial Number (from pump/WO) Jumper Box 2D Barcode Scanner Handheld 804. Pump with Serial Number Programmed and Initial Data Structure. | Wrong serial number programmed onto pump
Final Assembly and Qualification. Prompt the operator to verify purge fluid is dripping from each pump. AFQ system, Purge fluid, pump. System prompts the operator to verify purge fluid dripping for each pump. | System fails to instruct the Operator to verify purge fluid dripping for each pump
Final Assembly and Qualification. Pump detection on Pump Holder by Automated Central Station. Pump holder, Pump, Operator, LED, Pump Detector, Automated Central Station, ABC Software, ACS Software. Respective LED turned on. | Incorrect LED turns on
Final Assembly and Qualification. Pump detection on Pump Holder by Automated Central Station. Pump holder, Pump, Operator, LED, Pump Detector, Automated Central Station, ABC Software, ACS Software. Respective LED turned on. | LED does not turn on
Final Assembly and Qualification. Pump detection on Pump Holder by Automated Central Station. Pump holder, Pump, Operator, LED, Pump Detector, Automated Central Station, ABC Software, ACS Software. Respective LED turned on. | LED is not bright
Final Assembly and Qualification. Pump detection on Pump Holder by Automated Central Station. Pump holder, Pump, Operator, LED, Pump Detector, Automated Central Station, ABC Software, ACS Software. Respective LED turned on. | LED on without pump plugged in
Final Assembly and Qualification. Pump Holder present at staging area. Pump Holder, staging area. Pump holder present at staging area. | Pump holder at incorrect location
Final Assembly and Qualification. Pump holder present at the staging area. Pump Holder, staging area. Pump holder present at staging area. | Pump holder not present at the staging area
Final Assembly and Qualification. Pump lowered into beaker. Beaker, Holder, Switch, ABC software, ACS Software, Operator. Audible or Visual Indicator. | Incorrect Indicator
Final Assembly and Qualification. Pump lowered into beaker. Beaker, Holder, Switch, ABC software, ACS Software, Operator. Audible or Visual Indicator. | No Audible/ Visual Indication
Final Assembly and Qualification. Pump Run test. AIC, AFQ system, pumps. Run test operation started. | System fails to start run test operation
Final Assembly and Qualification. Pump Run test. Pump, Beaker, ACS Software, APC Software. Run test complete. | Pump run test not performed at P9
Final Assembly and Qualification. Pump Run test. Pump, Beaker, ACS Software, APC Software. Run test complete. | Pump run test performed for less than 20 seconds
Final Assembly and Qualification. Pump Run test. Pump, Beaker, ACS Software, APC Software. Run test complete. | Pump run test performed for more than 20 seconds
Final Assembly and Qualification. Pump Run test. Pump, Beaker, ACS Software, APC Software. Run test complete. | Pump run test performed for less than 20 seconds
Final Assembly and Qualification. Pump S/N and location displayed. Pumps, Central Station, Operator, ACS software. Pump S/N and location displayed. | Incorrect location or pair displayed
Final Assembly and Qualification. Pump S/N and location displayed. Pumps, Central Station, Operator, ACS software. Pump S/N and location displayed. | Incorrect S/N displayed
Final Assembly and Qualification. Pump S/N and location displayed. Pumps, Central Station, Operator, ACS software. Pump S/N confirmed. | Incorrect S/N displayed
Final Assembly and Qualification. Pump Unique Identifier programming. EEPROM, ACS software, APC software. Unique ID programmed. | Corrupted information
Final Assembly and Qualification. Pump Unique Identifier programming. EEPROM, ACS software, APC software. Unique ID programmed. | Unable to write
Final Assembly and Qualification. Pump Unique Identifier programming. EEPROM, ACS software, APC software. Unique ID programmed. | Unique ID erased
Final Assembly and Qualification. Purge pressure maintained. Purger, ACS/ ABC software. Purger pressure at 900 mm Hg. | Purger pressure incorrect
Final Assembly and Qualification. Record conductivity of NaHCO3 level in the rinse beaker. Rinse beaker, conductivity meter, Pyrogen-free water. Conductivity of NaHCO3 level in the rinse beaker recorded. | Incorrect conductivity reading
Final Assembly and Qualification. Recording pump programming status. EEPROM content verified and CRC checking performed. Programming and Write Protection reports generated. | No reports generated
Final Assembly and Qualification. Recording Run test data. Run test data, ACS software, external storage, internal storage, APC software, pumps. Maximum mean motor current recorded. | Data not received into ACS
Final Assembly and Qualification. Recording Run test data. Run test data, ACS software, external storage, internal storage, APC software, pumps. Maximum mean motor current recorded. | Data unable to be stored in External drive
Final Assembly and Qualification. Recording Run test data. Run test data, ACS software, external storage, internal storage, APC software, pumps. Maximum mean motor current recorded. | Incorrect data received
Final Assembly and Qualification. Recording Run test data. Run test data, ACS software, external storage, internal storage, APC software, pumps. Maximum mean motor current recorded. | Motor current incorrectly measured and fails specification (false positive)
Final Assembly and Qualification. Recording Run test data. Run test data, ACS software, external storage, internal storage, APC software, pumps. Pump run test graphs generated. | No graphs generated
Final Assembly and Qualification. Replacing the Rinse solution. Rinse solution, Beaker,. Rinse solution not over saturated. | Rinse solution is over saturted with Sodium Bicarbonate
Final Assembly and Qualification. Run pumps at P-5 P-level for 10 seconds for all pumps. Pumps, pump holder, rinse beaker. Pumps run at P-5 level for 5 seconds for all pumps. | Pumps run at P-5 level for less than 10 seconds for all pumps
Final Assembly and Qualification. S/N displayed on the Central Station. Pump S/N , Operator, Pump Holder. Correct Pump & Holder # displayed . | Incorrect Pump # and Holder # displayed
Final Assembly and Qualification. S/N displayed on the Central Station. SAP, Operator, Pump Holder. Correct Pump & Holder # displayed. | Incorrect Pump # and Holder # displayed
Final Assembly and Qualification. Save pump "current" profile (waveforms) (for all pumps). AFQ system, operator. Pump Current profile saved for all pumps. | Unable to save Pump current profile (waveforms) for all pumps
Final Assembly and Qualification. Select appropriate data structure bin from data storage location. Data Structure, Pump, APC software, ACS software. Appropriate Data structure selected from data storage location. | Unable to select Appropriate Data Structure from data storage location
Final Assembly and Qualification. Sensor Verification Test. Pump with G0-G3 loaded, serial number, and initial data structure, Deluxe Console Gen 3, Pressure Chamber Fixture. Pump with verified sensor, GO-G3 loaded, initial data structure and serial number. | Sensor not functioning
Final Assembly and Qualification. Setup ready. Central station, ABC, APC, Beaker setup, electrodes, operator, User interface. Setup complete. | Setup not complete
Final Assembly and Qualification. Setup ready. Central station, ABC, APC, Beaker setup, electrodes, operator. Setup complete. | Setup not complete
Final Assembly and Qualification. System ​detects that the pump holder is in the leakage beaker and starts the current leakage test. AFQ system, pump holder, pumps, leakage beaker. Leakage current test starts. | System fails to detect the pump holder in the leakage beaker
Final Assembly and Qualification. System boot-up. Central station, Controllers (Bench and pump), User. Turned on AFQ system. | System is not turned on
Final Assembly and Qualification. System boot-up. Central station, Controllers, AFQ system (all automated eq. ). Turned on AFQ system. | ACS is not turned on
Final Assembly and Qualification. System boot-up. Central station, Controllers, AFQ system (all automated eq. ). Turned on AFQ system. | System is not turned on
Final Assembly and Qualification. System initiation. AFQ system, ACS Software,. System Ready Acknowledge. | System is not acknowledged
Final Assembly and Qualification. System Login . Central station, Operator, Software, Credentials. Operator logged into the PC. | Operator unable to logon to the system
Final Assembly and Qualification. System Login. Central station, Operator, Software, Credentials. Operator logged into the PC. | Operator unable to logon to the system
Final Assembly and Qualification. System starts Rinse operation. AFQ system, pumps, pump holder, controller. Rinse Operation started. | System fails to start rinse operation
Final Assembly and Qualification. System to instruct operator to turn table to the Leakage Current test position. AFQ system, Rotating table. System instructs operator to turn the table to the leakage current test position. | System unable to instruct operator to turn the table to the leakage current test position
Final Assembly and Qualification. System turns on Leakage Current excitation voltage to verify Leakage Current test setup and indicates Red/Green. AFQ System. Leakage Current excitation voltage is turned on and Leakage Current test setup is verified by the system. | System fails to verify the Leakage current test setup
Final Assembly and Qualification. Test leakge at catheter junction. Pump assembly, glass beaker, ten port pressure regulator, leak test fixture, calibrated timer, qualified water. Leakage tested pump. | Pump with leak at catheter junction passes
Final Assembly and Qualification. Test motor leakage current of the pump. Pump with basic programming in it, high purity water, glass beaker, NAHCO3 solution, hot plate, Conductometer HANNA DiST3, purger, ten port pressure regulator, voltmeter, mounting device for lemo connector, leakage current measurement fixture, qualified water. Pump with measured leakage current. | Unable to detect pump with high leakage current
Final Assembly and Qualification. Turn ON the water heater. Water Heater, Controller, Central station. Heated water to correct temp.. | Water not heated to correct temp.
Final Assembly and Qualification. Verify EEPROM content. CRC check for all pumps. Inform the Operator that data is good.. Operator, EEPROM, APC software, ACS software, pumps. EEPROM content verified and CRC checking performed. | Unable to verify the EEPROM content and perform CRC check for all pumps
Final Assembly and Qualification. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Board corruption
Final Assembly and Qualification. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Damage pump surface leading to scrap
Final Assembly and Qualification. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | ESD damage to pump during handling leading to pump stop start issues
Final Assembly and Qualification. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | ESD damage to pump during handling leading to scrap
Final Assembly and Qualification. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. Pump, pump programmer, dynamic binning software, AIC/MPC Conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Sensor failure/low sensor signal
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Leakage current incorrectly recorded
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current failure pump identified. | False negative
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current failure pump identified. | Leakage current incorrectly recorded for wrong pump S/N
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current failure pump identified. | Leakage current tested and passes test even though the pump has leakage current (false negative)
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Leakage current not recorded
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | No leakage Current value recorded
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Incorrect Voltage to current conversion
Final Assembly and Qualification. Voltage from DMM read into ACS software . DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Leakage current not recorded in the database
Final Assembly and Qualification. Voltage from DMM read into ACS software. DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Incorrect Volatage to current translation
Final Assembly and Qualification. Voltage from DMM read into ACS software. DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | Leakage current not recorded in the database
Final Assembly and Qualification. Voltage from DMM read into ACS software. DMM, Controller, ABC software, ACS Software, local database, external database. Leakage current recorded. | No leakage Current value displayed
Final Assembly and Qualification. Wait for 2 mins. Prior to reading. DMM, ABC, ACS software. 2 min delay. | Incorrect delay in time
Final Assembly and Qualification. Write Data structure to EEPROM for all pumps except for gauge factor section​. ACS software, Pump, EEPROM, APC software. Data Structure written to EEPROM for all pumps except for gauge factor section. | Unable to write Data Structure to EEPROM for all pumps except for gauge factor section
Final Assembly and Qualification. Write the current value into the database. Operator, AFQ system. Current value stored in database. | Unable to save the current value into the database
Final Assembly and Qualification. Zero the DMM. DMM, Controller, ABC software, APC Softwre, ACS Software. DMM Zeroed. | Incorrect Leakage Current Reading
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm damaged
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge system occluded or break not detected
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection blocked
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection bond failure
Final Assm & Qualification (0550-7041). Assemble Purge Side Arm. Petroleum ether Pump PSS Banderole Glue Dispenser Adhesive Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
Final Assm & Qualification (0550-7041). Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Gage factors/G0 not loaded
Final Assm & Qualification (0550-7041). Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Low SNR leading to scrap
Final Assm & Qualification (0550-7041). Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Sensor failure/low sensor signal
Final Assm & Qualification (0550-7041). Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Wrong Gage factors/G0 offset loaded
Final Assm & Qualification (0550-7041). Dry in Oven. Oven Oven Rack Temperature Time Air manifold Air pressure/flow rate Pump with serial number programmed. Dry Pump. | Pump not fully dry, leads to bacterial growth
Final Assm & Qualification (0550-7041). Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, mulitmeter, holding fixture. Tested pump with Serial Number Programmed. | Acceptable pumps fail
Final Assm & Qualification (0550-7041). Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, multimeter, holding fixture. Tested pump with Serial Number Programmed. | Pumps with high leakage current pass
Final Assm & Qualification (0550-7041). Perform leakage test of the purge line. Pump with PSS attached, Tightness test fixture, USON. Tested pump with PSS attached. | Damage purge/PSS due to fixture and/or handling
Final Assm & Qualification (0550-7041). Perform leakage test of the purge line. Pump with PSS attached, Tightness test fixture, USON. Tested pump with PSS attached. | Scrapped unit due to damage
Final Assm & Qualification (0550-7041). Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Low SNR leading to scrap
Final Assm & Qualification (0550-7041). Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Sensor failure/low sensor signal
Final Assm & Qualification (0550-7041). Program EEPROM with final bin (2nd programming). Pump Deluxe Console Gen 3 Jumper Box. Pump with basic data structure. | 2nd Programming incomplete (CRC error)
Final Assm & Qualification (0550-7041). Program EEPROM with final bin (2nd programming). Pump Deluxe Console Gen 3 Jumper Box. Pump with basic data structure. | Pump loaded with incorrect bin file
Final Assm & Qualification (0550-7041). Program Serial Number. Pump Serial Number (from pump/WO) Deluxe Console Gen 3 Jumper Box 2D Barcode Scanner. Pump with Serial Number Programmed. | No Serial Number on pump
Final Assm & Qualification (0550-7041). Program Serial Number. Pump Serial Number (from pump/WO) Deluxe Console Gen 3 Jumper Box 2D Barcode Scanner. Pump with Serial Number Programmed. | Wrong serial number programmed onto pump
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Purge flow). | Damage catheter scrapping/kinking
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Purge flow). | Pass failed units for purge flow (high purge pressure)
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Inaccurate motor current measured
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Incorrect Average Max Motor Current
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Incorrect selection of bin
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Low SNR leading to scrap
Final Assm & Qualification (0550-7041). Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Sensor failure/low sensor signal
Final Assm & Qualification (0550-7041). Sensor Verification Test. Pump with G0-G3 loaded and serial number. Pump with verified sensor, GO-G3 loaded, and serial number. | Sensor not functioning
Final Assm & Qualification. 1st Programming. Pump with verified sensor, GO-G3 loaded, and serial number. Pump with 1st programming complete. | Damage sensor due to handling/fixture
Final Assm & Qualification. 1st Programming. Pump with verified sensor, GO-G3 loaded, and serial number. Pump with 1st programming complete. | Pump not programmed
Final Assm & Qualification. 1st Programming. Pump with verified sensor, GO-G3 loaded, and serial number. Pump with 1st programming complete. | Pump programmed incorrectly
Final Assm & Qualification. 1st Programming. Pump with verified sensor, GO-G3 loaded, and serial number. Pump with 1st programming complete. | Sensor not functioning
Final Assm & Qualification. Assemble Purge Side Arm (PSS). Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge side arm connection leaks
Final Assm & Qualification. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge side arm damaged
Final Assm & Qualification. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge system occluded or break not detected
Final Assm & Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge Side Arm connection blocked
Final Assm & Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge Side Arm connection bond failure
Final Assm & Qualification. Assemble Purge Side Arm. Petroleum ether Pump PSS Glue Dispenser Adhesive Holne UV Box Syringe Scissors/wire cutters. Pump with PSS attached. | Purge side arm connection leaks
Final Assm & Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Gage factors/G0 not loaded
Final Assm & Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Low SNR leading to scrap
Final Assm & Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Sensor failure/low sensor signal
Final Assm & Qualification. Calibrate sensor (i.e. load Gage Factors). Pump with serial number programmed Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded and serial number. | Wrong Gage factors/G0 offset loaded
Final Assm & Qualification. Dry in Oven. Oven Oven Rack Temperature Time Air manifold Air pressure/flow rate Pump with serial number programmed. Dry Pump. | Pump not fully dry, leads to bacterial growth
Final Assm & Qualification. Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, multimeter, holding fixture. Tested pump with Serial Number Programmed. | Acceptable pumps fail
Final Assm & Qualification. Measure motor leakage current. Pump with Serial Number Programmed, glass beaker, high purity water, 2% bicarbonate, hot plate, purger, power supply, multimeter, holding fixture. Tested pump with Serial Number Programmed. | Pumps with high leakage current pass
Final Assm & Qualification. Perform leakage test of the purge line. Pump with PSS attached, Tightness test fixture, USON. Tested pump with PSS attached. | Damage purge/PSS due to fixture and/or handling
Final Assm & Qualification. Perform leakage test of the purge line. Pump with PSS attached, Tightness test fixture, USON. Tested pump with PSS attached. | Scrapped unit due to damage
Final Assm & Qualification. Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Low SNR leading to scrap
Final Assm & Qualification. Prep pump and deluxe console for use. Pump Handle assembly, One-Click cleaner, . Prepped Pump handle assembly and deluxe console. | Sensor failure/low sensor signal
Final Assm & Qualification. Program EEPROM with final bin (2nd programming). Pump Deluxe Console Gen 3 Jumper Box. Pump with basic data structure. | 2nd Programming incomplete (CRC error)
Final Assm & Qualification. Program EEPROM with final bin (2nd programming). Pump Deluxe Console Gen 3 Jumper Box. Pump with basic data structure. | Pump loaded with incorrect bin file
Final Assm & Qualification. Program Serial Number. Pump Serial Number (from pump/WO) Deluxe Console Gen 3 Jumper Box 2D Barcode Scanner. Pump with Serial Number Programmed. | No Serial Number on pump
Final Assm & Qualification. Program Serial Number. Pump Serial Number (from pump/WO) Deluxe Console Gen 3 Jumper Box 2D Barcode Scanner. Pump with Serial Number Programmed. | Wrong serial number programmed onto pump
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Purge flow). | Damage catheter scrapping/kinking
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Purge flow). | Pass failed units for purge flow (high purge pressure)
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Inaccurate motor current measured
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Incorrect Average Max Motor Current
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Low SNR leading to scrap
Final Assm & Qualification. Purge flow run test. Purge rate test setup (pressure supply, tubings), finished pump assembly, console, flow sensors. Tested pump with Serial Number Programmed. | Sensor failure/low sensor signal
Final Assm & Qualification. Sensor Verification Test. Pump with G0-G3 loaded and serial number. Pump with verified sensor, GO-G3 loaded, and serial number. | Sensor not functioning
Final Inspection (0550-7042). Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage Impeller blades/tip
Final Inspection (0550-7042). Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage pump inner surfaces
Final Inspection (0550-7042). Install Sidearm Retainer. Tested full pump assembly, Retainer Top Shell, Retainer Bottom Shell. Tested full pump assembly with installed sidearm retainer. | retainer halves not fully engaged
Final Inspection (0550-7042). Install Sidearm Retainer. Tested full pump assembly, Retainer Top Shell, Retainer Bottom Shell. Tested full pump assembly with installed sidearm retainer. | retainer haves fail to engage
Final Inspection (0550-7042). Measure sensor SNR. SNR Test Fixture, Pump with PSS. Tested full pump assembly. | Sensor failure/low sensor signal
Final Inspection (0550-7042). Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Board corruption
Final Inspection (0550-7042). Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Damage pump surface leading to scrap
Final Inspection (0550-7042). Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | ESD damage to pump during handling leading to pump stop start issues
Final Inspection (0550-7042). Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | ESD damage to pump during handling leading to scrap
Final Inspection (0550-7042). Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Sensor failure/low sensor signal
Final Inspection. Air and O2 Function Test. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | With Air and O2 outlet tubing clamped, after two minutes, two high priority alarms are not reported: "Air Flow Out of Range" and "Oxygen Concentrator Flow Out of Range".
Final Inspection. Battery Disconnect Switch in "Disconnect" position. Test Completion. Test Completion Completed. | Battery Disconnect Switch not in "Disconnect" position
Final Inspection. Check Battery Charge Indicator (Blue LED). OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Battery Charge Indicator is illuminated when Console is not connected to A/C power
Final Inspection. Check Battery Charge Indicator (Blue LED). OXY-1 Functional Test 16 Hours. OXY-1 Functional Test 16 Hours Completed. | Battery Charge Indicator not illuminated when Console is connected to A/C power
Final Inspection. Check Battery Icon. OXY-1 Functional Test 16 Hours. OXY-1 Functional Test 16 Hours Completed. | Battery icon does not have 1-4 bars or lightning bolt present
Final Inspection. Check Console Date and time. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Date and time not verified
Final Inspection. Connect External Air. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Console Display indicated gas mode did not change from "Internal" to "External"
Final Inspection. Connect Tubing. OXY-1 ESS Test Preparation. OXY-1 ESS Test Preparation Completed. | Tube softening and deattachment from Console
Final Inspection. Console Burn In. OXY-1 End of 16 Hour Burn in. OXY-1 End of 16 Hour Burn in Completed. | "Data Export Complete" message not present
Final Inspection. Console Burn In. OXY-1 End of 16 Hour Burn in. OXY-1 End of 16 Hour Burn in Completed. | Operator does not recognize that other alarms are present (in addition to “Running on Battery” and “Powering Down –Press Power Button to Cancel”) during the burn in run time
Final Inspection. Console Burn In. OXY-1 End of 16 Hour Burn in. OXY-1 End of 16 Hour Burn in Completed. | Running on battery alarm did not occur in conjunction with loss of A/C power
Final Inspection. Console Burn In. OXY-1 End of 16 Hour Burn in. OXY-1 End of 16 Hour Burn in Completed. | SOC (State of Charge) < 30%
Final Inspection. Console Burn In. OXY-1 Start of 16 Hour Burn in. OXY-1 Start of 16 Hour Burn in Completed. | Operator does not recognize battery is not fully charged
Final Inspection. Console Burn In. OXY-1 Start of 16 Hour Burn in. OXY-1 Start of 16 Hour Burn in Completed. | Operator does not recognize alarms are present
Final Inspection. Console Burn In. OXY-1 Start of 16 Hour Burn in. OXY-1 Start of 16 Hour Burn in Completed. | Operator does not recognize console stops running or the display is inactive
Final Inspection. Console Burn In. OXY-1 Start of 16 Hour Burn in. OXY-1 Start of 16 Hour Burn in Completed. | Operator does not recognize flow rate and RPM strayed from their setpoints
Final Inspection. Enable write protection & Perform PPM Check. Write protection & PPM tool, pump (un-sterile), hand held programmer, AIC/ MPC conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Write protection not enabled
Final Inspection. Enable write protection & Perform PPM Check. Write protection tool & PPM front end tool, pump (un-sterile), hand held programmer, AIC/ MPC conversion box, computer, operator. Correctly Programmed pump with Write protection enabled. | Write protection not enabled
Final Inspection. ESS Room Phases 1-2. OXY-1 End of ESS Room. OXY-1 End of ESS Room Completed. | PLU unplugged from console when the unit is on
Final Inspection. ESS Room Phases 1-2. OXY-1 End of ESS Room. OXY-1 End of ESS Room Completed. | Pump RPM, O2, Air Flow not correctly set
Final Inspection. ESS Room Phases 1-2. OXY-1 Start of ESS Room. OXY-1 Start of ESS Room Completed. | Operator does not recognize battery is not fully charged
Final Inspection. ESS Room Phases 1-2. OXY-1 Start of ESS Room. OXY-1 Start of ESS Room Completed. | AC cord not connected to IEC Inlet on UUT
Final Inspection. ESS Room Phases 1-2. OXY-1 Start of ESS Room. OXY-1 Start of ESS Room Completed. | Operator does not recognize alarms are present
Final Inspection. ESS Room Phases 1-2. OXY-1 Start of ESS Room. OXY-1 Start of ESS Room Completed. | Operator does not recognize console stops running or the display is inactive
Final Inspection. ESS Room Phases 1-2. OXY-1 Start of ESS Room. OXY-1 Start of ESS Room Completed. | Operator does not recognize flow rate and RPM strayed from their setpoints
Final Inspection. Examine Front Shell, Right Shell, Left Shell, Rear Shell, Display panel surface, Casters, Exterior Surfaces for discoloration, markings, scratches. Final Cosmetic Review. Final Cosmetic Review Completed. | Operator does not recognize discoloration, markings, scratches are present
Final Inspection. Examine seam interface between shells. Final Cosmetic Review. Final Cosmetic Review Completed. | Operator does not recognize if gasket is present, seam interface is not uniform. Operator does not recognize if gasket is not present, seam gap is greater than 2mm
Final Inspection. Input Air and O2 Settings and Adjust Blood Flow Range and Pump Driver RPM. OXY-1 ESS Test Preparation. OXY-1 ESS Test Preparation Completed. | Alarms present after running for several minutes
Final Inspection. Input Air and O2 Settings and Adjust Blood Flow Range and Pump Driver RPM. OXY-1 Functional Test 16 Hours. OXY-1 Functional Test 16 Hours Completed. | Alarms present after running for several minutes
Final Inspection. Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage Impeller blades/tip
Final Inspection. Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage pump inner surfaces
Final Inspection. Install Battery Switch Plate. Test Completion. Test Completion Completed. | 4 fasteners not used
Final Inspection. Install PSS Top Housing. Tested full pump assembly, PSS Top Housing. Tested full pump assembly with installed PSS Housing. | Damage to PSS Housing
Final Inspection. Install PSS Top Housing. Tested full pump assembly, PSS Top Housing. Tested full pump assembly with installed PSS Housing. | PSS housing halves fail to engage
Final Inspection. Install PSS Top Housing. Tested full pump assembly, PSS Top Housing. Tested full pump assembly with installed PSS Housing. | PSS Housing halves not fully engaged
Final Inspection. Internal Stored Data Log. Test Completion. Test Completion Completed. | Internal stored data log not cleared
Final Inspection. Measure sensor SNR. SNR Test Fixture, Pump with PSS. Tested full pump assembly. | Sensor failure/low sensor signal
Final Inspection. Place Portable Mock Disposable loop in Pump Driver and connect Pump Driver to Console. OXY-1 ESS Test Preparation. OXY-1 ESS Test Preparation Completed. | Pump driver incorrectly connected to console
Final Inspection. Place Stationary Mock Disposable loop in Pump Driver and connect Pump Driver to Console. OXY-1 Functional Test Preparation. OXY-1 Functional Test Preparation Completed. | Pump driver incorrectly connected to console
Final Inspection. Prime Portable Mock Disposable Loop. OXY-1 ESS Test Preparation. OXY-1 ESS Test Preparation Completed. | Portable Mock Disposable Loop not primed
Final Inspection. Prime Stationary Mock Disposable Loop. OXY-1 Functional Test Preparation. OXY-1 Functional Test Preparation Completed. | Stationary Mock Disposable Loop not correctly primed
Final Inspection. Pump Driver LED Functional Check. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Pump Driver LED is off when Blood flow LPM is greater than zero
Final Inspection. Pump Driver LED Functional Check. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Pump Driver LED is on when Blood flow LPM is zero
Final Inspection. Set External Gas to 0 LPM. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Console Display indicated gas mode did not change from "External" to "Internal" and "Lost External Gas - Confirm Internal Settings" alarm not present
Final Inspection. Setup and Connect Outlet Timer to Console. OXY-1 Functional Test Preparation. OXY-1 Functional Test Preparation Completed. | Outlet Timer incorrectly setup
Final Inspection. Sigh Function Test. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | Alarms present after 5 minutes
Final Inspection. Transport of Console and Pump Driver. OXY-1 Functional Test Preparation. OXY-1 Functional Test Preparation Completed. | Pump driver damage
Final Inspection. Turn off Console and disconnect from AC wall power. Test Completion. Test Completion Completed. | Console not turned off
Final Inspection. Turn on Console and set Date and Time and Clear Logs. OXY-1 Functional Test 16 Hours. OXY-1 Functional Test 16 Hours Completed. | Date and time not correctly set
Final Inspection. Verify console gas mode. OXY-1 Functional Test 16 Hours. OXY-1 Functional Test 16 Hours Completed. | Console is in External O2 mode
Final Inspection. Verify Label Placement. Final Cosmetic Review. Final Cosmetic Review Completed. | Labels are not affixed
Final Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Board corruption
Final Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Damage pump surface leading to scrap
Final Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | ESD damage to pump during handling leading to pump stop start issues
Final Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | ESD damage to pump during handling leading to scrap
Final Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | Sensor failure/low sensor signal
Final Inspection. Volume Test. OXY-1 Additional Testing. OXY-1 Additional Testing Completed. | No audio output when volume level is adjusted
Final Inspection. Walkabout Phase 1-2. OXY-1 End of Walkabout. OXY-1 End of Walkabout Completed. | "Data Export Complete" message not present
Final Inspection. Walkabout Phase 1-2. OXY-1 End of Walkabout. OXY-1 End of Walkabout Completed. | Operator does not recognize that other alarms are present (in addition to “Running on Battery” and “System Shutting Down”) during the burn in run time
Final Inspection. Walkabout Phase 1-2. OXY-1 End of Walkabout. OXY-1 End of Walkabout Completed. | Running on battery alarm did not occur in conjunction with loss of A/C power
Final Inspection. Walkabout Phase 1-2. OXY-1 Start of Walkabout. OXY-1 Start of Walkabout Completed. | Cosmetic defect on display
Final Inspection. Walkabout Phase 1-2. OXY-1 Start of Walkabout. OXY-1 Start of Walkabout Completed. | O2 Concentrator Inactive or improper operation
Final Inspection. Walkabout Phase 1-2. OXY-1 Start of Walkabout. OXY-1 Start of Walkabout Completed. | O2, Air Flow not correctly set
Final Inspection. Walkabout Phase 1-2. OXY-1 Start of Walkabout. OXY-1 Start of Walkabout Completed. | Operator does not recognize alarms are present
Final Qualification, Final Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Domino Printer (ink), Finished pump assembly . Finished pump assembly with serial number. | Damage lemo connector pins (scrap)
Final Qualification, Final Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Domino Printer (ink), Finished pump assembly . Finished pump assembly with serial number. | missing/incorrect/smudged serial number leading to loss of traceability
Final Qualification, Final Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | Corrupted data leading to pump stop or failure to start
Final Qualification, Final Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | incorrect gains leading to pump stop or failure to start
Final Qualification, Final Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | Incorrect programmed serial number
Final Qualification, Final Visual Inspection. Connect to test setup, apply pressure level, remove pressure and verify pressure decay vs. time meets specification. Smart Monometer, finished pump assembly, pressure setup, Test Console (EEPROM programmer), Timer. Verified finished pump assembly. | Damage lemo connector pins (scrap)
Final Qualification, Final Visual Inspection. Connect to test setup, apply pressure level, remove pressure and verify pressure decay vs. time meets specification. Smart Monometer, finished pump assembly, pressure setup, Test Console (EEPROM programmer), Timer. Verified finished pump assembly. | pump with unacceptable pressure reading passes
Final Qualification, Final Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V, after that with +10V and -10V. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finishe pump assembly. | Contimnation of pump (sterility compromised)
Final Qualification, Final Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V, after that with +10V and -10V. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finishe pump assembly. | pump with leakage in purge line passes
Final Qualification, Final Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V, after that with +10V and -10V. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finished pump assembly. | Salt residue on the pump leading to corrosion
Final Qualification, Final Visual Inspection. Electrical test for sensor measurement system. Pump, dark box, electrical measurement setup (mulitmeter, power supply), software (ESD Measurement). Tested pump unit. | Damage sensor (ESD)
Final Qualification, Final Visual Inspection. Electrical test for sensor measurement system. Pump, dark box, electrical measurement setup (mulitmeter, power supply), software (ESD Measurement). Tested pump unit. | Failed sensor not detected
Final Qualification, Final Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | contamination of fluid path
Final Qualification, Final Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | Contamination of pump surface
Final Qualification, Final Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | corrosion of inner mechanical components leading to pump stop
Final Qualification, Final Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | Damage catheter scrapping/kinking
Final Qualification, Final Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | contamination of fluid path
Final Qualification, Final Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | contamination of fluid path
Final Qualification, Final Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too high)
Final Qualification, Final Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too high) - scrap
Final Qualification, Final Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too low)
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Damage impeller leading to no flow
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Damage Impeller/Pump housing internal surface/cannula internal surface
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump fails to start/corrupted data
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump with incorrect bins is accepted
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump with incorrect gains accepted
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Product free of particulates/debris. | Particluate accumulation
Final Qualification, Final Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Product free of particulates/debris. | Particluate migrate to other product
Final Qualification, Final Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM progrmmer), Formula, calculator, 20 degree pyrogen free water. Calibrated sensor on finished pump assembly. | Damage lemo connector pins (scrap)
Final Qualification, Final Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM progrmmer), Formula, calculator, 20 degree pyrogen free water. Calibrated sensor on finished pump assembly. | pump sensor incorrectly calibrated
Final Qualification, Final Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM progrmmer), Formula, calculator, 36.5 degree pyrogen free water. Calibrated sensor on finished pump assembly. | pump sensor incorrectly calibrated
Final Qualification, Final Visual Inspection. Position pump plug in print device; dial in pump serial number and perform printing. Domino Printer (ink), Finished pump assembly . Finished pump assembly with serial number. | missing/incorrect/smudged serial number leading to loss of traceability
Final Qualification, Final Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Damage lemo connector pins (scrap)
Final Qualification, Final Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Documenting bad pumps as good, sensor failure
Final Qualification, Final Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Documenting good pumps as bad (scrap)
Final Qualification, Final Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | incorrect motor current bin loaded leading to lower/inaccurate flows
Final Qualification, Final Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Unable to perform pump run test
Final Qualification, Final Visual Inspection. Purge flow rate test. Purge rate test setup (monometer, pressure supply, tubings), finished pump assembly. Measured assembly (Airflow). | Damage catheter scrapping/kinking
Final Qualification, Final Visual Inspection. Purge flow rate test. Purge rate test setup (monometer, pressure supply, tubings), finished pump assembly. Measured assembly (Airflow). | Pass failed units for airflow (high purge pressure)
Final Qualification. Clean 2x with petroleum ether. Then attach Pressure Storage Set with filter to pump assembly. Then dry in Oven. Then clean with petroleum ether again and apply UV adhesive.. THF, Pressure storage set, pump assembly, dosing needle, petroleum ether, oven (time and temperature). Pump assembly with Pressure Storage Set attached. | Bond Failure
Final Qualification. Clean 2x with petroleum ether. Then attach Pressure Storage Set with filter to pump assembly. Then dry in Oven. Then clean with petroleum ether again and apply UV adhesive.. THF, Pressure storage set, pump assembly, dosing needle, petroleum ether, oven (time and temperature). Pump assembly with Pressure Storage Set attached. | Kink/Bond failure
Final Qualification. Create work order. Print the pump serial number on handle. Finished AU pump plug, Domino Printer. Pump with serial number on handle. | incorrect serial/missing serial number
Final Qualification. Dry pumps in oven. Pump assembly, oven (time and temperature), hanging rank, air pressure. Dry pumps. | Pumps not dry leading to contamination issues
Final Qualification. Measure motor leakage current of the pump. Pump with basic programming in it, high purity water, glass beaker, NAHCO3 solution, hot plate, Conductometer HANNA DiST3, purger, ten port pressure regulator, voltmeter, mounting device for lemo connector, leakage current measurement fixture, high purity water. Pump with measured leakage current. | Unable to detect pump with high leakage current
Final Qualification. Perform functional testing of sensor. AIC, pump assembly, air pressure, gloves, run test plug holding fixture. Tested pump. | pump with malfunction sensor passes
Final Qualification. Perform passability test. Pump assembly with Pressure Storage Set attached, ten port pressure regulator, USON Flow Tester, Tightness & Passability Test Fixture, Input and output pressure gauge. Tested Pump assembly with Pressure Storage Set attached. | pump with occlusions passes
Final Qualification. Perform passability test. Pump assembly with Pressure Storage Set attached,Tightness & Passability Test Fixture, Input and output pressure gauge, smart monometer. Tested Pump assembly with Pressure Storage Set attached. | pump with occlusions passes
Final Qualification. Perform pump run test. AIC console, pump assembly, AIC Pump Test Run DAQ Software, purge cassette, high purity water, beaker, run test fixture, runtest plug holding fixture. Tested pump. | Pump data becomes corrupted during test
Final Qualification. Perform pump run test. AIC console, pump assembly, AIC Pump Test Run DAQ Software, purge cassette, high purity water, beaker, run test fixture, runtest plug holding fixture. Tested pump. | pump with errors passes (unable to run successfully)
Final Qualification. Perform purge, pressure, and PSS line pressure test. Pump assembly with Pressure Storage Set attached, USON Flow Tester, Tightness & Passability Test Fixture, Seal filter. Tested Pump assembly with Pressure Storage Set attached. | pump with weak leak passes
Final Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump assembly with correct motor characterisitc curve. | File corrupted
Final Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump assembly with correct motor characterisitc curve. | Incorrect curves uploaded leading to incorrect flow
Final Qualification. Perform second expanded programming and upload the motor characterisitc current curves. Handleheld programmer, pump assembly, PC, motor characteristic curve file. Pump with motor current curves programmed in it. | Corrupted data uploaded leading to device unable to run
Final Qualification. Pre-stress PSS and pump joint.. Pump assembly with Pressure Storage Set attached, ten port regulator. pre-stressed pump assembly with PSS attached. | pump with leak
Final Qualification. Program pump serial number onto pump using the handle held programmer and labview software.. Handleheld programmer, pump with serial number on handle, PC, labview program . Pump with programmed serial number. | incorrect serial/corrupted data in handle
Final Qualification. Test leakge at catheter junction. Pump assembly, glass beaker, ten port pressure regulator, leak test fixture, calibrated timer, high putiry water. Leakage tested pump. | Pump with leak at catheter junction passes
Final Qualification. Test motor leakage current of the pump. Pump with basic programming in it, high purity water, glass beaker, NAHCO3 solution, hot plate, Conductometer HANNA DiST3, purger, ten port pressure regulator, voltmeter, mounting device for lemo connector, leakage current measurement fixture, high purity water. Pump with measured leakage current. | Unable to detect pump with high leakage current
Final Visual Inspection. Check the Firmware and activate write protection. Pump assembly with pressure storage set, pump programmer. Write protected tested pump assembly. | write protection not activated
Final Visual Inspection. Inspect inner cannula surface and top of impellar. Pump assembly with pressure storage set, endoscope. Inspected pump assembly with pressure storage set attached. | Damage surface of cannula and tip of Impellar
Final Visual Inspection. Inspect inner cannula surface and top of impellar. Pump assembly with pressure storage set, endoscope. Inspected pump assembly with pressure storage set attached. | Damage surface of cannula and tip of Impellar are passed
Final Visual Inspection. Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage Impeller blades/tip
Final Visual Inspection. Inspect pump inner cannula surfaces, pump housing, and Impeller. Pump with PSS attached, Endoscope. Inspected pump with PSS attached. | Damage pump inner surfaces
Final Visual Inspection. Measure sensor SNR. SNR Test Fixture, Pump with PSS. Tested full pump assembly. | Sensor failure/low sensor signal
Final Visual Inspection. Perform visual inspection and then check data structure on pump. Pump assembly with pressure storage set, pump programmer. Tested Pump assembly with Pressure Storage Set attached. | Pump with incorrect data structure passes/data structure becomes corrupted
Final Visual Inspection. Verify, CRC check sum, Serial Number, G0-G3 values, Bin. ESD Wrist Band, powder free gloves, microscope, Deluxe console, Pump with PSS attached . Inspected pump with PSS attached. | ESD damage to pump during handling leading to pump stop start issues
Front Panel Installation Mfg Process (MFG-3018). Install Front Panel. . ASSY, Front Panel. | Cable strain
Front Panel Installation Mfg Process (MFG-3018). Install Front Panel. . ASSY, Front Panel. | Front panel covers battery light indicator
Front Panel Installation Mfg Process (MFG-3018). Install Front Panel. . ASSY, Front Panel. | Incorrect cable connection
Front Panel Installation Mfg Process (MFG-3018). Install Front Panel. . ASSY, Front Panel. | Insufficient display gasket contact
Front Panel Installation Mfg Process (MFG-3018). Install Front Panel. . ASSY, Front Panel. | Pinched concentrator tubing
Front Panel Installation Mfg Process (MFG-3018). Secure rear shell to front panel. . ASSY, Front Panel. | Excessive display gasket contact / Cosmetic defect on display
Front Panel Installation Mfg Process (MFG-3018). Secure rear shell to front panel. . ASSY, Front Panel. | Rear shell not properly fastened to front panel
Front panel shell Mfg Process (MFG-3016). Adhere Tape to Shroud.. . ASSY, Front Panel Shell. | Speaker Shroud is not flush to front panel
Front panel shell Mfg Process (MFG-3016). Assemble front panel shell. . ASSY, Front Panel Shell. | Occluded concentrator tubing
Front panel shell Mfg Process (MFG-3016). Install gasket on front panel. . ASSY, Front Panel Shell. | Insufficient display gasket contact
Front panel shell Mfg Process (MFG-3016). Install the front panel attachment block. . ASSY, Front Panel Shell. | Front panel attachment block not installed
Glass Cylinder UV Bonding. Allow assembly to cure for a minimum of 9 hours. assembly with stycast, holder. Glass Cylinder assembly. | under cured adhesive leading to leaking
Glass Cylinder UV Bonding. Apply glue at the end of the cylinder until flush with lip of Glass Cylinder. cured assembly, dispenser, dosing needle. glue filled assembly. | bond failure leading to leak
Glass Cylinder UV Bonding. Apply glue at the end of the cylinder until flush with lip of Glass Cylinder. cured assembly, dispenser, dosing needle. glue filled assembly. | piston is blocked unable to start purge system
Glass Cylinder UV Bonding. Apply glue at the end of the cylinder. positioned assembly on fixture, Loctite 3321, dosing needle, dispenser, lint free wipes. glued assembly. | bond failure leading to leak
Glass Cylinder UV Bonding. Apply glue between struts of the glass cylinder fixture. glass cylinder, glass cylinder housing, UV machine, holding fixture, Loctite 3321, dosing needle, dispenser. glued assembly. | bond failure leading to leak
Glass Cylinder UV Bonding. Apply Stycast between glass cylinder and glass cylinder housing. dispenser, cured assembly, stycast, syringe, precision glide needle, barrel adaptor, syringe piston, syringe barrel, safety goggles, holder. assembly with stycast. | bond failure leading to leak
Glass Cylinder UV Bonding. Apply Stycast between glass cylinder and glass cylinder housing. dispenser, cured assembly, stycast, syringe, precision glide needle, barrel adaptor, syringe piston, syringe barrel, safety goggles, holder. assembly with stycast. | inadequate seal leading to leak
Glass Cylinder UV Bonding. Clean inner surface of glass cylinder housing and outer surface of glass cylinder. swab, petroleum ether, glass cylinder housing, retainer plate. cleaned components. | Bonding failure leading to leaks
Glass Cylinder UV Bonding. Perform the curing process. glued assembly positioned correctly in machine, UV Machine (intensity, time). Cured assembly. | weak bond leading to leak
Glass Cylinder UV Bonding. Perform the curing process. glued filled assembly, UV curing machine (time and intensity). Cured assembly. | weak bond leading to leak
Glass Cylinder UV Bonding. Position glass cylinder and glass cylinder housing on holding fixture in UV machine. glass cylinder, glass cylinder housing, UV machine, holding fixture. components positioned in equipment. | Incorrect placement leading to blockage (unable to move pass case start)
Glass Cylinder UV Bonding. Position glass cylinder and glass cylinder housing on holding fixture in UV machine. glass cylinder, glass cylinder housing, UV machine, holding fixture. components positioned in equipment. | incorrect placement leading to leak
Glass Cylinder UV Bonding. Position glass cylinder fixture on adhesive filling holding fixture. Adhesive filling holding fixture, cured assembly. positioned assembly on fixture. | damage glass cylinder leading to leak
Glass Cylinder UV Bonding. Position glass cylinder with glass cylinder fixture in the curing device. Glued assembly, UV machine. glued assembly positioned correctly in machine. | damage glass cylinder leading to leak
Glass Cylinder UV Bonding. Position glass cylinder with glass cylinder housing in the curing device. Glued assembly, UV machine. glued assembly positioned correctly in machine. | damage glass cylinder leading to leak
Glass Cylinder UV Bonding. Prepare Stycast. Part A and Part B, scale, disposable cup, vacuum bell jar, vacuum pump. prepared stycast. | inadequate seal leading to leak
Glass Cylinder UV Bonding. Remove glass cylinder from the device. holding fixture, cured assembly. Cured assembly. | damage glass cylinder leading to leak
Glass Cylinder UV Bonding. Remove glass cylinder from the device. UV curing equipment, cured assembly. cured assembly. | damage glass cylinder leading to leak
Glass Cylinder UV Bonding. Set up machine for second UV curing . UV curing equipment, pressure regulator. setup equipment. | weak bond leading to leak
Glass Cylinder UV Bonding. Set up machine for UV curing. UV curing equipment, pressure regulator, adhesive dispenser. setup station. | bond failure leading to leak
Handle Installation Assy Mfg Process (MFG-3023) . Align hinge plates. . ASSY, Handle Installation. | Hinge plates not aligned
Handle Installation Assy Mfg Process (MFG-3023) . Assemble and attach Handle. . ASSY, Handle Installation. | Gap between two halves of the handle
Handle Installation Assy Mfg Process (MFG-3023) . Assemble and attach Handle. . ASSY, Handle Installation. | Handle does not rotate, handle does not lock, etc.
Handle Installation Assy Mfg Process (MFG-3023) . Insert O-ring, shaft, spring support, plunger, pin, knob, retaining ring. . ASSY, Handle Installation. | Handle does not rotate, handle does not lock, etc.
Impellatronic PCA Current Measurement Test (10003209). Impellatronic PMD ASYM S/W Programming. Impellatronic, Installation Dongle, PC, PMD ASYM S/W. PMD ASYM S/W installed on Impellatronic. | wrong S/W leads to higher measurement values
Inflow Cage Assembly (0550-7023). Apply adhesive to the inflow cage and place cannula over designated area. adhesive, cleaned and prepped cannula and inflow cage, timer. Cleaned and prepped cannula and inflow cage . | Bonding failure at inflow cage to cannula bond
Inflow Cage Assembly (0550-7023). Apply adhesive to the inflow cage and place cannula over designated area. adhesive, cleaned and prepped cannula and inflow cage, timer. Cleaned and prepped cannula and inflow cage . | Increase scrap
Inflow Cage Assembly (0550-7023). Clean outer surface of inflow cage and inner diameter of cannula with petroleum. Then apply HPA 41 to those same areas. Then dry.. Petroleum ether, Cannula, inflow cage, swab, HPA 41, Oven (time and temp). Cleaned and prepped cannula and inflow cage . | Bonding failure at inflow cage to cannula bond
Inflow Cage Assembly (0550-7023). Clean outer surface of inflow cage and inner diameter of cannula with petroleum. Then apply HPA 41 to those same areas.. Petroleum ether, Cannula, inflow cage, swab, HPA 41, Oven (time and temp). Cleaned and prepped cannula and inflow cage . | Bonding failure at inflow cage to cannula bond
Inflow Cage Assembly (0550-7023). Place assembly into fixture and then into oven for curing. Oven (time and temperature), glued together cannula and inflow cage, Inflow Cage Holding fixture. Inflow Cage assembly. | Bonding failure at inflow cage to cannula bond
Inflow Cage to Dual Stiff Pigtail. Evaluate pigtail for kinkability, clean inner and outer surface with petroleum ether, clean inflow cage with petroleum ether. Pigtail, inflow cage, petroleum ether, HFE wipe. Cleaned pigtail and inflow cage. | bond failure due to contiminated surface
Inflow Cage to Dual Stiff Pigtail. Flare end of pigtail. Hakko controller (teemp and time), pigtail flaring fixture, thread forming mandrel . Flared pigtail. | unable to assemble
Inflow Cage to Dual Stiff Pigtail. Flare end of pigtail. Hakko controller (temp and time), pigtail flaring fixture, thread forming mandrel . Flared pigtail. | Bond failure to end having poor contact with inflow cage
Inflow Cage to Dual Stiff Pigtail. Mix epotek 301, adhere pigtail to inflow cage top. Epo-Tek 301, sleeve tube, inflow cage, flared pigtail, inflow cage assembly fixture, dosing needle. Flared pigtail adhered to top of inflow cage. | incorrectly seated pigtail leading to bond failure
Inflow Cage to Dual Stiff Pigtail. Mix Epotek 302 together, apply to struts of Inflow cage, and cure in oven. cured pigtail to inflow cage, Epo-tek 302, holding fixture, oven (time and temperature), dosing needle, syringe, speed mixer. cured inflow cage to pigtail assembly. | rough struts and stagnat locations on inflow cage
Inflow Cage to Dual Stiff Pigtail. Place Epo-Tek 301 at transition area and place in oven for curing. Flared pigtail adhered to top of inflow cage, Epo-tek 301, dosing needle, Oven (time and temperature). cured pigtail to inflow cage with transition. | bond failure
Inflow Cage to Dual Stiff Pigtail. Place Epo-Tek 301 at transition area and place in oven for curing. Flared pigtail adhered to top of inflow cage, Epo-tek 301, dosing needle, Oven (time and temperature). cured pigtail to inflow cage with transition. | Gap between pigtail and inflow cage leading thrombus formation
Inflow Cage to Tapered Pigtail. Adhere pigtail and inflow cage using prepared EpoTek 301. Inflow cage, pigtail, EpoTek 301, yellow tipped dosing syringe. Pigtail adhered to the top of inflow cage. | Component with bent inflow cage struts
Inflow Cage to Tapered Pigtail. Adhere pigtail and inflow cage using prepared EpoTek 301. Inflow cage, pigtail, EpoTek 301, yellow tipped dosing syringe. Pigtail adhered to the top of inflow cage. | Crossthreading of pigtail to inflow cage, leading to transition issues, and bond failure
Inflow Cage to Tapered Pigtail. Adhere pigtail and inflow cage using prepared EpoTek 301. Inflow cage, pigtail, EpoTek 301, yellow tipped dosing syringe. Pigtail adhered to the top of inflow cage. | Incorrectly seated pigtail leading to bond failure
Inflow Cage to Tapered Pigtail. Applies to all subprocess steps in PF011. Pigtail. Unkinked pigtail component. | Pigtail kinking occurs during PF011
Inflow Cage to Tapered Pigtail. Apply EpoTek 302 to struts of inflow cage. EpoTek 302, assembled inflow cage-pigtail, yellow tipped dosing syringe, clean room swab. Smooth transition between struts and raindrop of inflow cage. | Assembly with bent inflow cage struts
Inflow Cage to Tapered Pigtail. Apply EpoTek 302 to struts of inflow cage. EpoTek 302, assembled inflow cage-pigtail, yellow tipped dosing syringe, clean room swab. Smooth transition between struts and raindrop of inflow cage. | Rough struts on inflow cage
Inflow Cage to Tapered Pigtail. Apply HPA bonding promoter. HPA, cleanroom swab, assembled inflow cage-pigtail. Inflow cage struts with bonding promoter . | EpoTek fails to bond to metal, exposing sharp edges on inflow cage
Inflow Cage to Tapered Pigtail. Apply HPA bonding promoter. HPA, cleanroom swab, assembled inflow cage-pigtail. Inflow cage struts with bonding promoter . | Promoter on wrong portion of inflow cage, Epotek is not effective smooth transition between struts and teardrop
Inflow Cage to Tapered Pigtail. Check patency of pigtail at adhesion with mandrel fixture. Inflow cage - pigtail assembly. Pigtail adhered to the top of inflow cage. | Untested with potential blockage leads to unable to insert Back Loading Lumen
Inflow Cage to Tapered Pigtail. Check patency of pigtail with mandrel fixture. Inflow cage - pigtail assembly. Pigtail adhered to the top of inflow cage. | Untested with potential blockage leads to unable to insert Back Loading Lumen
Inflow Cage to Tapered Pigtail. Check tear drop for adhesive on surface and remove. Inflow cage - pigtail assembly, acetone, cleanroom swap. Completed inflow cage to pigtail assembly . | Scratches on the metal of the inflow cage
Inflow Cage to Tapered Pigtail. Clean areas around end of strut and welded parts on teardrop with petroleum ether. Cleanroom swab, petroleum ether, assembled inflow cage - pigtail. Cleaned inflow cage struts. | EpoTek fails to bond to metal due to contaminated surface exposing sharp edges on inflow cage
Inflow Cage to Tapered Pigtail. Clean outer surface with HFE and inner surface of pigtail with acetone . Pigtail, acetone, HFE, wipe, cleanroom brush. Cleaned pigtail. | Bond failure due to contaminated surface
Inflow Cage to Tapered Pigtail. Dry HPA coated components in oven. Inflow cage - pigtail assembly, holding fixture. Inflow cage struts with dried bonding promoter . | Failure of Epotek to bond to metal, Epotek debris in body
Inflow Cage to Tapered Pigtail. Inspect pigtail . Cleaned Pigtail . Pigtail (inspected). | Pigtail with surface anomalies
Inflow Cage to Tapered Pigtail. Kink Test of Pigtail. Pigtail. Unkinked pigtail component. | Untested pigtail component
Inflow Cage to Tapered Pigtail. Mix EpoTek 301. EpoTek 301, clean container, stirrer, stopwatch. Mixed EpoTek 301. | EpoTek 301 past working life, bond failure
Inflow Cage to Tapered Pigtail. Mix EpoTek 301. EpoTek 301, clean container, stirrer, stopwatch. Mixed EpoTek 301. | EpoTek 302 used instead of EpoTek 301, leading to bond failure
Inflow Cage to Tapered Pigtail. Mix EpoTek 301. EpoTek 301, clean container, stirrer, stopwatch. Mixed EpoTek 301. | Inadequate EpoTek 301, bond failure
Inflow Cage to Tapered Pigtail. Mix EpoTek 302. EpoTek 302, beaker with cap, Speedmixer, Stopwatch. EpoTek 302 for smoothing tranistion between strut and tear drop. | Inadequate EpoTek 302, does not bond to metal exposing struts
Inflow Cage to Tapered Pigtail. Mix EpoTek 302. EpoTek 302, beaker with cap, Speedmixer, Stopwatch. EpoTek 302 for smoothing tranistion between strut and tear drop. | Inadequate EpoTek 302, does not bond to metal, Epotek debris in body
Inflow Cage to Tapered Pigtail. Mix EpoTek 302. EpoTek 302, beaker with cap, Speedmixer. EpoTek 302 for smoothing tranistion between strut and tear drop. | EpoTek 301 used in place of Epotek 302, debris in body
Inflow Cage to Tapered Pigtail. Place pigtail and inflow cage in oven and cure for at least 90 minutes. Inflow cage and pigtail, holding fixture, oven (time and temperature) . Cured pigtail to inflow cage with smooth transition. | Bond failure
Inflow Cage to Tapered Pigtail. Place pigtail and inflow cage in oven and cure for at least 90 minutes. Inflow cage and pigtail, holding fixture, oven (time and temperature) . Cured pigtail to inflow cage with smooth transition. | Bond failure
Inflow Cage to Tapered Pigtail. Place pigtail and inflow cage in oven and cure for at least 90 minutes. Inflow cage and pigtail, holding fixture, oven (time and temperature). Cured pigtail to inflow cage with smooth transition. | Bond failure
Inflow Cage to Tapered Pigtail. Smooth out transition area between pigtail and teardrop. Inflow cage and pigtail assembly, yellow dosing syringe, cleanroom wipe. Smooth transition between pigtail and inflow cage. | Gap between pigtail and inflow cage
Inflow Cage to Tapered Pigtail. Soak inflow cage with acetone. Inflow cage, absorbent cloth, acetone. Cleaned inflow cage. | Bond failure due to contaminated surface
Inflow Cage to Tapered Pigtail. Soak inflow cage with acetone. Inflow cage, absorbent cloth, acetone. Cleaned inflow cage. | Component with bent inflow cage struts
Inflow Cage to Tapered Pigtail. Soak inflow cage with acetone. Inflow cage, absorbent cloth, acetone. Cleaned inflow cage. | Crossthreading of pigtail to inflow cage, leading to bond failure
Inflow Cage to Tapered Pigtail. Soak inflow cage with acetone. Inflow cage, absorbent cloth, acetone. Cleaned inflow cage. | False positive for damaged threads
Insertion Pin Cleaning. Clean and dry insertion pin. Insertion pins, Micro-blasting chamber, Ultrasonic bath, Isopropanol, Petroleum Ether, Retention board, Oven (time and temperature). Cleaned insertion pins. | Contamination leading to bioburden issue
Insertion Pin Cleaning. Clean and dry insertion pin. Insertion pins, Micro-blasting chamber, Ultrasonic bath, Isopropanol, Petroleum Ether, Retention board, Oven (time and temperature). Cleaned insertion pins. | Contamination leading to improper installation of heat shrink
Insertion Pin Cleaning. Clean and dry insertion pin. Insertion pins, Micro-blasting chamber, Ultrasonic bath, Isopropanol, Petroleum Ether, Retention board, Oven (time and temperature). Cleaned insertion pins. | Undried insertion pins
Insertion Pin Shrink Tube. Clean shrink tube covered insertion pin. Insertion pin (w/shrink tube), wipes, Isopropanol. Cleaned insertion pins (w/ shrink tube). | Contamination leading to scrap
Insertion Pin Shrink Tube. Coat insertion pins with release agent and dry, then transport to cleanroom. Insertion pins, Fume hood, Acmos (release agent), Oven (time and temperature), plastic bags, paper towel. Coated and dried insertion pins. | Contamination leading to bioburden issue
Insertion Pin Shrink Tube. Coat insertion pins with release agent and dry, then transport to cleanroom. Insertion pins, Fume hood, Acmos (release agent), Oven (time and temperature), plastic bags, paper towel. Coated and dried insertion pins. | Too little release agent prevents removal of heat shrink from pin
Insertion Pin Shrink Tube. Cover insertion pin with shrink tube, cut off excess shrink tube. Insertion pins, shrink tube, Shrink tube process tray, Oven (time and temperature), Scissors. Insertion pin covered with shrink tube. | Scrap due to short shrink tube
Insertion Pin Shrink Tube. Cover insertion pin with shrink tube, cut off excess shrink tube. Insertion pins, shrink tube, Shrink tube process tray, Oven (time and temperature), Scissors. Insertion pin covered with shrink tube. | Scrap due to visual defects
Insertion Pin Shrink Tube. Inspect insertion pins, then place onto tranport fixture and double bag. Insertion pin (w/shrink tube), transport fixture, plastic bags. Insertion pins mounted and double bagged on transport fixure. | Contamination leading to bioburden issue
Insertion Pin Shrink Tube. Place insertion pins on retention board and place in oven, then remove excess heat shrink. Insertion pins, shrink tube, retention board, Oven (time and temperature), Scalpel. Insertion pin covered with shrink tube. | Shrink tube not thermoformed onto insertion pin
Installation for Backloading Lumen (10000176). load lumen until visible outside the pigtail. lumen, pump in fixture. pump with lumen loaded. | lumen not visible outside pigtail
Installation for Backloading Lumen (10002397). load lumen until visible outside the pigtail. lumen, pump in fixture. pump with lumen loaded. | lumen not visible outside pigtail
Installation for Backloading Lumen. place pump in backloading lumen loading fixture. pump, fixture. pump in fixture. | pump not placed in fixture correctly
Label Installation Mfg Process (MFG-3025). Affix labels. . CONSOLE. | Rear shell and left shell not cleaned
Label Installation Mfg Process (MFG-3025). Clean surface of rear shell and left shell.. . CONSOLE. | Rear shell not cleaned
Label Printing Procedure. Print labels. Label Printing Procedure. Label Printing Procedure. | Incorrect information on label
Laser welding PH/endcap (0046-7084). Clean components and inspect for damage. Petroleum ether, swabs, motor housing, endcap, peek sleeve. All components prepped for welding. | debry in weld area leading to incomplete weld/damaged weld
Laser welding PH/endcap (0046-7084). Clean components and inspect for damage. Petroleum ether, swabs, motor housing, endcap, peek sleeve. All components prepped for welding. | peak sleeve damage leads to damage on impeller
Laser welding PH/endcap (0046-7084). Load components into machine (manual position). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Low weld strength
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Damage to pump housing inner surface
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Damage to pump housing inner/outer surfaces
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Deformation of pump housing
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Low weld strength
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Pump housing out of concentricity
Laser welding PH/endcap (0046-7084). Load components into machine (manual positioning). Combi fixture, pump housing, pump assembly with endcap, Laser welder. Components aligned for welding. | Tip/Blades of impeller are damaged
Laser welding PH/endcap (0046-7084). Run machine and weld pump housing to motor housing. Combi fixture, assembly with welded end cap, Laser welder, pump houisng. Welded pump housing to motor housing as well as motor housing to endcap. | hemolysis
Laser welding PH/endcap (0046-7084). Run machine and weld pump housing to motor housing. Combi fixture, assembly with welded end cap, Laser welder, pump houisng. Welded pump housing to motor housing as well as motor housing to endcap. | Incomplete weld (ex. sink holes/distortions)
Laser welding PH/endcap (0046-7084). Run machine and weld pump housing to motor housing. Combi fixture, assembly with welded end cap, Laser welder, pump houisng. Welded pump housing to motor housing as well as motor housing to endcap. | Low weld strength
Laser welding PH/endcap (0046-7084). Run machine and weld pump housing to motor housing. Combi fixture, assembly with welded end cap, Laser welder, pump houisng. Welded pump housing to motor housing as well as motor housing to endcap. | window out of alignment with pressure port (scrap)
Laser welding PH/endcap (0046-7084). Run machine and weld endcap to motor housing. Combi fixture, pump housing, pump assembly with endcap, Laser welder. Welded endcap to motor housing. | Damage components (scrap)
Laser welding PH/endcap (0046-7084). Run machine and weld endcap to motor housing. Combi fixture, pump housing, pump assembly with endcap, Laser welder. Welded endcap to motor housing. | Incomplete weld (ex. sink holes/distortions)
Laser welding PH/endcap (0046-7084). Run machine and weld endcap to motor housing. Combi fixture, pump housing, pump assembly with endcap, Laser welder. Welded endcap to motor housing. | Low weld strength
Laser welding PH/endcap (0046-7084). Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program, combi fixture, combi fixture collet. Laser Welder operational and tested for power. | Damage to weld (thrombus formation)
Laser welding PH/endcap (0046-7084). Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program, combi fixture, combi fixture collet. Laser Welder operational and tested for power. | Energy meter not functional/providing incorrect output leading to low weld strength
Laser welding PH/endcap (0046-7084). Setup Equipment. Laser welder, laser welding program, energy meter, energy sensor, energy program, combi fixture, combi fixture collet. Laser Welder operational and tested for power. | Low weld strength
Laser welding PH/endcap (0046-7084). Trim PTFE mandrel, secure with O'ring and load peek sleeve. O'ring, assembly unit with mandrel, pump housing and peek sleeve. prepped unit ready for welding. | mandrel not aligned leading to damaged yoke/pump (scrap)
Laser welding PH/endcap (0046-7084). Unload pump from machine and perform eccentricity check. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap,. Welded assembly. | Deformation of pump housing
Laser welding PH/endcap (0046-7084). Unload pump from machine and perform eccentricity check. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap,. Welded assembly. | pump housing damaged
Laser welding PH/endcap (0046-7084). Unload pump from machine and perform eccentricity check. Feeler Gauge, Welded outflow cage to motor housing as well as motor housing to endcap,. Welded assembly. | Tip of impeller is damaged
Leak Test Procedure, Oxygenator. Perform leak test. Leak Test Procedure. Oxygenator Asm completed leak test. | Device contaminated
Leak Test Procedure, Oxygenator. Perform leak test. Leak Test Procedure. Oxygenator Asm completed leak test. | Excessive heat exposure
Leak Test Procedure, Oxygenator. Perform leak test. Leak Test Procedure. Oxygenator Asm completed leak test. | Test does not detect leaks
Left Shell Mfg Process (MFG-3019). Assemble Left Shell. . ASSY, Shell Left. | Left shell not properly assembled
Lower PCB Stack Mfg Process (MFG-3012) Upper PCB Stack Mfg Process (MFG-3013) PCB Installation Mfg Process (MFG-3015) Cable Harness #:. Install PCB. . ASSY, PCB Installation. | Incorrect cable connection Cable strain Cable damage
Lower PCB Stack Mfg Process (MFG-3012) Upper PCB Stack Mfg Process (MFG-3013) PCB Installation Mfg Process (MFG-3015) Cable Harness #:. Install PCB. . ASSY, PCB Installation. | Incorrect cable connection Cable strain Cable damage
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | contamination of rotor
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | damage/ deformation of rotor wings
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Damaging of tip during tumbling
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Damaging of wingduring tumbling
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Loss of tube leading to damage of rotor
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Protective tube not added to shaft; damage to DLC coating
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Protective tube not added to shaft; damage to DLC coating during tumbling
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Protective tube not removed from shaft
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Rotor damaged
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Rotor geometry incorrect
Manual deburring and grinding of rotor. Manual grinding of rotor. Internally- unprocessed rotor. Deburred rotor. | Rotor not clean
Molding Motor Housing (0550-7029). Apply primer on distal face of molded motor housing. Molded motor hosuing, and primer. Molder motor housing with primer. | Insufficient coating
Molding Motor Housing (0550-7029). Apply primer on distal face of molded motor housing. Molded motor housing, and primer. Molder motor housing with primer. | Insufficient coating
Molding Motor Housing (0550-7029). Check for stiff rotor using test rotor, remove test rotor. Molded motor housing, test rotor. AU Motor Housing. | stiff rotor
Molding Motor Housing (0550-7029). Clean assembly with cleanroom swab and pertroleum ether, wait 2 minutes. Molded Motor housing, cleanroom swab, petroleum ether. Cleaned Molded Motor Housing. | loose particles in motor housing/purge issue
Molding Motor Housing (0550-7029). Clean assembly with cleanroom swab and petroleum ether, wait 2 minutes. Molded Motor housing, cleanroom swab, petroleum ether. Cleaned Molded Motor Housing. | loose particles in motor housing/purge issue
Molding Motor Housing (0550-7029). close chamber and evacuate to -92 kPa. Reservoir of mold chamber filled with epotek. Evacuated chamber. | Partial molding
Molding Motor Housing (0550-7029). close chamber and evacuate to -92 kPa. Reservoir of mold chamber filled with epotek. Evacuated chamber. | unable to continue
Molding Motor Housing (0550-7029). close chamber and push close chamber soft button. Mold installed in chamber with tubing. Mold installed in chamber with tubing locked. | unable to continue
Molding Motor Housing (0550-7029). close pinch valves as epotek is seen above mold. molding in process. filled mold. | thermal properties out of spec - bubbles in epotek
Molding Motor Housing (0550-7029). cut tubing below syringe/reservoir. released vacuum from chamber, scissors. opened chamber with cut mold tubing. | unable to continue
Molding Motor Housing (0550-7029). dry pumps using compressed air and oven. Molded Motor housing, compressed air, oven (time and temp). Dried molded motor housing. | bioburden/contamination
Molding Motor Housing (0550-7029). fill 2 ml of epotek into syringe (down side of syringe) allow to pass pinch valve, Close pinch valves using buttons. Epotek, Mold installed in chamber. Reservoir of mold chamber filled with epotek. | thermal properties out of spec - bubbles in epotek
Molding Motor Housing (0550-7029). fill 2 ml of epotek into syringe (down side of syringe) allow to pass pinch valve, pinch silicone tubing by closing pneumatic valve. Epotek, Mold installed in chamber. Reservoir of mold chamber filled with epotek. | insufficient epotek - air bubbles in mold
Molding Motor Housing (0550-7029). fill 2 ml of epotek into syringe (down side of syringe) allow to pass pinch valve, pinch silicone tubing by closing pneumatic valve. Epotek, Mold installed in chamber. Reservoir of mold chamber filled with epotek. | thermal properties out of spec - bubbles in epotek
Molding Motor Housing (0550-7029). Fill any voids and coat distal face of molded motor housing with adhesive using a needle. Molded motor housing, adhesive, needle. Molded motor housing with voids filled, distal face covered with adhesive. | leakage current
Molding Motor Housing (0550-7029). Glue tubing into luer, harden adhesive.. mold tubing, transparent luer, UV adhesive, UV lamp. Glued mold tubing. | blocked luer - unable to mold
Molding Motor Housing (0550-7029). Glue tubing into luer, harden adhesive.. mold tubing, transparent luer, UV adhesive, UV lamp. Glued mold tubing. | epotek leak - unable to mold
Molding Motor Housing (0550-7029). guide mold tubing out of mold chamber and connect to reservoir using silicone tube. Mold installed in chamber, silicone tubing, reservoir syringe. Mold installed in chamber with tubing. | epotek leak - unable to mold
Molding Motor Housing (0550-7029). harden adhesive in oven. Molded motor housing with voids filled and distal face covered with adhesive, oven (time, temp). Molded motor housing with voids filled, distal face covered with adhesive. | thermal properties out of spec - bubbles in epotek
Molding Motor Housing (0550-7029). lock locknut on mold chamber. Mold installed in chamber with tubing. Mold installed in chamber with tubing locked. | vacuum not stable - unable to mold
Molding Motor Housing (0550-7029). once all molds are filled, press chamber open button. filled mold, evacuated vacuum chamber. released vacuum from chamber. | unable to continue
Molding Motor Housing (0550-7029). open all pinch valves. Evacuated chamber and reservoir filled with epotek. molding in process. | unable to continue
Molding Motor Housing (0550-7029). Open chamber and locknut, cut tubing below syringe/reservoir. released vacuum from chamber, scissors. opened chamber with cut mold tubing. | unable to continue
Molding Motor Housing (0550-7029). open clamp of mold, remove the clamp. cured mold. cured mold with clamp removed. | unable to continue
Molding Motor Housing (0550-7029). open main switch and pinch valves. Evacuated chamber and reservoir filled with epotek. molding in process. | unable to continue
Molding Motor Housing (0550-7029). place extra epotek into jar to ensure hardness. epotek in mixing jar. hardened epotek/passed crack test. | unable to continue without crack test
Molding Motor Housing (0550-7029). Place manifold into chamber, place mold into manifold, connect mold tubing to pink luer. AU Stator with glued luer, mold manifold, mold chamber, mold tubing. Mold installed in chamber. | epotek leak - unable to mold
Molding Motor Housing (0550-7029). Plug pink luer connector into yoke, add UV glue, harden. AU Stator, pink luer, UV adhesive, UV lamp. AU Stator with glued luer. | blocked luer - unable to mold
Molding Motor Housing (0550-7029). Plug pink luer connector into yoke, add UV glue, harden. AU Stator, pink luer, UV adhesive, UV lamp. AU Stator with glued luer. | epotek leak - unable to mold
Molding Motor Housing (0550-7029). prepare epotek. Epotek, speed mixer, mixing jar, filler. prepared Epotek. | incorrect mixture - thermal and chemical properties out of spec
Molding Motor Housing (0550-7029). prepare epotek. Epotek, speed mixer, mixing jar, filler. prepared Epotek. | incorrect mixture - thermal properties out of spec
Molding Motor Housing (0550-7029). press mold with a pin on the bottom onto a hard surface to remove the part from the mold. cured mold with clamp removed. cured mold with mold plate removed. | unable to continue
Molding Motor Housing (0550-7029). release vacuum from chamber. filled mold, evacuated vacuum chamber. released vacuum from chamber. | unable to continue
Molding Motor Housing (0550-7029). remove inlay from motor housing. cured mold with mold plate removed. Motor housing without inlay. | unable to continue
Molding Motor Housing (0550-7029). remove manifold from chamber, leave mold in manifold, place in oven for 10 hrs at 35 C+/- 2. opened chamber with cut mold tubing, oven (time, temp), timer. cured mold. | mechanical properties out of specification - uncured or partially cured mold
Molding Motor Housing (0550-7029). remove manifold from chamber, leave mold in manifold, place in oven for at lease 10 hrs at 35 C+/- 2. opened chamber with cut mold tubing, oven (time, temp), timer. cured mold. | mechanical properties out of specification, incomplete mold
Molding Motor Housing (0550-7029). remove mold flash with scalpel. Molded Motor housing, scalpel. Molded Motor housing with flash removed. | damage to assembly
Molding Motor Housing (0550-7029). remove mold ring and o-ring . Motor housing without inlay. Molded Motor housing. | unable to continue
Molding Motor Housing (0550-7029). Remove shim (if necessary). AU Stator with glued luer. AU Stator with glued luer with shim removed. | unable to see proper filling - insufficient molding
Molding Motor Housing (0550-7029). Start vacuum pump to 93hPa +/- 5. Reservoir of mold chamber filled with epotek. Evacuated chamber. | unable to continue
Molding Motor Housing (0550-7029). test leakage current using wet test. Molded motor housing, leakage current tester. leakage current tested molded motor housing. | high motor leakage current
Molding Motor Housing (0550-7029). test leakage current using dry test. Molded motor housing, leakage current tester. leakage current tested molded motor housing. | high motor leakage current
Molding Motor Housing (0550-7029). Test resistance and inductivitiy. Molded motor housing, multimeter (which can test inductivity). Electrically tested molded motor housing. | resistance and/or inductivity out of spec
Molding Motor Housing (0550-7029). Test resistance and inductivity. Molded motor housing, multimeter (which can test inductivity). Electrically tested molded motor housing. | resistance and/or inductivity out of spec
Molding Motor Housing (0550-7029). Unplug luer connectors. cured mold. cured mold with luer connectors removed. | unable to continue
Molding Motor Housing (0550-7029). wait for 2 minutes (use timer). Evacuated chamber, timer. Evacuated chamber. | Partial molding
Molding Motor Housing (0550-7029). wait for 2 minutes (use timer). Evacuated chamber, timer. Evacuated chamber. | unable to continue
Molding Motor Housing (0550-7029). wait until you see epotek in upper reservoir of mold, then close pinch valve. molding in process. filled mold. | empty syringe - air bubbles - thermal properties out of spec
Molding Motor Housing (0550-7029). wait until you see epotek in upper reservoir of mold, then close pinch valve. molding in process. filled mold. | Partial molding
Nitinol Wire Assembly (0550-7021). Bake coil in oven. Coiled nitinol with coating, oven (time and temp). baked Coiled nitinol with coating. | cannula not dried leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). clean nitinol coil in iso-propanol in a ultrasonic bath; swirl in acetone; swirl in petroleum ether. Purple nitinol coil on mandrel, isopropyl alcohol, ultrasonic bath, acetone, and petroleum ether. Cleaned Coiled nitinol (only). | components not cleaned leading to increased reject rate
Nitinol Wire Assembly (0550-7021). cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | Cannula is structurally weak
Nitinol Wire Assembly (0550-7021). cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | parts not stored in oven leading to weak structure
Nitinol Wire Assembly (0550-7021). immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B. Cilbond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | cannula not dried leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B. Cilbond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect material utilized leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B. Cilbond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect mixture of cilcure leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B. Cilbond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | not coated properly leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil. | incorrect water is selected leading to contamination (bioburden)
Nitinol Wire Assembly (0550-7021). quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil. | not enough water is utilized leading to structurally weak unit
Nitinol Wire Assembly (0550-7021). Setup coiling machine, select appropriate program for coil size. Coiling winder, coil winding program. Setup equipment. | increased reject rate
Nitinol Wire Assembly (0550-7021). Setup coiling machine, select appropriate program for coil size. Coiling winder, coil winding program. Setup equpiment. | increased reject rate
Nitinol Wire Assembly (0550-7021). start automatic mode until program end, fix clamp ring on overlapping winding. Threaded Nitinol on mandrel in coil winding machine. Coiled nitinol on mandrel . | coil not winded according to specification leading to weak structure
Nitinol Wire Assembly (0550-7021). thread Nitinol wire in coil bolt and fix it with clamp ring, fix coat bolt with tailstock . Nitinol wire, clamp ring, tailstock, mandrel, coil winder. Threaded Nitinol on mandrel in coil winding machine. | increased reject rate
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle diameter too large
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle diameter too large (won't fit into assembly)
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle diameter too small
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle too narrow
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle too wide
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Bundle too wide (will not fit into assembly)
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Fiber Deformation
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Fiber deformation and/or damage
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Fiber Not Tight at Splice
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Fiber not tight at termination
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Glue extends beyond potting line
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Glue not fully cured
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | ID termination end not attached to winding
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Major broken fibers
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Minor broken fibers
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Not enough glue used on tape
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | OD termination end not attached to winding
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Over cure of glue
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Splice Termination End not attached to winding.
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Tape extends beyond potting line
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Too little tension of membrane
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Too much tension of membrane
OXY XL Winder Operation. Wind fibers. Winder Operation. Oxy Winding. | Winding bundle not installed in proper quadrants
OXY-1 Battery Conditioning . Condition battery packs. . . | Batteries not properly discharged or charged
OXY-1 Console Software Update Procedure. Update OXY-1 Console software. . . | Incorrect or improper software installation
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Blood tubing not fully assembled to collar of port
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Cable tie in wrong location
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Glue not evenly dispensed/complete around OXY exit barb
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Glue not evenly dispensed/complete around OXY inlet barb
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Kinks in blood tubing
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Kinks in gas line
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Loose exit tubing after assembly to OXY exit
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Loose inlet tubing after assembly to OXY inlet
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Missing/loose cable tie on blood tubing/port connection
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Asm. | Red Stripe blood tubing is connected to the barb on the Pump Inlet Port and/or blue stripe blood tubing is connected to the barb on the OXY Exit port.
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Blood exit tubing off of exit port bundled to coiled tubing
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Blood tubing/connector damaged during installation
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Gas line under retaining straps
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Lifted foam adhesive strip/strap mount
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing condensation tray pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing corrugated support ring
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing hemostasis valve pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing IFU
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing in-line access port pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing lid protection insert
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing luer cap pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing PLU Labels pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing recirculation bag
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing recirculation bag hook
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing tether cover pouch assembly
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | PLU disposable not secured to tub
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Tubing under retaining straps
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure. PLU Blood Tubing Asm. | Pump retaining washer dislodged
OXY-1 Disposable Kit Packaging Procedure. Package disposable kit. Oxy-1 Disposable Kit Packaging Procedure; Belco Tray Sealing Procedure. Oxy-1 PLU Disposable Kit. | Incorrect Belco machine sealing settings
OXY-1 Disposable Kit Packaging Procedure. Place label on packaging. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing label on outer pack
OXY-1 Disposable Kit Packaging Procedure. Place label on packaging. Oxy-1 Disposable Kit Packaging Procedure. Oxy-1 PLU Disposable Kit. | Missing label on sealed tub
OXY-1 Disposable Kit Packaging Procedure. Print product label. Oxy-1 Disposable Kit Packaging Procedure. Label Printing Procedure. PLU Final Assembly. Oxy-1 PLU Disposable Kit. | Incorrect information on label
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Gas line caps not attached
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Incomplete bag seal
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Incorrect information on label
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Tubing cut too long
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Tubing cut too short
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Tubing loops not secure
OXY-1 Oxygen Tubing Assembly Process. Assemble Oxygen Tubing. . OXY-1 Tubing with Filter. | Tubing not connected to filter securely
OXY-1 Oxygen Tubing Final Assembly. Package Oxygen Tubing. . Packaged OXY-1 Tubing with Filter. | Filter Assembly incorrectly labeled
OXY-1 Post Sterilization Release. PLU Seal Integrity Test. OXY-1 PLU Release. Released Pallet of Product. | Failure of PLU Seal Integrity Test
OXY-1 System Packaging. Attach Arrow, Battery, and Shipping labels. Materials in MFG-4012. . | Labels not attached
OXY-1 System Packaging. Bag console. Materials in MFG-4012. . | Console not bagged properly
OXY-1 System Packaging. Bag E-drive. Materials in MFG-4012. . | E-drive not bagged properly
OXY-1 System Packaging. Bag Pump driver. Materials in MFG-4012. . | Pump driver not bagged properly
OXY-1 System Packaging. Close Case/box. Materials in MFG-4012. . | Lid does not shut without compressing foam more than 1/4"
OXY-1 System Packaging. Obtain packaging materials and OXY-1 system components. Materials in MFG-4012. . | Missing Case, Foam, Console, PLU driver, E-drive, memory token, UFX RUGGEDrive PC interface adapter, IV pole mount
OXY-1 System Packaging. Place OXY-1 System Operating Manual and unpackaging procedure in foam pocket. Materials in MFG-4012. . | Missing Operator Manual or unpackaging procedure
OXY-1 System Packaging. Place protective cap over pump driver cable connector. Materials in MFG-4012. . | PLU pump driver cable protective cap not placed
OXY-1 System Packaging. Print System packaging label. Materials in MFG-4012. . | Labels not printed or incorrect serial number on label
OXY-1 System Packaging. Replace Battery switch cover. Materials in MFG-4012. . | Battery switch cover not replaced
OXY-1 System Packaging. Verify console does not turn on when power switch on. Materials in MFG-4012. . | Power switch turns on console
OXY-1 Disposable Kit Palletization Procedure. Palletize disposable kit. See OXY-1 Disposable Kit Palletization Procedure. Oxy-1 Pallet Asm. | BET test units not properly labeled
Oxygen Concentrator Disassembly Procedure. Adjusting molecular sieve tanks position. O2 Concepts Oxlife INDEPENDENCE oxygen concentrator.. . | Oxygen concentration below the acceptance limit
Oxygen Concentrator Disassembly Procedure. Drilling hole in pump driver base. . Pump driver base with hole. | Hole is drilled in incorrect location
Oxygen Concentrator Disassembly Procedure. Drilling hole in pump driver base. . Pump driver base with hole. | LED can not fit into hole
Oxygen Concentrator Disassembly Procedure. Drilling hole in pump driver base. . Pump driver base with hole. | Plastic contamination in Pump driver
Oxygen Concentrator Disassembly Procedure. Drilling hole in pump driver base. . Pump driver base with hole. | Water ingress
Oxygen Manifold Mfg Process (MFG-3033). Assemble Oxygen Manifold. . ASSY, Oxygen Manifold. | Oxygen Manifold not properly assembled
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Blood exit port plug not assembled
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Bullet/mandrel snaps not fully engaged
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Bullet/mandrel snaps not fully engaged (potting leak)
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Major broken fibers
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Minor broken fibers
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | O-ring incorrectly assembled/not assembled (potting leak)
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Potting cap orifice too big
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Potting cap orifice too small/closed off (potting short-shot)
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Winding crushed during installation
Oxygenator Assembly Procedure. Assemble Oxygenator. Oxygenator Asm Procedure. Oxygenator Asm. | Winding termination disconnects during assembly
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Centrifuge stopped during oven cycle
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Centrifuge stopped during potting cycle
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Potting line too low
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Potting lines too high
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Potting overflow feature on blood side - ID
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Potting overflow feature on blood side - OD
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Potting overflows into blood path
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Uneven potting line (higher on one side than the other)
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Urethane not fully cured post leak test
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Urethane not fully cured prior to leak test
Oxygenator Urethane Potting Procedure. Apply potting. Oxygenator Asm Procedure. Oxygenator Asm. | Winding termination disconnects during potting
Packaging PSS Retainer. Labeling. sealed pouch product label. LN on label matches LN of Retainers. | LN on label does not match LN of Retainers
Packaging PSS Retainer. Packaging Inspection. sealed pouch magnifier light. no particles/fibers in pouch. | particles/fibers in pouch
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | creases/leak paths in seal
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | debris/fibers in pouch
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | debris/fibers in seal
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | inadequate seal strength
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | incomplete seal
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging PSS Retainer. Seal Pouch . pouch with retainer temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging Pump un-sterile. Apply ESD Label. sealed tray ESD label. ESD label applied to tray. | ESD label missing
Packaging Pump un-sterile. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins leading to off axis sealing
Packaging Pump un-sterile. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins, unable to operate equipment
Packaging Pump un-sterile. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | dirty sealing die, poor sealing
Packaging Pump un-sterile. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | worn Teflon cloths, inadequate seal strength
Packaging Pump un-sterile. Labeling. sealed pouch product label. SN on label matches SN on pump. | SN on label does not match SN on pump
Packaging Pump un-sterile. Load Tray. pouch sealed tray . tray placed in pouch correctly . | tray placed in pouch backwards, damage seal on pouch
Packaging Pump un-sterile. Load Tray. pouch sealed tray. tray placed in pouch correctly . | tray placed in pouch upside down, can't read pump SN
Packaging Pump un-sterile. Machine Setup. user ID password recipe ejection cylinder working hand temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect die temperature, inadequate seal strength
Packaging Pump un-sterile. Machine Setup. user ID password recipe hand correct temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect recipe loaded, causing sealing issues
Packaging Pump un-sterile. Packaging Inspection. sealed pouch, product (Pump, purge cassette, blister, inlays) . Tray free of particulate/ debris. | particles/fibers on product
Packaging Pump un-sterile. Pump cleaning. alcohol pump. cleaned pumps. | particles/fibers on product
Packaging Pump un-sterile. Pump Packaging. Pump, tray, pouch, ionized air gun. Tray free of particulate/ debris. | particles/fibers migrate to other product
Packaging Pump un-sterile. Pump Packaging. Pump, tray, pouch, ionized air gun. Tray free of particulate/ debris. | particles/fibers on product
Packaging Pump un-sterile. Pump placement. Cable wrapping fixture pump, spacer. plug in tray. | improper position of patient cable plug
Packaging Pump un-sterile. Pump placement. pump paper tape. paper tape on patient cable. | paper tape not secured
Packaging Pump un-sterile. Pump placement. pump tray cannula insert cover label. cannula insert cover clipped-in w/label. | cannula insert cover label in wrong location, difficulty removing pump from tray
Packaging Pump un-sterile. Pump placement. pump tray cannula insert cover label. cannula insert cover clipped-in w/label. | missing cannula insert cover label, difficulty removing pump from tray
Packaging Pump un-sterile. Pump placement. pump tray cannula insert cover. cannula insert cover clipped-in w/label. | cannula cover insert not clipped in to tray, sealing issue
Packaging Pump un-sterile. Pump placement. pump tray cannula insert cover. cannula insert cover clipped-in w/label. | missing cannula insert cover, cannula damage
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | Damage sensor face
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | foam not seated in plug cover, leading to inadequate seal
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | missing pump plug cover label, difficulty removing pump from tray
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | missing pump plug cover, leading to pump damage
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | pump plug cover label in wrong location, difficulty removing pump from tray
Packaging Pump un-sterile. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | pump plug cover not clipped-in, leading to inadequate seal
Packaging Pump un-sterile. Pump placement. pump tray. cannula in position. | cannula not in position, kink in pigtail or cannula bend radius
Packaging Pump un-sterile. Pump placement. pump tray. patient cable plug in position. | plug not in position, damage to patient plug
Packaging Pump un-sterile. Pump placement. pump tray. pump in tray. | pump not in position, causing pump damage
Packaging Pump un-sterile. Pump placement. pump tray. pump in tray. | red handle not clipped in to tray, damage to pump
Packaging Pump un-sterile. Pump placement. pump tray. pump in tray. | Repositioning unit not at the correct location on the catheter and/or the designated location in the packaging, causing the repositioning unit to interact with the introducer during pump insertion
Packaging Pump un-sterile. Pump placement. tray butterfly clip label. butterfly clip label applied. | butterfly clip label not properly aligned, difficulty opening package
Packaging Pump un-sterile. Pump placement. tray butterfly clip label. butterfly clip label applied. | missing butterfly clip label, difficulty opening package
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting. sealed pouch. | creases/leak paths in seal
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting. sealed pouch. | incomplete seal
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting time film thickness setting ionizer on hi. sealed pouch. | particles/fibers on product
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | inadequate seal strength
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging Pump un-sterile. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | creases/leak paths in seal
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | inadequate seal strength
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | incomplete seal
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | particles/fibers on product
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | Tyvek not properly aligned with tray
Packaging Pump un-sterile. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no fibers/debris in tray. | less than 6mm seal width, loss of seal integrity
Packaging Pump un-sterile. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | contaminated trays/Tyvek leading to particles/fibres left on product
Packaging Pump un-sterile. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | ionizer on wrong setting leading to particles/fibres left on product
Packaging. Apply EO indicator tape. sealed tray EO indicator tape. EO indicator tape applied to tray. | EO indicator tape missing, unable to indicate sterility
Packaging. Apply ESD Label. sealed tray ESD label. ESD label applied to tray. | ESD label missing
Packaging. Apply ESD Label. sealed tray ESD label. ESD label applied to tray. | ESD label missing
Packaging. Apply the butterfly clip and label. pump tray pump layout in tray inserts. butterfly clip label applied. | butterfly clip label not properly aligned, difficulty opening package
Packaging. Apply the butterfly clip and label. pump tray pump layout in tray inserts. butterfly clip label applied. | missing butterfly clip label, difficulty opening package
Packaging. Apply the butterfly clip and label. pump tray pump layout in tray inserts. butterfly clip label applied. | butterfly clip label not properly aligned, difficulty opening package
Packaging. Apply the butterfly clip and label. pump tray pump layout in tray inserts. butterfly clip label applied. | missing butterfly clip label, difficulty opening package
Packaging. Clean catheter, motor housing and workplace with isopropanol; switch on ionizer blower; lay hard shell and seal foil under blower. Finished pump assembly, Isopropanol, ionizer, hard shell, seal foil. cleaned pump and area, hard shell and seal foil under ionizer. | particles/fibers on product
Packaging. Clean Packaging. Hard shell, seal foil, Air gun. Tray free of particulate/debris. | particles/fibers migrate to other product
Packaging. Clean Packaging. Hard shell, seal foil, Air gun. Tray free of particulate/debris. | particles/fibers on product
Packaging. Clean pump. alcohol pump wipe. cleaned pumps. | debris/fibers on pump
Packaging. Clip in cannula insert cover and label. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | missing cannula insert cover, cannula damage
Packaging. Clip in cannula insert cover and label. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | cannula cover insert not clipped in to tray, sealing issue
Packaging. Clip in cannula insert cover and label. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | cannula insert cover label in wrong location, difficulty removing pump from tray
Packaging. Clip in cannula insert cover and label. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | missing cannula insert cover label, difficulty removing pump from tray
Packaging. Clip in cannula insert cover and label. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | missing cannula insert cover, cannula damage
Packaging. ensure backloading lumen label is located in the correct position. pump in butterfly clip, label. pump in butterfly clip. | backloading lumen damaged
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins leading to off axis sealing
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins, unable to operate equipment
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | dirty sealing die, poor sealing
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | worn Teflon cloths, inadequate seal strength
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins leading to off axis sealing
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | bent centering pins, unable to operate equipment
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | dirty sealing die, poor sealing
Packaging. Equipment Checks. Teflon cloths sealing die centering pins. machines ready for use. | worn Teflon cloths, inadequate seal strength
Packaging. Heat foil welding device for the second packaging and switch on ionizer. Foil welding device (temperature), ionizer. Heated foil welding device and ionizer switched on. | Poor seal of package leading sterility compromise
Packaging. Insert Cannula cover. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | cannula cover insert not clipped in to tray, sealing issue
Packaging. Insert Cannula cover. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | cannula insert cover label in wrong location, difficulty removing pump from tray
Packaging. Insert Cannula cover. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | missing cannula insert cover label, difficulty removing pump from tray
Packaging. Insert Cannula cover. pump tray pump layout in tray inserts. cannula insert cover clipped-in w/label. | missing cannula insert cover, cannula damage
Packaging. Insert pump plug cover insert with label and foam. pump tray pump layout in tray inserts. pump plug cover clipped-in w/ label and foam. | foam not seated in plug cover, leading to inadequate seal
Packaging. Insert pump plug cover insert with label and foam. pump tray pump layout in tray inserts. pump plug cover clipped-in w/ label and foam. | missing pump plug cover label, difficulty removing pump from tray
Packaging. Insert pump plug cover insert with label and foam. pump tray pump layout in tray inserts. pump plug cover clipped-in w/ label and foam. | missing pump plug cover, causing sterile barrier breach during shipping
Packaging. Insert pump plug cover insert with label and foam. pump tray pump layout in tray inserts. pump plug cover clipped-in w/ label and foam. | pump plug cover label in wrong location, difficulty removing pump from tray
Packaging. Insert pump plug cover insert with label and foam. pump tray pump layout in tray inserts. pump plug cover clipped-in w/ label and foam. | pump plug cover not clipped-in, leading to inadequate seal
Packaging. Inspect Packaging. Hard shell, seal foil. Tray free of particulate/debris. | particles/fibers on product
Packaging. Label pouch. Then load Tray into pouch. pouch sealed tray . tray placed in pouch correctly . | tray placed in pouch backwards, damage seal on pouch
Packaging. Label pouch. Then load Tray into pouch. pouch sealed tray. tray placed in pouch correctly . | tray placed in pouch upside down, can't read pump SN
Packaging. Labeling. sealed pouch product label. SN on label matches SN on pump. | SN on label does not match SN on pump
Packaging. Labeling. sealed pouch product label. SN on label matches SN on pump. | SN on label does not match SN on pump
Packaging. Laminate sterile bag with patient label. Sterile bag, patient label. Laminated sterile bag. | Label missing or incorrectly applied
Packaging. Lay packaging bag in the foil welding device; insert tray into the bag and weld it. laminated sterile bag, foil welding device (temperature), tray with pump. Final packaged device (unsterile). | Poor seal of package leading sterility compromise
Packaging. Lay packed pump under the ionizer; then laminate ESD label. sealed tray with pump, ionizer, ESD label. Sealed tray with pump with ESD label. | ESD label missing or incorrect label applied
Packaging. Load Tray. pouch sealed tray . tray placed in pouch correctly . | tray placed in pouch backwards, damage seal on pouch
Packaging. Load Tray. pouch sealed tray. tray placed in pouch correctly . | tray placed in pouch upside down, can't read pump SN
Packaging. Load Tray. pouch sealed tray . tray placed in pouch correctly . | tray placed in pouch backwards, damage seal on pouch
Packaging. Load Tray. pouch sealed tray. tray placed in pouch correctly . | tray placed in pouch upside down, can't read pump SN
Packaging. Machine Setup. user ID password program recipe gloves Tray sealer (time, pressure, and temperature). correct recipe correct temperature, time and pressure tray sealer. | incorrect die temperature, inadequate seal strength
Packaging. Machine Setup. user ID password program recipe gloves Tray sealer (time, pressure, and temperature). correct recipe correct temperature, time and pressure tray sealer. | incorrect recipe loaded, causing sealing issues
Packaging. Machine Setup. user ID password program recipe gloves Tray sealer (time, pressure, and temperature). correct recipe correct temperature, time and pressure tray sealer. | incorrect die temperature, inadequate seal strength
Packaging. Machine Setup. user ID password program recipe gloves Tray sealer (time, pressure, and temperature). correct recipe correct temperature, time and pressure tray sealer. | incorrect recipe loaded, causing sealing issues
Packaging. Machine Setup. user ID password recipe ejection cylinder working hand temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect die temperature, inadequate seal strength
Packaging. Machine Setup. user ID password recipe hand correct temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect recipe loaded, causing sealing issues
Packaging. Packaging Inspection. sealed pouch magnifier light. no particles/fibers in tray or pouch. | particles/fibers in tray or pouch
Packaging. Packaging Inspection. sealed pouch magnifier light. no particles/fibers in tray or pouch. | particles/fibers in tray or pouch
Packaging. Packaging Inspection. sealed pouch magnifier light. no particles/fibers in tray or pouch. | particles/fibers in tray or pouch not found
Packaging. Packaging Inspection. sealed pouch magnifier light. no particles/fibers in tray or pouch. | particles/fibers on product
Packaging. Packaging Inspection. sealed pouch, product (Pump, purge cassette, blister, inlays) . Tray free of particulate/ debris. | particles/fibers on product
Packaging. Place cannula inserts into tray. pump tray pump layout in tray inserts. cannula in position. | cannula not in position, kink in pigtail or cannula bend radius
Packaging. Place cannula inserts into tray. pump tray pump layout in tray inserts. cannula in position. | cannula not in position, kink in pigtail or cannula bend radius
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | cannula cover insert not clipped in to tray, sealing issue
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | cannula not in position, kink in pigtail or cannula bend radius
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | foam not seated in plug cover, leading to inadequate seal
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | missing cannula insert cover, cannula damage
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | missing pump plug cover, causing sterile barrier breach during shipping
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | pump not in position, causing pump damage
Packaging. Place pump into tray with covers. pump tray Cannula Covers Foam Inserts Plug cover insert Catheter Butterfly. pump in tray with covers in correct position. | pump plug cover not clipped-in, leading to inadequate seal
Packaging. Prep workstation. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | contaminated trays/Tyvek
Packaging. Prep workstation. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | ionizer on wrong setting
Packaging. Pump cleaning. alcohol pump. cleaned pumps. | Damaged Contamination Sleeve
Packaging. Pump cleaning. alcohol pump. cleaned pumps. | debris/fibers on pump
Packaging. Pump cleaning. alcohol pump. cleaned pumps. | particles/fibers on product
Packaging. Pump cleaning. alcohol pump. cleaned pumps. | particles/fibers on product
Packaging. Pump Packaging. Pump, tray, pouch, air gun. Tray free of particulate/ debris. | particles/fibers migrate to other product
Packaging. Pump Packaging. Pump, tray, pouch, air gun. Tray free of particulate/ debris. | particles/fibers on product
Packaging. Pump placement. pump tray pump layout in tray inserts. pump in tray. | pump not in position, causing pump damage
Packaging. Pump placement. pump tray pump layout in tray inserts. pump in tray. | red plug not clipped in to tray, damage to pump
Packaging. Pump placement. pump tray, operator pump layout in tray, inserts, pump. pump in tray. | Incorrectly placed Repositioning Unit on catheter in the tray
Packaging. Pump placement. pump paper tape. paper tape on patient cable. | paper tape not secured, causing seal breach during shipping
Packaging. Pump placement. pump tray cannula insert cover label. cannula insert cover clipped-in w/label. | cannula insert cover label in wrong location, difficulty removing pump from tray
Packaging. Pump placement. pump tray cannula insert cover label. cannula insert cover clipped-in w/label. | missing cannula insert cover label, difficulty removing pump from tray
Packaging. Pump placement. pump tray cannula insert cover. cannula insert cover clipped-in w/label. | cannula cover insert not clipped in to tray, sealing issue
Packaging. Pump placement. pump tray cannula insert cover. cannula insert cover clipped-in w/label. | missing cannula insert cover, cannula damage
Packaging. Pump placement. pump tray pump layout in tray inserts. pump in tray. | pump not in position, causing pump damage
Packaging. Pump placement. pump tray pump layout in tray inserts. pump in tray. | red plug not clipped in to tray, damage to pump
Packaging. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | foam not seated in plug cover, leading to inadequate seal
Packaging. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | missing pump plug cover label, difficulty removing pump from tray
Packaging. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | missing pump plug cover, causing sterile barrier breach during shipping
Packaging. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | pump plug cover label in wrong location, difficulty removing pump from tray
Packaging. Pump placement. pump tray pump plug cover pump plug cover label pump plug cover foam. pump plug cover clipped-in w/ label and foam. | pump plug cover not clipped-in, leading to inadequate seal
Packaging. Pump placement. pump tray. cannula in position. | cannula not in position, kink in cannula bend radius
Packaging. Pump placement. pump tray. cannula in position. | cannula not in position, kink in pigtail or cannula bend radius
Packaging. Pump placement. pump tray. patient cable plug in position. | plug not in position, damage to patient plug
Packaging. Pump placement. pump tray. pump in tray. | pump not in position, causing pump damage
Packaging. Pump placement. pump tray. pump in tray. | red plug not clipped in to tray, damage to pump
Packaging. Pump placement. tray butterfly clip label. butterfly clip label applied. | butterfly clip label not properly aligned, difficulty opening package
Packaging. Pump placement. tray butterfly clip label. butterfly clip label applied. | missing butterfly clip label, difficulty opening package
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | creases/leak paths in seal
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | inadequate seal strength
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | incomplete seal
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | creases/leak paths in seal
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | inadequate seal strength
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | incomplete seal
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | particles/fibers on product
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging. Seal Pouch . pouch with sealed tray temperature setting dwell time film thickness setting Pouch Sealer. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | creases/leak paths in seal
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | debris/fibers in pouch
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | debris/fibers in seal
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | incomplete seal
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting Pump label. sealed pouch. | weak seal strength
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting. sealed pouch. | creases/leak paths in seal
Packaging. Seal Pouch . pouch sealed tray temperature setting dwell time film thickness setting. sealed pouch. | incomplete seal
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting ionizer on hi. sealed pouch. | debris/fibers in pouch
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting ionizer on hi. sealed pouch. | debris/fibers in seal
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting ionizer on hi. sealed pouch. | particles/fibers on product
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | inadequate seal strength
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | pouch seal not properly aligned leading to breach in sterile barrier
Packaging. Seal Pouch . pouch sealed tray temperature setting time film thickness setting. sealed pouch. | pouch seal too close to tray, pre-load on pouch seal, leading to breach unsterile barrier
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | creases/leak paths in seal
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | inadequate seal strength
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | incomplete seal
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | less than 6mm seal width, loss of seal integrity
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | Tyvek not properly aligned with tray
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | creases/leak paths in seal
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | debris/fibers in tray, contamination of product
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | inadequate seal strength
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | Tyvek not properly aligned with tray
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | Particles/fibers on product
Packaging. Seal Tray. tray assembly Tyvek alignment pins on sealing die time temperature pressure. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | less than 6mm seal width, loss of seal integrity
Packaging. Setup Packaging Machine Setup. user ID password recipe correct temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect die temperature, inadequate seal strength
Packaging. Setup Packaging Machine Setup. user ID password recipe correct temperature @ 4 points. correct recipe correct temperature @ 4 points. | incorrect recipe loaded, causing sealing issues
Packaging. Start program. Place in tray. Click in cover and lay white foil on the hard shell; start sealing; afterwards hard shell is pushed automatically out of the cavity . Tray, Cover, white foil, packaging program (time, temperature, and pressure), Nelipak packaging machine. Sealed tray with pump. | Damage pump side arm
Packaging. Start program. Place in tray. Click in cover and lay white foil on the hard shell; start sealing; afterwards hard shell is pushed automatically out of the cavity . Tray, Cover, white foil, packaging program (time, temperature, and pressure), Nelipak packaging machine. Sealed tray with pump. | Poor seal of package leading sterility compromise
Packaging. Start program. Place in tray. Click in cover and lay white foil on the hard shell; start sealing; afterwards hard shell is pushed automatically out of the cavity . Tray, Cover, white foil, packaging program (time, temperature, and pressure), Nelipak packaging machine. Sealed tray with pump. | Unable to seal package leading sterility compromise
Packaging. Switch on packaging device; remove contaminants with Isopropanol. Isopropanol, Nelipak packaging machine. Prepped Nelipak packaging machine. | Contamination in package seal area impacting sterility
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | creases/leak paths in seal
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | debris/fibers in tray, contamination of product
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | inadequate seal strength
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | incomplete seal
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | particles/fibers on product
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure hands. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | Tyvek not properly aligned with tray
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | creases/leak paths in seal
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | inadequate seal strength
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | incomplete seal
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | less than 6mm seal width, loss of seal integrity
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | particles/fibers on product
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure Tray sealer. Sealed package with pump. | Tyvek not properly aligned with tray
Packaging. Tray sealing. tray assembly Tyvek alignment pins on sealing die time temperature pressure. 6mm seal width Tyvek bonded to internal points no creases in seal no breaks in seal no debris in seal no fibers/debris in tray. | less than 6mm seal width, loss of seal integrity
Packaging. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | contaminated trays/Tyvek
Packaging. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | ionizer on wrong setting
Packaging. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | contaminated trays/Tyvek
Packaging. Work station prep. alcohol ionizers on hi packaging trays Tyvek film. work station ready for use. | ionizer on wrong setting
PCB Installation Mfg Process (MFG-3015). Attach System Power Enable Harness. . ASSY, PCB Installation. | Incorrect cable connection Cable strain Cable damage
PCB Installation Mfg Process (MFG-3015). Connect Breethe Connect Cables. . ASSY, PCB Installation. | Incorrect cable connection Cable strain Cable damage
PCB Installation Mfg Process (MFG-3015). Data transfer export. . ASSY, PCB Installation. | Data not transferred
PCB Installation Mfg Process (MFG-3015). Set Date and time. . ASSY, PCB Installation. | Date and time reset every restart
Pick to Order (PTO). Apply label on set. Label with lot and SN, operator. Proper pump set label applied. | Incorrect SN label used on part
Pick to Order (PTO). Apply label on set. Label with lot and SN, operator. Proper pump set label applied. | Incorrect/imprecise label application
Pick to Order (PTO). Apply label on set. Label with lot and SN, operator. Proper pump set label applied. | labels on pump set box are mismatched
Pick to Order (PTO). Apply label on set. Label with lot and SN, operator. Proper pump set label applied. | Labels on pump sets are applied incorrectly
Pick to Order (PTO). Apply label on set. Label with lot and SN, operator. Proper pump set label applied. | Wrong label used on pump set
Pick to Order (PTO). Assemble set. Work order, comissioned material, operator. Pumps & accessories placed in the package. | Incorrect items packed into set
Pick to Order (PTO). Assemble set. Work order, comissioned material, operator. Pumps & accessories placed in the package. | Incorrect Lot number of items kitted
Pick to Order (PTO). Assemble set. Work order, comissioned material, operator. Pumps & accessories placed in the package. | Incorrect quantity assembled
Pick to Order (PTO). Assemble set. Work order, comissioned material, operator. Pumps & accessories placed in the package. | Items excluded from Pump set
Pick to Order (PTO). Assemble set. work order, comissioned material, operator. Pumps & accessories placed in the package. | Items not placed in a defined configuration
Pick to Order (PTO). Create Work Order(WO). Bill of material, lot Number,order specifications(quantity and type), Operator. Complete and print work order sheet. | Pumps with multiple SLED dates added to same Work Order
Pick to Order (PTO). Create Work Order(WO). Bill of material, lot Number,order specifications(quantity and type), Operator. Complete and print Work Order sheet. | Wrong components selected
Pick to Order (PTO). Create Work Order(WO). Bill of material, lot Number,order specifications(quantity and type), Operator. Complete and print Work Order sheet. | Wrong location picked
Pick to Order (PTO). Create Work Order(WO). Bill of material, lot Number,order specifications(quantity and type), Operator. Complete work order, printed packing list. | Expired pump selected
Pick to Order (PTO). Create Work Order(WO). Bill of material, lot Number,order specifications(quantity and type),Operator. Complete and print Work Order sheet. | Wrong components selected
Pick to Order (PTO). End. Scan Work Order into Drive. Verification of pump set labels and box sets. | Work Order is lost/ incompletly scanned
Pick to Order (PTO). Goods issue(MIGO :261). Work order, Quality Inspector,comissioned material. Successful goods issue . | Part number overwritten manually or unverified
Pick to Order (PTO). Items placement. Packaged Purge Cassette, White Carton (0505-4200), Incision, templates, silicone plugs, Axillary insertion clamps, Sterilization cycle 11- 2ABM2. Packaged Incision templates, axillary insertion clamps or Silicone plugs into White Carton. | Incorrect number of items packaged
Pick to Order (PTO). Line Clearance . SOP-MF11, Work Area, Operators. Clean Work Area. | Foreign Material/ particulate present in Work Area
Pick to Order (PTO). Line Clearance . SOP-MF11, Work Area, Operators. Clean Work Area. | Mechanical failure of motor
Pick to Order (PTO). Line Clearance . SOP-MF11, Work Area, Operators. Correct materials placed in Work Area. | Incorrect Material placed in Pump Set
Pick to Order (PTO). Print Patient Tracking form. Operator, shipping information. Paitent tracking form . | Incorrect Patient Tracking form P/N used.
Pick to Order (PTO). Purge Cassette placement. Packaged Purge Cassette, White Carton (0505-4200), Incision, templates, silicone plugs, Axillary insertion clamps, Sterilization cycle 11- 2ABM1. Packaged Purge Cassette into White Carton. | Incorrect number of Purge cassettes packaged
Pick to Order (PTO). Q.Check. Pump and pump set label. Verification of pump set labels and box sets. | Incorrect Pump serial number
Pick to Order (PTO). Q.check. Pump and pump set label. Verification of pump set labels and box sets. | Label Revision is incorrect
Pick to Order (PTO). Q.Check. Pump and pump set label. Verification of pump set labels and box sets. | Labels are illegible
Pick to Order (PTO). Q.Check. Pump and pump set label. Verification of pump set labels and box sets. | labels are printed incorrectly
Pick to Order (PTO). Q.check. Pump and pump set label. Verification of pump set labels and box sets. | Work Order is incomplete/ Incorrect
Pick to Order (PTO). Quality Check. Work order, Quality Inspector,comissioned material. Verification of correct parts on cart against work order. | Shopped items are not matching Work Order
Pick to Order (PTO). Quality Check. Work order, Quality Inspector,comissioned material. Verification of correct parts on cart against work order. | Shopped items are not matching Work Order (unnoticed)
Pick to Order (PTO). Quality Check. Work order, Quality Inspector,shopped cart. Verification of correct parts on cart against work order. | Mistake in count
Pick to Order (PTO). Scan Barcode on pump. Work order barcode, pump barcode, scanner, operator. Label printed of the scanned pump. | Labels are printed incorrectly
Pick to Order (PTO). Scan Barcode on pump. Work order barcode, pump barcode, scanner, operator. Successful scan of work order and pump. | Incorrect pump scanned
Pick to Order (PTO). Scan Barcode on Work Order. Work order barcode, pump barcode, scanner, operator. Successful scan of work order and pump. | Barcode is not readable
Pick to Order (PTO). Scan Barcode on Work Order. Work order barcode, pump barcode, scanner, operator. Successful scan of work order and pump. | Incorrect WO scanned
Pick to Order (PTO). Second Inspection. Pump and pump set label. Verification of pump set labels and box sets. | Improper pump set verified
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components on Work Order collected. | Wrong revision pulled
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | First in First Out not followed
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | Lowex pump selected (<18 months)
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | Missing items from Work Order
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | Wrong item pulled
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | Wrong quantity pulled
Pick to Order (PTO). Shop. Work Order, Operator, Shopping cart. All components onWork Order collected. | Wrong Serial no. pulled of the same part
Pick to Order (PTO). Transfer. Move pumps to Shipping. Pumps transferred from PTO to FG physically. | Pumps physically moved to incorrect location
Pick to Order (PTO). Transfer. Scan label. Work order automatically closes to FGC(in SAP). | Pumps are not transferred into FGC location
PLU Durable Base Assy Mfg Process (MFG-0005). Assemble PLU Durable Base. . PLU Durb Base Asm. | Breach in sealant
PLU Durable Base Top Asm Mfg Process (MFG-0004). Assemble PLU Durable Base Top. . Base Top Asm. | Motor locked
PLU Durable Base Top Asm Mfg Process (MFG-0004). Assemble PLU Durable Base Top. . Base Top Asm. | No FBS communication
PLU Durable Base Top Asm Mfg Process (MFG-0004). Assemble PLU Durable Base Top. . Base Top Asm. | No Flow bubble sensor communication
PLU Durable Base Top Asm Mfg Process (MFG-0004). Plug in LED Cable. . PLU Durable Base Top Assembly. | Improper LED harness connection
PLU Durable Base Top Asm Mfg Process (MFG-0004). Route LED Cable. . PLU Durable Base Top Assembly. | controller harness damage
PLU Durable Base Top Asm Mfg Process (MFG-0004). Route LED Cable. . PLU Durable Base Top Assembly. | LED Harness damage
PLU Durable Base Top Asm Mfg Process (MFG-0004). Route LED Cable. . PLU Durable Base Top Assembly. | Top Housing damage
PLU Durable Base Top Chassis Assy Mfg Process (MFG-0003) Cable harness #:. Assemble Durable Base Top Chassis. Multiple inputs - see MFG-0003 PLU Durable Base Top Chassis. Base Chassis Asm. | No controller communication
PLU Durable Base Top Chassis Assy Mfg Process (MFG-0003) Cable harness #:. Assemble Durable Base Top Chassis. Multiple inputs - see MFG-0003 PLU Durable Base Top Chassis. Base Chassis Asm. | Wires rubbing on sheet metal edge
PLU Durable Base Top Chassis Assy Mfg Process (MFG-0003). Assemble Durable Base Top Chassis. Multiple inputs - see MFG-0003 PLU Durable Base Top Chassis. Base Chassis Asm. | No controller communication
PLU Durable Base Top Chassis Assy Mfg Process (MFG-0003). Secure cables. . Durable Base Top Chassis Assembly. | Cable strain Cable damage
PLU Durable Base Top Latch Assy Mfg Process (MFG-0007). Assemble PLU Durable Base Top Latch. . Top Latch Asm. | Top Latch Pin and Hinge not properly assembled into Latch Arm and Latch Base Back
PLU Durable Base Top Latch Lever Assy Mfg Process (MFG-0006). Assemble PLU Durable Base Top Latch Lever. . Top Latch Lever Asm. | Latch pin loose
PLU Durable Lid Assy Mfg Process (MFG-0009). Assemble PLU Durable Lid Assy. . PLU Durable Lid Assembly. | PLU Durable window not fastened properly
PLU Durable Lid Hinge Pin Assy Mfg Process (MFG-0008) PLU Durable Asm Mfg Process (MFG-0010). Assemble PLU Durable Lid Hinge Pin Assy. . PLU Durb Lid Hinge Pin Asm. | Lid wont stay open, hinge stack will not go together
PLU Durable Top Housing Assy Mfg Process (MFG-0002). Assemble PLU Durable Top Housing. . PLU Durable Base Top Housing Asm. | Bubble sensor not properly assembled
PLU Durable Top Housing Assy Mfg Process (MFG-0002). Install connector to LED harness. . Top Housing Assembly. | Circuit failure
PLU Durable Top Housing Assy Mfg Process (MFG-0002). Install green LED indicator. . Top Housing Assembly. | Water ingress
PLU Final Assembly Procedure. Assemble final disposable product. PLU Final Asm Procedure. PLU Disposable Asm. | Missing or incorrectly installed gasket/o-ring
PLU Final Assembly Procedure. Assemble final disposable product. PLU Final Asm Procedure. PLU Disposable Asm. | Missing or incorrectly installed gasket/o-ring in blood path
PLU Final Assembly Procedure. Assemble final disposable product. PLU Final Asm Procedure. PLU Disposable Asm. | PLU Pouch assembly missing labels
PLU Final Assembly Procedure. Assemble final disposable product. PLU Final Asm Procedure. PLU Disposable Asm. | Stripped out screw threads in plastic components
PLU Final Assembly Procedure. Assemble final disposable product. PLU Final Asm Procedure. PLU Disposable Asm. | Wrong screw used or missing screw
PLU Final Assembly Test. Perform final assembly test. PLU Final Asm Test. PLU Disposable Asm completed final test. | Device contaminated
PLU Final Assembly Test. Perform final assembly test. PLU Final Asm Test. PLU Disposable Asm completed final test. | Device damaged from over-pressure
PLU Pump Housing Lot Integrity Test. Pump Housing Integrity Test Set up. PLU Pump Housing Lot Integrity Test. Verified Glue Joint. | Test Leaks
PLU Pump Housing Lot Integrity Test. Pump Housing Integrity Test. PLU Pump Housing Lot Integrity Test. Verified Glue Joint. | Induced leaks
PLU Pump Housing Lot Integrity Test. Pump Housing Integrity Test. PLU Pump Housing Lot Integrity Test. Verified Glue Joint. | Leaks not observed
Pneumatic Circuit Enclosure Assy (MFG-3026) Cable harness #:. Assemble Pneumatic Circuit Enclosure. . ASSY, Pneumatic Circuit Enclosure. | Loose sensor connection
Pneumatic Circuit Enclosure Mfg Process (MFG-3026) Cable harness #:. Assemble Pneumatic Circuit Enclosure. . ASSY, Pneumatic Circuit Enclosure. | Loose sensor connection
Pneumatic Circuit Enclosure Mfg Process (MFG-3026) Cable harness #:. Assemble Pneumatic Circuit Enclosure. . ASSY, Pneumatic Circuit Enclosure. | Switched solenoid connectors
Pneumatic Circuit Enclosure Mfg Process (MFG-3026). Assemble Pneumatic Circuit Enclosure. . ASSY, Pneumatic Circuit Enclosure. | Kinked Tubing
Pneumatic Circuit Enclosure Mfg Process (MFG-3026). Assemble Pneumatic Circuit Enclosure. . ASSY, Pneumatic Circuit Enclosure. | Oxygen/Air Leak
Pneumatic Circuit Functional Test. Test pneumatic circuit for functionality. . . | Incorrect test setup and test fails (Tubing connection, fixture leak)
Pneumatic Circuit Functional Test. Test pneumatic circuit for functionality. . . | Incorrect test setup and test passes (Leak in setup, switches incorrectly setup)
Pneumatic Circuit Installation Mfg Process (MFG-3021) Cable Harness #:. Assemble Pneumatic Circuit. . . | Incorrect cable connection Cable strain
Pneumatic Circuit Installation Mfg Process (MFG-3021). Assemble Pneumatic Circuit. . ASSY, Pneumatic Circuit Installation. | Pinched concentrator tubing
Pneumatic Circuit Leak Test. Test pneumatic circuit for leaks. . . | Incorrect test setup and test fails (Tubing connection, fixture leak)
Pneumatic Circuit Leak Test. Test pneumatic circuit for leaks. . . | Incorrect test setup and test passes (Leak in setup, switches incorrectly setup)
Pneumatic Circuit Lid Mfg Process (MFG-3027). Assemble Pneumatic Circuit Lid. . ASSY, Pneumatic Circuit Lid. | Pneumatic circuit lid not properly assembled
Pneumatic Circuit Mfg Process (MFG-3028). Assemble Pneumatic Circuit. . ASSY, Pneumatic Circuit. | Oxygen/Air Leak
Pneumatic Circuit Mfg Process (MFG-3028). Assemble Pneumatic Circuit. . ASSY, Pneumatic Circuit. | Stripped threadforming screw
Pouch Label. Attach label to pouch. Label (ICU Sheath Lock SPEC) . Labeled pouch. | Missing label
Pouch Label. Attach label to pouch. Label (ICU Sheath Lock SPEC) . Labeled pouch. | Wrong label
Pouch Product. Close and latch ICU Sheath Lock and place inside pouch. ICU Sheath Lock, Outer Pouch. ICU Sheath Lock closed and latched inside labeled pouch. | Damage to Pouch
Pouch Product. Close and latch ICU Sheath Lock and place inside pouch. ICU Sheath Lock, Outer Pouch. ICU Sheath Lock closed and latched inside labeled pouch. | Damage to Sheath Lock
Pouch Product. Close and latch ICU Sheath Lock and place inside pouch. ICU Sheath Lock, Outer Pouch. ICU Sheath Lock closed and latched inside labeled pouch. | Incorrect pouch
Pouch Product. Close and latch ICU Sheath Lock and place inside pouch. ICU Sheath Lock, Outer Pouch. ICU Sheath Lock closed and latched inside labeled pouch. | Missing ICU Sheath Lock
Pouch Seal. Place open end of pouch in bar sealer with Tyvek facing downwards and depress foot pedal. ICU Sheath Lock contained in pouch. Correctly sealed pouch containing ICU Sheath Lock accessory. | Inadequate pouch seal
Power Supply Mounting Mfg Process (MFG-3003). Mount Power Supply. . ASSY, Power Supply Mounting. | DC cable not properly connected
Power Supply Mounting Mfg Process (MFG-3003). Mount Power Supply. . ASSY, Power Supply Mounting. | IEC cable insulating material improperly installed
Power Supply Mounting Mfg Process (MFG-3003). Mount Power Supply. . ASSY, Power Supply Mounting. | IEC inlet cable not properly connected
Power Supply Mounting Mfg Process (MFG-3003). Mount Power Supply. . ASSY, Power Supply Mounting. | Power supply inlet cable not properly connected
Preparations of Guiding Lumens. Cut guide lumen to length and attach label. Lumen label application fixture, backloading lumen cutter/flexible tubing cutter, tweexers, rounded end spatula, bar sealer . Backloading lumen with label in bag. | Damage lumen leading to break upon removal leading to pump stop
Preparations of Guiding Lumens. Cut guide lumen to length and attach label. Lumen label application fixture, backloading lumen cutter/flexible tubing cutter, tweexers, rounded end spatula, bar sealer . Backloading lumen with label in bag. | Damage lumen leading to scrap
Pressure Lumen Assembly. Apply glue to lumen, position inside pressure cube and cure under UV spot light . clear catheter, butterfly flushing system, mandrel, loctite 3321, UV spot light (Time and distance), swab. butterfly flushing system with glued in clear catheter. | Leaking at the butterfly flushing system
Pressure Lumen Assembly. Apply Loctite 3321 to red lumen and slide into the butterfly valve, fill the gap with Loctite 3321 and cure . pressure cube, red catheter, Loctite 3321, UV spot light (time and distance), butterfly valve. Pressure Lumen Kink Protector Assembly. | leaking at bond site
Pressure Lumen Assembly. Clean wire mandrel, pressure cube, and red lumen. wire mandrel, pressure cube, red lumen, petroleum ether. cleaned components. | loss of pressure signal due to weak seal
Pressure Lumen Assembly. cut connection clear catheter, clean catheter and junction point on butterfly flushing system valve with petroleum ether. petroleum ether, scalpel or razer, calibrated ruler or flexible tubing cutter, butterfly flushing system valve, swab. cut and cleaned catheter, cleaned butterfly flushing system. | Bonding failure leading to leak
Pressure Lumen Assembly. Cut the red lumen to length. red lumen, red lumen fixture, blade. red lumen cut to length. | unable to assesmble
Pressure Lumen Assembly. Place benderole and kink protector on red lumen. butterfly flushing system with glued in clear catheter, benderole, kink protector. butterfly flushing system with glued in clear catheter with benderole label. | incorrect position of component (cosmetic defect)
Pressure Lumen Assembly. Place benderole and kink protector on red lumen. butterfly flushing system with glued in clear catheter, benderole, kink protector. butterfly flushing system with glued in clear catheter with benderole label. | incorrect position of component (unable to assemble)
Pump Assembly Procedure. Apply glue to suction line. Suction tubing, Loctite 3321, needle. suction line with glue. | EQT failure
Pump Assembly Procedure. Apply glue to suction line. Suction tubing, Loctite 3321, needle. suction line with glue. | EQT failure due to blocked lumen
Pump Assembly Procedure. Apply glue to suction line. Suction tubing, Loctite 3321, needle. suction line with glue. | weak bond leading to loss of purge pressure
Pump Assembly Procedure. Cure bonded area. LED curing system, cylinder and suction line with Loctite 3321. Cured bond. | EQT Failure
Pump Assembly Procedure. Cure bonded area. LED curing system, cylinder and suction line with Loctite 3321. Cured bond. | Incomplete curing
Pump Assembly Procedure. Cure bonded area. LED curing system, cylinder and suction line with Loctite 3321. Cured bond. | weak bond leading to loss of purge pressure
Pump Assembly Procedure. Insert rack and piston into the cylinder such that the teeth of the rack are facing the Suction Port. Rack & Piston Assembly and Cylinder . Rack & Piston Assembly inserted into Cylinder . | Unable to assemble rack and driveshaft (WI 0043-7015)
Pump Assembly Procedure. Insert tubing into suction port of the cylinder. Suction tubing with Loctite 3321 and cylinder. suction line inserted into cylinder. | weak bond leading to loss of purge pressure
Pump Assembly Procedure. Issue label with sub-assembly serial number and temporairly attach to pump assembly. Label and Pump Assembly. Label with manually input information attached to pump assembly. | Did not attach label to Pump Assembly
Pump Assembly Procedure. Issue label with sub-assembly serial number and temporairly attach to pump assembly. Label and Pump Assembly. Label with manually input information attached to pump assembly. | Incorrectly serialized label
Pump Assembly Procedure. Plasma treat suction line. Plasma chamber, Suction line assembly. Plasma treated suction line assembly. | Insufficent plasma treatment
Pump Assembly Procedure. Push open end of black piston onto rack. Piston and Rack. Piston fully seated on rack head. | Piston not fully seated
Pump Assembly Procedure. Wipe end of suction line. Lint free wipes, Suction tubing, IPA. Cleaned suction line. | contaminated surface leading to weak bond
Pump Assembly. Insert a Rack and Piston Assembly into the upper nest. Insert Pump Body Assembly onto lower nest of the test fixture. Slide lower nest backward on the fixture to initiate the test. . Rack & Piston Assembly, Pump Body Assembly, Pump Assembly Functional Test Fixture. Pump Assembly in fixture. | Damage to piston leading leak
Pump Assembly. Insert a Rack and Piston Assembly into the upper nest. Insert Pump Body Assembly onto lower nest of the test fixture. Slide lower nest backward on the fixture to initiate the test. . Rack & Piston Assembly, Pump Body Assembly, Pump Assembly Functional Test Fixture. Pump Assembly in fixture. | piston and rack overshoots causing damage to piston and/or cylinder leading to leak
Pump Assembly. Insert a Rack and Piston Assembly into the upper nest. Insert Pump Body Assembly onto lower nest of the test fixture. Slide lower nest backward on the fixture to initiate the test. . Rack & Piston Assembly, Pump Body Assembly, Pump Assembly Functional Test Fixture. Pump Assembly in fixture. | poor insertion depth leading unacceptable component passes EQ test
Pump Assembly. Insert a Rack and Piston Assembly into the upper nest. Insert Pump Body Assembly onto lower nest of the test fixture. Slide lower nest backward on the fixture to initiate the test. . Rack & Piston Assembly, Pump Body Assembly, Pump Assembly Functional Test Fixture. Pump Assembly in fixture. | too much insertion leading to damage of piston and/or glass leading to leak
Pump Assembly. Check pressure regulator and label maker. Pump Assembly Functional Test Fixture, label maker. Setup Fixture and label maker. | Unable to assemble
Pump Assembly. Ensure correct Piston and Glass Cylinder matching. piston, glass cylinder. correct piston and glass cylinder. | incorrect match of components causing damage of piston leading to leak
Pump Assembly. Ensure correct Piston and Glass Cylinder matching. piston, glass cylinder. correct piston and glass cylinder. | incorrect matching of components leading to blocked piston (unable to move pass case start)
Pump Assembly. Remove glass cylinder from the device and store in carrier. Pump Assembly in fixture, Pump Assembly Functional Test Fixture. Pump Assembly. | damage glass cylinder leading to leak
Pump Drive Assy Mfg Process (MFG-0025). Assemble Pump Drive. . Pump Drive Asm. | Incorrect polarity of fly wheel
Pump Drive Assy Mfg Process (MFG-0025). Assemble Pump Drive. . Pump Drive Asm. | Motor locked
Pump Drive Wheel Assembly (MFG-0028). Assemble Pump Drive Wheel. . . | Incorrect magnetic orientation
Pump Drive Wheel Assembly (MFG-0028). Assemble Pump Drive Wheel. . . | Interference with pump drive
Pump Drive Wheel Assembly (MFG-0028). Assemble Pump Drive Wheel. . . | Magnets shifting
Pump Handle Assembly (0550-7033). Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage sensor wires
Pump Handle Assembly (0550-7033). Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage to patient cable wires
Pump Handle Assembly (0550-7033). Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage unit/leading scrap
Pump Handle Assembly (0550-7033). Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
Pump Handle Assembly (0550-7033). Cover solder joints with UV glue. UV wand, Adhesive, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Improper sealing of board/joints leading to shorts
Pump Handle Assembly (0550-7033). Cover solder joints with UV glue. UV wand, Adhesive, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Insufficient straining of wires leading to them coming loosing causing pump stop
Pump Handle Assembly (0550-7033). Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Increased processing time
Pump Handle Assembly (0550-7033). Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Low bond strength leading to component separation
Pump Handle Assembly (0550-7033). Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Scrap product due to visual defect
Pump Handle Assembly (0550-7033). Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Damage component/unusable
Pump Handle Assembly (0550-7033). Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Logo and/or serial number printed on incorrect handle half
Pump Handle Assembly (0550-7033). Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Smugged/illegible printing
Pump Handle Assembly (0550-7033). Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Damage components/unusable
Pump Handle Assembly (0550-7033). Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Loss of traceability
Pump Handle Assembly (0550-7033). Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Placement signal failure
Pump Handle Assembly (0550-7033). Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Pump halves come apart causing damage to internal components
Pump Handle Assembly (0550-7033). Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Adhesive, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | kink protector window orientation to catheter incorrect
Pump Handle Assembly (0550-7033). Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Adhesive, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | Low bond strength leading to component separation
Pump Handle Assembly (0550-7033). Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Excessive torque applied - Catheter damaged/deformed
Pump Handle Assembly (0550-7033). Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Excessive torque applied - catheter to kink protector bond broken
Pump Handle Assembly (0550-7033). Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Insufficient torque applied
Pump Handle Assembly (0550-7033). Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Damage pump/unusable
Pump Handle Assembly (0550-7033). Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Low motor current(Low flow)
Pump Handle Assembly (0550-7033). Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Short of PCB board
Pump Handle Assembly (0550-7033). Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Wires unsecured and come loose causing pump stop
Pump Handle Assembly (0550-7033). Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | parts printed with incorrect logo
Pump Handle Assembly (0550-7033). Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | Unable to print components with logo
Pump Handle Assembly (0550-7033). Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Low SNR leading to inaccurate placement signal information
Pump Handle Assembly (0550-7033). Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Placement signal failure
Pump Handle Assembly (0550-7033). Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Scrap product due to low SNR failure
Pump Handle Assembly (0550-7033). Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Sensor failure/low sensor signal
Pump Handle Assembly. Bond Purge Lumens. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Stainless Steel Connector, Petroleum Ether, Cleanroom swab, SS Connector Loading Fixture, 30GA Dosing Needle, Loctite 3321. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumens connected. | Low bond strength leading to component separation
Pump Handle Assembly. Bond Purge Lumens. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Stainless Steel Connector, Petroleum Ether, Cleanroom swab, SS Connector Loading Fixture, 30GA Dosing Needle, Loctite 3321. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumens connected. | Purge Lumen trimmed too short
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage sensor wires
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage to patient cable wires
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Damage unit/leading scrap
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Adhesive, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Adhesive. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3325. Handle half with serial number and glued in patient cable assembly. | Damage sensor wires
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3324. Handle half with serial number and glued in patient cable assembly. | Damage to patient cable wires
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3326. Handle half with serial number and glued in patient cable assembly. | Damage unit/leading scrap
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3322. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
Pump Handle Assembly. Clean strain relief, cable and red handle, then glue strain relief end and board to pump handle half. Petroleum ether, Loctite 4062, Handle with serial number, Patient cable assembly, 25GA dosing needle, 23GA dosing needle, UV cure spotlight, Loctite 3323. Handle half with serial number and glued in patient cable assembly. | Low bond strength leading to component separation
Pump Handle Assembly. Cover solder joints with UV glue. UV wand, Adhesive, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Improper sealing of board/joints leading to shorts
Pump Handle Assembly. Cover solder joints with UV glue. UV wand, Adhesive, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Insufficient straining of wires leading to them coming loosing causing pump stop
Pump Handle Assembly. Cover solder joints with UV glue. UV wand, Loctite 3321, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Improper sealing of board/joints leading to shorts
Pump Handle Assembly. Cover solder joints with UV glue. UV wand, Loctite 3321, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue. | Insufficient straining of wires leading to them coming loosing causing pump stop
Pump Handle Assembly. Cut microtube, slide microtube into catheter, put protection lumen above nitinol wire and glue the opening with Loctite. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, microtube, protection lumen, UV light source. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, entrance to catheter with microtube glued in and protection lumen above nitinol wire. | Microtube occlude leading insufficient sterilization of internal catheter
Pump Handle Assembly. Cut microtube, slide microtube into catheter, put protection lumen above nitinol wire and glue the opening with Loctite. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, microtube, protection lumen, UV light source. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, entrance to catheter with microtube glued in and protection lumen above nitinol wire. | protection lumen does not fit on nitinol wire
Pump Handle Assembly. Cut microtube, slide microtube into catheter, put protection lumen above nitinol wire and glue the opening with Loctite. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, microtube, protection lumen, UV light source. Loctite, 23GA dosing needle, Distance and Time, Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, entrance to catheter with microtube glued in and protection lumen above nitinol wire. | protection lumen fits loose on nitinol wire
Pump Handle Assembly. Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Increased processing time
Pump Handle Assembly. Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Low bond strength leading to component separation
Pump Handle Assembly. Glue purge kink protector into top shell handle. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, glue, Kink protector. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in. | Scrap product due to visual defect
Pump Handle Assembly. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Damage component/unusable
Pump Handle Assembly. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Logo and/or serial number printed on incorrect handle half
Pump Handle Assembly. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Smudged/illegible printing
Pump Handle Assembly. Load parts into laser printer and print logo/serial number onto handle half. Laser printer on with logo/serial number file loaded, handle tray. Handle half with logo/serial number. | Smugged/illegible printing
Pump Handle Assembly. Microtube Install. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumens connected. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped and microtube extended from catheter. | Unable to sterilize
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Damage components/unusable
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Loss of traceability
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Placement signal failure
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped with purge lumen kink protector glued in, Printed top shell half, Press Fixture . Pump handle assembly. | Pump halves come apart causing damage to internal components
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Printed top shell half, Press Fixture . Pump handle assembly. | Damage components/unusable
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Printed top shell half, Press Fixture . Pump handle assembly. | Loss of traceability
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Printed top shell half, Press Fixture . Pump handle assembly. | Placement signal failure
Pump Handle Assembly. Place other half of shell on top of assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped, Printed top shell half, Press Fixture . Pump handle assembly. | Pump halves come apart causing damage to internal components
Pump Handle Assembly. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Adhesive, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | kink protector window orientation to catheter incorrect
Pump Handle Assembly. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Adhesive, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | Low bond strength leading to component separation
Pump Handle Assembly. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Loctite 4062, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | kink protector window orientation to catheter incorrect
Pump Handle Assembly. Place unit into fixture and clean gluing areas, then glue catheter/kink protector end of pump into handle half with serial number. Fixture nest, Petroleum ether, Loctite 4062, 25GA dosing needle, Handle half with serial number and glued in patient cable assembly, pump catheter assembly. Handle half with serial number and glued in patient cable assembly and pump catheter assembly. | Low bond strength leading to component separation
Pump Handle Assembly. Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Excessive torque applied - Catheter damaged/deformed
Pump Handle Assembly. Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Excessive torque applied - catheter to kink protector bond broken
Pump Handle Assembly. Rotational Proof Loading of Catheter to Kink Protector Bond. Red handle (Closed), Coil reinforced catheter, red stopper, Pump Cannula Assembly. Pump handle assembly. | Insufficient torque applied
Pump Handle Assembly. Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Damage pump/unusable
Pump Handle Assembly. Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Low motor current(Low flow)
Pump Handle Assembly. Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Short of PCB board
Pump Handle Assembly. Solder the motor cable wires into PCB board. Red handle with patient cable and pump assembly glued in, Assembly Fixture , Solder material, Soldering iron, Handle half with serial number and glued in patient cable assembly and pump catheter assembly. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board. | Wires unsecured and come loose causing pump stop
Pump Handle Assembly. Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | parts printed with incorrect logo
Pump Handle Assembly. Turn on laser printer and select designated pump logo/serial number. Laser printer, laser printer logo/serial number file. Laser printer on with logo/serial number file loaded. | Unable to print components with logo
Pump Handle Assembly. Unpack Patient Cable and inspect orientation of Plug relative to PSS Housing . Visual . Patient Cable with confirmed proper orientation. | Plug orientation unacceptable
Pump Handle Assembly. Unpack Patient Cable and inspect purge lumen . Visual . Patient Cable with confirmed acceptable purge lumen. | Purge lumen unacceptable
Pump Handle Assembly. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Low SNR leading to inaccurate placement signal information
Pump Handle Assembly. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Placement signal failure
Pump Handle Assembly. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Scrap product due to low SNR failure
Pump Handle Assembly. Wrap fibers around fiber guides on PCB board, measure SNR, and then make connection via ferrule sleeve and mount sleeve on clip. Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue, SNR Tester. Tested Red handle with patient cable and pump assembly glued in, with motor cable wires soldered on board with UV glue with fibers wrapped. | Sensor failure/low sensor signal
Pump Housing Welding (0550-7025). Activate Argon supply. Argon supply. Activated argon supply to correctly shield weld seam. | Corroded weld seam (too little or no argon)
Pump Housing Welding (0550-7025). Activate Argon supply. Argon supply. Activated argon supply to correctly shield weld seam. | Sputtering of weld seam (too much argon)
Pump Housing Welding (0550-7025). Activate Argon supply. Argon supply. Activted argon supply to correctly shield weld seam. | Corroded weld seam (too little or no argon)
Pump Housing Welding (0550-7025). Activate Argon supply. Argon supply. Activted argon supply to correctly shield weld seam. | Sputtering of weld seam (too much argon)
Pump Housing Welding (0550-7025). Align starting point of weld by means of crosshairs (re-aligning at each strut). Parts loaded in welding fixture, welding fixture in correct position, laser. Properly aligned starting point. | weld strength comprimised
Pump Housing Welding (0550-7025). Align starting point of weld by means of crosshairs (re-aligning at each strut). Parts loaded in welding fixture, welding fixture in correct position, laser. Properly aligned starting point. | weld strength compromised
Pump Housing Welding (0550-7025). Apply Epotek and cure. mixed epotek, assembly with primer, dispenser, dispensing needle, oven (time and temperature), timer. Welded assembly with cured epotek. | inadequate coverage leading to thrombus formation
Pump Housing Welding (0550-7025). Apply Epotek and cure. mixed epotek, assembly with primer, dispenser, dispensing needle, oven (time and temperature), timer. Welded assembly with cured epotek. | uncured epotek falls off leading to inadequate coverage causing thrombus formation
Pump Housing Welding (0550-7025). Apply UV Adhesive at the slot transition and then cure. Adhesive, cleaned assembly, honle blue point UV energy source, dispenser, dispensing needle. assembly with block transition between strut to hub sensor slot. | epotek flows into struct (unable to assemble sensor)
Pump Housing Welding (0550-7025). Assemble pump housing onto rotor. Pump housing with PEEK sleeve, Rotor. Assembled pump housing to rotor. | Eccentric parts after welding
Pump Housing Welding (0550-7025). Check for correct welding fixture and if incorrect, change. Laser, fixture. correct welding fixture installed in laser. | Unable to assemble
Pump Housing Welding (0550-7025). Check, clean and prepare parts. AU rotor, peek sleeve, pump housing, petroleum ether, cleanroom swabs. Cleaned parts. | Corrosion of welding site leading to weak weld
Pump Housing Welding (0550-7025). Check, clean and prepare parts. AU rotor, peek sleeve, pump housing, petroleum ether, cleanroom swabs. Cleaned parts. | Imperfect weld seam appearance - rough surface/holes in seam
Pump Housing Welding (0550-7025). Clean rotor struts and hub with petroleum ether. welded pump housing assembly, petroleum ether, swabs. Cleaned assembly. | Poor UV adhesion, epotek flows into struts (unable to assemble sensor)
Pump Housing Welding (0550-7025). Let parts dry. Cleaned parts. Parts ready for installation in laser. | Imperfect weld seam appearance - rough surface/holes in seam
Pump Housing Welding (0550-7025). Load parts into welding fixture. Welding fixture, cleaned parts. Fixture with parts loaded. | damage to magnet coating - corrosion
Pump Housing Welding (0550-7025). Load parts into welding fixture. Welding fixture, cleaned parts. Fixture with parts loaded. | Parts not aligned - eccentricity
Pump Housing Welding (0550-7025). load PEEK sleeve onto pump housing. PEEK sleeve, pump housing. PEEK sleeve on Pump housing. | Eccentric parts after welding
Pump Housing Welding (0550-7025). Mix epotek. epotek 302-3M, speed mixer, scale. mixed epotek. | white spots of epotek
Pump Housing Welding (0550-7025). Open program on laser welder. Laser welder, program. laser ready for welding. | Incorrect weld - unable to continue
Pump Housing Welding (0550-7025). Open program on laser welder. Laser welder, program. laser ready for welding. | Program incorrect - insufficient weld strength
Pump Housing Welding (0550-7025). Open program on laser welder. Laser welder, program. laser ready for welding. | Program incorrect - unable to proceed
Pump Housing Welding (0550-7025). Prime the surface of the struct and hub of the welded assembly, then cure in oven. Primer HPA41, swab, oven (time and temperature), timer. welded assembly with primer. | poor adhesion of epotek (thrombus formation)
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scalpel. finished assembly. | Damage to pump housing, scratches (hemolysis)
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scalpel. finished assembly. | Damage yoke, scratches (thrombus formation)
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scalpel. finished assembly. | unable to assemble sensor to slot
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scapel. finished assembly. | Damage to pump housing, scratches (hemolysis)
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scapel. finished assembly. | Damage yoke, scratches (thrombus formation)
Pump Housing Welding (0550-7025). Remove excesses UV Adhesive from sensor slot. welded assembly with cured epotek, scapel. finished assembly. | unable to assemble sensor to slot
Pump Housing Welding (0550-7025). switch on laser. Laser. activated laser. | Unable to weld
Pump Housing Welding (0550-7025). Visually inspect welds. Welded AU. inspected AU. | weld strength comprimised
Pump Housing Welding (0550-7025). Visually inspect welds. Welded AU. inspected AU. | weld strength compromised
Pump Housing Welding (0550-7025). weld PH to AU rotor. Laser, assembled parts. AU rotor pump housing. | weld strength comprimised
Pump Housing Welding (0550-7025). weld PH to AU rotor. Laser, assembled parts. AU rotor pump housing. | weld strength compromised
Purge Cassette Assembly. Adjust pressure regulator to acceptable pressure. Blow-Out Air Fixture, Pressure Regulator. Setup fixture. | damage to components leading leak and loss of purge pressure
Purge Cassette Assembly. Adjust pressure regulator to acceptable pressure. THF dispenser, THF, THF bottle. Filled dispenser. | occlusion of tubing (unable to move pass case start)
Purge Cassette Assembly. Apply half arc of glue to Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | bond failure of cartridge (unable to utilize device)
Purge Cassette Assembly. Apply half arc of glue to Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | damage to suction line (unable to move pass case start)
Purge Cassette Assembly. Assembly gear box and drive shaft onto rack assembly. Gear box, Drive Shaft, rack assesmbly. Assembled components. | Rack position incorrect, failure at final qualification
Purge Cassette Assembly. Assembly gear box and drive shaft onto rack assembly. Gear box, Drive Shaft, rack assesmbly. Assembled components. | Unable to assemble
Purge Cassette Assembly. Clean pressure line and glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | damage tubing leading to leak and loss of purge pressure
Purge Cassette Assembly. Clean pressure line and glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | occlusion of tubing (unable to move pass case start)
Purge Cassette Assembly. Clean pressure line and glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | poor engagement of tubing
Purge Cassette Assembly. Clean pressure line and glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | pressure transmitter not in line with drive shaft causing tubing kink (unable to move pass case start)
Purge Cassette Assembly. Clean pressure line and glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | weak bond leading to loss of purge pressure
Purge Cassette Assembly. Clean pressure line then glue with THF into the Luer Connection of the cartridge cap. Tubing assembly, lint free wipe, THF, THF dispenser, Glued assembly. Completed assembly. | contaminated surface leading to weak bond causing purge pressure loss
Purge Cassette Assembly. Clean tubing and bond tubing to Pump Assembly with THF. Lint-free wipe, alcohol, suction line, THF dispenser, THF, pump assembly. Bonded assembly. | contaminated surface leading to weak bond causing purge pressure loss
Purge Cassette Assembly. Clean tubing and bond tubing to Pump Assembly with THF. Lint-free wipe, alcohol, suction line, THF dispenser, THF, pump assembly. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Purge Cassette Assembly. Clean tubing and bond tubing to Pump Assembly with THF. Lint-free wipe, alcohol, suction line, THF dispenser, THF, pump assembly. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Purge Cassette Assembly. Clean tubing and bond tubing to Pump Assembly with THF. Lint-free wipe, alcohol, suction line, THF dispenser, THF, pump assembly. Bonded assembly. | poor engagement of tubing
Purge Cassette Assembly. Clean tubing and bond tubing to Pump Assembly with THF. Lint-free wipe, alcohol, suction line, THF dispenser, THF, pump assembly. Bonded assembly. | weak bond leading to loss of purge pressure
Purge Cassette Assembly. Clear tubing of residual bonding agent using blow-out fixture and allow for proper cure time. Bonded assembly, blow-out fixture. Cured assembly. | occlusion of tubing (unable to move pass case start)
Purge Cassette Assembly. Clear tubing of residual bonding agent using blow-out fixture and allow for proper cure time. Bonded assembly, blow-out fixture. Cured assembly. | weak bond leading to loss of purge pressure
Purge Cassette Assembly. Connect Suction line to fixture, place pressure transmitter in holding block, and use a quick burst of air to clear tubing. Completed assembly, tubing blow-out air fixture. Completed assembly. | damage transmitter (unable to move pass case start)
Purge Cassette Assembly. Connect Suction line to fixture, place pressure transmitter in holding block, and use a quick burst of air to clear tubing. Completed assembly, tubing blow-out air fixture. Completed assembly. | occlusion of tubing (unable to move pass case start)
Purge Cassette Assembly. Fill THF dispenser until indicator stops rising. THF dispenser, THF, THF bottle. Filled dispenser. | weak bond leading to loss of purge pressure
Purge Cassette Assembly. Place a drop of adhesive inside Cartridge Bottom insert and place cylinder pin . Loctite 4601, Cartridge Bottom, cylinder pin, dosing needle 22 gauge, Pin setting tool, swab. Cartridge Bottom bonded to pin. | bond failure leading to detachement of cylinder pin and blockage of piston (unable to move pass case start)
Purge Cassette Assembly. Place a drop of adhesive inside Cartridge Bottom insert and place cylinder pin . Loctite 4601, Cartridge Bottom, cylinder pin, dosing needle 22 gauge, Pin setting tool, swab. Cartridge Bottom bonded to pin. | glue migrates to rack, blocking piston (unable to move pass case start)
Purge Cassette Assembly. Remove serial number and place on Cartridge Top. Cartridge top, pump assembly. serial number on cartridge top. | Unable to assemble
Purge Cassette Assembly. Using fixture, press RFID chip into cartridge top in correct placement. RFID Chip Press Fixture, RFID label, cartridge top. RFID chip assembled into cartridge top. | damage RFID chip (unable to move pass case start)
Purge Cassette Assembly. Using fixture, press RFID chip into cartridge top in correct placement. RFID Chip Press Fixture, RFID label, cartridge top. RFID chip assembled into cartridge top. | unable to read RFID chip (unable to move pass case start)
Purge Cassette Final Assy Gen 2. Apply glue to outflow valve. Outflow valve, Loctite 3321. Outflow valve with glue. | EQT failure
Purge Cassette Final Assy Gen 2. Apply glue to outflow valve. Outflow valve, Loctite 3321. Outflow valve with glue. | weak bond leading to loss of purge pressure
Purge Cassette Final Assy Gen 2. Apply half arc of glue to Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | bond failure of cartridge (unable to utilize device)
Purge Cassette Final Assy Gen 2. Apply half arc of glue to either end of Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | bond failure of cartridge (unable to utilize device)
Purge Cassette Final Assy Gen 2. Apply half arc of glue to either end of Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | damage to suction line (unable to move past case start)
Purge Cassette Final Assy Gen 2. Apply half arc of glue to either end of Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | Kinked Suction Line
Purge Cassette Final Assy Gen 2. Apply half arc of glue to either end of Cartridge Bottom. Insert components into Cartridge Bottom and add Ventilation Valve. Apply glue into interlocking groove of Cartridge Bottom and to the half arc of the Cartridge top. Then attach Cartridge Top onto the Cartridge bottom and compress the parts in fixture. Loctite 4601, Cartridge Bottom with cyinder pin, dosing needle 22 gauge, Cartridge top, ventilation valve, pump assembly with suction line, Cartridge Assembly Fixture. Glued assembly. | Tightness test failure due to kinking
Purge Cassette Final Assy Gen 2. Cure bonded area. LED curing system, cylinder and outflow valve with Loctite 3321. Cured bond. | Incomplete curing
Purge Cassette Final Assy Gen 2. Insert outflow valve into cylinder until blue end bottoms out at end of bonding region. Outflow valve, Loctite 3321 and cylinder. Outflow valve to cylinder assembly. | EQT failure
Purge Cassette Final Assy Gen 2. Insert outflow valve into cylinder until blue end bottoms out at end of bonding region. Outflow valve, Loctite 3321 and cylinder. Outflow valve to cylinder assembly. | weak bond leading to loss of purge pressure
Purge Cassette Final Assy Gen 2. Insert Rack of Pump Assembly into Gearbox and Insert Driveshaft into larger remaining hole of Gearbox). Gear box, Drive Shaft, Cylinder, Suction Line, rack and piston assesmbly. Assembled components. | EQT Failure
Purge Cassette Final Assy Gen 2. Insert Rack of Pump Assembly into Gearbox and Insert Driveshaft into larger remaining hole of Gearbox). Gear box, Drive Shaft, Cylinder, Suction Line, rack and piston assesmbly. Assembled components. | Unable to assemble
Purge Cassette Final Assy Gen 2. Place a drop of adhesive inside Cartridge Bottom insert and place cylinder pin . Loctite 4601, Cartridge Bottom, cylinder pin, dosing needle 22 gauge, Pin setting tool, swab. Cartridge Bottom bonded to pin. | glue migrates to rack, blocking piston (unable to move past case start)
Purge Cassette Final Assy Gen 2. Place a drop of adhesive inside Cartridge Bottom insert and place cylinder pin . Loctite 4601, Cartridge Bottom, cylinder pin, dosing needle 22 gauge, Pin setting tool, swab. Cartridge Bottom bonded to pin. | Mechanical Integrity Test Failure (EQT)
Purge Cassette Final Assy Gen 2. Using fixture, press RFID chip into cartridge top in correct placement. RFID Chip Press Fixture, RFID label, cartridge top. RFID chip assembled into cartridge top. | damage RFID chip (unable to move past case start)
Purge Cassette Final Assy Gen 2. Using fixture, press RFID chip into cartridge top in correct placement. RFID Chip Press Fixture, RFID label, cartridge top. RFID chip assembled into cartridge top. | EQT Failure
Purge Cassette Final Assy Gen 2. Using fixture, press RFID chip into cartridge top in correct placement. RFID Chip Press Fixture, RFID label, cartridge top. RFID chip assembled into cartridge top. | unable to read RFID chip (unable to move past case start)
Purge Cassette Final Assy Gen 2. Wipe outflow check valve of pressure line. Lint free wipes, Outflow valve. Cleaned outflow valve. | contaminated surface leading to weak bond causing purge pressure loss
Purge Cassette Final Assy Gen 2. Wipe outflow check valve of pressure line. Lint free wipes, Outflow valve. Cleaned outflow valve. | EQT failure
Purge Cassette Packaging. Connect Quick Setup Bifurcated Assembly to Pressure Line tubing . Purge Cassette in inner tray, Inner Tray, Quick Setup Bifurcated Assembly. Unit in Tray, Quick Setup Bifurcated Assembly connected to pressure line. | Quick Setup Bifurcated Assembly not included (non-quick setup configuration instead of quick setup configuration)
Purge Cassette Packaging. Connect Quick Setup Bifurcated Assembly to Pressure Line tubing . Purge Cassette in inner tray, Inner Tray, Quick Setup Bifurcated Assembly. Unit in Tray, Quick Setup Bifurcated Assembly connected to pressure line. | parts not connected properly (failure to move pass case start)
Purge Cassette Packaging. Connect Quick Setup Bifurcated Assembly to Pressure Line tubing . Purge Cassette in inner tray, Inner Tray, Quick Setup Bifurcated Assembly. Unit in Tray, Quick Setup Bifurcated Assembly connected to pressure line. | parts not connected properly (failure to move past case start)
Purge Cassette Packaging. Place other components in tray and connect appropriate parts. Purge Cassette in inner tray, Inner Tray. Unit in Tray. | Quick Setup Bifurcated Assembly included (quick setup configuration instead of non-quick setup configuration)
Purge Cassette Packaging. Place other components in tray and connect appropriate parts. Purge Cassette in inner tray, Inner Tray. Unit in Tray. | Kink tubing leading to occlusion (failure to move pass case start)
Purge Cassette Packaging. Place other components in tray and connect appropriate parts. Purge Cassette in inner tray, Inner Tray. Unit in Tray. | Kink tubing leading to occlusion (failure to move past case start)
Purge Cassette Packaging. Place other components in tray and connect appropriate parts. Purge Cassette in inner tray, Inner Tray. Unit in Tray. | parts not connected properly (failure to move pass case start)
Purge Cassette Packaging. Place other components in tray and connect appropriate parts. Purge Cassette in inner tray, Inner Tray. Unit in Tray. | parts not connected properly (failure to move past case start)
Purge Cassette Packaging. Place Tyvek lid over Blister and Align within white locating pins. Press and hold Sealer Activation Buttons to initiate sealing cycle. Visually inspect seal. Assembly in inner tray, tyvek cover, Nelipak Tray Sealer (time, pressure, temperature). Packaged unit. | inadequate package seal
Purge Cassette Packaging. Put green and tray cover on tray and place unit in tray into blue tray. Place Tyvek lid over Blister and Align within white locating pins. Press and hold Sealer Activation Buttons to initiate sealing cycle. Visually inspect seal. Green tray cover, blue tray cover, assembly in inner tray, tyvek cover, Nelipak Tray Sealer (time, pressure, temperature). Packaged unit. | inadequate package seal
Purge Cassette Packaging. Slide tray with components in place, green tray cover, and sealing blister through the Static Clean MCS. Static Clean MCS vacuum hood, ionizing bar, packaged assembly, green tray. Unit in Tray. | unable to seal blister (scrap)
Purge Cassette Packaging. Use Ionized air gun to blow off inner tray. Then place Purge Cassette in inner tray with printed symbols to the left. Ionzer air gun, Inner Tray, purge cassette final assembly. Purge Cassette in inner tray. | Kink tubing leading to occlusion (failure to move pass case start)
Purge Cassette Packaging. Use Ionized air gun to blow off inner tray. Then place Purge Cassette in inner tray with printed symbols to the left. Ionzer air gun, Inner Tray, purge cassette final assembly. Purge Cassette in inner tray. | Kink tubing leading to occlusion (failure to move past case start)
Purge Cassette Packaging. Use Ionized air gun to blow off inner tray. Then place Purge Cassette in inner tray with printed symbols to the left. Ionzer air gun, Inner Tray, purge cassette final assembly. Purge Cassette in inner tray. | Particles in tray/contamination of unit in tray
Purge Cassette Pressure Transmitter Component Test. Place tubing with connections in housing and perform test. Connected pressure transmitter to Sprint IQ tester tubing, Pressure Transmitter Housing, purge cassette pressure transmitter. Tested Transmitter. | damage transmitter (unable to move pass case start)
Purge Cassette Pressure Transmitter Component Test. Place tubing with connections in housing and perform test. Connected pressure transmitter to Sprint IQ tester tubing, Pressure Transmitter Housing, purge cassette pressure transmitter. Tested Transmitter. | damage transmitter (unable to move past case start)
Purge Cassette Pressure Transmitter Component Test. Place tubing with connections in housing and perform test. Connected pressure transmitter to Sprint IQ tester tubing, Pressure Transmitter Housing, purge cassette pressure transmitter. Tested Transmitter. | tubing dislodges leading to failure of good components
Purge Cassette Pressure Transmitter Component Test. Place tubing with connections in housing and perform test. Connected pressure transmitter to Sprint IQ tester tubing, Pressure Transmitter Housing, purge cassette pressure transmitter. Tested Transmitter. | Unacceptable units are passed leading case start failures in field
Purge Cassette Pressure Transmitter Component Test. Place tubing with connections in housing and perform test. Connected pressure transmitter to Sprint IQ tester tubing, Pressure Transmitter Housing, purge cassette pressure transmitter. Tested Transmitter. | Unacceptable units are passed leading to case start failures in field
Purge Cassette Pressure Transmitter Component Test. Set up dead end tubing. Closed Valve, Cap. setup dead end tubing. | tubing not on valve leading to rejection of good parts
Purge Cassette Pressure Transmitter Component Test. Set up Sprint IQ tester tubing. Then confirm Program #11 is selected.. Sprint IQ Tester, Cap. Setup system. | Unacceptable units are passed leading case start failures in field
Purge Cassette Pressure Transmitter Component Test. Set up Sprint IQ tester tubing. Then confirm Program #11 is selected.. Sprint IQ Tester, Cap. Setup system. | Unacceptable units are passed leading to case start failures in field
Purge Cassette Pressure Transmitter Component Test. Visual Inspection of transmitter for defects. Connect Pressure Transmitter to Sprint IQ tester tubing and dead-ended tubing. Setup system, Sprint IQ Tester Tubing, Dead-ended tubing. Connected pressure transmitter to Sprint IQ tester tubing. | tubing not tight leading to failure of good components
Purge Cassette Pressure Transmitter Component Test. Visual Inspection of transmitter for defects. Connect Pressure Transmitter to Sprint IQ tester tubing and dead-ended tubing. Setup system, Sprint IQ Tester Tubing, Dead-ended tubing. Connected pressure transmitter to Sprint IQ tester tubing. | Unacceptable units are passed leading case start failures in field
Purge Cassette Pressure Transmitter Component Test. Visual Inspection of transmitter for defects. Connect Pressure Transmitter to Sprint IQ tester tubing and dead-ended tubing. Setup system, Sprint IQ Tester Tubing, Dead-ended tubing. Connected pressure transmitter to Sprint IQ tester tubing. | Unacceptable units are passed leading to case start failures in field
Purge Lumen Kink Protector. Cut protective catheter and clean. protective catheter, cutting block aide, cutting board, flexible tubing cutter, petroleum ether, clean wipes. cut and cleaned protective catheter. | incorrect length leading to unable to assemble unit
Purge Lumen Kink Protector. Cut protective catheter and clean. protective catheter, cutting block aide, cutting board, flexible tubing cutter, petroleum ether, wipes. cut and cleaned protective catheter. | not cleaned catheter leading to bond failure at connection between catheter and kink protector leading to compromised lumen bond
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 2. Purge lumen Kink protector sub-assembly unit. | Assembly not correct length leading to scrap
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 3. Purge lumen Kink protector sub-assembly unit. | Assembly not correct length leading to scrap
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 2. Purge lumen Kink protector sub-assembly unit. | bond failure at connection between catheter and kink protector leading to compromised lumen bond
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 3. Purge lumen Kink protector sub-assembly unit. | bond failure at connection between catheter and kink protector leading to compromised lumen bond
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 2. Purge lumen Kink protector sub-assembly unit. | bond failure at connection between catheter and kink protector leading to excessive stress on lumen bond
Purge Lumen Kink Protector. Insert protective catheter into kink protector, cut to length, apply glue, and cure using UV. Loctite 3321, Kink protector, UV flood lamp (intensity, time, and distance), Honle Bluepoint 2, Loctite UV adhesive dispensor, Assembly aid, Purge lumen kink protector tubing height gauge, Assembly Aid 1, Assembly Aid 2. Purge lumen Kink protector sub-assembly unit. | Glue migrates into catheter
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether . Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | bond failure at connection between catheter and kink protector leading to compromised lumen bond
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether. Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | bond failure at connection between catheter and kink protector leading to stress on lumen bond
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether. Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | bond failure at connection between lumen and kink protector leading to stress on lumen bond
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether. Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | bond failure at connection between kink protector and handle compromising lumen bond
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether. Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | scrap due to visual defects
Purge Lumen Kink Protector. Trim flash around kink protector then clean surfaces with petroleum ether. Kink protector, scapel or surgical scissors, petroleum ether, swab. Trimmed kink protector. | unable to assemble/difficulty assembling
QSK Bifurcated Tubing Bond. Adjust pressure regulator to acceptable pressure. Tubing Blow-Out Air Fixture, pressure regulator. Setup fixture. | damage tubing leading to leak and loss of purge pressure
QSK Bifurcated Tubing Bond. Adjust pressure regulator to acceptable pressure. Tubing Blow-Out Air Fixture, pressure regulator. Setup fixture. | occlusion of tubing (unable to move pass case start)
QSK Bifurcated Tubing Bond. Adjust pressure regulator to acceptable pressure. Tubing Blow-Out Air Fixture, pressure regulator. Setup fixture. | occlusion of tubing (unable to move past case start)
QSK Bifurcated Tubing Bond. Bond components in appropriate order, using blow-out fixture to blow air through assembly after each bonding step. Y-tubing Inflow Luer, 65mm tubing, Y-Joint, 125mm tubing, 175mm tubing, tethered cap, red luer, yellow luer, Y-tubing clamp, drying rack, timer, THF, THF dispenser. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
QSK Bifurcated Tubing Bond. Bond components in appropriate order, using blow-out fixture to blow air through assembly after each bonding step. Y-tubing Inflow Luer, 65mm tubing, Y-Joint, 125mm tubing, 175mm tubing, tethered cap, red luer, yellow luer, Y-tubing clamp, drying rack, timer, THF, THF dispenser. Bonded assembly. | occlusion of tubing (unable to move pass case start)
QSK Bifurcated Tubing Bond. Bond components in appropriate order, using blow-out fixture to blow air through assembly after each bonding step. Y-tubing Inflow Luer, 65mm tubing, Y-Joint, 125mm tubing, 175mm tubing, tethered cap, red luer, yellow luer, Y-tubing clamp, drying rack, timer, THF, THF dispenser. Bonded assembly. | occlusion of tubing (unable to move past case start)
QSK Bifurcated Tubing Bond. Bond components in appropriate order, using blow-out fixture to blow air through assembly after each bonding step. Y-tubing Inflow Luer, 65mm tubing, Y-Joint, 125mm tubing, 175mm tubing, tethered cap, red luer, yellow luer, Y-tubing clamp, drying rack, timer, THF, THF dispenser. Bonded assembly. | weak bond leading to loss of purge pressure
QSK Bifurcated Tubing Bond. Place assembly on tubing setup, attach filter cap, and test for pressure decay and flow. Cured assembly, Sprint LC-PF Tester, Sprint iQ Tester, QSK Test Station Tubing. Tested assembly. | damage leur connection leading to leak and loss of purge pressure leading test failure
QSK Bifurcated Tubing Bond. Place completed assemblies on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Compressed air, bonded assembly, drying fixture, manifold. Cured assembly. | occlusion of tubing (unable to move pass case start)
QSK Bifurcated Tubing Bond. Place completed assemblies on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Compressed air, bonded assembly, drying fixture, manifold. Cured assembly. | occlusion of tubing (unable to move past case start)
QSK Bifurcated Tubing Bond. Place completed assemblies on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Compressed air, bonded assembly, drying fixture, manifold. Cured assembly. | weak bond leading to loss of purge pressure
QSK Bifurcated Tubing Bond. Place completed assembly on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Drying fixture, manifold, timer, bonded assembly, luer connector. Cured assembly. | occlusion of tubing (unable to move pass case start)
QSK Bifurcated Tubing Bond. Place completed assembly on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Drying fixture, manifold, timer, bonded assembly, luer connector. Cured assembly. | occlusion of tubing (unable to move past case start)
QSK Bifurcated Tubing Bond. Place completed assembly on manifold and mount on drying fixture, open air valves, and dry for appropriate cure time. Drying fixture, manifold, timer, bonded assembly, luer connector. Cured assembly. | weak bond leading to loss of purge pressure
QSK Bifurcated Tubing Bond. Remove assembly from drying rack and bond filter to QSK Setup Kit. Filter, cured assembly, THF, THF dispenser, Tubing Blow-out Air Fixture. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
QSK Bifurcated Tubing Bond. Remove assembly from drying rack and bond filter to QSK Setup Kit. Filter, cured assembly, THF, THF dispenser, Tubing Blow-out Air Fixture. Bonded assembly. | occlusion of tubing (unable to move pass case start)
QSK Bifurcated Tubing Bond. Remove assembly from drying rack and bond filter to QSK Setup Kit. Filter, cured assembly, THF, THF dispenser, Tubing Blow-out Air Fixture. Bonded assembly. | occlusion of tubing (unable to move past case start)
QSK Bifurcated Tubing Bond. Remove assembly from drying rack and bond filter to QSK Setup Kit. Filter, cured assembly, THF, THF dispenser, Tubing Blow-out Air Fixture. Bonded assembly. | weak bond leading to loss of purge pressure
QSK Bifurcated Tubing Bond. Turn on fume hood. Then fill THF dispenser until indicator stops rising.. Fume Hood, THF dispenser, THF. Turned on hood and filled THF dispenser. | weak bond leading to loss of purge pressure
QSK Bifurcated Tubing Bond. Verify parts cleaning of all applicable components, clean bonding regions of tubing, and attach non-vented luer cap. lint-free cloth, alcohol, Non-Vented Luer Cap, Tubing components. cleaned components. | contaminated surface leading to weak bond causing purge pressure loss
Qualification, Visual Inspection. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection blocked
Qualification, Visual Inspection. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge Side Arm connection bond failure
Qualification, Visual Inspection. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm connection leaks
Qualification, Visual Inspection. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge side arm damaged
Qualification, Visual Inspection. Assemble Purge Side Arm and stress test system. Petroleum ether Pump PSS Banderole Glue Dispenser Loctite Holne UV Box THF Syringe. Pump with PSS attached. | Purge system occluded or break not detected
Qualification, Visual Inspection. Calibrate optical sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Gage factors/G0 not loaded
Qualification, Visual Inspection. Calibrate optical sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Low SNR leading to scrap
Qualification, Visual Inspection. Calibrate optical sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Sensor failure/low sensor signal
Qualification, Visual Inspection. Calibrate optical sensor (i.e. load Gage Factors). Pump with serial number programmed and initial data structure Deluxe Console Gen 3 2D Barcode Scanner. Pump with G0-G3 loaded, serial number, initial data structure. | Wrong Gage factors/G0 offset loaded
Qualification, Visual Inspection. Connect pump plug to pump programmer and program pump with serial number. EEPROM Programmer, finished pump assembly, software . Finished pump assembly with serial number. | mismatch serial number leading to loss of traceability
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. offset and gain). Test Console, finished pump assembly, software . Programmed pump. | Corrupted data leading to pump stop or failure to start
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. offset and gain). Test Console, finished pump assembly, software . Programmed pump. | incorrect gains leading to pump stop or failure to start
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Finished pump assembly with serial number. | Damage lemo connector pins (scrap)
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Finished pump assembly with serial number. | mismatch serial number leading to loss of traceability
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | Corrupted data leading to pump stop or failure to start
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | incorrect gains leading to pump stop or failure to start
Qualification, Visual Inspection. Connect pump plug to test console and program pump (incl. serial number, offset and gain). Test Console (EEPROM Programmer), finished pump assembly, software . Programmed pump. | Incorrect programmed serial number
Qualification, Visual Inspection. Connect to test setup, apply pressure level, remove pressure and verify pressure decay vs. time meets specification. Smart Manometer, finished pump assembly, pressure setup with nozzle, Test Console (EEPROM programmer), Timer. Verified finished pump assembly. | pump with unacceptable pressure reading passes
Qualification, Visual Inspection. Connect to test setup, apply pressure level, remove pressure and verify pressure decay vs. time meets specification. Smart Monometer, finished pump assembly, pressure setup with nozzle, Test Console (EEPROM programmer), Timer. Verified finished pump assembly. | Damage lemo connector pins (scrap)
Qualification, Visual Inspection. Connect to test setup, apply pressure level, remove pressure and verify pressure decay vs. time meets specification. Smart Monometer, finished pump assembly, pressure setup with nozzle, Test Console (EEPROM programmer), Timer. Verified finished pump assembly. | pump with unacceptable pressure reading passes
Qualification, Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V for motor leakage current, after that with +10V and -10V for sensor leakage current. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finished pump assembly. | Contimnation of pump (sterility compromised)
Qualification, Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V for motor leakage current, after that with +10V and -10V for sensor leakage current. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finished pump assembly. | Leakage current from motor leading to damage to mechanical structures causing pump stop
Qualification, Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V for motor leakage current, after that with +10V and -10V for sensor leakage current. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finished pump assembly. | pump with leakage current passes
Qualification, Visual Inspection. Create solution NaHCO3 and place electrodes in bath. Connect purge lumen with the compressed air supply, position pump in the bath and perform the leakage current inspection with +20V and -20V for motor leakage current, after that with +10V and -10V for sensor leakage current. Pre-cleaned beakers, leakage current test setup, solution NaHCO3, air supply, digital system multimeter, electrodes. Tested finished pump assembly. | Salt residue on the pump leading to corrosion
Qualification, Visual Inspection. Electrical test for sensor measurement system. Pump, dark box, electrical measurement setup (mulitmeter, power supply), software (ESD Measurement). Tested pump unit. | Damage sensor (ESD)
Qualification, Visual Inspection. Electrical test for sensor measurement system. Pump, dark box, electrical measurement setup (mulitmeter, power supply), software (ESD Measurement). Tested pump unit. | Failed sensor not detected
Qualification, Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | contamination of fluid path
Qualification, Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | Contamination of pump surface
Qualification, Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | corrosion of inner mechanical components leading to pump stop
Qualification, Visual Inspection. Hang pumps in bracket in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | Damage catheter scrapping/kinking
Qualification, Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | contamination of fluid path
Qualification, Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | Contamination of pump surface
Qualification, Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | corrosion of inner mechanical components leading to pump stop
Qualification, Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Drying Rack fixture. Reduced humidity pumps. | Damage catheter scrapping/kinking
Qualification, Visual Inspection. Hang pumps in rack in an oven and connect them with compressed air. Oven (time and temperature), clean compressed air, pump assembly, Rack fixture. Reduced humidity pumps. | contamination of fluid path
Qualification, Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | contamination of fluid path
Qualification, Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too high)
Qualification, Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too high) - scrap
Qualification, Visual Inspection. Peform pump run test . Pump with motor current bin, console, run test sofware, PC, test plug, beaker with pyrgon free water (room temperature), Run test fixture. Characterized pump motor current and report for run test. | Incorrect characterized pump motor current (too low)
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Damage impeller flow issues
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Damage Impeller/Pump housing internal surface/cannula internal surface
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Particluate accumulation cause thrombus formation
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump fails to start/corrupted data
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump with incorrect bins is accepted
Qualification, Visual Inspection. Perform final visual inspection, evaluating pump for correct programming, damage, and configuration. Endoscope, microscope, Ispropyl, wipes, console, pump programmer, ESD bands. Finshed pump assembly. | Pump with incorrect gains accepted
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | contamination of fluid path
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Damage catheter scrapping/kinking
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Incorrect characterized pump motor current (too high)
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Incorrect characterized pump motor current (too high) - scrap
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Incorrect characterized pump motor current (too low)
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Pass failed units for airflow (high purge pressure)
Qualification, Visual Inspection. Perform purge flow run test. Calibrated pump, console, run test sofware, PC, test plug, beaker with pyrogen free water (room temperature), Run test fixture, purge flow sensor. Characterized pump motor current and purge flow, and report for run test. | Pass failed units for purge flow
Qualification, Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM programmer), Formula, calculator, 36.5 degree pyrogen free water. Calibrated sensor on finished pump assembly. | pump sensor incorrectly calibrated
Qualification, Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM progrmmer), Formula, calculator, 20 degree pyrogen free water. Calibrated sensor on finished pump assembly. | Damage lemo connector pins (scrap)
Qualification, Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console (EEPROM progrmmer), Formula, calculator, 20 degree pyrogen free water. Calibrated sensor on finished pump assembly. | pump sensor incorrectly calibrated
Qualification, Visual Inspection. Perform Sensor Calibration. Finished pump assembly, beaker, test setup, Test console, Formula, calculator, 36.5 degree pyrogen free water. Calibrated sensor on finished pump assembly. | pump sensor incorrectly calibrated
Qualification, Visual Inspection. Perform Tightness test at hub. time, test setup, glass beaker, lumen, air pressure, pump assembly. Tested finished pump assembly. | Untight pump passes (impacting purge pressure)
Qualification, Visual Inspection. Position pump plug in print device; dial in pump serial number and perform printing. Domino Printer (ink), Finished pump assembly . Finished pump assembly with serial number. | missing/incorrect/smudged serial number leading to loss of traceability
Qualification, Visual Inspection. Program pump with final programming file. Pump, Handheld programmer, file, PC, cable. Programmed pump and report for programming. | incorrect program (wrong pump)
Qualification, Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Damage lemo connector pins (scrap)
Qualification, Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | incorrect motor current bin loaded leading to lower/inaccurate flows/no flow
Qualification, Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Unable to perform pump run test
Qualification, Visual Inspection. Program motor current bin into pump. Pump, Handheld programmer, bin file, PC, cable. Pump with motor current bin programmed and report for programming. | Unable to perform purge rate test
Qualification, Visual Inspection. Purge flow rate test. Purge rate test setup (monometer, pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Airflow). | Damage catheter scrapping/kinking
Qualification, Visual Inspection. Purge flow rate test. Purge rate test setup (monometer, pressure supply, tubings), finished pump assembly, console, flow sensors. Measured assembly (Airflow). | Pass failed units for airflow (high purge pressure)
Rack and Piston Assembly. Ensure pressure regulator is set to acceptable pressure. Rack & Piston Assembly Fixture. Rack & Piston Assembly Fixture with correct pressure. | Damage to piston leading leak
Rack and Piston Assembly. Insert piston into the right cradle with the help of a supporting tool and lock the cradle. Slide carriage to the working position and use hand crack to move piston into rack.. Piston, Rack, Rack & Piston Assembly Fixture, piston pre-assembly tool. components positioned in equipment. | incorrect position of components casusing damage of piston and rack leading to leak
Rack and Piston Assembly. Insert rack into the left cradle and lock the cradle. Rack & Piston Assembly Fixture, Rack. Rack inside cradle of fixture. | incorrect insertion into rack leading to damage of rack (unable to move pass case start)
Rack and Piston Assembly. Place Split pin holder over cradle. Then stamp the split pin into the Rack and Piston. Remove assembly from fixture. Split pin holder, split pin, Piston, Rack, Rack & Piston Assembly Fixture, split pin assembly tool. Rack and Piston Assembly. | incorrect assembly leading to damage of piston causing leak
Rack and Piston Assembly. Place Split pin holder over cradle. Then stamp the split pin into the Rack and Piston. Remove assembly from fixture. Split pin holder, split pin, Piston, Rack, Rack & Piston Assembly Fixture, split pin assembly tool. Rack and Piston Assembly. | incorrect position of components leading to misalignment of piston and rack causing leak
Rack and Piston Assembly. Place Split pin holder over cradle. Then stamp the split pin into the Rack and Piston. Remove assembly from fixture. Split pin holder, split pin, Piston, Rack, Rack & Piston Assembly Fixture, split pin assembly tool. Rack and Piston Assembly. | shallow split pin leading to block piston (unable to move pass case start)
Rear Shell Mfg Process (MFG-3006). Attached panel mount receptable. . ASSY, Expansion Panel. | Panel Mount is not secured
Rear Shell Mfg Process (MFG-3006). Insert Handle Dampening Foam. . ASSY, Rear Shell. | Dampening foam not installed in the handle indentation
Rear Shell Mfg Process (MFG-3006). Install PLU Outlet Cable . . ASSY, Rear Shell. | LEMO GND (Grounding) Strap harness is not attached to PLU output harness
Rear Shell Mfg Process (MFG-3006). Power Cord Strap. . ASSY, Rear Shell. | Power Cord Tangling
Relief Valve Mfg Process (MFG-3031). Assemble Relief Valve. . ASSY, Relief Valve. | Relief valve not properly assembled
Reposition Unit Assembly. Insert dialator into unit. Then apply loctite to the butterfly and mount together with the wound closure device. Loctite 4061, butterfly, wound closure device, 9Fr dialator, dosing needle. Bonded together unit. | Damage tip of wound closure/repo unit device leading to bleeding
Reposition Unit Assembly. Insert dialator into unit. Then apply loctite to the butterfly and mount together with the wound closure device. Loctite 4061, butterfly, wound closure device, 9Fr dialator, dosing needle. Bonded together unit. | Positioning issues with pump (i.e. pump moves)
Reposition Unit Assembly. Insert dialator into unit. Then apply loctite to the butterfly and mount together with the wound closure device. Loctite 4061, butterfly, wound closure device, 9Fr dialator, dosing needle. Bonded together unit. | Sealing in forward position leading leak
Reposition Unit Assembly. Insert dialator into unit. Then apply loctite to the butterfly and mount together with the wound closure device. Loctite 4061, butterfly, wound closure device, 9Fr dialator, dosing needle. Bonded together unit. | Sealing in forward position leading positioning issues
Reposition Unit Assembly. Insert dialator into unit. Then apply loctite to the butterfly and mount together with the wound closure device. Loctite 4061, butterfly, wound closure device, 9Fr dialator, dosing needle. Bonded together unit. | weak bond leading to leakage at repo unit
Reposition Unit Assembly. Add contamination sleeve and store unit. Repositioning Unit, contamination sleeve, plastic bags. Stored assembly. | Damage of contamination sleeve leading contamination
Reposition Unit Assembly. Add contamination sleeve and store unit. Repositioning Unit, contamination sleeve, plastic bags. Stored assembly. | Damage of the tip of the assembly leading to bleeding
Reposition Unit Assembly. Apply primer to the wound closure device. Wound closure, Primer (7701 Loctite), primer bottle with brush. Wound closure with primer. | Damage tip of wound closure/repo unit device leading to bleeding
Reposition Unit Assembly. Apply primer to the wound closure device. Wound closure, Primer (7701 Loctite), primer bottle with brush. Wound closure with primer. | too much solution leading to weak bond causing to leakage at repo unit
Reposition Unit Assembly. Apply primer to the wound closure device. Wound closure, Primer (7701 Loctite), primer bottle with brush. Wound closure with primer. | weak bond leading to leakage at repo unit
Reposition Unit Assembly. Clean butterfly and apply primer. Repositioning Unit, Petroleum ether, Primer (7701 Loctite), Oven (time and temperature), swabs, Primer bottle with brush. Cleaned and primered repositioning unit. | insufficient drying of unit leading to bond failure causing leakage at repo unit
Reposition Unit Assembly. Clean butterfly and apply primer. Repositioning Unit, Petroleum ether, Primer (7701 Loctite), Oven (time and temperature), swabs, Primer bottle with brush. Cleaned and primered repositioning unit. | too much solution leading to weak bond causing to leakage at repo unit
Reposition Unit Assembly. Clean butterfly and apply primer. Repositioning Unit, Petroleum ether, Primer (7701 Loctite), Oven (time and temperature), swabs, Primer bottle with brush. Cleaned and primered repositioning unit. | weak bond leading to leakage at repo unit
Reposition Unit Assembly. Prepare wound closure device removing all burrs. . AAC - Rework of Tapered wound closure, cutting fixture, length test fixture, scapel, 13Fr dialator . Prepared Tappered wound closure device. | Bleeding due to burrs
Reposition Unit Assembly. Prepare wound closure device removing all burrs. . AAC - Rework of Tapered wound closure, cutting fixture, length test fixture, scapel, 13Fr dialator . Prepared Tappered wound closure device. | Bleeding due to incorrect length
Reposition Unit Assembly. Prepare wound closure device removing all burrs. . AAC - Rework of Tapered wound closure, cutting fixture, length test fixture, scapel, 13Fr dialator . Prepared Tappered wound closure device. | Damage tip of wound closure device leading to bleeding
Reposition Unit Assembly. Then place in chamber for plasma activation of Tapered wound closure. Tapered wound closure, plasma chamber (program BG002164), oxygen. Plasma activated tappered wounder closure. | Damage tip of wound closure device leading to bleeding
Reposition Unit Assembly. Then place in chamber for plasma activation of Tapered wound closure. Tapered wound closure, plasma chamber (program BG002164), oxygen. Plasma activated tappered wounder closure. | weak bond leading to leakage at repo unit
Right Shell Mfg Process (MFG-3020). Assemble Right Shell. . ASSY, Shell Right. | Right shell not properly assembled
Rotor Assembly (0550-7024). Clean sleeve bearing and hub. Sleeve bearing, hub, swab, petroleum ether. Cleaned parts. | insufficient bond strength - pump stop
Rotor Assembly (0550-7024). Clean thrust washer, magnet, and Impeller shaft. thrust washer, Impeller, magnet, wipe, swab, petroleum ether. Cleaned thrust washer, magnet, and impeller shaft. | insufficient bond strength - pump stop
Rotor Assembly (0550-7024). cure adhesive. glued and cleaned assembly, oven (time and temp), timer, rotor assembly fixture. glued assembly - cured. | insufficient bond strength
Rotor Assembly (0550-7024). cure adhesive. glued and cleaned assembly, oven (time and temp), timer, sleeve bearing alignment fixture. glued assembly - cured. | insufficient bond strength - pump stop
Rotor Assembly (0550-7024). cure adhesive. uncured AU rotor, oven (time, temp), timer, backfill fixture. AU Rotor. | corrosion of magnet - pump stop
Rotor Assembly (0550-7024). Fill chamfer of the magnet with adhesive. assembly, adhesive, dosing needle. glued assembly. | expired adhesive used
Rotor Assembly (0550-7024). Fill chamfer of the magnet with adhesive. assembly, adhesive, dosing needle. glued assembly. | insufficient adhesive - corrosion of magnet - pump stop
Rotor Assembly (0550-7024). Fill chamfer of the magnet with adhesive. assembly, adhesive, dosing needle. glued assembly. | too much adhesive - pump stop
Rotor Assembly (0550-7024). Glue magnet to thrust washer and shaft. magnet, thrust washer, shaft, adhesive, dosing needle. glued assembly. | expired adhesive used
Rotor Assembly (0550-7024). Glue magnet to thrust washer and shaft. magnet, thrust washer, shaft, adhesive, dosing needle. glued assembly. | insufficient adhesive - insufficient bond strength
Rotor Assembly (0550-7024). Glue magnet to thrust washer and shaft. magnet, thrust washer, shaft, adhesive, dosing needle. glued assembly. | too much adhesive - pump stop
Rotor Assembly (0550-7024). Glue sleeve bearing to hub. Sleeve bearing, hub, adhesive, dosing needle, sleeve bearing alignment fixture. glued assembly. | bearing not flush with hub hemolysis, thrombus
Rotor Assembly (0550-7024). Glue sleeve bearing to hub. Sleeve bearing, hub, adhesive, dosing needle, sleeve bearing alignment fixture. glued assembly. | expired adhesive used
Rotor Assembly (0550-7024). Glue sleeve bearing to hub. Sleeve bearing, hub, adhesive, dosing needle, sleeve bearing alignment fixture. glued assembly. | insufficient adhesive - insufficient bond strength
Rotor Assembly (0550-7024). Glue sleeve bearing to hub. Sleeve bearing, hub, adhesive, dosing needle, sleeve bearing alignment fixture. glued assembly. | too much adhesive - unable to assemble
Rotor Assembly (0550-7024). Mount glued sleeve bearing and hub and the thrust washer on shaft. Glued sleeve bearing to hub, thrust washer, Impeller. assembly. | mounted incorrectly - shaft does not engage
Rotor Assembly (0550-7024). Remove excessive adhesive. AU rotor, scalpel. AU Rotor. | adhesive not cleaned - pump stop
Rotor Assembly (0550-7024). Remove excessive adhesive. AU rotor, scalpel. AU Rotor. | damage to assembly
Rotor Assembly (0550-7024). Remove excessive adhesive. glued assembly, swab, sleeve bearing alignment fixture. glued and cleaned assembly. | adhesive not cleaned - hemolysis or pump stop
Rotor Assembly (0550-7024). Remove excessive adhesive. glued assembly, swab. glued and cleaned assembly. | adhesive not cleaned - pump stop
Rotor Assembly (0550-7024). Remove excessive adhesive. uncured AU rotor, swab. uncured and cleaned AU rotor. | adhesive not cleaned - pump stop
Rotor Assembly (0550-7024). Store assembly in shipping tray. AU Rotor, shipping tray. AU Rotor in shipping tray. | corrosion
Rotor Assembly (0550-7024). Store assembly in shipping tray. AU Rotor, shipping tray. AU Rotor in shipping tray. | damage to impeller - hemolysis
RP IJ Cannula Coil. clean nitinol coil in iso-propanol in a ultrasonic bath; pivot it in benzene. Purple nitinol coil on mandrel, isopropyl alcohol, ultrasonic batch. Cleaned Coiled nitinol (only). | components not cleaned leading to increased reject rate
RP IJ Cannula Coil. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | Cannula is structurally weak
RP IJ Cannula Coil. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | parts not stored in oven leading to weak structure
RP IJ Cannula Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | cannula not dried leading to structurally weak unit
RP IJ Cannula Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect material utilized leading to structurally weak unit
RP IJ Cannula Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect mixture of cilcure leading to structurally weak unit
RP IJ Cannula Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | not coated properly leading to structurally weak unit
RP IJ Cannula Coil. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | incorrect water is selected leading to structurally weak unit
RP IJ Cannula Coil. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | not enough water is utilized leading to structurally weak unit
RP IJ Cannula Coil. Setup coiling machine, select appropriate program for coil size. Coiling winder, coil winding program. Setup equpiment. | increased reject rate
RP IJ Cannula Coil. start automatic mode until program end, fix clamp ring on overlapping winding. Threaded Nitinol on mandrel in coil winding machine. Coiled nitinol on mandrel . | coil not winded according to specification leading to weak structure
RP IJ Cannula Coil. thread Nitinol wire in coil bolt and fix it with clamp ring, fix coat bolt with tailstock . Nitinol wire, clamp ring, tailstock, mandrel, coil winder. Threaded Nitinol on mandrel in coil winding machine. | increased reject rate
RP TFC Cannula Bend Forming. Clean cannula in ultrasonic bath with 80% alcohol and 20% non-pyrogenic water solution for 2 minutes minimum Blow cannula off with compressed air . Ultrasonic bath (80% alcohol/20% non-pyrogenic water), Time, compressed air, annealed shapped cannula. cleaned component. | contimnation from process leading sterility issue
RP TFC Cannula Bend Forming. Evaporate residual moisture by placing cannulae in tray into oven set to 50°C for 30 minutes minimum. Tray, shapped cannula, oven (time and temperature), gloves. Completed Cannula assembly. | weak cannula tensile strength
RP TFC Cannula Bend Forming. Lay cannula flat into bin and place into oven set to 70ºC on middle of rack Anneal for 8 hours minimum Remove from oven and allow to cool to room temperature. Bin, shapped cannula, oven (time and temperature). annealed shapped cannula. | Cannula not with low rigidity leading to delivery issues
RP TFC Cannula Bend Forming. Lay cannula flat into bin and place into oven set to 70ºC on middle of rack Anneal for 8 hours minimum Remove from oven and allow to cool to room temperature. Bin, shapped cannula, oven (time and temperature). annealed shapped cannula. | Improper shape leading to placement issues
RP TFC Cannula Bend Forming. Place into oven set to 160°C for 10 ± ½ minutes. Cannula on Bend Mandrel in fixture, Oven (time and temperature). Formed cannula on bend mandrel in fixture. | Improper shape leading to placement issues
RP TFC Cannula Bend Forming. Preheat straight cannula and bend mandrel in oven set to 160°C for 1 minute. Oven (time and temperature), cannula on mandrel, bend mandrel, RP TFC Bend Mandrel Holding Fixture. preheated components. | unable to hold shape leading to placement issues
RP TFC Cannula Bend Forming. Remove bushings on both sides of mandrel Pull cannula off mandrel in direction of PH side while rotating slightly . Formed cannula on bend mandrel in fixture. Shapped cannula off bend mandrel. | distorted shape, leading to placement issues
RP TFC Cannula Bend Forming. Remove cannula on bend mandrel from 160°C oven Immediately place cannula into oven set to 70°C for 10 ± ½ minutes . Cannula on Bend Mandrel in fixture, Oven (time and temperature). Formed cannula on bend mandrel in fixture. | delamination of cannula
RP TFC Cannula Bend Forming. Remove pump housing bushing and carefully slide cannula on bending mandrel starting at outflow end Push cannula down mandrel until outflow bushing is fully seated on mandrel Insert bending mandrel with mounted cannula into perforated aluminum plate. Cannula on mandrel assembly, bend mandrel, RP TFC Bend Mandrel Holding Fixture. Cannula on Bend Mandrel in fixture. | damaged cannula surface leading
RP TFC Cannula Bend Forming. Remove pump housing bushing and carefully slide cannula on bending mandrel starting at outflow end Push cannula down mandrel until outflow bushing is fully seated on mandrel Insert bending mandrel with mounted cannula into perforated aluminum plate. Cannula on mandrel assembly, bend mandrel, RP TFC Bend Mandrel Holding Fixture. Cannula on Bend Mandrel in fixture. | distorted shape due to missing bushings, leading to placement issues
RP TFC Cannula Bend Forming. Remove pump housing bushing and carefully slide cannula on bending mandrel starting at outflow end Push cannula down mandrel until outflow bushing is fully seated on mandrel Insert bending mandrel with mounted cannula into perforated aluminum plate. Cannula on mandrel assembly, bend mandrel, RP TFC Bend Mandrel Holding Fixture. Cannula on Bend Mandrel in fixture. | Improper shape leading to placement issues
RP TFC Cannula Bend Forming. Slide outflow bushing into outflow end of cannula and pump housing bushing into pump housing end of cannula. Trimmed cannula on mandrel, bushings. Cannula on mandrel assembly. | Missing busshings, Unable to assemble device (scrap)
RP TFC Cannula NiTi Coil. clean nitinol coil in iso-propanol in a ultrasonic bath; pivot it in benzene. Purple nitinol coil on mandrel, isopropyl alcohol, ultrasonic batch. Cleaned Coiled nitinol (only). | components not cleaned leading to increased reject rate
RP TFC Cannula NiTi Coil. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | Cannula is structurally weak
RP TFC Cannula NiTi Coil. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | parts not stored in oven leading to weak structure
RP TFC Cannula NiTi Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | cannula not dried leading to structurally weak unit
RP TFC Cannula NiTi Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect material utilized leading to structurally weak unit
RP TFC Cannula NiTi Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect mixture of cilcure leading to structurally weak unit
RP TFC Cannula NiTi Coil. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | not coated properly leading to structurally weak unit
RP TFC Cannula NiTi Coil. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | incorrect water is selected leading to structurally weak unit
RP TFC Cannula NiTi Coil. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | not enough water is utilized leading to structurally weak unit
RP TFC Cannula NiTi Coil. Setup coiling machine, select appropriate program for coil size. Coiling winder, coil winding program. Setup equpiment. | increased reject rate
RP TFC Cannula NiTi Coil. start automatic mode until program end, fix clamp ring on overlapping winding. Threaded Nitinol on mandrel in coil winding machine. Coiled nitinol on mandrel . | coil not winded according to specification leading to weak structure
RP TFC Cannula NiTi Coil. thread Nitinol wire in coil bolt and fix it with clamp ring, fix coat bolt with tailstock . Nitinol wire, clamp ring, tailstock, mandrel, coil winder. Threaded Nitinol on mandrel in coil winding machine. | increased reject rate
RP TFC Cannula Straight. apply release agent Acmos 70-2406 and cure. Cannula mandrel with grease, ACMOS, oven (time and temperature). Cured mandrel with grease and ACMOS. | Increased scrap rate
RP TFC Cannula Straight. cool the cannulas; clamp cannulas in 4mm chuck and turn it with a cutting tool . chuck, cutting tool, heat shrink on coiled nitinol on mandrel. Cooled heat shrinked on coiled nitinol on mandrel. | Cannula is structurally weak
RP TFC Cannula Straight. cool the cannulas; clamp cannulas in 4mm chuck and turn it with a cutting tool . chuck, cutting tool, heat shrink on coiled nitinol on mandrel. Cooled heat shrinked on coiled nitinol on mandrel. | Cannula surfaces rough leading to thrombus formation/hemolysis
RP TFC Cannula Straight. cut off coil and remove tailstock; pull end of coil out of the bolt boring, cut coil . Coiled nitinol on mandrel, tweezers, clamps, tailstock . Coiled nitinol on mandrel . | Cannula is structurally weak
RP TFC Cannula Straight. dip and turn cannula mandrel in the exhaust system in THF, shake off THF and cure assembly in oven. THF, Cannula on mandrel, Oven (time and temperature), exhaust system. cured cannula on mandrel assembly. | Cannula is structurally weak
RP TFC Cannula Straight. dip and turn cannula mandrel in the exhaust system in THF, shake off THF and cure assembly in oven. THF, Cannula on mandrel, Oven (time and temperature), exhaust system. cured cannula on mandrel assembly. | Surface imperfections
RP TFC Cannula Straight. Insert urethane tubing on top of coil assembly, then insert heat shrink and cure in oven. heat shrink, oven (time and temperature), coiled nitinol on mandrel. Heat shrink on coiled nitinol on mandrel. | Cannula is structurally weak
RP TFC Cannula Straight. remove cannula bolt and clamp in infrared oven, select program Ref PU. Cured mandrel with grease and ACMOS and loctite, IR Oven (Time and Temperature). Cured mandrel with grease and ACMOS and loctite. | Cannula is structurally weak
RP TFC Cannula Straight. remove cannula out of oven and cool down, remove shrink tube from cannula with scalpel. scapel, heat shrink on coiled mandrel. Cannula on mandrel. | Damage to cannula surface
RP TFC Cannula Straight. remove cannula out of oven and cool down, remove shrink tube from cannula with scalpel. scapel, heat shrink on coiled mandrel. Cannula on mandrel. | Increased scrap rate
RP TFC Cannula Straight. select appropriate program per WI, flush & prime dispensing valve, and start automatic dosing program. Loctite Dispenser, syringe tip, Cured mandrel with grease and ACMOS. Mandrel with grease and ACMOS and loctite. | Cannula is structurally weak
RP TFC Cannula Straight. select appropriate program per WI, flush & prime dispensing valve, and start automatic dosing program. Loctite Dispenser, syringe tip, Cured mandrel with grease and ACMOS. Mandrel with grease and ACMOS and loctite. | Increased scrap rate
RP TFC Cannula Straight. separate cannula mandrel, lube with silicone grease, . cannula mandrel and silicone grease. Cannula mandrel with grease. | Cannula is structurally weak
RP TFC Cannula Straight. start coil winding machine, select program “Coil On”, clamp cannula mandrel, start automatic winding mode; dose Loctite 4061 on the shoulder. coil winder, Cured mandrel with grease and ACMOS and loctite, with coiled nitinol. Coiled nitinol on mandrel . | Cannula is structurally weak
RP TFC Cannula Straight. start coil winding machine, select program “Coil On”, clamp cannula mandrel, start automatic winding mode; dose Loctite 4061 on the shoulder. coil winder, Cured mandrel with grease and ACMOS and loctite. completed winding cycle. | Cannula is structurally weak
RP TFC Cannula Straight. store cannula mandrel assembly in an oven. Oven (time and temperature), coiled nitinol on mandrel. Cured coiled nitinol on mandrel . | Cannula is structurally weak
RP TFC Cannula Straight. trim cannula. cannula, cannula trim fixture. trimmed cannula . | Insertion/placement issues
RP TFC Cannula Straight. trim cannula. cannula, cannula trim fixture. trimmed cannula . | unable to assesmble
SBC SMARC Module. Clean flux residue. . . | Flux residue not cleaned
SBC SMARC Module. Ensure no shorts between pads. . . | Shorts between pads
SBC SMARC Module. ESD Control. . . | SMARC Module board damage
SBC SMARC Module. Install coin cell battery. . . | Coin cell battery not installed
SBC SMARC Module. Remove coin cell battery. . . | Coin cell battery not removed
SBC SMARC Module. Remove excess solder. . . | Excess solder not removed
SBC SMARC Module. Remove Q502. . . | Unexpected system restart alarm
SBC SMARC Module. Remove SMARC Module. . . | SMARC Module board damage
SBC SMARC Module. Remove SMARC Module. . . | Thermal pad damage
SBC SMARC Module. Remove U904 and replace with 300464. . . | Coin cell battery drains quickly
Shipping. Add Customs Form. File for Customs for product value > $2500. Attached appropriate customs form. | Missing Customs information
Shipping. Add delivery note. Box, pump set, delivery note. Undamaged pump set and delivery note. | Damaged delivery note used
Shipping. Close box, apply green label. Box, green label. Packed pumpset. | Label Applied to wrong box
Shipping. Delivery note print. Customer data and order data. Legible delivery note. | Incorrect data input (part, cost, material, batch, serial number)
Shipping. Delivery note print. Customer data and order data. Legible delivery note. | Printer malfunction
Shipping. Fedex station . Delivery note, customer data. Address verified. | Incorrect address verified
Shipping. Fedex station . Delivery note, customer data. Auto-population of data. | Incorrect data
Shipping. Fedex station . Delivery note, customer data. Scanned delivery note. | Wrong delivery note scanned
Shipping. Fedex station . Tracking tag, delivery note. Tracking tag attached to delivery note. | Wrong tag applied
Shipping. Fedex station . Tracking tag. Tag scanned. | Wrong tag scanned
Shipping. Place in box. Box, pump set. Undamaged pump set in box. | Unapproved box used
Shipping. Place in box. Box, pump set. Undamaged pump set packed in box. | Box with damage selected
Shipping. Pull material from Finished Goods Location. Delivery note data. Correct item/quantity pulled. | Wrong item or quantity pulled
Shipping. Scan delivery Note . Operator, Barcode scanner, Delivery Note. Proper pump set confirmed with customer order. | Incorrect note scanned
Shipping. Scan delivery Note . Operator, Barcode scanner, Delivery Note. Proper pump set confirmed with customer order. | Wrong line scanned
Shipping. Second delivery note. Delivery note, box. Packing slip in envelope on box. | Applied to wrong box
Side Panel Mfg Process (MFG-3022). Attach Left Shell and Right Shell to Pneumatic Circuit Installation Assy. . ASSY, Side Panels. | Left shell and right shell not properly fastened to pneumatic circuit assembly
Spine Assy Mfg Process (MFG-3007). Assemble Spine. . ASSY, Spine. | DC cable and ferrite not properly routed or captured
Spine Assy Mfg Process (MFG-3007). Assemble Spine. . ASSY, Spine. | DC cable not properly strain-relieved
Spine Assy Mfg Process (MFG-3007). Assemble Spine. . ASSY, Spine. | IEC inlet cable not properly connected
Spine Assy Mfg Process (MFG-3007). Assemble Spine. . ASSY, Spine. | Rear shell not properly fastened to Rear bracket or Bottom shell
Spine Assy Mfg Process (MFG-3007). Attach Handle bracket to rear support bracket. . ASSY, Spine. | Hand Bracket is not fastened to rear support bracket
Spine Assy Mfg Process (MFG-3007). Attach rear support bracket to bottom bracket. . ASSY, Spine. | Rear support bracket not properly fastened to Bottom shell
Spine Assy Mfg Process (MFG-3007). Breethe Connect Cable. . ASSY, Spine. | Breethe connect cable not installed
Spine Assy Mfg Process (MFG-3007). Install Spring Plunger. . ASSY, Spine. | Spring plunger over tightened and broken
Spine Assy Mfg Process (MFG-3007). Install Spring Plunger. . ASSY, Spine. | Spring plunger under tightened
Spine Assy Mfg Process (MFG-3007). Secure handle pivot and hinge . . ASSY, Spine. | Pivot and hinge not secure
SSD Mounting Mfg Process (MFG-3008). Assemble SSD Mount. . ASSY, Subassemblies. | Incorrect or improper software installation
Stator Soldering (0550-7028). Apply primer on sleeve. Assembly with glued purge lumen, primer, oven (time and temp). Assembly with primer on sleeve. | insufficient coating
Stator Soldering (0550-7028). Applying primer to yoke. Yoke, primer, cleanroom swab, oven (time temp). Yoke coated with primer. | insufficient coating
Stator Soldering (0550-7028). Clean purge lumen with petroleum ether. Purge lumen, petroleum ether, lint free wipe. cleaned purge lumen. | Insufficient bond strength - epotek flash - unable to demold
Stator Soldering (0550-7028). Clean soldered area with petroleum ether. assembly with soldered winding wires on PCB, petroleum ether, swabs. assembly with soldered winding wires on PCB - cleaned soldered area. | Excess flux - insufficient gluing of strain relief
Stator Soldering (0550-7028). Connect multimeter with motor cable, measure resistance and inductance. AU stator, multimeter. Electrically inspected assembly. | Failure in downstream test
Stator Soldering (0550-7028). Cover PCB with adhesive. Stator assemble, adhesive, dispensing tip. Stator assembly with PCB covered. | solder joint fails
Stator Soldering (0550-7028). Curing of adhesive. Stator assembly with PCB covered, oven (time and temp). Stator assembly with PCB covered - cured. | solder joint fails
Stator Soldering (0550-7028). Curing of glue. AU Stator - uncured, oven (time and temp). AU Stator. | epotek flash - unable to assemble
Stator Soldering (0550-7028). Fill gap between purge lumen and bearing with adhesive, harden adhesive. Assembly with glued purge lumen. Assembly with glued purge lumen with gap filled. | epotek flash - unable to demold
Stator Soldering (0550-7028). Fill opening of yoke with adhesive. Assembly with insertion pin and coil locked on backing plate and locked with shim with yoke in place, adhesive, dispenser. AU Stator - uncured. | epotek flash - unable to assemble
Stator Soldering (0550-7028). Insert sheath on to yoke. Yoke, sheath. Yoke sheath assembly. | unable to assemble
Stator Soldering (0550-7028). Mount backing ring and o-ring over coil onto yoke. Assembly with glued purge lumen with gap filled, backing ring, o-ring. Assembly with backing ring and o-ring on yoke. | damage to coil
Stator Soldering (0550-7028). Place assembly onto soldering fixture. Insertion pin with AU coil sleeve and o-ring, soldering fixture. Insertion pin with AU coil sleeve and o-ring on soldering fixture. | placed upside down - unable to assemble
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and O-ring on soldering fixture, PCB, fixture, solder, soldering iron, cutter. assembly with soldered winding wires on PCB. | temp too high - damaged PCB - motor leakage current too high
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and o-ring on soldering fixture, PCB, fixture, soldering iron, cutter. assembly with soldered winding wires on PCB. | damaged winding
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and o-ring on soldering fixture, PCB, fixture, soldering iron, cutter. assembly with soldered winding wires on PCB. | short circuit
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and o-ring on soldering fixture, PCB, fixture, soldering iron, cutter. assembly with soldered winding wires on PCB. | Solder Temperature out of specification
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and o-ring on soldering fixture, PCB, fixture, soldering iron, cutter. assembly with soldered winding wires on PCB. | temp too low - insufficient soldering - low resistance
Stator Soldering (0550-7028). position PCB on bearing, solder winding wires to PCB. Insertion pin with AU coil sleeve and o-ring on soldering fixture, PCB, fixture,solder, soldering iron, cutter. assembly with soldered winding wires on PCB. | temp too high - damaged PCB - motor leakage current too high
Stator Soldering (0550-7028). remove foil from PCB. PCB, tweezer. PCB with foil removed. | PCB cannot be soldered
Stator Soldering (0550-7028). Reposition motor cable leads away from purge lumen, apply adhesive to lumen, advance lumen into bearing, harden adhesive. Assembly with purge lumen inserted to yoke, adhesive, dispensing tip, UV lamp. Assembly with glued purge lumen. | purge lumen interferes with shaft
Stator Soldering (0550-7028). Reposition motor cable leads away from purge lumen, apply adhesive to lumen, advance lumen into bearing, harden adhesive. Assembly with purge lumen inserted to yoke, adhesive, dispensing tip, UV lamp. Assembly with glued purge lumen. | purge lumen interferes with shaft - unable to assemble
Stator Soldering (0550-7028). Slide assembly with insertion pin onto black backing plate, lock with shim. Assembly with backing ring and o-ring on yoke, backing plate, shim. Assembly with insertion pin and coil locked on backing plate and locked with shim. | incorrect position of ceramic sleeve with respect to yoke - unable to assemble to hub
Stator Soldering (0550-7028). Slide AU coil sleeve onto insertion pin. Au coil sleeve, insertion pin with o-ring. Insertion pin with AU coil sleeve and o-ring. | ceramics damaged - epotek flash - high motor current
Stator Soldering (0550-7028). Slide AU coil sleeve onto insertion pin. Au coil sleeve, insertion pin with o-ring. Insertion pin with AU coil sleeve and o-ring. | pump stop
Stator Soldering (0550-7028). Slide purge lumen on proximal end of insertion pin. Purge lumen, AU stator. Assembly with purge lumen inserted to yoke. | purge lumen interferes with shaft - unable to assemble
Stator Soldering (0550-7028). Slide silicone ring onto insertion pin. insertion pin, silicone O-ring. insertion pin with O-ring. | Epotek leaks to inside of motor housing - cannot demold
Stator Soldering (0550-7028). Slide silicone ring onto insertion pin. insertion pin, silicone O-ring. insertion pin with O-ring. | O-ring on wrong side - unable to assemble
Stator Soldering (0550-7028). Slide silicone ring onto insertion pin. insertion pin, solicone o-ring. insertion pin with o-ring. | Epotek leaks to inside of motor housing - cannot demold
Stator Soldering (0550-7028). Slide silicone ring onto insertion pin. insertion pin, solicone o-ring. insertion pin with o-ring. | o-ring on wrong side - unable to assemble
Stator Soldering (0550-7028). Slide yoke onto coil and connect multimeter with motor cable, measure resistance and inductance. Electrically inspected assembly. Electrically inspected assembly with yoke. | Failure in downstream test
Stator Soldering (0550-7028). Slide yoke, backing ring and o-ring assembly over the coil into the black backing plate. Assembly with insertion pin and coil locked on backing plate and locked with shim. Assembly with insertion pin and coil locked on backing plate and locked with shim with yoke in place. | incorrect position of ceramic sleeve with respect to yoke - unable to assemble to hub
Stator Soldering (0550-7028). Solder leads of motor cable to PCB. . motor cable leads tinned inserted into yoke, assembly with soldered, trimmed and cleaned leads, solder, soldering iron. Stator assembly. | fails electrical check
Stator Soldering (0550-7028). Solder leads of motor cable to PCB. . motor cable leads tinned inserted into yoke, assembly with soldered, trimmmed and cleaned leads, solder, soldering iron. Stator assembly. | fails electrical check
Stator Soldering (0550-7028). Trim excess wire . assembly with soldered winding wires on PCB - cleaned soldered area. assembly with soldered, trimmed and cleaned leads. | too long - interference with assembly into yoke
Stator Soldering (0550-7028). Trim excess wire . assembly with soldered winding wires on PCB - cleaned soldered area. assembly with soldered, trimmed and cleaned leads. | too short - no electrical contact
Stator Soldering (0550-7028). Trim excess wire . assembly with soldered winding wires on PCB - cleaned soldered area. assembly with soldered, trimmmed and cleaned leads. | too long - interference with assembly into yoke
Stator Soldering (0550-7028). Trim excess wire . assembly with soldered winding wires on PCB - cleaned soldered area. assembly with soldered, trimmmed and cleaned leads. | too short - no electrical contact
Stator Soldering. Clean insertion pin and let dry. Petroleum ether, insertion pin, swabs. Cleaned insertion pin. | Pin not cleaned/debry still present leading to damage of the coil
Stator Soldering. Deburr the soldering PCB, slide on the PCB on the engage pin . PCB, side cutter. PCB on engage pin. | damage to PCB during deburring leading to pump stop
Stator Soldering. Deburr the soldering PCB, slide on the PCB on the insertion pin . PCB, side cutter, PCB holder, PCB press. PCB on insertion pin. | PCB with burrs, leading to increase reject rate
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | damage causing short circuit leading to pump stop
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | Damage insulation leading to leakage current
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | failure to solder all ends of motor cable leading to pump failure to start
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | Flux debris not removed leading to increased scrap
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | leakage current
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | solder ends in wrong direction leading to impeller rotating in wrong direction, no flow
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | unable to solder correctly due to low iron temperature leading to pump stop
Stator Soldering. Position stator in rotating device and sold the motor cables. Motor cables, soldered assembly. Assembled Stator unit. | unable to solder correctly leading to pump stop
Stator Soldering. Solder the coil leads to PCB solder pads and then clean. PCB on insertion pin, coil winding, petroleum ether, swabs, scapel. Soldered assembly unit. | not all joints soldered leading to pump stop/start issues
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | damage insulation leading to leakage current
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | Damage to PCB causing short occurs leading to pump stop
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | Flux debris not removed leading to increased scrap
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | leakage current
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | not all joints soldered leading to pump stop/start issues
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | solder ends in wrong direction leading to impeller rotating in wrong direction, no flow
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | unable to solder correctly due to low iron temperature leading to pump stop
Stator Soldering. Solder the ends of the winding with the copper pads and remove the flux leavings . Engagement pin, PCB, coil winding, petroleum ether. Soldered assembly unit. | unable to solder correctly leading to pump stop
Strain Relief PSS. Cut protection catheter for purge lumen, then clean with petroleum ether. protection catheter, petroleum ether, wipe, catheter cutting fixture, scapel, cutting tray, pressurized air. cleaned and cut protection catheter. | bond failure leading to leak
Strain Relief PSS. Cut protection catheter for purge lumen, then clean with petroleum ether. protection catheter, petroleum ether, wipe, catheter cutting fixture, scapel, cutting tray, pressurized air. cleaned and cut protection catheter. | cut length incorrect leading to scrap
Strain Relief PSS. Cut protection catheter for purge lumen, then clean with petroleum ether. protection catheter, petroleum ether, wipe, catheter cutting fixture, scapel, cutting tray, pressurized air. cleaned and cut protection catheter. | misalignment leading to unable to assemble
Strain Relief PSS. Remove from curing fixture. Apply Loctite 3321 around the kink protection and in the second rectangular opening. Place back into fixture and cure under direct UV lamp. bonded together kink protector and protection catheter, loctite 3321, curing fixture, UV flood lamp. Bonded together kink protector and protection catheter. | Bond Failure leading strain on inner lumen (leak)
Strain Relief PSS. Remove from curing fixture. Apply Loctite 3321 around the kink protection and in the second rectangular opening. Place back into fixture and cure under direct UV lamp. bonded together kink protector and protection catheter, loctite 3321, curing fixture, UV flood lamp. Bonded together kink protector and protection catheter. | Incorrect amount of adhesive leading strain on inner lumen (leak)
Strain Relief PSS. Remove from curing fixture. Apply Loctite 3321 to front face of kink protector and place into front face curing fixture and cure under UV lamp. bond together kink protector and protection catheter, loctite 3321, UV flood lamp, front face curing fixture. Bonded together kink protector and protection catheter with formed fit adhesive. | Bond Failure leading strain on inner lumen (leak)
Strain Relief PSS. Remove from curing fixture. Apply Loctite 3321 to front face of kink protector and place into front face curing fixture and cure under UV lamp. bond together kink protector and protection catheter, loctite 3321, UV flood lamp, front face curing fixture. Bonded together kink protector and protection catheter with formed fit adhesive. | Incorrect amount of adhesive leading strain on inner lumen (leak)
Strain Relief PSS. slide in protection catheter into kink protector and fill window with loctite, then position in fixture and cure. cut protection catheter, loctite 3321, curing fixture, UV flood lamp, kink protector. Bonded together kink protector and protection catheter. | Bond Failure leading strain on inner lumen (leak)
Strain Relief PSS. slide in protection catheter into kink protector and fill window with loctite, then position in fixture and cure. cut protection catheter, loctite 3321, curing fixture, UV flood lamp, kink protector. Bonded together kink protector and protection catheter. | Incorrect amount of adhesive leading strain on inner lumen (leak)
Subassemblies Mfg Process (MFG-3011). Install exhaust duct bumper. . ASSY, Power Supply Mounting. | Console noise
Subassemblies Mfg Process (MFG-3011). Join subassemblies. . ASSY, Subassemblies. | Corrupted hard drive
Subassemblies Mfg Process (MFG-3011). Join subassemblies. . ASSY, Subassemblies. | Oxygen flow blocked/leaking
Subassemblies Mfg Process (MFG-3011). Join subassemblies. . ASSY, Subassemblies. | Pinched concentrator cables
Subassemblies Mfg Process (MFG-3011). Join subassemblies. . ASSY, Subassemblies. | Unconstrained battery bracket assembly
Thermoformed Cannula. apply release agent Acmos 70-2406 and cure. Cannula mandrel with grease, ACMOS, oven (time and temperature). Cured mandrel with grease and ACMOS. | Increased scrap rate
Thermoformed Cannula. clean nitinol coil in iso-propanol in a ultrasonic bath; pivot it in benzene. Purple nitinol coil on mandrel, isopropyl alcohol, ultrasonic batch. Cleaned Coiled nitinol (only). | components not cleaned leading to increased reject rate
Thermoformed Cannula. cool the cannulas; clamp cannulas in 4mm chuck and turn it with a cutting tool . chuck, cutting tool, heat shrink on coiled nitinol on mandrel. Cooled heat shrinked on coiled nitinol on mandrel. | Cannula is structurally weak
Thermoformed Cannula. cool the cannulas; clamp cannulas in 4mm chuck and turn it with a cutting tool . chuck, cutting tool, heat shrink on coiled nitinol on mandrel. Cooled heat shrinked on coiled nitinol on mandrel. | Cannula surfaces rough leading to thrombus formation/hemolysis
Thermoformed Cannula. cut off coil and remove tailstock; pull end of coil out of the bolt boring, cut coil . Coiled nitinol on mandrel, tweezers, clamps, tailstock . Coiled nitinol on mandrel . | Cannula is structurally weak
Thermoformed Cannula. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | Cannula is structurally weak
Thermoformed Cannula. cut wire, withdraw coil bolt and temper it. Snips, Coiled nitinol wire on mandrel, oven (time and temperature), heat gloves, pliers. Tempered Coiled nitinol on mandrel. | parts not stored in oven leading to weak structure
Thermoformed Cannula. dip and turn cannula mandrel in the exhaust system in THF, shake off THF and cure assembly in oven. THF, Cannula on mandrel, Oven (time and temperature), exhaust system. cured cannula on mandrel assembly. | Cannula is structurally weak
Thermoformed Cannula. dip and turn cannula mandrel in the exhaust system in THF, shake off THF and cure assembly in oven. THF, Cannula on mandrel, Oven (time and temperature), exhaust system. cured cannula on mandrel assembly. | Surface imperfections
Thermoformed Cannula. fill 800 grams of DMF into fixture. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Weighed out Desmopan and DMF. | contaminated cannula leading to it being weak/brittle
Thermoformed Cannula. fill 800 grams of DMF into fixture. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Weighed out Desmopan and DMF. | degradation of material during leading to particles in blood stream
Thermoformed Cannula. fill 800 grams of DMF into fixture. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Weighed out Desmopan and DMF. | increased reject rate
Thermoformed Cannula. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | cannula not dried leading to structurally weak unit
Thermoformed Cannula. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect material utilized leading to structurally weak unit
Thermoformed Cannula. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | incorrect mixture of cilcure leading to structurally weak unit
Thermoformed Cannula. immerse nitinol coil in a bath with primer cilbond 49SF / Cilcure B and dry coil. Cibond mixture, beaker, Cleaned Coiled nitinol (only), oven (time and temperature), gloves . Coiled nitinol with coating. | not coated properly leading to structurally weak unit
Thermoformed Cannula. Insert urethane tubing on top of coil assembly, then insert heat shrink and cure in oven. heat shrink, oven (time and temperature), coiled nitinol on mandrel. Heat shrink on coiled nitinol on mandrel. | Cannula is structurally weak
Thermoformed Cannula. load the marker band onto the cannula with the cannula marker inspection fixture. cannula on mandrel assembly with marker band, cannula marker inspection fixture. inspected cannula with marker band. | difficulty placing pump in optimal position
Thermoformed Cannula. mix solution a minimum of six hours and let sit overnight. The following day mix an additional six hours. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Mixed together Desmopan and DMF. | solution not mixed consistently leading to weak/brittle cannula
Thermoformed Cannula. quench bending mandrel with cannula in pyrogenic free water; remove bending mandrel and cure cannulas in oven, clean cannula with isopropyl- panol in a ultrasonic bath and dry it with compressed air . cannula on bend mandrel, utltrasonic bath, isopropyl alcohol, pyrogen free water, overn (time and temperature). cured and cleaned cannula on bend mandrel. | Cannula is structurally weak
Thermoformed Cannula. refill solution in four 250ml containers. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Desmopan and DMF in 250ml containers. | contaminated cannula leading to it being weak/brittle
Thermoformed Cannula. refill solution in four 250ml containers. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Desmopan and DMF in 250ml containers. | degradation of material during leading to particles in blood stream
Thermoformed Cannula. refill solution in four 250ml containers. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Desmopan and DMF in 250ml containers. | increased reject rate
Thermoformed Cannula. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | incorrect water is selected leading to structurally weak unit
Thermoformed Cannula. remove bolt and quench in pyrogenic free water; remove clamp ring and peel away coil; cut over wounded end. pyrogen free water, tempered coiled nitinol on mandrel, heat gloves, pliers. Purple nitinol coil on mandrel. | not enough water is utilized leading to structurally weak unit
Thermoformed Cannula. remove cannula bolt and clamp in infrared oven, select program Ref PU. Cured mandrel with grease and ACMOS and loctite, IR Oven (Time and Temperature). Cured mandrel with grease and ACMOS and loctite. | Cannula is structurally weak
Thermoformed Cannula. remove cannula out of oven and cool down, remove shrink tube from cannula with scalpel. scapel, heat shrink on coiled mandrel. Cannula on mandrel. | Damage to cannula surface
Thermoformed Cannula. remove cannula out of oven and cool down, remove shrink tube from cannula with scalpel. scapel, heat shrink on coiled mandrel. Cannula on mandrel. | Increased scrap rate
Thermoformed Cannula. select appropriate program per WI, flush & prime dispensing valve, and start automatic dosing program. Loctite Dispenser, syringe tip, Cured mandrel with grease and ACMOS. Mandrel with grease and ACMOS and loctite. | Cannula is structurally weak
Thermoformed Cannula. select appropriate program per WI, flush & prime dispensing valve, and start automatic dosing program. Loctite Dispenser, syringe tip, Cured mandrel with grease and ACMOS. Mandrel with grease and ACMOS and loctite. | Increased scrap rate
Thermoformed Cannula. separate cannula mandrel, lube with silicone grease, . cannula mandrel and silicone grease. Cannula mandrel with grease. | Cannula is structurally weak
Thermoformed Cannula. Setup coiling machine, select appropriate program for coil size. Coiling winder, coil winding program. Setup equpiment. | increased reject rate
Thermoformed Cannula. slide cannula up to the marking on the bending mandrel, slide sleeves into cannula ends and temper it in an oven . bending mandrel, slide sleeves, oven (time and temperature). Cured cannula on bend mandrel. | Cannula is structurally weak
Thermoformed Cannula. start automatic mode until program end, fix clamp ring on overlapping winding. Threaded Nitinol on mandrel in coil winding machine. Coiled nitinol on mandrel . | coil not winded according to specification leading to weak structure
Thermoformed Cannula. start coil winding machine, select program “Coil On”, clamp cannula mandrel, start automatic winding mode; dose Loctite 4061 on the shoulder. coil winder, Cured mandrel with grease and ACMOS and loctite, with coiled nitinol. Coiled nitinol on mandrel . | Cannula is structurally weak
Thermoformed Cannula. start coil winding machine, select program “Coil On”, clamp cannula mandrel, start automatic winding mode; dose Loctite 4061 on the shoulder. coil winder, Cured mandrel with grease and ACMOS and loctite. completed winding cycle. | Cannula is structurally weak
Thermoformed Cannula. Start mixer, add Desmopan pellets at a rate of 50 grams every 15 minutes to DMF. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Mixed together Desmopan and DMF. | solution not mixed consistently leading to weak/brittle cannula
Thermoformed Cannula. store cannula mandrel assembly in an oven. Oven (time and temperature), coiled nitinol on mandrel. Cured coiled nitinol on mandrel . | Cannula is structurally weak
Thermoformed Cannula. Submerge cannula marker bands and place in ultrasonic bath, let the piece dry. cannula on mandrel assembly, ultrasonic bath, marker band, lint free wipes. prepped marker band. | degredation of material leading to particles
Thermoformed Cannula. thread Nitinol wire in coil bolt and fix it with clamp ring, fix coat bolt with tailstock . Nitinol wire, clamp ring, tailstock, mandrel, coil winder. Threaded Nitinol on mandrel in coil winding machine. | increased reject rate
Thermoformed Cannula. trim cannula. cannula with marker band, cannula trim fixture. trimmed cannula with marker band. | Insertion/placement issues
Thermoformed Cannula. trim cannula. cannula with marker band, cannula trim fixture. trimmed cannula with marker band. | unable to assesmble
Thermoformed Cannula. Weigh 200grams of Desmopan and 800 grams of DMF. Desmopan, DMF, glass bowl, mixing fixture, gloves, safety glasses, mask, exhaust hood, scale, graduated cylinder. Weighed out Desmopan and DMF. | Incorrect ration of Desmopan and DMF leading to weak/brittle cannula
Tubing THF Bonding. Adjust pressure regulator to acceptable pressure. Blow-Out Air Fixture, Pressure Regulator. Setup fixture. | damage of components leading to leak
Tubing THF Bonding. Adjust pressure regulator to acceptable pressure. Blow-Out Air Fixture, Pressure Regulator. Setup fixture. | residual agent leading to occlusion (unable to move pass case start)
Tubing THF Bonding. Adjust pressure regulator to acceptable pressure. Blow-Out Air Fixture, Pressure Regulator. Setup fixture. | residual agent leading to occlusion (unable to move past case start)
Tubing THF Bonding. Bond open Pressure Line end to Outflow Male Luer with THF. . Bonded assembly, Outflow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Tubing THF Bonding. Bond open Pressure Line end to Outflow Male Luer with THF. . Bonded assembly, Outflow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Bond open Pressure Line end to Outflow Male Luer with THF. . Bonded assembly, Outflow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Bond open Pressure Line end to Outflow Male Luer with THF. . Bonded assembly, Outflow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Bond open Pressure Line end to Outflow Male Luer with THF. . Bonded assembly, Outflow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Bond open Pressure Line end to Pressure Transmitter with THF.. Pressure transmitter, THF, Bonded assembly. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Tubing THF Bonding. Bond open Pressure Line end to Pressure Transmitter with THF.. Pressure transmitter, THF, Bonded assembly. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Bond open Pressure Line end to Pressure Transmitter with THF.. Pressure transmitter, THF, Bonded assembly. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Bond open Pressure Line end to Pressure Transmitter with THF.. Pressure transmitter, THF, Bonded assembly. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Bond open Pressure Line end to Pressure Transmitter with THF.. Pressure transmitter, THF, Bonded assembly. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Bond Pressure Line to Female Luer with THF on Extension Line. Bonded assembly, female luer, THF, THF dispenser, metric ruler. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Tubing THF Bonding. Bond Pressure Line to Female Luer with THF on Extension Line. Bonded assembly, female luer, THF, THF dispenser, metric ruler. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Bond Pressure Line to Female Luer with THF on Extension Line. Bonded assembly, female luer, THF, THF dispenser, metric ruler. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Bond Pressure Line to Female Luer with THF on Extension Line. Bonded assembly, female luer, THF, THF dispenser, metric ruler. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Bond Pressure Line to Yellow Male Luer with THF on Short Line. Bonded assembly, yellow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Tubing THF Bonding. Bond Pressure Line to Yellow Male Luer with THF on Short Line. Bonded assembly, yellow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Bond Pressure Line to Yellow Male Luer with THF on Short Line. Bonded assembly, yellow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Bond Pressure Line to Yellow Male Luer with THF on Short Line. Bonded assembly, yellow male luer, THF, THF dispenser, metric ruler. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | contaminated surface leading to weak bond causing purge pressure loss
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | contamination from particles
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | particles/fibers on product
Tubing THF Bonding. Clean other end of Suction tubing and bond tubing to white side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | contaminated surface leading to weak bond causing purge pressure loss
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | contamination from particles
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | particles/fibers on product
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Clean Pressure tubing and bond tubing to clear side of Outflow check valve with THF. Outflow check valve, cured assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Clean shorter Suction tubing and bond tubing to clear side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | contaminated surface leading to weak bond causing purge pressure loss
Tubing THF Bonding. Clean shorter Suction tubing and bond tubing to clear side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Clean shorter Suction tubing and bond tubing to clear side of valve with THF. Valve, bonded assembly, THF, metric ruler, THF dispenser, Alcohol, lint-free wipe. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | contaminated surface leading to weak bond causing purge pressure loss
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | contamination from particles
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | particles/fibers on product
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | poor engagement leading to leaks and loss of purge pressure
Tubing THF Bonding. Clean Suction tubing and bond tubing to Inflow Female Luer with THF. Lint-free wipe, alcohol, suction line, inflow female leur, tubing blow-out fixture, THF dispenser, THF. Bonded assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Clear tubing of residual bonding agent using blow-out fixture. Bonded assembly, blow-out fixture. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Clear tubing of residual bonding agent using blow-out fixture. Bonded assembly, blow-out fixture. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Clear tubing of residual bonding agent using blow-out fixture. Bonded assembly, blow-out fixture. Bonded assembly. | residual agent leading to occlusion (unable to move pass case start)
Tubing THF Bonding. Clear tubing of residual bonding agent using blow-out fixture. Bonded assembly, blow-out fixture. Bonded assembly. | residual agent leading to occlusion (unable to move past case start)
Tubing THF Bonding. Fill THF dispenser until indicator stops rising. THF dispenser, THF, THF bottle. Filled dispenser. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Luer Connection utilizing Modular Pressure Line Assembly Fixture. Yellow male luer on short line, female luer extension line, modular pressure line assembly fixture. Connected assembly. | Connection not tight enough
Tubing THF Bonding. Luer Connection utilizing Modular Pressure Line Assembly Fixture. Yellow male luer on short line, female luer extension line, modular pressure line assembly fixture. Connected assembly. | Connection too tight
Tubing THF Bonding. Place Pressure Transmitter in housing and clear tubing of residual bonding agent using blow-out fixture. PC Pressure Transmitter Holding Block, Pressure Transmitter. cleared pressure transmitter. | damage of transmitter (unable to move pass case start)
Tubing THF Bonding. Place Pressure Transmitter in housing and clear tubing of residual bonding agent using blow-out fixture. PC Pressure Transmitter Holding Block, Pressure Transmitter. cleared pressure transmitter. | damage of transmitter (unable to move past case start)
Tubing THF Bonding. Place Pressure Transmitter in housing and clear tubing of residual bonding agent using blow-out fixture. PC Pressure Transmitter Holding Block, Pressure Transmitter. cleared pressure transmitter. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Place Pressure Transmitter in housing and clear tubing of residual bonding agent using blow-out fixture. PC Pressure Transmitter Holding Block, Pressure Transmitter. cleared pressure transmitter. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Set aside for proper cure time. Bonded Assembly, tray. Cured assembly. | weak bond leading to loss of purge pressure
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser, Transmitter Reinforcemnet Length Gauge. Bonded assembly. | kink/damage of spike line
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser. Bonded assembly. | damage tubing leading to leak and loss of purge pressure
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser. Bonded assembly. | occlusion of tubing (unable to move pass case start)
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser. Bonded assembly. | occlusion of tubing (unable to move past case start)
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser. Bonded assembly. | poor engagement of tubing
Tubing THF Bonding. Slide small Spike Line tubing open long Pressure Line. Bond long Pressure Line end to open Pressure Transmitter end with THF. Slide Spike Line over Pressure Transmitter. Bonded assembly, spike line, THF, THF dispenser. Bonded assembly. | weak bond leading to loss of purge pressure
Ultrasonic Cleaning. Clean parts. . . | Contaminants on parts post cleaning
Ultrasonic Cleaning. Clean parts. Ultrasonic Cleaning. Various components. | Contaminants on parts post cleaning
Upper PCB Stack Mfg Process (MFG-3013). Attach Power Control Harness. . ASSY, PCB Installation. | Power contol harness not connected
Upper PCB Stack Mfg Process (MFG-3013). Attach RIB GPIO Harness. . ASSY, PCB Installation. | Ribbon cable not connected
Upper PCB Stack Mfg Process (MFG-3013). Attach RIB I2C Harness. . ASSY, PCB Installation. | Pneumatic sensor cable not connected
Upper PCB Stack Mfg Process (MFG-3013). Attach RIB SPI/RESET Harness. . ASSY, PCB Installation. | PCB interface connector not connected
Upper PCB Stack Mfg Process (MFG-3013). Attach System Power Enable Harness. . ASSY, PCB Installation. | Power enable cable not connected
Upper PCB Stack Mfg Process (MFG-3013). Attach thermistor harness. . ASSY, PCB Installation. | Thermistor not connected
Upper PCB Stack Mfg Process (MFG-3013). Route and Secure DC Cable and Ferrite.. . ASSY, PCB Installation. | Incorrect cable connection Cable strain Cable damage
Upper PCB Stack Mfg Process (MFG-3013). Set dip switches. . ASSY, PCB Installation. | Dip switches incorrectly set
Upper PCB Stack Mfg Process (MFG-3013). Set Jumper location. . ASSY, PCB Installation. | Jumper incorrectly set
UV Gluing Pumps Process. Glue Pump Impeller components. UV Gluing Pumps Process. Pump Impeller. | Impeller too tall / too short
UV Gluing Pumps Process. Glue Pump Impeller components. UV Gluing Pumps Process. Pump Impeller. | Incorrect magnet polarities
UV Gluing Pumps Process. Glue Pump Impeller components. UV Gluing Pumps Process. Pump Impeller. | Non-continuous glue joint
UV Gluing Pumps Process. Glue Pump Impeller components. UV Gluing Pumps Process. Pump Impeller. | Not all magnets installed
Vibratory grinding of rotors. Inspection of diameter. Pre-tumbled rotors w/o adequate diameter. Pre-tumbled rotors with adequate diameter. | Damaging of rotors
Vibratory grinding of rotors. Inspection of diameter. Pre-tumbled rotors w/o adequate diameter. Pre-tumbled rotors with adequate diameter. | Dimensional nonconformance
Vibratory grinding of rotors. Mixing of vibratory grinding additives. Unmixed grinding additives, rotors. Rotors & appropriate grinding additives ready for tumbling. | Mixing components characteristics not in specification, resulting in surface or dimensional nonconformities
Vibratory grinding of rotors. Mixing of vibratory grinding additives. Unmixed grinding additives, rotors. Rotors & appropriate grinding additives ready for tumbling. | Wrong granulates or additives used, resulting in surface- or dimensional non- conformities
Vibratory grinding of rotors. Mixing of vibratory grinding additives. Unmixed grinding additives, rotors. Rotors & appropriate grinding additives ready for tumbling. | Wrong mixing ratio used, resulting in surface or dimensional nonconformities
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Damaging of rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Machine characteristics wrong, resulting in damaged rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Not polished
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Polished without additives
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Polishing container not fixed, resulting in damage of rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Polishing duration too short / too long, resulting in damage of rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | polishing materials worn out, resulting in damaged rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | residues in the polishing container, resulting in damaged rotors
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | Sieves damage rotor
Vibratory grinding of rotors. Polishing of rotors. Rotors cleaned from injection molding edges & smoothed. Machine- polished rotors. | wrong sequence of work steps, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | Damaging of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | Machine characteristics wrong, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | residues in the polishing container, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | residues in the tumbling container, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | Rotor geometry incorrect (wing radius/ OD)
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | Sieves damage rotor
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | tumbling container not fixed, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | tumbling container not fixed, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | tumbling duration too short / too long, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors w/o adequate diameter. | Tumbling materials worn out, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | Damaging of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | Machine characteristics wrong, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | residues in the polishing container, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | residues in the tumbling container, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | Rotor geometry incorrect (wing radius/ OD)
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | Sieves damage rotor
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | tumbling container not fixed, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | tumbling container not fixed, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | tumbling duration too short / too long, resulting in damage of rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbled rotors. | Tumbling materials worn out, resulting in damaged rotors
Vibratory grinding of rotors. Pre-tumbling. Rotors cleaned from injection molding edges & smoothed. Pre-tumbPre-tumbled rotors w/o adequate diameterled rotors. | tumbling duration too short / too long, resulting in damage of rotors
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Brittle glue in glue joint
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Excessive glue at glue joint
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Negative glue joint
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Non-continuous glue joint
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Not fully cured glue in glue joint
XL UV Gluing Procedure, UV Gluing Inlet Process UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Glue Pump and Oxy components. See UV Gluing Procedure, UV Gluing Inlet Process, UV Gluing Exits Process, UV Gluing Mid Housings XL, UV Gluing Pumps Process. Pump Impeller, Pump Inlet, Pump, Oxy Mid, Oxy Exit. | Part deformation and/or damage
Note on B3824: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Welche Siebe?
Note on B3834: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Welche Siebe?

## Sheet1
Categories | rows
dFMEA | 2359
pFMEA | 3845
uFMEA | 954
sFMEA | 281
Impella_IFU | 290
Connect_IFU | 67
Connect_dFMEA | 31
total | 7827

## hazard_HS_Harm
Hazard | Hazardous Situation | Harm | S | Hazard (58) | Hazardous Situation (65) | Harm (29)
Active Pump Removed | Mechanical Damage To Cardiac Structure | Heart Injury | 5 | Active Pump Removed | AIC Purge Cassette Door Wont Stay Closed | 1st Degree Burn
AIC Alarm Failure/Delay | Unexpected Stop Of Pump, Due To AIC Failure | Loss Of Hemodynamic Support | 5 | AIC Alarm Failure/Delay | Air Within System, Introduced Within Vasculature | Additional Surgical (Case) Intervention
AIC Failure | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support | 3 | AIC Failure | Cart/Console Fall On Patient/ User/Bystander | Additional Surgical Intervention
AIC Failure | Non-Essential AIC Defect | User Inconvenience | 1 | AIC Inoperable | Compromised Product Packaging | Anemia
AIC Inoperable | Unable To Drive Pump | Loss Of Hemodynamic Support | 5 | AIC Purge Cassette Door | Excessive Current Delivered To Patient, Thru Pump | Arrhythmia
AIC Purge Cassette Door | AIC Purge Cassette Door Wont Stay Closed | User Inconvenience | 1 | AIC Startup Failure | Excessive Current Delivered To User | Burn
AIC Startup Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Allergen | Excessive Force On Ventricle Wall | Burn, 2nd Degree
Allergen | Incorrect Selection Of Purge Solution Additive | Heparin-Induced Thrombocytopenia (HIT) | 5 | Automated Flow Rate Adjustment | Excessive Force Required For Insertion | Cardiac Arrhythmia
Automated Flow Rate Adjustment | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Battery Overheating / Failure | Exposure To High Temperature Surfaces On AIC | Cardiac Perforation
Battery Overheating / Failure | Fire | Burn | 5 | Cannula/Cage Separation | Exposure To Sharp Edges On AIC | Disection
Cannula/Cage Separation | Loose Cannula/Cage In Great Vessel Or Heart | Thromboembolism | 5 | Catheter Separation | Fire | Heart Injury
Cannula/Cage Separation | Loose Cannula/Cage In Peripheral Vessels | Peripheral Vascular Occlusion | 4 | Device Compatibility | Hemolysis, Due To High Temperature | Hemorphagic Shock
Catheter Separation | Loose Device Segment In Vessel Or Heart | Surgical Intervention | 3 | Disruption of Blood Flow | Hemolysis, Due To Turbulent Flow | Heparin-Induced Thrombocytopenia (HIT)
Device Compatibility | Inability To Insert Concomitant Device Through Sheath | Procedural Delay | 1 | Disruption Of Blood Flow Near Access Site | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support
Device Compatibility | Inability To Remove Concomitant Device Through Sheath | Additional Surgical (Case) Intervention | 3 | Electrical - Leakage Current | Inability To Insert Concomitant Device Through Sheath | Ischemia
Disruption of Blood Flow | Release Of Thrombus Within Vessel | Thromboembolism | 5 | Electrical Pump Motor Failure | Inability To Remove Concomitant Device Through Sheath | Loss Of Hemodynamic Support
Disruption Of Blood Flow Near Access Site | Irreversible Obstruction Of Peripheral Vessel | Peripheral Vascular Occlusion | 4 | Falling Object | Inability To Reposition Without Advanced Imaging | Major Bleed
Disruption Of Blood Flow Near Access Site | Reversible Obstruction Of Peripheral Vessel | Ischemia | 3 | False Alarm/ Notification | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | Peripheral Vascular Occlusion
Electrical - Leakage Current | Excessive Current Delivered To Patient, Thru Pump | Cardiac Arrhythmia | 5 | Foreign Material | Incorrect Clinical Decision Made Based On System Provided Data | Peripheral Vessel Disection
Electrical - Leakage Current | Excessive Current Delivered To User | Burn, 2nd Degree | 4 | High Restistance Between Devices | Incorrect Selection Of Purge Solution Additive | Procedural Delay
Electrical Pump Motor Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | High Temperature | Increased Replacement Of Purge System | Sepsis
Electrical Pump Motor Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Impella Interaction With Other Device | Irreversible Obstruction Of Peripheral Vessel | Sepsis
Electrical Pump Motor Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Inadequate Priming Of Purge Line | Loose Cannula/Cage In Great Vessel Or Heart | Skin Laceration
Electrical Pump Motor Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Inadequate Pump Visualization | Loose Cannula/Cage In Peripheral Vessels | Soft Tissue Injury
Falling Object | Cart/Console Fall On Patient/ User/Bystander | Soft Tissue Injury | 4 | Incorrect Placement | Loose Component Due To Excessive Removal Force | Stroke
False Alarm/ Notification | Unnecessary Troubleshooting | User Inconvenience | 1 | Incorrect Placement Of Graft | Loose Device Segment In Vessel Or Heart | Surgical Intervention
Foreign Material | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5 | Incorrect Pump Insertion | Loose Pigtail In Great Vessel Or Heart | Thromboembolism
High Restistance Between Devices | Excessive Force Required For Insertion | User Inconvenience | 1 | Incorrect Pump Performance Metrics | Loose Pigtail In Peripheral Vessels | Toxic Reaction
High Restistance Between Devices | Loose Component Due To Excessive Removal Force | Additional Surgical (Case) Intervention | 3 | Incorrect Pump Setup | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience
High Restistance Between Devices | Vessel Trauma At Access Site | Major Bleed | 3 | Incorrect Purge Solution Selection | Loss Of Full Support Pump Therapy
High Temperature | Exposure To High Temperature Surfaces On AIC | 1st Degree Burn | 2 | Incorrect Repositioning | Loss Of Partial Support Pump Therapy
Impella Interaction With Other Device | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5 | Introducer /Sheath Mechanical Issue | Mechanical Damage To Cardiac Structure
Inadequate Priming Of Purge Line | Air Within System, Introduced Within Vasculature | Thromboembolism | 5 | Loss Of Battery Power | Non-Essential AIC Defect
Inadequate Pump Visualization | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Mechanical Pump Motor Failure | Non-Functional Device Identified Prior To Therapy Initiation
Incorrect Placement | Excessive Force On Ventricle Wall | Cardiac Perforation | 5 | Micro-Organism/Pathogen | Non-Sterile Device Used In Patient
Incorrect Placement | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Non-Biocompatibility | Overestimated Diagnostic Cardiac Performance Information
Incorrect Placement | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Obstruction Of Impella Flow Path | Patient Exposed To Non-Biocompatible Material
Incorrect Placement | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Optical Sensor Error | Patient/User/Bystander Trips On Console Cable
Incorrect Placement | Mechanical Damage To Cardiac Structure | Heart Injury | 5 | Optical Sensor Failure | Perforation Of Great Vessel
Incorrect Placement | Perforation Of Great Vessel | Hemorphagic Shock | 5 | Partial Obstruction Of Impella Flow Path | Perforation Of Peripheral Vessel
Incorrect Placement | Perforation Of Peripheral Vessel | Major Bleed | 3 | Pigtail Separation | Perforation On Great Vessel
Incorrect Placement | Pump Function Degraded, Due To Suction | Inadequate Hemodynamic Support | 3 | Pressure Storage Set Leak/Failure | Pump Fails To Initiate Prior To Therapy Initiation
Incorrect Placement | Sharp Surface Skives Great Vessel | Disection | 5 | Product Packaging Integrity | Pump Function Degraded
Incorrect Placement | Sharp Surface Skives Peripheral Vessel | Peripheral Vessel Disection | 3 | Programming / Software Error | Pump Function Degraded, Due To Suction
Incorrect Placement | Unintended Stimulation Of Cardiac Wall | Arrhythmia | 5 | Pump Failure | Release of Thrombus From Device
Incorrect Placement Of Graft | Unable To Effectively Place Pump | Additional Surgical Intervention | 3 | Purge Cassette Failure | Release Of Thrombus Within Vessel
Incorrect Pump Insertion | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Purge System Failure | Reversible Obstruction Of Peripheral Vessel
Incorrect Pump Insertion | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Recoverable AIC Failure | Sharp Surface Skives Great Vessel
Incorrect Pump Performance Metrics | Incorrect Clinical Decision Made Based On System Provided Data | Inadequate Hemodynamic Support | 3 | Rough Surfaces | Sharp Surface Skives Peripheral Vessel
Incorrect Pump Setup | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Sharp Edges | Sheath Removal Compromised
Incorrect Pump Setup | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Small Component Dislodge From Device | Small Component(S) Left Within Patient Great Vessel Or Heart
Incorrect Pump Setup | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Software / Algorithm Error | Small Component(S) Left Within Patient Peripheral Vessels
Incorrect Pump Setup | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Software Error | Soft Tissue Trauma At Access Site
Incorrect Purge Solution Selection | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Thermal - High Temperature | Temporary Pump Stop
Incorrect Purge Solution Selection | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Thrombus | Thermal Damage To Ventricle/Vasculature
Incorrect Purge Solution Selection | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Tripping | Unable To Drive Pump
Incorrect Repositioning | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Unsecured Device | Unable To Effectively Place Pump
Incorrect Repositioning | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Use Error | Unable To Successfully Place Sheath
Incorrect Repositioning | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5 |  | Underestimated Diagnostic Cardiac Performance Information
Incorrect Repositioning | Perforation On Great Vessel | Hemorphagic Shock | 5 |  | Unexpected Stop Of Full Support Pump
Incorrect Repositioning | Sharp Surface Skives Great Vessel | Disection | 5 |  | Unexpected Stop Of Partial Support Pump
Introducer /Sheath Mechanical Issue | Sheath Removal Compromised | Procedural Delay | 1 |  | Unexpected Stop Of Pump, Due To AIC Failure
Introducer /Sheath Mechanical Issue | Soft Tissue Trauma At Access Site | Minor Bleed | 2 |  | Unintended Stimulation Of Cardiac Wall
Introducer /Sheath Mechanical Issue | Unable To Successfully Place Sheath | Procedural Delay | 1 |  | Unintentional Loss Of Sheath Integrity Or Placement
Introducer /Sheath Mechanical Issue | Unintentional Loss Of Sheath Integrity Or Placement | Hemorphagic Shock | 5 |  | Unnecessary Troubleshooting
Introducer /Sheath Mechanical Issue | Vessel Trauma At Access Site | Major Bleed | 3 |  | User Selects Incorrect P-Level
Loss Of Battery Power | Unable To Drive Pump | Loss Of Hemodynamic Support | 5 |  | Vessel Trauma At Access Site
Mechanical Pump Motor Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3
Mechanical Pump Motor Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5
Mechanical Pump Motor Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5
Micro-Organism/Pathogen | Non-Sterile Device Used In Patient | Sepsis | 5
Non-Biocompatibility | Patient Exposed To Non-Biocompatible Material | Toxic Reaction | 5
Obstruction Of Impella Flow Path | Pump Function Degraded | Inadequate Hemodynamic Support | 3
Obstruction Of Impella Flow Path | Release Of Thrombus Within Vessel | Thromboembolism | 5
Obstruction Of Impella Flow Path | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5
Obstruction Of Impella Flow Path | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5
Optical Sensor Error | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | User Inconvenience | 1
Optical Sensor Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1
Optical Sensor Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1
Optical Sensor Failure | Inability To Reposition Without Advanced Imaging | User Inconvenience | 1
Optical Sensor Failure | Non-Functional Device Identified Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3
Partial Obstruction Of Impella Flow Path | Hemolysis, Due To Turbulent Flow | Anemia | 3
Pigtail Separation | Loose Pigtail In Great Vessel Or Heart | Thromboembolism | 5
Pigtail Separation | Loose Pigtail In Peripheral Vessels | Peripheral Vascular Occlusion | 4
Pressure Storage Set Leak/Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3
Product Packaging Integrity | Compromised Product Packaging | User Inconvenience | 1
Programming / Software Error | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3
Programming / Software Error | Pump Function Degraded | Inadequate Hemodynamic Support | 3
Programming / Software Error | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5
Programming / Software Error | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5
Pump Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3
Purge Cassette Failure | Increased Replacement Of Purge System | User Inconvenience | 1
Purge System Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3
Purge System Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3
Purge System Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5
Purge System Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5
Recoverable AIC Failure | Temporary Pump Stop | Inadequate Hemodynamic Support | 3
Rough Surfaces | Hemolysis, Due To Turbulent Flow | Anemia | 3
Sharp Edges | Exposure To Sharp Edges On AIC | Skin Laceration | 2
Small Component Dislodge From Device | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5
Small Component Dislodge From Device | Small Component(S) Left Within Patient Peripheral Vessels | Peripheral Vascular Occlusion | 4
Software / Algorithm Error | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | User Inconvenience | 1
Software / Algorithm Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1
Software / Algorithm Error | Overestimated Diagnostic Cardiac Performance Information | User Inconvenience | 1
Software / Algorithm Error | Underestimated Diagnostic Cardiac Performance Information | User Inconvenience | 1
Software Error | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support | 3
Thermal - High Temperature | Hemolysis, Due To High Temperature | Anemia | 3
Thermal - High Temperature | Thermal Damage To Ventricle/Vasculature | Heart Injury | 5
Thrombus | Release of Thrombus From Device | Stroke | 5
Tripping | Patient/User/Bystander Trips On Console Cable | Soft Tissue Injury | 4
Unsecured Device | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5
Unsecured Device | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5
Unsecured Device | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3
Unsecured Device | Pump Function Degraded, Due To Suction | Inadequate Hemodynamic Support | 3
Use Error | Non-Sterile Device Used In Patient | Sepsis | 5
Use Error | Unnecessary Troubleshooting | User Inconvenience | 1
Use Error | User Selects Incorrect P-Level | Inadequate Hemodynamic Support | 3