## HA_Pool
| Hazard | Hazardous Situation | Harm | S | Hazard (58) | Hazardous Situation (65) | Harm (29)
P2017-131_rY | Active Pump Removed | Mechanical Damage To Cardiac Structure | Heart Injury | 5 | Active Pump Removed | AIC Purge Cassette Door Wont Stay Closed | 1st Degree Burn
P2017-131_rY | AIC Alarm Failure/Delay | Unexpected Stop Of Pump, Due To AIC Failure | Loss Of Hemodynamic Support | 5 | AIC Alarm Failure/Delay | Air Within System, Introduced Within Vasculature | Additional Surgical (Case) Intervention
P2017-131_rY | AIC Failure | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support | 3 | AIC Failure | Cart/Console Fall On Patient/ User/Bystander | Additional Surgical Intervention
P2017-131_rY | AIC Failure | Non-Essential AIC Defect | User Inconvenience | 1 | AIC Inoperable | Compromised Product Packaging | Anemia
P2017-131_rY | AIC Inoperable | Unable To Drive Pump | Loss Of Hemodynamic Support | 5 | AIC Purge Cassette Door | Excessive Current Delivered To Patient, Thru Pump | Arrhythmia
P2017-131_rY | AIC Purge Cassette Door | AIC Purge Cassette Door Wont Stay Closed | User Inconvenience | 1 | AIC Startup Failure | Excessive Current Delivered To User | Burn
P2017-131_rY | AIC Startup Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Allergen | Excessive Force On Ventricle Wall | Burn, 2nd Degree
P2017-131_rY | Allergen | Incorrect Selection Of Purge Solution Additive | Heparin-Induced Thrombocytopenia (HIT) | 5 | Automated Flow Rate Adjustment | Excessive Force Required For Insertion | Cardiac Arrhythmia
P2017-131_rY | Automated Flow Rate Adjustment | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Battery Overheating / Failure | Exposure To High Temperature Surfaces On AIC | Cardiac Perforation
P2017-131_rY | Battery Overheating / Failure | Fire | Burn | 5 | Cannula/Cage Separation | Exposure To Sharp Edges On AIC | Disection
P2017-131_rY | Cannula/Cage Separation | Loose Cannula/Cage In Great Vessel Or Heart | Thromboembolism | 5 | Catheter Separation | Fire | Heart Injury
P2017-131_rY | Cannula/Cage Separation | Loose Cannula/Cage In Peripheral Vessels | Peripheral Vascular Occlusion | 4 | Device Compatibility | Hemolysis, Due To High Temperature | Hemorphagic Shock
P2017-131_rY | Catheter Separation | Loose Device Segment In Vessel Or Heart | Surgical Intervention | 3 | Disruption of Blood Flow | Hemolysis, Due To Turbulent Flow | Heparin-Induced Thrombocytopenia (HIT)
P2017-131_rY | Device Compatibility | Inability To Insert Concomitant Device Through Sheath | Procedural Delay | 1 | Disruption Of Blood Flow Near Access Site | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support
P2017-131_rY | Device Compatibility | Inability To Remove Concomitant Device Through Sheath | Additional Surgical (Case) Intervention | 3 | Electrical - Leakage Current | Inability To Insert Concomitant Device Through Sheath | Ischemia
P2017-131_rY | Disruption of Blood Flow | Release Of Thrombus Within Vessel | Thromboembolism | 5 | Electrical Pump Motor Failure | Inability To Remove Concomitant Device Through Sheath | Loss Of Hemodynamic Support
P2017-131_rY | Disruption Of Blood Flow Near Access Site | Irreversible Obstruction Of Peripheral Vessel | Peripheral Vascular Occlusion | 4 | Falling Object | Inability To Reposition Without Advanced Imaging | Major Bleed
P2017-131_rY | Disruption Of Blood Flow Near Access Site | Reversible Obstruction Of Peripheral Vessel | Ischemia | 3 | False Alarm/ Notification | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | Peripheral Vascular Occlusion
P2017-131_rY | Electrical - Leakage Current | Excessive Current Delivered To Patient, Thru Pump | Cardiac Arrhythmia | 5 | Foreign Material | Incorrect Clinical Decision Made Based On System Provided Data | Peripheral Vessel Disection
P2017-131_rY | Electrical - Leakage Current | Excessive Current Delivered To User | Burn, 2nd Degree | 4 | High Restistance Between Devices | Incorrect Selection Of Purge Solution Additive | Procedural Delay
P2017-131_rY | Electrical Pump Motor Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | High Temperature | Increased Replacement Of Purge System | Sepsis
P2017-131_rY | Electrical Pump Motor Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Impella Interaction With Other Device | Irreversible Obstruction Of Peripheral Vessel | Sepsis
P2017-131_rY | Electrical Pump Motor Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Inadequate Priming Of Purge Line | Loose Cannula/Cage In Great Vessel Or Heart | Skin Laceration
P2017-131_rY | Electrical Pump Motor Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Inadequate Pump Visualization | Loose Cannula/Cage In Peripheral Vessels | Soft Tissue Injury
P2017-131_rY | Falling Object | Cart/Console Fall On Patient/ User/Bystander | Soft Tissue Injury | 4 | Incorrect Placement | Loose Component Due To Excessive Removal Force | Stroke
P2017-131_rY | False Alarm/ Notification | Unnecessary Troubleshooting | User Inconvenience | 1 | Incorrect Placement Of Graft | Loose Device Segment In Vessel Or Heart | Surgical Intervention
P2017-131_rY | Foreign Material | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5 | Incorrect Pump Insertion | Loose Pigtail In Great Vessel Or Heart | Thromboembolism
P2017-131_rY | High Restistance Between Devices | Excessive Force Required For Insertion | User Inconvenience | 1 | Incorrect Pump Performance Metrics | Loose Pigtail In Peripheral Vessels | Toxic Reaction
P2017-131_rY | High Restistance Between Devices | Loose Component Due To Excessive Removal Force | Additional Surgical (Case) Intervention | 3 | Incorrect Pump Setup | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience
P2017-131_rY | High Restistance Between Devices | Vessel Trauma At Access Site | Major Bleed | 3 | Incorrect Purge Solution Selection | Loss Of Full Support Pump Therapy | Arrhythmia
P2017-131_rY | High Temperature | Exposure To High Temperature Surfaces On AIC | 1st Degree Burn | 2 | Incorrect Repositioning | Loss Of Partial Support Pump Therapy | Bleeding lead to medical intervention which only requires a vasculature closure device (ex. Balloon Tamponade)
P2017-131_rY | Impella Interaction With Other Device | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5 | Introducer /Sheath Mechanical Issue | Mechanical Damage To Cardiac Structure | Cardiac/Vasculature Surgery
P2017-131_rY | Inadequate Priming Of Purge Line | Air Within System, Introduced Within Vasculature | Thromboembolism | 5 | Loss Of Battery Power | Non-Essential AIC Defect | Death
P2017-131_rY | Inadequate Pump Visualization | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Mechanical Pump Motor Failure | Non-Functional Device Identified Prior To Therapy Initiation | Delay of Therapy/Inconvenience
P2017-131_rY | Incorrect Placement | Excessive Force On Ventricle Wall | Cardiac Perforation | 5 | Micro-Organism/Pathogen | Non-Sterile Device Used In Patient | Embolism/Stroke
P2017-131_rY | Incorrect Placement | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Non-Biocompatibility | Overestimated Diagnostic Cardiac Performance Information | Excess bleeding requires an unplanned invasive or operable percutaneous intervention (ex. Vasculature surgery to repair femoral artery)
P2017-131_rY | Incorrect Placement | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Obstruction Of Impella Flow Path | Patient Exposed To Non-Biocompatible Material | Inconvenience or temporary discomfort
P2017-131_rY | Incorrect Placement | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Optical Sensor Error | Patient/User/Bystander Trips On Console Cable | Infection/Sepsis
P2017-131_rY | Incorrect Placement | Mechanical Damage To Cardiac Structure | Heart Injury | 5 | Optical Sensor Failure | Perforation Of Great Vessel | Loss of hemodynamic support leading to medical intervention
P2017-131_rY | Incorrect Placement | Perforation Of Great Vessel | Hemorphagic Shock | 5 | Partial Obstruction Of Impella Flow Path | Perforation Of Peripheral Vessel | Loss of hemodynamic support leading to medical intervention (ex. Premature removal of device)
P2017-131_rY | Incorrect Placement | Perforation Of Peripheral Vessel | Major Bleed | 3 | Pigtail Separation | Perforation On Great Vessel | Loss of hemodynamic support leading to medical intervention (ex: Use of Snare)
P2017-131_rY | Incorrect Placement | Pump Function Degraded, Due To Suction | Inadequate Hemodynamic Support | 3 | Pressure Storage Set Leak/Failure | Pump Fails To Initiate Prior To Therapy Initiation | Loss of hemodynamic support leading to medical intervention (New pump insertion)
P2017-131_rY | Incorrect Placement | Sharp Surface Skives Great Vessel | Disection | 5 | Product Packaging Integrity | Pump Function Degraded | Medical intervention (ex: Use of Snare)
P2017-131_rY | Incorrect Placement | Sharp Surface Skives Peripheral Vessel | Peripheral Vessel Disection | 3 | Programming / Software Error | Pump Function Degraded, Due To Suction | Medical intervention to install bypass or remove blockage.
P2017-131_rY | Incorrect Placement | Unintended Stimulation Of Cardiac Wall | Arrhythmia | 5 | Pump Failure | Release of Thrombus From Device | Minor Inconvenience/Delay in Therapy
P2017-131_rY | Incorrect Placement Of Graft | Unable To Effectively Place Pump | Additional Surgical Intervention | 3 | Purge Cassette Failure | Release Of Thrombus Within Vessel | Organ damage
P2017-131_rY | Incorrect Pump Insertion | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Purge System Failure | Reversible Obstruction Of Peripheral Vessel | Procedure disruption causing longer than expected delays
P2017-131_rY | Incorrect Pump Insertion | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Recoverable AIC Failure | Sharp Surface Skives Great Vessel | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products)
P2017-131_rY | Incorrect Pump Performance Metrics | Incorrect Clinical Decision Made Based On System Provided Data | Inadequate Hemodynamic Support | 3 | Rough Surfaces | Sharp Surface Skives Peripheral Vessel | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products, applying manual pressure)
P2017-131_rY | Incorrect Pump Setup | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Sharp Edges | Sheath Removal Compromised | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex: Femstop or additionally blood products)
P2017-131_rY | Incorrect Pump Setup | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Small Component Dislodge From Device | Small Component(S) Left Within Patient Great Vessel Or Heart | Reduced blood flow throughout the body
P2017-131_rY | Incorrect Pump Setup | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Software / Algorithm Error | Small Component(S) Left Within Patient Peripheral Vessels | Reduced hemodynamic support
P2017-131_rY | Incorrect Pump Setup | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Software Error | Soft Tissue Trauma At Access Site | Skin burn (1st degree)
P2017-131_rY | Incorrect Purge Solution Selection | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Thermal - High Temperature | Temporary Pump Stop | Stroke
P2017-131_rY | Incorrect Purge Solution Selection | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Thrombus | Thermal Damage To Ventricle/Vasculature | Surgical Removal of Extremity
P2017-131_rY | Incorrect Purge Solution Selection | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Tripping | Unable To Drive Pump | Systemic toxic reaction
P2017-131_rY | Incorrect Repositioning | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Unsecured Device | Unable To Effectively Place Pump | Unplanned invasive operable or percutaneous intervention (surgical removal)
P2017-131_rY | Incorrect Repositioning | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Use Error | Unable To Successfully Place Sheath | Vasculature/Cardiac surgery
P2017-131_rY | Incorrect Repositioning | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Access Site Bleeding | Underestimated Diagnostic Cardiac Performance Information | Ventricular Fibrillation/Ventricular tachycardia/Asystole
P2017-131_rY | Incorrect Repositioning | Perforation On Great Vessel | Hemorphagic Shock | 5 | Air in lumen | Unexpected Stop Of Full Support Pump | Delay of Therapy/ Inconvenience
P2017-131_rY | Incorrect Repositioning | Sharp Surface Skives Great Vessel | Disection | 5 | Allergen | Unexpected Stop Of Partial Support Pump | Inconvenience
P2017-131_rY | Introducer /Sheath Mechanical Issue | Sheath Removal Compromised | Procedural Delay | 1 | Biological blockage of the system | Unexpected Stop Of Pump, Due To AIC Failure | Loss of hemodynamic support leading to medical intervention
P2017-131_rY | Introducer /Sheath Mechanical Issue | Soft Tissue Trauma At Access Site | Minor Bleed | 2 | Blockage of the system | Unintended Stimulation Of Cardiac Wall | No patient harm is associated with this use step
P2017-131_rY | Introducer /Sheath Mechanical Issue | Unable To Successfully Place Sheath | Procedural Delay | 1 | Blood Clotting | Unintentional Loss Of Sheath Integrity Or Placement | Reduced hemodynamic support
P2017-131_rY | Introducer /Sheath Mechanical Issue | Unintentional Loss Of Sheath Integrity Or Placement | Hemorphagic Shock | 5 | Blood Damage | Unnecessary Troubleshooting | Cardiac/Vascular Surgery
P2017-131_rY | Introducer /Sheath Mechanical Issue | Vessel Trauma At Access Site | Major Bleed | 3 | Bond failures | User Selects Incorrect P-Level | Death
P2017-131_rY | Loss Of Battery Power | Unable To Drive Pump | Loss Of Hemodynamic Support | 5 | Clot on pump head | Vessel Trauma At Access Site | Delay of Therapy/Inconvenience
P2017-131_rY | Mechanical Pump Motor Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Components damaged during insertion/removal | AIC shuts down/freezes/ inoperable | Elevation of the leg or device removal
P2017-131_rY | Mechanical Pump Motor Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Console malfunction | Air Embolism | Embolism/Stroke
P2017-131_rY | Mechanical Pump Motor Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Electrical failure of motor | Clot formation/Thrombosis | Excess bleeding requires an unplanned invasive or operable percutaneous intervention (ex. Vasculature surgery to repair vein)
P2017-131_rY | Micro-Organism/Pathogen | Non-Sterile Device Used In Patient | Sepsis | 5 | Electrical failure of the pump | Clot prevents pump head from collapsing | Inconvenience or temporary discomfort
P2017-131_rY | Non-Biocompatibility | Patient Exposed To Non-Biocompatible Material | Toxic Reaction | 5 | High Temperature | Components break inside body | Inconvenience/Pump remains functional
P2017-131_rY | Obstruction Of Impella Flow Path | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Incorrect Algorithm | Components left inside patient | Infection/Sepsis
P2017-131_rY | Obstruction Of Impella Flow Path | Release Of Thrombus Within Vessel | Thromboembolism | 5 | Incorrect description/specification for use of device | Damage to vasculature | Loss of hemodynamic support leading to medical intervention
P2017-131_rY | Obstruction Of Impella Flow Path | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Incorrect Placement | Damage to ventricle/vasculature | Loss of hemodynamic support leading to medical intervention (ex. Premature removal of device)
P2017-131_rY | Obstruction Of Impella Flow Path | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Incorrect Programing/Insufficient Power | Excessive Bleeding | Loss of hemodynamic support leading to medical intervention (ex: Use of Snar)
P2017-131_rY | Optical Sensor Error | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | User Inconvenience | 1 | Incorrect Programing/Insufficient Power/ Communication Error | Excessive Bleeding due to damage to vasculature | Loss of hemodynamic support leading to medical intervention (New pump insertion)
P2017-131_rY | Optical Sensor Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1 | Incorrect Repositioning | Excessive current delivered to patient | Medical interention required
P2017-131_rY | Optical Sensor Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1 | Leakage Current | False positive alarm | Medical intervention (ex: Use of Snare)
P2017-131_rY | Optical Sensor Failure | Inability To Reposition Without Advanced Imaging | User Inconvenience | 1 | Mechanical blockage of the system | Hemolysis | Minor Inconvenience/Delay in Therapy
P2017-131_rY | Optical Sensor Failure | Non-Functional Device Identified Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Mechanical failure of motor | Heparin-induced thrombocytopenia (HIT) | Organ damage
P2017-131_rY | Partial Obstruction Of Impella Flow Path | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Micro-organism | Hot part in skin contact | Procedure disruption causing longer than expected delays
P2017-131_rY | Pigtail Separation | Loose Pigtail In Great Vessel Or Heart | Thromboembolism | 5 | Optimal Purge solution not utilized | Interference with cardiac conduction | Pulmonary Edema
P2017-131_rY | Pigtail Separation | Loose Pigtail In Peripheral Vessels | Peripheral Vascular Occlusion | 4 | Outflow in ventricle | Ischemia | Pulmonary Embolism
P2017-131_rY | Pressure Storage Set Leak/Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Particulates generation | Leachable compound in design path | Reduced blood flow throughout the body
P2017-131_rY | Product Packaging Integrity | Compromised Product Packaging | User Inconvenience | 1 | Pump contacting LV wall | Low Flow | Reduced hemodynamic support
P2017-131_rY | Programming / Software Error | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Pump head stuck | Malposition of the device | Results in patient death
P2017-131_rY | Programming / Software Error | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Pump Setup incorrectly | Obstruction of vessels/Thrombosis | Systemic toxic reaction
P2017-131_rY | Programming / Software Error | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Purge System Failure | Pathogens present after sterilization | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure)
P2017-131_rY | Programming / Software Error | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Reduced Blood Flow to Leg | Procedure aborted | Transient Arrythmia
P2017-131_rY | Pump Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Toxicity | Prolong skin and vasculature contact with non biocompatible material | Unplanned invasive operable or percutaneous intervention (surgical removal)
P2017-131_rY | Purge Cassette Failure | Increased Replacement Of Purge System | User Inconvenience | 1 | AIC Cart falls over | Pump Fails to Start | Ventricular Fibrillation/Ventricular tachycardia/Asystole
P2017-131_rY | Purge System Failure | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Data not transmitted to cloud based server | Pump head does not collapse
P2017-131_rY | Purge System Failure | Pump Function Degraded | Inadequate Hemodynamic Support | 3 | Incorrect Display (Abi) | Pump provides lower than expected support
P2017-131_rY | Purge System Failure | Unexpected Stop Of Full Support Pump | Loss Of Hemodynamic Support | 5 | Incorrect Display (AIC) | Pump Stop/Failed to start after implantation
P2017-131_rY | Purge System Failure | Unexpected Stop Of Partial Support Pump | Loss Of Hemodynamic Support | 5 | Incorrect Output (avg. flow defict and minutes in suction) | Pump Stop/Removal
P2017-131_rY | Recoverable AIC Failure | Temporary Pump Stop | Inadequate Hemodynamic Support | 3 | Integrating new information beyond AIC Screen | Unable to insert device leading to procedure being aborted/second pump utilized
P2017-131_rY | Rough Surfaces | Hemolysis, Due To Turbulent Flow | Anemia | 3 | Interference | AIC reboots due to interference
P2017-131_rY | Sharp Edges | Exposure To Sharp Edges On AIC | Skin Laceration | 2 | LED blink codes | AIC shuts down/freezes/ inoperable
P2017-131_rY | Small Component Dislodge From Device | Small Component(S) Left Within Patient Great Vessel Or Heart | Thromboembolism | 5 | Loss of Privacy & Security | Pump provides inadequate support
P2017-131_rY | Small Component Dislodge From Device | Small Component(S) Left Within Patient Peripheral Vessels | Peripheral Vascular Occlusion | 4 | Loss or deterioration of AIC's function | User misintrepets information
P2017-131_rY | Software / Algorithm Error | Inaccurate Or Loss Of Non-Diagnostic Hemodynamic Measurement | User Inconvenience | 1 | Loss or deterioration of RLM's user interface functionality | User unable to access accurate information on IC
P2017-131_rY | Software / Algorithm Error | Loss Of Diagnostic Cardiac Performance Information | User Inconvenience | 1 | Overheating of the RLM | User unable to access information on IC
P2017-131_rY | Software / Algorithm Error | Overestimated Diagnostic Cardiac Performance Information | User Inconvenience | 1 | Web-browser Incompatability | User unable to configure RLM
P2017-131_rY | Software / Algorithm Error | Underestimated Diagnostic Cardiac Performance Information | User Inconvenience | 1 | Access Site Bleeding | Abortion of operation
P2017-131_rY | Software Error | Inability To Adjust Pump Flow Rate | Inadequate Hemodynamic Support | 3 | Air in lumen | Clot formation/Thrombosis
P2017-131_rY | Thermal - High Temperature | Hemolysis, Due To High Temperature | Anemia | 3 | Allergen | Components break inside body
P2017-131_rY | Thermal - High Temperature | Thermal Damage To Ventricle/Vasculature | Heart Injury | 5 | Biological blockage of the system | Components left inside patient
P2017-131_rY | Thrombus | Release of Thrombus From Device | Stroke | 5 | Blockage of the system | Damage to blood cells
P2017-131_rY | Tripping | Patient/User/Bystander Trips On Console Cable | Soft Tissue Injury | 4 | Blood Clotting | Damage to cardiac or vascular structures
P2017-131_rY | Unsecured Device | Loss Of Full Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Blood Damage | Damage to heart/vasculature/tricuspid valve/pulmonary valve
P2017-131_rY | Unsecured Device | Loss Of Partial Support Pump Therapy | Loss Of Hemodynamic Support | 5 | Bond failures | Damage to vasculature
P2017-131_rY | Unsecured Device | Pump Fails To Initiate Prior To Therapy Initiation | Additional Surgical (Case) Intervention | 3 | Components damaged during insertion/removal | Embolism
P2017-131_rY | Unsecured Device | Pump Function Degraded, Due To Suction | Inadequate Hemodynamic Support | 3 | Electrical failure of motor | Excessive Bleeding
P2017-131_rY | Use Error | Non-Sterile Device Used In Patient | Sepsis | 5 | High interaction forces | Excessive Bleeding due to damage to vasculature
P2017-131_rY | Use Error | Unnecessary Troubleshooting | User Inconvenience | 1 | High Temperature | Excessive current delivered to patient
P2017-131_rY | Use Error | User Selects Incorrect P-Level | Inadequate Hemodynamic Support | 3 | Incorrect description/specification for use of device | Hemolysis
P2019-585_rB ECP | Purge System Failure | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Incorrect Placement | Heparin-induced thrombocytopenia (HIT)
P2019-585_rB ECP | Mechanical failure of motor | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Incorrect Programing/Insufficient Power | Interference with cardiac conduction
P2019-585_rB ECP | Electrical failure of motor | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Incorrect Repositioning | Leachable compound in design path
P2019-585_rB ECP | Mechanical blockage of the system | Low Flow | Reduced blood flow throughout the body | 2 | Incorrect system removal | Loss of Optical Sensor signal
P2019-585_rB ECP | Mechanical blockage of the system | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Incorrect system setup | Loss of sensor signal
P2019-585_rB ECP | Blockage of the system | Low Flow | Reduced blood flow throughout the body | 2 | Leakage Current | Low Flow
P2019-585_rB ECP | Biological blockage of the system | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Mechanical blockage of the system | Lymphedema/DVT
P2019-585_rB ECP | Leakage Current | Excessive current delivered to patient | Ventricular Fibrillation/Ventricular tachycardia/Asystole | 4 | Mechanical failure of motor | Malposition of the device
P2019-585_rB ECP | High Temperature | Hot part in skin contact | Skin burn (1st degree) | 3 | Micro-organism | Outlet falls below Pulmonary Valve
P2019-585_rB ECP | Incorrect Placement | Hemolysis | Organ damage | 4 | Misuse of pump metrics | Pathogens present after sterilization
P2019-585_rB ECP | Incorrect Placement | Damage to ventricle/vasculature | Vasculature/Cardiac surgery | 4 | Motor Vibrates during the usage of the device | Patient is weaned too early
P2019-585_rB ECP | Incorrect Placement | Malposition of the device | Loss of hemodynamic support leading to medical intervention | 2 | Optical Sensor System Failure | Placement Signal Failure
P2019-585_rB ECP | Bond failures | Damage to vasculature | Vasculature/Cardiac surgery | 4 | Optimal Purge solution not utilized | Procedure aborted
P2019-585_rB ECP | Bond failures | Components left inside patient | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4 | Pump Setup incorrectly | Prolong skin and vasculature contact with non biocompatible material
P2019-585_rB ECP | Blood Clotting | Clot formation/Thrombosis | Embolism/Stroke | 4 | Purge System Failure | Pump Fails to Start
P2019-585_rB ECP | Blood Clotting | Clot formation/Thrombosis | Loss of hemodynamic support leading to medical intervention | 3 | Reduced Blood Flow from Leg | Pump operates with a purge system leak
P2019-585_rB ECP | Pump Setup incorrectly | Pump Fails to Start | Minor Inconvenience/Delay in Therapy | 1 | Rough Surface | Pump provides greater flow than displayed
P2019-585_rB ECP | Pump Setup incorrectly | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention | 3 | Sensor System Failure | Pump provides lower flow than displayed
P2019-585_rB ECP | Pump Setup incorrectly | Low Flow | Reduced hemodynamic support | 2 | Smart Assist Metrics | Pump provides lower than expected support
P2019-585_rB ECP | Optimal Purge solution not utilized | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 | Stability Failure | Pump Stop/Failed to start after implantation
P2019-585_rB ECP | Air in lumen | Air Embolism | Stroke | 4 | Toxicity | Pump Stop/Removal
P2019-585_rB ECP | Purge System Failure | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 |  | Reversal of forward blood flow
P2019-585_rB ECP | Purge System Failure | Procedure aborted | Delay of Therapy/Inconvenience | 1 |  | Smart Assist metrics are not accurate
P2019-585_rB ECP | Allergen | Heparin-induced thrombocytopenia (HIT) | Organ damage | 4 |  | Unable to insert device leading to procedure being aborted/second pump utilized
P2019-585_rB ECP | Components damaged during insertion/removal | Pump Stop/Failed to start after implantation | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3 |  | User is unable to position/reposition the pump
P2019-585_rB ECP | Toxicity | Leachable compound in design path | Systemic toxic reaction | 4
P2019-585_rB ECP | Toxicity | Prolong skin and vasculature contact with non biocompatible material | Systemic toxic reaction | 4
P2019-585_rB ECP | Micro-organism | Pathogens present after sterilization | Infection/Sepsis | 4
P2019-585_rB ECP | Incorrect description/specification for use of device | Malposition of the device | Reduced hemodynamic support | 2
P2019-585_rB ECP | Incorrect description/specification for use of device | Damage to ventricle/vasculature | Cardiac/Vasculature Surgery | 4
P2019-585_rB ECP | Incorrect description/specification for use of device | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention | 3
P2019-585_rB ECP | Access Site Bleeding | Excessive Bleeding | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex: Femstop or additionally blood products) | 2
P2019-585_rB ECP | Access Site Bleeding | Excessive Bleeding | Bleeding lead to medical intervention which only requires a vasculature closure device (ex. Balloon Tamponade) | 3
P2019-585_rB ECP | Access Site Bleeding | Excessive Bleeding | Excess bleeding requires an unplanned invasive or operable percutaneous intervention (ex. Vasculature surgery to repair femoral artery) | 4
P2019-585_rB ECP | Access Site Bleeding | Excessive Bleeding | Death | 5
P2019-585_rB ECP | Reduced Blood Flow to Leg | Ischemia | Medical intervention to install bypass or remove blockage. | 3
P2019-585_rB ECP | Reduced Blood Flow to Leg | Ischemia | Surgical Removal of Extremity | 4
P2019-585_rB ECP | Incorrect Repositioning | Hemolysis | Organ damage | 4
P2019-585_rB ECP | Incorrect Repositioning | Damage to ventricle/vasculature | Vasculature/Cardiac surgery | 4
P2019-585_rB ECP | Incorrect Repositioning | Malposition of the device | Loss of hemodynamic support leading to medical intervention | 2
P2019-585_rB ECP | Bond failures | Components left inside patient | Medical intervention (ex: Use of Snare) | 3
P2019-585_rB ECP | Incorrect Placement | Damage to ventricle/vasculature | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products) | 2
P2019-585_rB ECP | Incorrect Repositioning | Damage to ventricle/vasculature | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products) | 2
P2019-585_rB ECP | Incorrect description/specification for use of device | Damage to ventricle/vasculature | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products) | 2
P2019-585_rB ECP | Blood Damage | Hemolysis | Organ damage | 4
P2019-585_rB ECP | Blood Damage | Hemolysis | Loss of hemodynamic support leading to medical intervention (ex. Premature removal of device) | 3
P2019-585_rB ECP | Components damaged during insertion/removal | Components break inside body | Loss of hemodynamic support leading to medical intervention (ex: Use of Snare) | 3
P2019-585_rB ECP | Components damaged during insertion/removal | Components break inside body | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4
P2019-585_rB ECP | Components damaged during insertion/removal | Excessive Bleeding | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products, applying manual pressure) | 2
P2019-585_rB ECP | Components damaged during insertion/removal | Excessive Bleeding due to damage to vasculature | Cardiac/Vasculature Surgery | 4
P2019-585_rB ECP | Components damaged during insertion/removal | Unable to insert device leading to procedure being aborted/second pump utilized | Inconvenience or temporary discomfort | 1
P2019-585_rB ECP | Incorrect Programing/Insufficient Power | Pump provides lower than expected support | Procedure disruption causing longer than expected delays | 2
P2019-585_rB ECP | Incorrect Programing/Insufficient Power | Pump Fails to Start | Minor Inconvenience/Delay in Therapy | 1
P2019-585_rB ECP | Blood Damage | Hemolysis | Product Malfunction leading to a temporary injury or impairment and require non-invasive (either operable or percutaneous) treatment. (ex Blood products, applying manual pressure) | 2
P2019-585_rB ECP | Electrical failure of the pump | Pump Fails to Start | Minor Inconvenience/Delay in Therapy | 1
P2019-585_rB ECP | Incorrect Programing/Insufficient Power/ Communication Error | Pump Stop/ Removal | Loss of hemodynamic support leading to medical intervention | 3
P2019-585_rB ECP | Particulates generation | Obstruction of vessels/Thrombosis | Embolism/Stroke | 4
P2019-585_rB ECP | Outflow in ventricle | Pump provides lower than expected support | Procedure disruption causing longer than expected delays | 2
P2019-585_rB ECP | Clot on pump head | Clot prevents pump head from collapsing | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4
P2019-585_rB ECP | Pump contacting LV wall | Interference with cardiac conduction | Arrhythmia | 3
P2019-585_rB ECP | Pump head stuck | Pump head does not collapse | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4
P2019-585_rB ECP | Incorrect Algorithm | False positive alarm | Minor Inconvenience/Delay in Therapy | 1
P2019-585_rB ECP | Console malfunction | AIC shuts down/freezes/ inoperable | Delay of Therapy/Inconvenience | 1
0042-9817_rD Connect | Interference | AIC reboots due to interference | Delay of Therapy/ Inconvenience | 1
0042-9817_rD Connect | Overheating of the RLM | AIC shuts down/freezes/ inoperable | Loss of hemodynamic support leading to medical intervention | 3
0042-9817_rD Connect | AIC Cart falls over | AIC shuts down/freezes/ inoperable | Loss of hemodynamic support leading to medical intervention | 3
0042-9817_rD Connect | Loss or deterioration of AIC's function | AIC shuts down/freezes/ inoperable | Loss of hemodynamic support leading to medical intervention | 3
0042-9817_rD Connect | Loss or deterioration of RLM's user interface functionality | User unable to configure RLM | Inconvenience | 1
0042-9817_rD Connect | LED blink codes | User unable to configure RLM | Inconvenience | 1
0042-9817_rD Connect | Data not transmitted to cloud based server | User unable to access information on IC | Inconvenience | 1
0042-9817_rD Connect | Web-browser Incompatability | User unable to access information on IC | Inconvenience | 1
0042-9817_rD Connect | Loss of Privacy & Security | Users unable to access information on IC. | Inconvenience | 1
0042-9817_rD Connect | Incorrect Display (AIC) | User unable to access information on IC | Inconvenience | 1
0042-9817_rD Connect | Incorrect Display (Abi) | Pump provides inadequate support | Reduced hemodynamic support | 2
0042-9817_rD Connect | Incorrect Output (avg. flow defict and minutes in suction) | User unable to access accurate information on IC | Inconvenience | 1
0042-9817_rD Connect | Integrating new information beyond AIC Screen | User misintrepets information | No patient harm is associated with this use step | N/A
P2018-169_rH RP | Purge System Failure | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Mechanical failure of motor | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Electrical failure of motor | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Mechanical blockage of the system | Low Flow | Reduced blood flow throughout the body | 2
P2018-169_rH RP | Mechanical blockage of the system | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Blockage of the system | Low Flow | Reduced blood flow throughout the body | 2
P2018-169_rH RP | Biological blockage of the system | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Leakage Current | Excessive current delivered to patient | Results in patient death | 5
P2018-169_rH RP | Leakage Current | Excessive current delivered to patient | Ventricular Fibrillation/Ventricular tachycardia/Asystole | 4
P2018-169_rH RP | High Temperature | Damage to blood cells | Organ damage | 4
P2018-169_rH RP | Incorrect Placement | Hemolysis | Organ damage | 4
P2018-169_rH RP | Incorrect Placement | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Incorrect Placement | Malposition of the device | Loss of hemodynamic support leading to medical intervention | 2
P2018-169_rH RP | Bond failures | Damage to vasculature | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Bond failures | Components left inside patient | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4
P2018-169_rH RP | Blood Clotting | Clot formation/Thrombosis | Embolism/Stroke | 4
P2018-169_rH RP | Blood Clotting | Clot formation/Thrombosis | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Pump Setup incorrectly | Pump Fails to Start | Minor Inconvenience/Delay in Therapy | 1
P2018-169_rH RP | Pump Setup incorrectly | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Pump Setup incorrectly | Placement Signal Failure | Reduced hemodynamic support | 2
P2018-169_rH RP | Optimal Purge solution not utilized | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Purge System Failure | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Purge System Failure | Procedure aborted | Delay of Therapy/Inconvenience | 1
P2018-169_rH RP | Allergen | Heparin-induced thrombocytopenia (HIT) | Organ damage | 4
P2018-169_rH RP | Components damaged during insertion/removal | Pump Stop/Failed to start after implantation | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Sensor System Failure | Malposition of the device | Reduced hemodynamic support | 2
P2018-169_rH RP | Sensor System Failure | Loss of sensor signal | Loss of hemodynamic support leading to medical intervention (New pump insertion) | 3
P2018-169_rH RP | Sensor System Failure | Abortion of operation | Delay of Therapy/Inconvenience | 1
P2018-169_rH RP | Toxicity | Leachable compound in design path | Systemic toxic reaction | 4
P2018-169_rH RP | Toxicity | Prolong skin and vasculature contact with non biocompatible material | Systemic toxic reaction | 4
P2018-169_rH RP | Micro-organism | Pathogens present after sterilization | Infection/Sepsis | 4
P2018-169_rH RP | Incorrect description/specification for use of device | Malposition of the device | Reduced hemodynamic support | 2
P2018-169_rH RP | Incorrect description/specification for use of device | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Incorrect description/specification for use of device | Pump Stop/Removal | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Access Site Bleeding | Excessive Bleeding | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Access Site Bleeding | Excessive Bleeding | Death | 5
P2018-169_rH RP | Reduced Blood Flow from Leg | Lymphedema/DVT | Elevation of the leg or device removal | 3
P2018-169_rH RP | Incorrect Repositioning | Hemolysis | Organ damage | 4
P2018-169_rH RP | Incorrect Repositioning | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Incorrect Repositioning | Malposition of the device | Loss of hemodynamic support leading to medical intervention | 2
P2018-169_rH RP | Bond failures | Components left inside patient | Medical intervention (ex: Use of Snare) | 3
P2018-169_rH RP | Incorrect Placement | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Incorrect Repositioning | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Incorrect description/specification for use of device | Damage to heart/vasculature/tricuspid valve/pulmonary valve | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Blood Damage | Hemolysis | Organ damage | 4
P2018-169_rH RP | Blood Damage | Hemolysis | Loss of hemodynamic support leading to medical intervention (ex. Premature removal of device) | 3
P2018-169_rH RP | Components damaged during insertion/removal | Components break inside body | Loss of hemodynamic support leading to medical intervention (ex: Use of Snar) | 3
P2018-169_rH RP | Components damaged during insertion/removal | Components break inside body | Unplanned invasive operable or percutaneous intervention (surgical removal) | 4
P2018-169_rH RP | Components damaged during insertion/removal | Excessive Bleeding | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Components damaged during insertion/removal | Excessive Bleeding due to damage to vasculature | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Components damaged during insertion/removal | Unable to insert device leading to procedure being aborted/second pump utilized | Inconvenience or temporary discomfort | 1
P2018-169_rH RP | Incorrect Programing/Insufficient Power | Pump provides lower than expected support | Procedure disruption causing longer than expected delays | 2
P2018-169_rH RP | Incorrect Programing/Insufficient Power | Pump Fails to Start | Minor Inconvenience/Delay in Therapy | 1
P2018-169_rH RP | Blood Damage | Hemolysis | Temporary injury or impairment and require non-invasive treatment. (ex 1 unit of Blood products, applying manual pressure) | 2
P2018-169_rH RP | Incorrect Placement | Reversal of forward blood flow | Death | 5
P2018-169_rH RP | Misuse of pump metrics | Patient is weaned too early | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Misuse of pump metrics | Pump provides lower than expected support | Procedure disruption causing longer than expected delays | 2
P2018-169_rH RP | Pump Setup incorrectly | Pump provides lower than expected support | Procedure disruption causing longer than expected delays | 2
P2018-169_rH RP | Optical Sensor System Failure | Loss of Optical Sensor signal | Inconvenience or temporary discomfort | 1
P2018-169_rH RP | Sensor System Failure | Pump provides lower flow than displayed | Medical interention required | 3
P2018-169_rH RP | Sensor System Failure | Pump provides greater flow than displayed | Pulmonary Edema | 3
P2018-169_rH RP | Smart Assist Metrics | Smart Assist metrics are not accurate | Inconvenience or temporary discomfort | 1
P2018-169_rH RP | Air in lumen | Embolism | Pulmonary Embolism | 2
P2018-169_rH RP | Access Site Bleeding | Excessive Bleeding | Excess bleeding requires an unplanned invasive or operable percutaneous intervention (ex. Vasculature surgery to repair vein) | 3
P2018-169_rH RP | Incorrect system setup | Abortion of operation | Delay of Therapy/Inconvenience | 1
P2018-169_rH RP | Incorrect system removal | Damage to cardiac or vascular structures | Cardiac/Vascular Surgery | 4
P2018-169_rH RP | Motor Vibrates during the usage of the device | Hemolysis | Organ damage | 3
P2018-169_rH RP | Motor Vibrates during the usage of the device | Interference with cardiac conduction | Transient Arrythmia | 1
P2018-169_rH RP | Stability Failure | Outlet falls below Pulmonary Valve | Loss of hemodynamic support leading to medical intervention | 3
P2018-169_rH RP | Rough Surface | Damage to cardiac or vascular structures | Cardiac/Vascular surgery | 4
P2018-169_rH RP | High interaction forces | User is unable to position/reposition the pump | Inconvenience or temporary discomfort | 1
P2018-169_rH RP | Purge System Failure | Pump operates with a purge system leak | Inconvenience/Pump remains functional | 1

## Hazard_Uni
Hazard (total 109) | Hazard (simply 101)
Access Site Bleeding | Access Site Bleeding
Active Pump Removed | Active Pump Removed
AIC Alarm Failure/Delay | AIC Alarm Failure/Delay
AIC Cart falls over | AIC Cart falls over
AIC Failure | AIC Inoperable (failure, startup)
AIC Inoperable | AIC Purge Cassette Door
AIC Purge Cassette Door | Air in lumen
AIC Startup Failure | Algorithm Error
Air in lumen | Allergen
Allergen | Automated Flow Rate Adjustment
Automated Flow Rate Adjustment | Battery Overheating / Failure
Battery Overheating / Failure | Biological blockage of the system
Biological blockage of the system | Blockage of the system
Blockage of the system | Blood Clotting (e.g. pump head)
Blood Clotting | Blood Damage
Blood Damage | Bond failures
Bond failures | Cage Separation
Cannula/Cage Separation | Cannula Separation
Catheter Separation | Catheter Separation
Clot on pump head | Communication Error
Components damaged during insertion/removal | Components damaged during insertion/removal
Console malfunction | Console malfunction
Data not transmitted to cloud based server | Data not transmitted to cloud based server
Device Compatibility | Device Compatibility
Disruption of Blood Flow | Disruption of Blood Flow
Disruption Of Blood Flow Near Access Site | Disruption Of Blood Flow Near Access Site
Electrical - Leakage Current | Electrical - Leakage Current
Electrical failure of motor | Electrical failure of motor
Electrical failure of the pump | Electrical failure of the pump
Electrical Pump Motor Failure | Electrical Pump Motor Failure
Falling Object | Falling Object
False Alarm/ Notification | False Alarm/ Notification
Foreign Material | Foreign Material
High interaction forces | High interaction forces
High Restistance Between Devices | High Restistance Between Devices
High Temperature | High Temperature
Impella Interaction With Other Device | Impella Interaction With Other Device
Inadequate Priming Of Purge Line | Inadequate Priming Of Purge Line
Inadequate Pump Visualization | Inadequate Pump Visualization
Incorrect Algorithm | Incorrect Algorithm
Incorrect description/specification for use of device | Incorrect description/specification for use of device
Incorrect Display (Abi) | Incorrect Display (Abi, AIC)
Incorrect Display (AIC) | Incorrect Output (avg. flow defict and minutes in suction)
Incorrect Output (avg. flow defict and minutes in suction) | Incorrect Placement (e.g. graft)
Incorrect Placement | Incorrect Programing
Incorrect Placement Of Graft | Incorrect Pump Insertion
Incorrect Programing/Insufficient Power | Incorrect Pump Performance Metrics
Incorrect Programing/Insufficient Power/ Communication Error | Incorrect Pump Setup
Incorrect Pump Insertion | Incorrect Purge Solution Selection
Incorrect Pump Performance Metrics | Incorrect Repositioning
Incorrect Pump Setup | Incorrect system removal
Incorrect Purge Solution Selection | Incorrect system setup
Incorrect Repositioning | Insufficient Power
Incorrect system removal | Integrating new information beyond AIC Screen
Incorrect system setup | Interference
Integrating new information beyond AIC Screen | Leakage Current
Interference | LED blink codes
Introducer /Sheath Mechanical Issue | Loss Of Battery Power
Leakage Current | Loss of Privacy & Security
LED blink codes | Loss or deterioration of AIC's function
Loss Of Battery Power | Loss or deterioration of RLM's user interface functionality
Loss of Privacy & Security | Mechanical blockage of the system
Loss or deterioration of AIC's function | Mechanical failure of motor
Loss or deterioration of RLM's user interface functionality | Mechanical Issue (e.g. introducer, sheath)
Mechanical blockage of the system | Mechanical Pump Motor Failure
Mechanical failure of motor | Micro-organism (e.g. Pathogen)
Mechanical Pump Motor Failure | Misuse of pump metrics
Micro-organism | Motor Vibrates during the usage of the device
Micro-Organism/Pathogen | Non-Biocompatibility
Misuse of pump metrics | Obstruction Of Impella Flow Path
Motor Vibrates during the usage of the device | Optical Sensor Failure (e.g.error, system failure)
Non-Biocompatibility | Optimal Purge solution not utilized
Obstruction Of Impella Flow Path | Outflow in ventricle
Optical Sensor Error | Overheating of the RLM
Optical Sensor Failure | Partial Obstruction Of Impella Flow Path
Optical Sensor System Failure | Particulates generation
Optimal Purge solution not utilized | Pigtail Separation
Outflow in ventricle | Pressure Storage Set Leak/Failure
Overheating of the RLM | Product Packaging Integrity
Partial Obstruction Of Impella Flow Path | Programming Error
Particulates generation | Pump contacting LV wall
Pigtail Separation | Pump Failure
Pressure Storage Set Leak/Failure | Pump head stuck
Product Packaging Integrity | Pump Setup incorrectly
Programming / Software Error | Purge Cassette Failure
Pump contacting LV wall | Purge System Failure
Pump Failure | Recoverable AIC Failure
Pump head stuck | Reduced Blood Flow (e.g. to or from Leg)
Pump Setup incorrectly | Rough Surface
Purge Cassette Failure | Sharp Edges
Purge System Failure | Small Component Dislodge From Device
Recoverable AIC Failure | Smart Assist Metrics
Reduced Blood Flow from Leg | Software Error
Reduced Blood Flow to Leg | Stability Failure
Rough Surface | Thermal - High Temperature
Sensor System Failure | Thrombus
Sharp Edges | Toxicity
Small Component Dislodge From Device | Tripping
Smart Assist Metrics | Unsecured Device
Software / Algorithm Error | Use Error
Software Error | Web-browser Incompatability
Stability Failure
Thermal - High Temperature
Thrombus
Toxicity
Tripping
Unsecured Device
Use Error
Web-browser Incompatability

## Abimed
hazard_abiomed
Access Site Bleeding
Active Pump Removed
AIC Alarm Failure/Delay
AIC Inoperable (failure, startup)
AIC Purge Cassette Door
Air in lumen
Algorithm Error
Automated Flow Rate Adjustment
Battery Overheating / Failure
Communication Error
Console malfunction
Data not transmitted to cloud based server
Device Compatibility
Electrical - Leakage Current
Electrical failure of motor
Electrical failure of the pump
Electrical Pump Motor Failure
Falling Object
False Alarm/ Notification
High Temperature
Inadequate Priming Of Purge Line
Inadequate Pump Visualization
Incorrect Algorithm
Incorrect Display (Abi, AIC)
Incorrect Output (avg. flow defict and minutes in suction)
Incorrect Placement (e.g. graft)
Incorrect Programing
Incorrect Purge Solution Selection
Incorrect system removal
Incorrect system setup
Insufficient Power
Integrating new information beyond AIC Screen
Leakage Current
LED blink codes
Loss Of Battery Power
Loss of Privacy & Security
Loss or deterioration of AIC's function
Loss or deterioration of RLM's user interface functionality
Optimal Purge solution not utilized
Overheating of the RLM
Programming Error
Purge Cassette Failure
Purge System Failure
Recoverable AIC Failure
Software Error
Thermal - High Temperature
Tripping
Unsecured Device
Use Error
Web-browser Incompatability

## FDA
hazard_FDA
device alarm system issue
application interface becomes nonfunctional or program exits abnormally
application program issue
application program version or upgrade problem
application security issue
battery issue
low battery
communication or transmission issue
component or accessory incompatibility
computer software issue
computer system security issue
data issue
detachment of device component
device alarm system issue
device inoperable
device maintenance issue
device markings issue
device or device component damaged by another device
device unsafe to use in environment
disconnection
electrical issue
electrical shorting
electromagnetic interference (EMI)
failure to flush
failure to power-up
failure to prime
failure to reset
failure to service
false alarm
human-device interface issue
image display error
improper alarm
inadequate user interface
incorrect display
installation-related problem
invalid sensing
loss of data
no apparent adverse event
no display or display failure
output issue
problem with software installation
protective measure issue
pumping issue
radio signal problem
radiofrequency interference (RFI)
unauthorized access to computer system
use of device issue
use of incorrect control settings
user used incorrect product for intended use
wireless communication issue

## ABIOMED->FDA
hazard_abiomed | hazard_FDA
AIC Inoperable (failure, startup)/Recoverable AIC Failure (device inoperable, failure to power-up, failure to reset) | application interface becomes nonfunctional or program exits abnormally
AIC Purge Cassette Door (break, difficult to open or close, device or device component damaged by another device) | application program issue
Air in lumen (obstruction within device) | application program version or upgrade problem
Algorithm Error/Incorrect Algorithm (parameter calculation error due to software problem, application program issue) | application security issue
Automated Flow Rate Adjustment (improper flow or infusion, application program issue) | battery issue
Battery Overheating / Failure (battery issue, low battery) | break
Communication Error (communication or transmission issue, wireless communication issue) | chemical issue
Console malfunction (device inoperable, failure to service) | communication or transmission issue
Data not transmitted to cloud based server (data issue, wireless communication issue, loss of data) | component falling
Device Compatibility (device-device incompatibility, component or accessory incompatibility) | component or accessory incompatibility
Electrical - Leakage Current/Leakage Current (electrical issue) | computer operating system issue
Falling Object (component falling) | computer software issue
False Alarm/ Notification (device alarm system issue, false alarm, improper alarm) | computer system security issue
High Temperature (overheating of device or device component) | data issue
Inadequate Priming Of Purge Line (failure to prime, failure to flush) | detachment of device component
Inadequate Pump Visualization (image display error) | device alarm system issue
Incorrect Display (Abi, AIC) (incorrect display, no display or display failure) | device inoperable
Incorrect Output (avg. flow defict and minutes in suction) (improper flow or infusion, improper device output) | device maintenance issue
Incorrect Placement (e.g. graft) (malposition of device) | device markings issue
Incorrect Programing/Programming Error (application program issue) | device or device component damaged by another device
Incorrect Purge Solution Selection (chemical issue) | device unsafe to use in environment
Insufficient Power (loss of power, failure to power-up) | device-device incompatibility
Integrating new information beyond AIC Screen (incorrect display) | difficult to open or close
LED blink codes (incorrect display) | disconnection
Loss Of Battery Power (low battery) | electrical issue
Loss of Privacy & Security (computer system security issue, application security issue, unauthorized access to computer system) | electrical shorting
Loss or deterioration of AIC's function (application interface becomes nonfunctional or program exits abnormally) | electromagnetic interference (EMI)
Loss or deterioration of RLM's user interface functionality (application interface becomes nonfunctional or program exits abnormally) | failure to flush
Optimal Purge solution not utilized (improper flow or infusion, failure to prime, failure to flush, use of incorrect control settings) | failure to power-up
Overheating of the RLM (overheating of device or device component) | failure to prime
Purge Cassette Failure (break, fluid leak) | failure to reset
Purge System Failure (pressure issue) | failure to service
Software Error (computer software issue, application program version or upgrade problem) | false alarm
Thermal - High Temperature (overheating of device or device component) | fluid leak
Tripping (tip-over) | human-device interface issue
Unsecured Device (unstable, detachment of device component, disconnection) | image display error
Use Error (use of device issue, human-device interface issue, inadequate user interface) | improper alarm
Web-browser Incompatability (computer operating system issue) | improper device output
improper flow or infusion
inadequate user interface
incorrect display
incorrect display
installation-related problem
electrical shorting | loss of data
electromagnetic interference (EMI) | loss of power
installation-related problem | low battery
problem with software installation | malposition of device
pumping issue | no display or display failure
radio signal problem | obstruction within device
radiofrequency interference (RFI) | overheating of device or device component
device maintenance issue | parameter calculation error due to software problem
device markings issue | pressure issue
device unsafe to use in environment | problem with software installation
pumping issue
radio signal problem
radiofrequency interference (RFI)
tip-over
unauthorized access to computer system
unstable
use of device issue
use of incorrect control settings
wireless communication issue

## FDA->ABIOMED
FDA | Abiomed
application interface becomes nonfunctional or program exits abnormally | Loss or deterioration of AIC's function, Loss or deterioration of RLM's user interface functionality
application program issue | Algorithm Error, Incorrect Algorithm, Automated Flow Rate Adjustment, Incorrect Programing, Programming Error
application program version or upgrade problem | Software Error
application security issue | Loss of Privacy & Security
battery issue | Battery Overheating, Battery Overheating Failure
break | AIC Purge Cassette Door, Purge Cassette Failure
chemical issue | Incorrect Purge Solution Selection
communication or transmission issue | Communication Error
component falling | Falling Object
component or accessory incompatibility | Device Compatibility
computer operating system issue | Web-browser Incompatability
computer software issue | Software Error
computer system security issue | Loss of Privacy & Security
data issue | Data not transmitted to cloud based server
detachment of device component | Unsecured Device
device alarm system issue | False Alarm/ Notification
device inoperable | AIC Inoperable (failure, startup), Recoverable AIC Failure, Console malfunction
device or device component damaged by another device | AIC Purge Cassette Door
device-device incompatibility | Device Compatibility
difficult to open or close | AIC Purge Cassette Door
disconnection | Unsecured Device
electrical issue | Leakage Current
failure to flush | Inadequate Priming Of Purge Line, Optimal Purge solution not utilized
failure to power-up | AIC Inoperable (failure, startup), Recoverable AIC Failure, Insufficient Power
failure to prime | Inadequate Priming Of Purge Line, Optimal Purge solution not utilized
failure to reset | AIC Inoperable (failure, startup), Recoverable AIC Failure
failure to service | Console malfunction
false alarm | False Alarm/ Notification
fluid leak | Purge Cassette Failure
human-device interface issue | Use Error
image display error | Inadequate Pump Visualization
improper alarm | False Alarm/ Notification
improper device output | Incorrect Output (avg. flow defict and minutes in suction)
improper flow or infusion | Automated Flow Rate Adjustment, Incorrect Output (avg. flow defict and minutes in suction), Optimal Purge solution not utilized
inadequate user interface | Use Error
incorrect display | Incorrect Display (Abi, AIC), Integrating new information beyond AIC Screen, LED blink codes
loss of data | Data not transmitted to cloud based server
loss of power | Insufficient Power
low battery | Battery Overheating, Battery Overheating Failure, Loss Of Battery Power
malposition of device | Incorrect Placement (e.g. graft)
no display or display failure | Incorrect Display (Abi, AIC)
obstruction within device | Air in lumen
overheating of device or device component | High Temperature, Overheating of the RLM, Thermal - High Temperature
parameter calculation error due to software problem | Algorithm Error, Incorrect Algorithm
pressure issue | Purge System Failure
tip-over | Tripping
unauthorized access to computer system | Loss of Privacy & Security
unstable | Unsecured Device
use of device issue | Use Error
use of incorrect control settings | Optimal Purge solution not utilized
wireless communication issue | Communication Error, Data not transmitted to cloud based server